https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140623&pc=MTA | 2023826-2020-01194 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND NARROWING OF THE ANTERIOR CHAMBER ANGLE. THE PROBLEM WAS RESOLVED. PRIOR TO EXCHANGE THE LENS HAD BEEN REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4) Manufacturer Narrative: CORRECTED DATA: D9- "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2021" IN THE PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10168247&pc=MTA | 2023826-2020-01306 | 2020-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242535&pc=MTA | 2023826-2020-01430 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10245041&pc=MTA | 2023826-2020-01401 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION- LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1-ADVERSE EVENT. B2- REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON THE SAME DAY (B)(6) 2020 BUT SECONDARY SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. H1- MALFUNCTION CORRECTED TO SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398561&pc=MTA | 2023826-2020-01749 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403845&pc=MTA | 2023826-2020-01736 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403911&pc=MTA | 2023826-2020-01735 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY, WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEFORMED WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10236679&pc=MTA | 2023826-2020-01442 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -13.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254213&pc=MTA | 2023826-2020-01449 | 2020-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10224702&pc=MTA | 2023826-2020-01393 | 2020-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10236322&pc=MTA | 2023826-2020-01441 | 2020-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, INSIDE A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, RESIDUE ON THE LENS SURFACE WAS OBSERVED. CLAIM# (B)(4). Manufacturer Narrative: B5 - EXCHANGE DATE SHOULD BE (B)(6) 2020, IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10224706&pc=MTA | 2023826-2020-01394 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242539&pc=MTA | 2023826-2020-01429 | 2020-06-02 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10245037&pc=MTA | 2023826-2020-01437 | 2020-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10247427&pc=MTA | 2023826-2020-01467 | 2020-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS DEFORMED AND RESIDUE ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189471&pc=MTA | 2023826-2020-01381 | 2020-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. THE LENS WAS A FUSCHIA COLOR. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINTS TYPE EVENTS FOUND WITHIN ASSOCIATED LOTS. # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189567&pc=MTA | 2023826-2020-01382 | 2020-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. THE LENS IS A FUSCHIA COLOR. CLAIM# (B)(4). Manufacturer Narrative: B3- CORRECTED TO UNK IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10247547&pc=MTA | 2023826-2020-01468 | 2020-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242516&pc=MTA | 2023826-2020-01432 | 2020-06-05 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10246120&pc=MTA | 2023826-2020-01463 | 2020-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285769&pc=MTA | 2023826-2020-01489 | 2020-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285773&pc=MTA | 2023826-2020-01490 | 2020-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451134&pc=MTA | 2023826-2020-01863 | 2020-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426600&pc=MTA | 2023826-2020-01764 | 2020-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335310&pc=MTA | 2023826-2020-01556 | 2020-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392499&pc=MTA | 2023826-2020-01674 | 2020-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476195&pc=MTA | 2023826-2020-01931 | 2020-06-20 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10277692&pc=MTA | 2023826-2020-01499 | 2020-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON (B)(6) 2020 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED, AND THE PROBLEM RESOLVED. PATIENT IS HAPPY. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354274&pc=MTA | 2023826-2020-01613 | 2020-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUAGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10369902&pc=MTA | 2023826-2020-01679 | 2020-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -2.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO EXCESSIVE VAULT. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER REPORTED "TEMPORAL CHAFING/TRANSILLUMINATION DEFECT (FROM IRIS PROLAPSE) AND STATED THAT THE CAUSE OF IT WAS THAT THE LENS WAS TOO BIG FOR THE PATIENTS EYE, THE LENS APPEARED BULGING OUT AFTER IMPLANTATION AND PATIENT WAS NAUSEATING. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED ON THE SAME DAY. PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -2.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE REPORTER REPORTED "TEMPORAL CHAFING/TRANSILLUMINATION DEFECT (FROM IRIS PROLAPSE) AND STATED THAT THE CAUSE OF IT WAS THAT THE LENS WAS TOO BIG FOR THE PATIENTS EYE, THE LENS APPEARED BULGING OUT AFTER IMPLANTATION AND PATIENT WAS NAUSEATING. THE LENS WAS EXCHANGED ON THE SAME DAY DIFFERENT SURGERY FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332982&pc=MTA | 2023826-2020-01577 | 2020-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476250&pc=MTA | 2023826-2020-01937 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354270&pc=MTA | 2023826-2020-01611 | 2020-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429123&pc=MTA | 2023826-2020-01840 | 2020-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10366023&pc=MTA | 2023826-2020-01657 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED WITH MOISTURE, IN A MICROCENTRIFUGE VIAL. THERE WAS DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457349&pc=MTA | 2023826-2020-01879 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398470&pc=MTA | 2023826-2020-01658 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429078&pc=MTA | 2023826-2020-01835 | 2020-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REFERENCE (B)(4) FOR REPLACEMENT LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10433118&pc=MTA | 2023826-2020-01800 | 2020-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445903&pc=MTA | 2023826-2020-01820 | 2020-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476283&pc=MTA | 2023826-2020-01924 | 2020-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10492093&pc=MTA | 2023826-2020-02017 | 2020-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -12.6 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE IN A STERILITY POUCH. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10480713&pc=MTA | 2023826-2020-02002 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10480716&pc=MTA | 2023826-2020-02001 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491980&pc=MTA | 2023826-2020-01962 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH AN ARTISAN LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND DEBRIS ON THE LENS. VISUAL INPSECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CORRECTED DATA: H6 - 3191: LENS EXCHANGE SHOULD BE CORRECTED TO 4627 LENS EXPLANT. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K-THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4820211&pc=MTA | 2023826-2015-00642 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). NEW INCISION. SECONDARY SURGERY. EXPLANTED. VAULTING. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - MOSTLY LIKELY THE EVENT WAS CAUSED BY WRONG LENS SIZING. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4820234&pc=MTA | 2023826-2015-00643 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - MOSTLY LIKELY THE EVENT WAS CAUSED BY WRONG LENS SIZING. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: EVALUATION: RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND BOTH HAPTICS TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885853&pc=MTA | 2023826-2015-00795 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT AGE AND WEIGHT: UNK. EVENT DATE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS DRIED SURGICAL RESIDUE ON THE LENS SURFACE. CONCLUSIONS: UNABLE TO CONFIRM COMPLAINT, BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - ACCURATE DETERMINATION OF ANTERIOR SEGMENT DIMENSIONS IS KEY TO MAINTENANCE OF ADEQUATE VAULT. THE METHOD OF ICL SIZING FOR THIS PATIENT WAS BASED ON ACD AND WTW RECOMMENDATIONS PER DFU. THE REPORT PROVIDES NO SPECIFIC INFORMATION ON THE DEGREE OF VAULTING, OR THE CAUSE OF EXCESSIVE VAULT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5985008&pc=MTA | 2023826-2016-01332 |  | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: IMPLANT DATE WAS NOT VERIFIED. DEVICE EVALUATION: VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S.(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5987383&pc=MTA | 2023826-2016-01338 |  | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS - -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4) AS PER DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIAL TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM IS CONTRAINDICATED. DEVICE EVALUATION-LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND DRIED DISCOLORED FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6001928&pc=MTA | 2023826-2016-01397 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +7.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND UNSPECIFIED FOREIGN MATERIAL ON SURFACE AND HAPTIC TORN). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041624&pc=MTA | 2023826-2016-01494 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS - -11.50/+3.0/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6091604&pc=MTA | 2023826-2016-01602 |  | Injury | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS, -9.5/+2.50/091 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (NOT POSSIBLE TO SPECIFY) ON LENS SURFACE. CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094684&pc=MTA | 2023826-2016-01586 |  | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+2.0/086 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FIBERS ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6146489&pc=MTA | 2023826-2016-01719 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN AS WELL AS FOREIGN MATERIAL ON LENS SURFACE (FIBERS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6146492&pc=MTA | 2023826-2016-01720 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6156877&pc=MTA | 2023826-2016-01713 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S EYE. THE REPORTER INDICATED THE LENS HAD LITTLE TO NO VAULT AND WILL BE EXPLANTED AND EXCHANGED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6156880&pc=MTA | 2023826-2016-01714 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S EYE. THE REPORTER INDICATED THE LENS HAD LITTLE TO NO VAULT AND WILL BE EXPLANTED AND EXCHANGED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6233642&pc=MTA | 2023826-2017-00035 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: AFTER EXCHANGE OF LENS, THE VAULT VALUE CHANGED FROM 680 MICROMETERS TO 350 MICROMETERS. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. PERFORMED VISUAL INSPECTION AND FOUND THE HAPTIC TO BE BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6234140&pc=MTA | 2023826-2017-00037 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.5/111 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE VAULT VALUE WENT FROM 1215 MICROMETERS TO 452 MICROMETERS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DRIED, DISCOLORED MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6246839&pc=MTA | 2023826-2017-00048 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+0.5/59 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2015, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE LENS HAS BEEN RETURNED HOWEVER, AS OF REPORT SUBMISSION DATE, IT HAS NOT YET BEEN EVALUATED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+0.5/59 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT (VAULT VALUE: 1182 UM) AND WAS EXCHANGED ON (B)(6) 2016 FOR A SHORTER SAME MODEL, SIMILAR DIOPTER LENS. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS REPORTED AS 20/25 WITH VAULT VALUE = 450 UM. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6260605&pc=MTA | 2023826-2017-00100 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. AS OF THE DATE OF MDR REPORT SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6326258&pc=MTA | 2023826-2017-00103 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS "EXCHANGED FOR LONGER." NO OTHER INFORMATION IS AVAILABLE AT THIS TIME, MULTIPLE ATTEMPTS MADE TO CONTACT CUSTOMER. Manufacturer Narrative: ADVERSE EVENT WAS "OTHER SERIOUS (IMPORTANT MEDICAL EVENTS), NOT "LIFE-THREATENING". (B)(4). Manufacturer Narrative: MANUFACTURED IN U.S. BUT NOT DISTRIBUTED IN U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6370477&pc=MTA | 2023826-2017-00307 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON 12/13/2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION-LENS RETURNED IN A SMALL CENTRIFUGE TUBE WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. FOREIGN MATERIAL (CLEAR AND RED RESIDUE) WAS FOUND ON LENS SURFACE. LENS IS CURVED DUE TO THE POSITION IT DRIED IN THE CENTRIFUGE TUBE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGICAL INTERVENTION, LENS WAS EXCHANGED FOR SHORTER LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6426165&pc=MTA | 2023826-2017-00472 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE PATIENT HAD A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IMPLANTED IN HER RIGHT EYE (OD), ON (B)(6) 2016, BY A SURGEON IN (B)(6). THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, BY A SURGEON IN (B)(6). THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A TEAR ON THE LENS OPTIC. A PIECE OF LENS OPTIC AND HAPTIC WERE TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID AND THERE WAS CLEAR RESIDUE ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6432222&pc=MTA | 2023826-2017-00438 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11/1.0/139 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6521050&pc=MTA | 2023826-2017-00694 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER INDICATED A PROBABLE CAUSE OF THE EVENT WAS INACCURATE WHITE-TO-WHITE MEASUREMENT. THE PATIENTS POST-OP BSCVA WAS 20/20. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE REPORTER INDICATED THE PATIENT POST-OP WAS EXPERIENCING GLARE/HALO ISSUES. THE PATIENT WAS MEDICALLY TREATED WITH ALPHAGAN, WHICH FAILED TO HELP THE PATIENT. THE PATIENT WAS THEN TREATED WITH PILOCARPINE 1%, WHICH HAD A MIOTIC EFFECT. (B)(4). Manufacturer Narrative: (B)(4): SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6527428&pc=MTA | 2023826-2017-00722 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/76 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PIGMENT DISPERSION, AND UNDERWENT THE SECONDARY YAG. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH OCULAR HYPERTENSION IN EITHER EYE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546100&pc=MTA | 2023826-2017-00766 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED EXCESSIVE VAULTING, NARROWING OF THE IRIDO-CORNEAL ANGLES. THE DOCTOR PLANNED TO EXCHANGE TO A SMALLER LENS AND SAID THAT THE PROBLEM IS DUE TO W-T-W MEASUREMENT. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. PATIENT DOES NOT UNDERSTAND WHY EXCHANGE IS NECESSARY WHEN VISION IS GOOD. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546101&pc=MTA | 2023826-2017-00767 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/X008 IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED EXCESSIVE VAULTING, NARROWING OF THE IRIDO-CORNEAL ANGLES. THE DOCTOR PLANS TO EXCHANGE TO A SMALLER LENS AND SAID THAT THE PROBLEM IS DUE TO W-T-W MEASUREMENT. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. PATIENT DOES NOT UNDERSTAND WHY EXCHANGE IS NECESSARY WHEN VISION IS GOOD. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577113&pc=MTA | 2023826-2017-00851 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN. THE REPORTER INDICATED THE LENS HAD AN OVER VAULT OF 2 1/2 WITH TEMPORAL IRIS DEFECT. THE LENS WILL BE EXPLANTED AND EXCHANGED FOR A 12.6MM, ALSO HE WILL ATTEMPT TO REPAIR THE TORN IRIS. THE VAULT WAS ONLY HIGH AT THE IRIS DEFECT. THERE WERE NO SIGNS OF ANY CYST OR PROBLEM WITH THE ICL WITH UBM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6714543&pc=MTA | 2023826-2017-01189 |  | Injury | STAAR SURGICAL COMPANY | 2017-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON EXPLANTED A MICL13.7 IMPLANTABLE COLLAMER LENS DUE TO INCORRECT SIZING. THE REPORTER INDICATED THAT THE 13.7MM MICL CAUSED THE PATIENT TO GO INTO ANGLE CLOSURE. THE SURGEON PERFORMED A SURGICAL PI ( PERIPHERAL IRIDECTOMY) WHICH DID NOT REMEDY THE SITUATION. THE WHITE TO WHITE WAS REMEASURED AND A SHORTER LENS LENGTH (12.6MM) WAS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6767809&pc=MTA | 2023826-2017-01271 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER. IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ONE WEEK AFTER THE LENS WAS IMPLANTED, THE LENS HAD SOME RESIDUAL VAULTING. THE PRESSURE WAS GOOD, SO THE SURGEON AUGMENTED THE PERIPHERAL IRIDECTOMIES. THE VAULT AND PATIENT WERE FINE AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6811211&pc=MTA | 2023826-2017-01341 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 , -7.5/+1.5/112 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO PATIENT EXPERIENCED EXCESSIVE VAULT. ATTEMPTS TO GET MORE INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832555&pc=MTA | 2023826-2017-01386 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: ON (B)(6) 2017 THE REPORTER INDICATED THAT THE SURGEON PERFORMED A LENS EXCHANGE. THE ORIGINALLY IMPLANTED LENS (12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS) HAD BEEN REMOVED AND A "RESIZED" LENS WAS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS WERE FOUND. CLAIM# (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832594&pc=MTA | 2023826-2017-01387 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 IT WAS REPORTED THAT THE LENS WAS EXHIBITING EXCESSIVE VAULT AND THE PATIENT EXPERIENCING A SHALLOW ANTERIOR CHAMBER. THE LENS IS PLANNED TO BE EXCHANGED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND DEFORMED WITH CLEAR RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.884MM)]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836404&pc=MTA | 2023826-2017-01384 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/161 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED LENS HAD EXCESSIVE VAULTING, CROWDED ANGLES AND THE ANTERIOR CHAMBER WAS VERY SHALLOW. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/161 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED LENS HAD EXCESSIVE VAULTING, CROWDED ANGLES AND THE ANTERIOR CHAMBER WAS VERY SHALLOW. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY NOW. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEBRIS ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955558&pc=MTA | 2023826-2017-01580 |  | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO LOW VAULT. Manufacturer Narrative: ADDITIONAL INFORMATION: THE ORIGINAL LENS WAS EXCHANGED WITH AN ARTISAN LENS. DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE:OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.92MM)] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955563&pc=MTA | 2023826-2017-01581 |  | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO LOW VAULT. Manufacturer Narrative: LENS WAS EXCHANGED WITH AN ARTISAN LENS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: 24-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.89MM)] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6998023&pc=MTA | 2023826-2017-01684 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.5/+4.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF LENS HAPTIC BROKEN AND BENT. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7112915&pc=MTA | 2023826-2017-01858 |  | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND THE LENS HAD A CURVED SHAPE. THERE WAS EVIDENCE OF DEBRIS AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7177073&pc=MTA | 2023826-2018-00005 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS FOUND TO HAVE AN EXCESSIVE VAULT. THE SURGEON DILATED THE PATIENT POST-OPERATIVELY BUT THERE WAS NO CHANGE TO THE VAULT AND THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE SURGEON DILATED THE PATIENT POST-OPERATIVELY BUT THERE WAS NO CHANGE IN VAULT. THE LENS WAS SUBSEQUENTLY EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7211131&pc=MTA | 2023826-2018-00118 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -7.5/+1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SHALLOW AC, AND THE IRIS BEING PUSHED FORWARD. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7308076&pc=MTA | 2023826-2018-00309 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR LONGER LENS (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7322043&pc=MTA | 2023826-2018-00333 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A (B)(4) BAG. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7346769&pc=MTA | 2023826-2018-00419 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, DUE TO HIGH VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7508610&pc=MTA | 2023826-2018-00721 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. REPORT SOURCE: USER FACILITY IS INCORRECT AND SHOULD BE DISTRIBUTOR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535055&pc=MTA | 2023826-2018-00784 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS THAT, AT THE 1 WEEK POST-OP FOLLOW-UP VISIT, THE ICL NEEDED TO HAVE THE POSITION ADJUSTED DUE TO HIGH VAULT (>1100UM). REPORTEDLY, AFTER THE ADJUSTMENT, THE VAULT DECREASED TO LESS THAN 1000UM. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535082&pc=MTA | 2023826-2018-00786 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS THAT, AT THE 1 WEEK POST-OP FOLLOW-UP VISIT, THE ICL NEEDED TO HAVE THE POSITION ADJUSTED DUE TO HIGH VAULT (>1100UM). REPORTEDLY, AFTER THE ADJUSTMENT, THE VAULT DECREASED TO LESS THAN 1000UM. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557876&pc=MTA | 2023826-2018-00815 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -15.00, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES A "HIGH VAULT AND OVERSIZED LENS." THE LENS WAS EXPLANTED (DATE OF EXPLANT UNKNOWN) AND REPLACED WITH AN IOL LENS. Manufacturer Narrative: PREVIOUSLY STATED THE LENS EXPLANT DATE IS UNKNOWN. ADDITIONAL INFORMATION-THE LENS WAS EXPLANTED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7558966&pc=MTA | 2023826-2018-00833 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED DUE TO BEING TOO LONG. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4) Manufacturer Narrative: PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7562916&pc=MTA | 2023826-2018-00837 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE: SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS PATIENT CODE: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7629864&pc=MTA | 2023826-2018-00948 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE REPORTER INDICATED THEY FELT THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7631786&pc=MTA | 2023826-2018-00949 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON HAS HAD TO EXPLANT 3 - 4 IMPLANTABLE COLLAMER LENSES OVER THE LAST COUPLE OF YEARS DUE TO LENS SIZING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7786610&pc=MTA | 2023826-2018-01217 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/5.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO ASYMMETRICAL VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7903841&pc=MTA | 2023826-2018-01391 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/+3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8030339&pc=MTA | 2023826-2018-01588 |  | Injury | STAAR SURGICAL COMPANY | 2018-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE "SUREGON" ATTEMPTED TO IMPLANT A 13.7MM, MICL13.7, -6.5 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) AND IT DID NOT FIT (TOO LARGE) SO HE EXPLANTED THE LENS AND IMPLANTED A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8040072&pc=MTA | 2023826-2018-01610 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-11-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -13.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AS REPLACEMENT LENS BUT IT DID NOT RESOLVE THE PROBLEM COMPLETELY. THE AC (ANTERIOR CHAMBER) REMAINS VERY SHALLOW, BUT IT IS NO LONGER EXTREMELY SHALLOW AND THREATENING THE CORNEAL ENDOTHELIUM. EVEN WITH A SMALLER LENS THAN CALL FOR (DESPITE REPEAT CONFIRMATION OF CORRECT WTW) THE NEW SMALLER LENS REMAIN EXCESSIVELY VAULTED. THE PATIENT IS DOING VERY WELL. PATIENT IS HAPPY WITH HER VISION AND HER IOP (INTRAOCULAR PRESSURE) IS NORMAL. SURGEON WILL CONTINUE TO MONITOR THE LENS POSITION, CORNEAL HEALTH AND IOP (INTRAOCULAR PRESSURE). Manufacturer Narrative: PT INFO: UNK. EVENT DATE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8040111&pc=MTA | 2023826-2018-01611 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-11-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -13.5 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AS A REPLACEMENT LENS BUT IT DID NOT RESOLVE THE PROBLEM COMPLETELY. THE AC (ANTERIOR CHAMBER) REMAINS VERY SHALLOW, BUT IT IS NO LONGER EXTREMELY SHALLOW AND THREATENING THE CORNEAL ENDOTHELIUM. EVEN WITH A SMALLER LENS THAN CALL FOR (DESPITE REPEAT CONFIRMATION OF CORRECT WTW) THE NEW SMALLER LENS REMAIN EXCESSIVELY VAULTED. THE PATIENT IS DOING VERY WELL. PATIENT IS HAPPY WITH HER VISION AND HER IOP (INTRAOCULAR PRESSURE) IS NORMAL. SURGEON WILL CONTINUE TO MONITOR THE LENS POSITION, CORNEAL HEALTH AND IOP (INTRAOCULAR PRESSURE). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129039&pc=MTA | 2023826-2018-01783 |  | Injury | STAAR SURGICAL COMPANY | 2018-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: BMC "OPTHALMOLOGY" ARTICLE WAS RECEIVED ON 07-NOV-2018 WHICH CITED A RETROSPECTIVE STUDY TITLED, 'IMPLANTABLE COLLAMER LENS SURGERY IN PATIENTS WITH PRIMARY IRIS AND/OR CILIARY BODY CYSTS.' THE ARTICLE STATES THAT IN ONE EYE THE POSTOPERATIVE ANTERIOR CHAMBER ANGLE WAS CLOSED IN THE 3 O'CLOCK POSITION; IT WAS CLOSED AT THE LOCATION CORRESPONDING TO THE ICL HAPTIC. ALSO OBSERVED WAS THAT THE CENTRAL VAULT WAS MORE THAN 1.25MM. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8337619&pc=MTA | 2023826-2019-00253 |  | Injury | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A SHORTER LENS WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8417792&pc=MTA | 2023826-2019-00421 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | MTA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND DUE TO LOW VAULT THE SURGEON FELT COMPELLED TO EXCHANGE THE LENS. THE REPORTER STATED THAT THE SURGEON EXCHANGED THE LENS FOR A 13.2 TMICL (LONGER LENGTH LENS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8471358&pc=MTA | 2023826-2019-00425 |  | Injury | STAAR SURGICAL COMPANY | 2019-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018 AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492947&pc=MTA | 2023826-2019-00567 |  | Injury | STAAR SURGICAL COMPANY | 2019-04-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) IN (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND NARROWING OF THE ANGLE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609183&pc=MTA | 2023826-2019-00832 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DATE OF EVENT-UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8841847&pc=MTA | 2023826-2019-01365 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -4.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. " NO COMPLAINTS" WAS REPORTER FOR STATUS OF THE EYE ON PATIENTS LAST FOLLOW UP VISIT. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS; RESIDUE ON LENS. (B)(4). Manufacturer Narrative: EXPLANTED DATE: CORRECTED THE YEAR REPORTED IN INITIAL MDR FROM 2109 TO 2019. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871949&pc=MTA | 2023826-2019-01388 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED HIGH VAULTING WITH SOME MICL IMPLANTABLE COLLAMER LENSES AND HAS HAD TO EXPLANT AND EXCHANGE FOR SHORTER LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871950&pc=MTA | 2023826-2019-01389 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED HIGH VAULTING WITH SOME MICL IMPLANTABLE COLLAMER LENSES AND HAS HAD TO EXPLANT AND EXCHANGE FOR SHORTER LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871952&pc=MTA | 2023826-2019-01387 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED HIGH VAULTING WITH SOME MICL IMPLANTABLE COLLAMER LENSES AND HAS HAD TO EXPLANT AND EXCHANGE FOR SHORTER LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871957&pc=MTA | 2023826-2019-01386 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED HIGH VAULTING WITH SOME MICL IMPLANTABLE COLLAMER LENSES AND HAS HAD TO EXPLANT AND EXCHANGE FOR SHORTER LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8977984&pc=MTA | 2023826-2019-01694 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, MICL13.7, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND A 13.2MM LENS WAS IMPLANTED. ANGLE CLOSURE WAS REPORTED AS AN OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: ANGLE CLOSURE SHOULD HAVE BEEN INCLUDED IN INITIAL MDR CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN LENS CASE. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN AND RESIDUE ON THE LENS SURFACE. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9070258&pc=MTA | 2023826-2019-01705 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN A ZIPLOC BAG, IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125393&pc=MTA | 2023826-2019-01790 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, SCRATES ON THE LENS, RESIDUE ON LENS SURFACE, AND A TORN AND BENT HAPTIC. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9282305&pc=MTA | 2023826-2019-02072 |  | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL, WITH FIBER ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9293575&pc=MTA | 2023826-2019-02100 |  | Injury | STAAR SURGICAL COMPANY | 2019-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PROBLEMS WITH VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC BENT. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9416987&pc=MTA | 2023826-2019-02169 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -14.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3:DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753513&pc=MTA | 2023826-2020-00377 |  | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD A HIGH VAULT AND BOWED FORWARD IRIS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED NO EXPLANTATION IS PLANNED FOR NOW AND THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9844048&pc=MTA | 2023826-2020-00546 |  | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE, " BUT EXPLANTED AFTER SEEING EXCESSIVE VAULT. HE THEN IMPLANTED A 12.6 SIZE AND WAS MUCH BETTER SIZE FOR PATIENT." Manufacturer Narrative: CORRECTED DATA: D7 - EXPLANTED DATE: (B)(6) 2020 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965119&pc=MTA | 2023826-2020-00893 |  | Injury | STAAR SURGICAL COMPANY | 2020-04-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS EYE. SEVERAL MONTHS POST-OP ICL SURGERY, THE LENS WAS REPORTED AS HAVING HIGH VAULTING AND THE PATIENT EXPERIENCED A MONOCULAR LARGE PUPIL. THE LENS WAS EXPLANTED. UPON INVESTIGATION, THE SURGEON DISCOVERED THE TECHNICIAN HAD USED THE WTW FROM LENSTAR INSTEAD OF THE CALIPER MEASUREMENT HE WANTED TO BE USED FOR THE CALCULATION, RESULTING IN AN ICL BEING IMPLANTED THAT WAS LARGER THAN APPROPRIATE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476052&pc=MTA | 2023826-2020-01921 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS EXPLANTED AND EXCHANGED DUE TO HIGH VAULTING AND PUPILLARY BLOCK. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1: PRODUCT PROBLEM. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS IN THE PATIENTS RIGHT EYE (OD). THE SURGEON THOUGHT HE MAY HAVE HAD TO EXPLANT THE LENS DUE TO HIGH VAULTING AND PUPILLARY BLOCK; HOWEVER, THE RIGHT EYE SETTLED AND DID NOT HAVE TO BE REMOVED. THE LENS REMAINED IMPLANTED. H1: MALFUNCTION. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4922988&pc=MTA | 2023826-2015-00820 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773049&pc=MTA | 2023826-2018-01190 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO VAULTING AND TO PREVENT CATARACT FORMATION. THE LENS WAS REPLACED WITH A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS HAPTICS TORN, WITH A PIECE MISSING. (B)(4) Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6199638&pc=MTA | 2023826-2016-01813 | 2016-11-20 | Injury | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Activation, Positioning or Separation Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND THE LENS WAS NOTED TO BE TORN DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND LENS TEAR. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND A PIECE OF HAPTIC MISSING. THERE WAS CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7253318&pc=MTA | 2023826-2018-00150 | 2017-12-04 | Injury | STAAR SURGICAL COMPANY | 2018-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+4.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE LENS EXCHANGE RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED THAT "THE DOCTOR'S REPLY WAS A SPARE LENS FOR THE SAME OPERATION AND RESERVED." THE LENS WAS EXPLANTED AND EXCHANGED PRESUMABLY DUE TO OBSERVATION OF AN EXCESSIVE VAULT. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS), AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7253307&pc=MTA | 2023826-2018-00149 | 2017-12-11 | Injury | STAAR SURGICAL COMPANY | 2018-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE LENS EXCHANGE RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED DUE TO "HIGH ARCHES AND THERE WERE NO OTHER COMPLICATIONS. THE PATIENT WAS IN GOOD CONDITION AFTER THE REPLACEMENT." (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9595578&pc=MTA | 2023826-2020-00078 | 2019-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO LOW VAULT. REMOVAL OF LENS RESOLVED PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, DUE TO PATIENT'S "HIGH DEMAND FOR NEAR VISION AND POOR ADJUSTMENT, THE PATIENT STRONGLY REQUESTED TO TAKE OUT." Manufacturer Narrative: CORRECTED DATA: B3- DATE OF EVENT IS CORRECTED TO UNK. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-DEVICE CODE 2993 IN INITIAL MDR IS CORRECTED TO 1583. CLAIM# : (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9664308&pc=MTA | 2023826-2020-00233 | 2019-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PER THE REPORTER ON (B)(6) 2020, "THEY IMPLANTED AND EXPLANTED (REMOVED)" LENS AND IMPLANTED A 13.2MM, MICL 13.2, -9.0 DIOPTER, LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND EXCHANGED IN A SECOND SURGERY ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS LOW VAULT. REPORTEDLY, "PATIENT'S LAST VISIT ON (B)(6) 2020 UCVA 20/20-2 IOP 14. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- DEVICE CODE 12993 SHOULD BE CORRECTED TO 1583 IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4471676&pc=MTA | 2023826-2015-00104 | 2008-10-15 | Injury | STAAR SURGICAL COMPANY | 2015-01-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: PT AGE: - UNK. PT WEIGHT: UNK. IMPLANT DATE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS WAS RETURNED IN LIQUID AND THERE WAS NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RE-HYDRATED AND BOTH THE LENGTH AND WIDTH WERE RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS A LENS WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). IF ROTATION: SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ECT.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4476459&pc=MTA | 2023826-2015-00108 | 2008-10-15 | Injury | STAAR SURGICAL COMPANY AG | 2015-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. SEE MANUFACTURER REPORT #2023826-2015-00107 FOR THE OTHER EYE. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS WAS RETURNED IN LIQUID AND THERE WAS NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RE-HYDRATED AND BOTH THE LENGTH AND WIDTH WERE RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS A LENS WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). IF ROTATION: SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ECT.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: THIS IS A DUPLICATE REPORT OF MFR# 2023826-2015-00104 AND WAS SUBMITTED IN ERROR. THIS IS NOT A COMPLAINT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592743&pc=MTA | 2023826-2015-00230 | 2011-03-14 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) - (NEW INCISION); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592543&pc=MTA | 2023826-2015-00274 | 2011-07-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST UCVA WAS 20/14. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY AND A HAPTIC WAS TORN. THERE WAS A CLEAR SURGICAL RESIDUE ON THE LENS. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870309&pc=MTA | 2023826-2016-01050 | 2011-08-09 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS, -5.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE CHANGE OVERTIME. THE LENS WAS PLANNED TO EXCHANGE FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/40. DEVICE EVALUATION - THE LENS WAS RETURNED BENT AND DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE PRODUCT WAS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870312&pc=MTA | 2023826-2016-01049 | 2011-08-09 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE CHANGE OVERTIME. THE LENS WAS PLANNED TO EXCHANGE FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING NO DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THIS LENS MODEL IS CONTRAINDICATED DUE TO THE ANTERIOR CHAMBER DEPTH (ACD) BEING LOWER THAN WHAT WAS INDICATED FOR CORRECT USE IN THE DFU. CLAIM #: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4962378&pc=MTA | 2023826-2015-00932 | 2012-01-14 | Injury | STAAR SURGICAL COMPANY | 2015-08-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012 AND LOW VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH ON (B)(6) 2012 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4973824&pc=MTA | 2023826-2015-00939 | 2012-02-03 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. LOW VAULT WAS OBSERVED AND THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2012. THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4595898&pc=MTA | 2023826-2015-00288 | 2012-02-11 | Injury | STAAR SURGICAL COMPANY | 2015-03-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LET EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST BCVA WAS 20/20 ON (B)(6) 2014. Manufacturer Narrative: PT AGE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY AND A HAPTIC WAS TORN. THERE WAS A CLEAR SURGICAL RESIDUE ON THE LENS. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4596168&pc=MTA | 2023826-2015-00292 | 2012-02-28 | Injury | STAAR SURGICAL COMPANY | 2015-03-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: PT AGE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC AND THERE WAS A CLEAR SURGICAL RESIDUE ON THE LENS. A LENS WORK ORDER SEARCH REVEALED THERE WAS ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4889285&pc=MTA | 2023826-2015-00764 | 2012-02-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED WITH A BROKEN HAPTIC. A LENS WORK ORDER SEARCH REVEALED THERE WAS ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578360&pc=MTA | 2023826-2015-00222 | 2012-05-29 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4).: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS.. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4784826&pc=MTA | 2023826-2015-00592 | 2012-05-29 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012 AND LOW VAULT WAS NOTED IN THE POST-OPERATIVE PERIOD. THE LENS WAS EXCHANGED FOR A LONGER LENGTH ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY WITH A TORN HAPTIC. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830544&pc=MTA | 2023826-2016-00989 | 2012-07-06 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. THE PATIENT EXPERIENCED A REFRACTIVE CHANGE OVERTIME. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592353&pc=MTA | 2023826-2015-00280 | 2012-07-11 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC AND THERE WAS A CLEAR SURGICAL RESIDUE ON THE LENS. A LENS WORK ORDER SEARCH REVEALED THERE WAS ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4872767&pc=MTA | 2023826-2015-00720 | 2012-08-01 | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885591&pc=MTA | 2023826-2015-00721 | 2012-08-01 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4596538&pc=MTA | 2023826-2015-00294 | 2012-09-22 | Injury | STAAR SURGICAL COMPANY | 2015-03-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012 . THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA 20/20. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN OPTIC. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4596082&pc=MTA | 2023826-2015-00290 | 2012-12-13 | Injury | STAAR SURGICAL COMPANY | 2015-03-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20 ON (B)(6) 2014. SEE MFR REPORT #2023826-2015-00289 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: VISUAL INPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE OBSERVED. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4595986&pc=MTA | 2023826-2015-00289 | 2012-12-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20 ON (B)(6) 2014. SEE MFR REPORT #2023826-2015-00290 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY AND A HAPTIC WAS TORN. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923018&pc=MTA | 2023826-2015-00821 | 2013-03-04 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5183958&pc=MTA | 2023826-2015-01330 | 2013-03-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER -02.00 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS NOTED (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01322 FOR LEFT EYE (OD). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DIOPTER: -02.00. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5183969&pc=MTA | 2023826-2015-01322 | 2013-03-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER -02.25 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LOW VAULT WAS NOTED (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01330 FOR RIGHT EYE (OD). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DIOPTER: -02.25. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642888&pc=MTA | 2023826-2016-00589 | 2013-04-28 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4, -20.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4870129&pc=MTA | 2023826-2015-00714 | 2013-05-17 | Injury | STAAR SURGICAL COMPANY | 2015-06-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE PATIENT IS HAPPY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DATE OF BIRTH AND WEIGHT - UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4577723&pc=MTA | 2023826-2015-00255 | 2013-06-13 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2013 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THERE WAS NO VISIBLE DAMAGE TO THE LENS. A LENS WORK ORDER SEARCH REVEALED THAT THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4577445&pc=MTA | 2023826-2015-00254 | 2013-06-21 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2013 DUE TO EXCESSIVE VAULTING. A YAG WAS PERFORMED AND LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP BCVA WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THERE WAS NO VISIBLE DAMAGE TO THE LENS. A LENS WORK ORDER SEARCH REVEALED THAT THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574309&pc=MTA | 2023826-2015-00216 | 2013-08-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786542&pc=MTA | 2023826-2015-00570 | 2013-08-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION - BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4740739&pc=MTA | 2023826-2015-00490 | 2013-09-04 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) - (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVALUATION METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A HAPTIC WAS TORN AND THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600883&pc=MTA | 2023826-2015-00306 | 2013-10-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW -ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786551&pc=MTA | 2023826-2015-00572 | 2013-10-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2014, AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN BCVA WAS 20/20 AS OF (B)(6) 2014. REFERENCE MFR. REPORT #2023826-2015-00571 FOR RIGHT EYE (OD) OF SAME PATIENT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY WITH A TORN HAPTIC. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION - BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4787597&pc=MTA | 2023826-2015-00571 | 2013-10-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-21 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN BCVA WAS 20/20 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY WITH A TORN HAPTIC. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161951&pc=MTA | 2023826-2015-01275 | 2013-11-02 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. LOW VAULT WAS OBSERVED ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 AND WAS EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2014 - UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -12.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954776&pc=MTA | 2023826-2015-00963 | 2013-11-07 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855101&pc=MTA | 2023826-2015-00686 | 2013-11-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954764&pc=MTA | 2023826-2015-00961 | 2013-11-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592742&pc=MTA | 2023826-2015-00231 | 2013-11-26 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/18. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND BOTH HAPTICS TORN. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5384458&pc=MTA | 2023826-2016-00021 | 2013-12-01 | Injury | STAAR SURGICAL COMPANY | 2016-01-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -12.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2006. LOW VAULT WAS OBSERVED ON (B)(6) 2013. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST OFFICE VISIT ON (B)(6) 2015, BCVA WAS 20/20. SEE MFR. #2023826-2016-00022 FOR LEFT EYE. Manufacturer Narrative: DIOPTER: -12.00. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266055&pc=MTA | 2023826-2015-01604 | 2014-01-09 | Injury | STAAR SURGICAL COMPANY | 2015-12-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -13.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. LOW VAULT WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP, PATIENT'S UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -13.00. (B)(4). CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806954&pc=MTA | 2023826-2015-00584 | 2014-01-11 | Injury | STAAR SURGICAL COMPANY | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786497&pc=MTA | 2023826-2015-00566 | 2014-01-15 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2014, AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOW UCVA WAS 20/20 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY AND THE OPTIC WAS TORN. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885632&pc=MTA | 2023826-2015-00727 | 2014-02-12 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4968447&pc=MTA | 2023826-2015-00935 | 2014-02-12 | Injury | STAAR SURGICAL COMPANY | 2015-08-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS OBSERVED AND THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. PATIENT LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293051&pc=MTA | 2023826-2015-01648 | 2014-02-23 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 -10.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. THE PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -10.50. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266246&pc=MTA | 2023826-2015-01598 | 2014-02-25 | Injury | STAAR SURGICAL COMPANY | 2015-12-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -13.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/20. SEE MFR. #2023826-2015-01599 FOR LEFT EYE. Manufacturer Narrative: DIOPTER: -13.50. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266250&pc=MTA | 2023826-2015-01599 | 2014-02-25 | Injury | STAAR SURGICAL COMPANY | 2015-12-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -17.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LOW VAULT WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/20. SEE MFR. #2023826-2015-01598 FOR RIGHT EYE. Manufacturer Narrative: DIOPTER: -17.00. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4601238&pc=MTA | 2023826-2015-00300 | 2014-02-28 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4973612&pc=MTA | 2023826-2015-00936 | 2014-03-04 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. EXCESSIVE VAULT WAS OBSERVED AND THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. PATIENT LAST UCVA WAS 20/25. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592747&pc=MTA | 2023826-2015-00237 | 2014-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592749&pc=MTA | 2023826-2015-00236 | 2014-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885809&pc=MTA | 2023826-2015-00734 | 2014-04-11 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4974539&pc=MTA | 2023826-2015-00945 | 2014-04-11 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND EXCESSIVE VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5535265&pc=MTA | 2023826-2016-00257 | 2014-04-19 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4, -17.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2011. DECREASED VISION OVER TIME WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015; UCVA WAS 20/16. SEE MFR. #2023826-2016-00258 FOR RIGHT EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. DECREASED VISION OVER TIME WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO REFRACTIVE SURPRISE OVER TIME. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2015, THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/16. THIS PRODUCT IS NOT MANUFACTURED IN THE U.S. AND NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: DIOPTER: -17.00. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4489013&pc=MTA | 2023826-2015-00113 | 2014-05-05 | Injury | STAAR SURGICAL COMPANY | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULTS: THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4489014&pc=MTA | 2023826-2015-00112 | 2014-05-05 | Injury | STAAR SURGICAL COMPANY | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266251&pc=MTA | 2023826-2015-01600 | 2014-05-10 | Injury | STAAR SURGICAL COMPANY | 2015-12-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -07.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/15. Manufacturer Narrative: DIOPTER: -07.00. (B)(4). CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786529&pc=MTA | 2023826-2015-00569 | 2014-05-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. EXCESSIVE VAULT AND CLOSURE OF IRIDO-CORNEAL ANGLES ASSOCIATED WITH ELEVATED IOP WERE NOTED, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2014, AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN UCVA WAS 20/20 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW & WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITION MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FOR OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. THERE IS NOT ENOUGH INFORMATION TO DETERMINE THE EXACT CAUSE OF THIS EVENT, HOWEVER, WE CANNOT RULE OUT EXCESSIVE VAULTING SECONDARY TO A LONG LENS AS A CONTRIBUTING FACTOR. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5185673&pc=MTA | 2023826-2015-01328 | 2014-05-19 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS +6.5/+1.0/X150 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01325 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: +6.5/+1.0/X150. (B)(4). CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5187573&pc=MTA | 2023826-2015-01325 | 2014-05-19 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS +6.0/+2.0/X155 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SMALLER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01328 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: +6.0/+2.0X155. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4966354&pc=MTA | 2023826-2015-00933 | 2014-05-23 | Injury | STAAR SURGICAL COMPANY | 2015-08-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS OBSERVED AND THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. PATIENT LAST BCVA WAS 20/50. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773338&pc=MTA | 2023826-2015-00546 | 2014-05-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULTING WAS NOTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THIS RESOLVED THE PROBLEM. THE PATIENTS LAST BCVA WAS 20/20 ON (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4573899&pc=MTA | 2023826-2015-00194 | 2014-05-30 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4599024&pc=MTA | 2023826-2015-00308 | 2014-06-02 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC AND THERE WAS A CLEAR SURGICAL RESIDUE ON THE LENS. A LENS WORK ORDER SEARCH REVEALED THERE WAS ONCE SIMILAR COMPLAINT WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4555206&pc=MTA | 2023826-2015-00195 | 2014-06-05 | Injury | STAAR SURGICAL COMPANY | 2015-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4599019&pc=MTA | 2023826-2015-00305 | 2014-06-05 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4599012&pc=MTA | 2023826-2015-00304 | 2014-06-13 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4973800&pc=MTA | 2023826-2015-00938 | 2014-06-20 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS OBSERVED AND THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4615063&pc=MTA | 2023826-2015-00331 | 2014-06-21 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE, RESULTING IN AN EXPLANT OF THE LENS ON (B)(6) 2014 AND EXCHANGE FOR A LENS WITH A SMALLER VAULT. THE ISSUE WAS RESOLVED WITH THE IMPLANTATION OF THE NEW LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) - (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592748&pc=MTA | 2023826-2015-00234 | 2014-06-25 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4573903&pc=MTA | 2023826-2015-00191 | 2014-07-02 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4573891&pc=MTA | 2023826-2015-00197 | 2014-07-08 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4555213&pc=MTA | 2023826-2015-00189 | 2014-07-15 | Injury | STAAR SURGICAL COMPANY | 2015-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) -(NEW INCISION); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574465&pc=MTA | 2023826-2015-00192 | 2014-07-15 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4816489&pc=MTA | 2023826-2015-00636 | 2014-07-21 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12MM ICM120V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2007. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A CLEAR SURGICAL RESIDUE. NO VISIBLE DAMAGE WAS OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC, POSITION ECT.). ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4552424&pc=MTA | 2023826-2015-00133 | 2014-07-23 | Injury | STAAR SURGICAL COMPANY | 2015-02-26 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SIZE INCORRECT FOR PATIENT, UNLABELED. NO CODE AVAILABLE (LENS EXPLANTED), LABELED. NO CODE AVAILABLE (VAULTING), UNLABELED. : METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4542632&pc=MTA | 2023826-2015-00130 | 2014-07-25 | Injury | STAAR SURGICAL COMPANY | 2015-02-24 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION: METHOD - (OTHER): WORK ORDER SEARCH. RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614651&pc=MTA | 2023826-2015-00340 | 2014-08-01 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) - (NEW INCISION); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING); DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4555259&pc=MTA | 2023826-2015-00208 | 2014-08-02 | Injury | STAAR SURGICAL COMPANY | 2015-02-27 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). IF ROTATION: SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ECT.). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4577134&pc=MTA | 2023826-2015-00253 | 2014-08-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. SIGNIFICANT REDUCTION OF ICA WERE NOTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP BCVA WAS 20/66. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975314&pc=MTA | 2023826-2015-00951 | 2014-08-10 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574310&pc=MTA | 2023826-2015-00219 | 2014-08-11 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578985&pc=MTA | 2023826-2015-00218 | 2014-08-11 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378612&pc=MTA | 2023826-2015-01672 | 2014-08-13 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4839509&pc=MTA | 2023826-2015-00630 | 2014-08-16 | Injury | STAAR SURGICAL COMPANY | 2015-06-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH, MEDICAL REVIEW RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS -(OTHER): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592561&pc=MTA | 2023826-2015-00275 | 2014-08-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER/HIGH POWER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT;S LAST UCVA WAS 20/14. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786479&pc=MTA | 2023826-2015-00567 | 2014-08-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2014, AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN UCVA WAS 20/20 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY AND A HAPTIC WAS TORN. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287686&pc=MTA | 2023826-2015-01607 | 2014-08-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592744&pc=MTA | 2023826-2015-00228 | 2014-08-28 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) - (NEW INCISION); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592737&pc=MTA | 2023826-2015-00232 | 2014-08-30 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH AND DIOPTER POWER WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5164296&pc=MTA | 2023826-2015-01268 | 2014-08-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/X079 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 WAS 20/16. SEE MFR. #2023826-2015-01274 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5164307&pc=MTA | 2023826-2015-01274 | 2014-08-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/X094 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 WAS 20/13. SEE MFR. #2023826-2015-01268 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614648&pc=MTA | 2023826-2015-00343 | 2014-09-02 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER: NO - LENS NOT RETURNED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULC.US TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975063&pc=MTA | 2023826-2015-00949 | 2014-09-02 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4736021&pc=MTA | 2023826-2015-00480 | 2014-09-12 | Injury | STAAR SURGICAL COMPANY | 2015-04-29 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578363&pc=MTA | 2023826-2015-00221 | 2014-09-19 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS MOST PROBABLY A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592738&pc=MTA | 2023826-2015-00229 | 2014-09-19 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT ANTERIOR SUBCAPSULAR CATARACT OCCURS IN 1.3% OF PATIENTS WHERE AN ICL HAS BEEN IMPLANTED, FOLLOWING THE RECOMMENDATIONS FOR USE. THIS USUALLY OCCURS WITHIN THE 1ST 6 MONTHS POST-OP. THE MOST PROBABLE CAUSE HAS BEEN DETERMINED TO BE CRYSTALLINE LENS TOUCH DURING SURGERY WHICH COULD HAVE BEEN EITHER FROM THE INSTRUMENTS USED OR THE ICL ITSELF. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR SULCUS OR CILIARY SULCUS CYST, ETC). STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUBCAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/OR REPLACE THE ICL INCREASES THE RISK FOR ANTERIOR SUBCAPSULAR CATARACT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975024&pc=MTA | 2023826-2015-00948 | 2014-09-21 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND LOW VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/40. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5082837&pc=MTA | 2023826-2015-01224 | 2014-09-24 | Injury | STAAR SURGICAL COMPANY | 2015-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -12.0/+3.0/X176 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.0/+3.0/X176. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4784694&pc=MTA | 2023826-2015-00591 | 2014-09-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULT WAS NOTED IN THE POST-OPERATIVE PERIOD. THE LENS WAS EXCHANGED FOR A LONGER LENGTH ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4958310&pc=MTA | 2023826-2015-00979 | 2014-10-01 | Injury | STAAR SURGICAL COMPANY | 2015-07-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5168232&pc=MTA | 2023826-2015-01271 | 2014-10-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.0/+4.5/X097 DIOPTER IN THE PATIENT'S RIGHT EYE (OD). LOW VAULT AND LENS ROTATION WAS OBSERVED ON (B)(6) 2014. THE LENS WAS REPOSITIONED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773352&pc=MTA | 2023826-2015-00552 | 2014-10-03 | Injury | STAAR SURGICAL COMPANY | 2015-05-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161949&pc=MTA | 2023826-2015-01276 | 2014-10-03 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULT AND WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2014 - UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -17.00. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5209471&pc=MTA | 2023826-2015-01435 | 2014-10-04 | Injury | STAAR SURGICAL COMPANY | 2015-11-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -09.00/+06.00/X088 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023829-2015-01431 FOR RIGHT EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED (DATE UNKNOWN) FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE LENS WAS DEFORMED. HOWEVER, THERE WAS NO OTHER VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION OF THE RETURNED LENS FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -09.00/+06.00/X088. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210090&pc=MTA | 2023826-2015-01431 | 2014-10-04 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -8.50/+6.00/X094 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. THE LENS WAS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4576766&pc=MTA | 2023826-2015-00250 | 2014-10-05 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW & LENS WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY AND A HAPTIC WAS TORN. A LENS WORK ORDER SEARCH REVEALED THAT THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4576823&pc=MTA | 2023826-2015-00251 | 2014-10-05 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY AND A HAPTIC WAS TORN. A LENS WORK ORDER SEARCH REVEALED THAT THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629118&pc=MTA | 2023826-2015-00372 | 2014-10-06 | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) -(NEW INCISION); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD): WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4745849&pc=MTA | 2023826-2015-00500 | 2014-10-06 | Injury | STAAR SURGICAL COMPANY | 2015-05-04 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT AGE AND WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OS THE RETURNED LENS SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885281&pc=MTA | 2023826-2015-00754 | 2014-10-09 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/22. SEE MFR# 2023826-2015-00755 FOR THE OTHER EYE. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. (EVALUATION ANTICIPATED BUT NOT YET DONE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5168268&pc=MTA | 2023826-2015-01288 | 2014-10-10 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS WITH -12.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2014, PATIENT'S BCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -12.50. (B)(4).CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806959&pc=MTA | 2023826-2015-00586 | 2014-10-11 | Injury | STAAR SURGICAL COMPANY | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(6): EVALUATION: METHOD - WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614657&pc=MTA | 2023826-2015-00339 | 2014-10-14 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) - (NEW INCISION); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4596593&pc=MTA | 2023826-2015-00295 | 2014-10-18 | Injury | STAAR SURGICAL COMPANY | 2015-03-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA 20/20. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161952&pc=MTA | 2023826-2015-01280 | 2014-10-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT AND WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -11.50 DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5163903&pc=MTA | 2023826-2015-01279 | 2014-10-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT AND WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/18. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600865&pc=MTA | 2023826-2015-00313 | 2014-10-21 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE ICL WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DATE OF BIRTH - UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629140&pc=MTA | 2023826-2015-00373 | 2014-10-21 | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4596617&pc=MTA | 2023826-2015-00298 | 2014-10-28 | Injury | STAAR SURGICAL COMPANY | 2015-03-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013 . THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA 20/20. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855108&pc=MTA | 2023826-2015-00692 | 2014-10-28 | Injury | STAAR SURGICAL COMPANY | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. PER MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4863670&pc=MTA | 2023826-2015-00709 | 2014-10-28 | Malfunction | STAAR SURGICAL COMPANY | 2015-06-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULT WAS NOTED. THE LENS WAS REMOVED FROM THE EYE AND A LONGER LENS HAS BEEN ORDERED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). RESULTS: (OTHER) A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (OTHER) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4489047&pc=MTA | 2023826-2015-00111 | 2014-11-04 | Injury | STAAR SURGICAL COMPANY | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210650&pc=MTA | 2023826-2015-01430 | 2014-11-10 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, 8.00/+2.50/X103 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED AND THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -8.0/+2.5/X103. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629117&pc=MTA | 2023826-2015-00374 | 2014-11-14 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? LENS NOT RETURNED. EVALUATION METHOD : WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574342&pc=MTA | 2023826-2015-00220 | 2014-11-15 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614656&pc=MTA | 2023826-2015-00342 | 2014-11-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER: NO -LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614661&pc=MTA | 2023826-2015-00341 | 2014-11-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016358&pc=MTA | 2023826-2015-01055 | 2014-11-17 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS WITH -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS OBSERVED ON (B)(6) 2014 AND THE LENS WAS EXCHANGED FOR A LONGER LEN ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST UCVA WAS 20/13. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5079600&pc=MTA | 2023826-2015-01226 | 2014-11-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -20.00/+2.50/X097 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. SEE MFR #2023826-2015-01225 FOR LEFT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -20.00/+2.5/X097. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5088287&pc=MTA | 2023826-2015-01225 | 2014-11-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -22.00/+3.00/X083 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25. SEE MFR #2023826-2015-01226 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -22.00/+3.00/X083. DEVICE EVALUATED BY MANUFACTURER? NO-LENS NOT RETURNED. (B)(4). METHOD CODE: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4966366&pc=MTA | 2023826-2015-00934 | 2014-11-19 | Injury | STAAR SURGICAL COMPANY | 2015-08-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS OBSERVED AND THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON 11/19/2014. THIS RESOLVED THE PROBLEM. PATIENT LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975673&pc=MTA | 2023826-2015-00953 | 2014-11-19 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578608&pc=MTA | 2023826-2015-00259 | 2014-11-20 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW & LENS WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY AND A HAPTIC WAS TORN. LENS WORK ORDER SEARCH REVEALED THAT THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629093&pc=MTA | 2023826-2015-00369 | 2014-11-24 | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - 81 (OTHER): LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) . Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806336&pc=MTA | 2023826-2015-00581 | 2014-11-26 | Injury | STAAR SURGICAL COMPANY | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH, MEDICAL REVIEW RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786483&pc=MTA | 2023826-2015-00564 | 2014-11-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED NO VISIBLE DAMAGE WAS OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629137&pc=MTA | 2023826-2015-00376 | 2014-11-29 | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) -(SECONDARY SURGERY); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629138&pc=MTA | 2023826-2015-00377 | 2014-11-30 | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (NEW INCISION); INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4650378&pc=MTA | 2023826-2015-00405 | 2014-12-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED IN LIQUID AND NO VISIBLE DAMAGE WAS OBSERVED. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5011306&pc=MTA | 2023826-2015-01038 | 2014-12-02 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND EXCESSIVE VAULT WAS NOTED THE NEXT DAY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: -11.50. EVENT PROBLEM AND EVALUATION CODES: PATIENT CODE: (B)(4)- NO CODE AVAILABLE (NEW INCISION), LABELED (B)(4)- NO CODE AVAILABLE (SECONDARY SURGERY), UNLABELED DEVICE CODE: (B)(4)- NO CODE AVAILABLE (LENS EXPLANTED), LABELED (B)(4)- NO CODE AVAILABLE (VAULTING), UNLABELED EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574469&pc=MTA | 2023826-2015-00217 | 2014-12-03 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4) Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885630&pc=MTA | 2023826-2015-00722 | 2014-12-03 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954756&pc=MTA | 2023826-2015-00957 | 2014-12-03 | Injury | STAAR SURGICAL COMPANY | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5178938&pc=MTA | 2023826-2015-01329 | 2014-12-05 | Injury | STAAR SURGICAL COMPANY | 2015-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -13.0/+2.0/X103 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.0/+2.0/X103. DEVICE EVALUATED BY MANUFACTURER? LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5182257&pc=MTA | 2023826-2015-01327 | 2014-12-05 | Injury | STAAR SURGICAL COMPANY | 2015-10-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER -08.00 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULTING WAS NOTED N (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT IS NOT MARKETED IN THE U.S. DIOPTER: -08.00. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5189894&pc=MTA | 2023826-2015-01446 | 2014-12-05 | Injury | STAAR SURGICAL COMPANY | 2015-10-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH . RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4577024&pc=MTA | 2023826-2015-00252 | 2014-12-06 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). NEW INCISION. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. LENS EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY AND THE OPTIC WAS TORN. A LENS WORK ORDER SEARCH REVEALED THAT THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4642445&pc=MTA | 2023826-2015-00399 | 2014-12-06 | Injury | STAAR SURGICAL COMPANY | 2015-03-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266212&pc=MTA | 2023826-2015-01605 | 2014-12-06 | Injury | STAAR SURGICAL COMPANY | 2015-12-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -12.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP, PATIENT'S LAST VISIT, UCVA WAS 20/13. Manufacturer Narrative: DIOPTER: -12.00. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629108&pc=MTA | 2023826-2015-00370 | 2014-12-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - 81 (OTHER): LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855100&pc=MTA | 2023826-2015-00687 | 2014-12-08 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287996&pc=MTA | 2023826-2015-01608 | 2014-12-08 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/072 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4920281&pc=MTA | 2023826-2015-00818 | 2014-12-10 | Injury | STAAR SURGICAL COMPANY | 2015-07-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4949711&pc=MTA | 2023826-2015-00905 | 2014-12-10 | Injury | STAAR SURGICAL COMPANY | 2015-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4642574&pc=MTA | 2023826-2015-00400 | 2014-12-11 | Injury | STAAR SURGICAL COMPANY | 2015-03-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED AS WELL AS LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/20 AS OF (B)(6) 2014. Manufacturer Narrative: (B)(4) - (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING); UNINTENDED MOVEMENT; (ROTATION). DEVICE EVALUATED BY THE MANUFACTURER? NO. : LENS WAS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). IF ROTATION, SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ECT.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885646&pc=MTA | 2023826-2015-00725 | 2014-12-11 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885651&pc=MTA | 2023826-2015-00726 | 2014-12-11 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171636&pc=MTA | 2023826-2015-01388 | 2014-12-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -03.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4641288&pc=MTA | 2023826-2015-00398 | 2014-12-13 | Injury | STAAR SURGICAL COMPANY | 2015-03-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED AND THE LENS WAS REPOSITIONED ON (B)(6) 2014. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST KNOW UCVA IS 20/22 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY AND A HAPTIC WAS TORN. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN (B)(4) OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION,ECT). THERE IS NOT ENOUGH CLINICAL INFORMATION TO DETERMINE THE EXACT ROOT CAUSE; HOWEVER, IT SEEMS THE ICL WAS SHORT FOR THE EYE. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533613&pc=MTA | 2023826-2016-00345 | 2014-12-13 | Injury | STAAR SURGICAL COMPANY | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5079611&pc=MTA | 2023826-2015-01223 | 2014-12-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -13.5/+1.5/X081 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -13.5/+1.5/X081. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4646535&pc=MTA | 2023826-2015-00403 | 2014-12-14 | Injury | STAAR SURGICAL COMPANY | 2015-03-31 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20 AS OF (B)(6) 2015. (REFER MFR REPORT# 2023826-2015-00402 FOR OD) Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO - LENS WAS NOT RETURNED. EVALUATION: METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - NO CONCLUSION CAN BE DRAWN. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). IF ROTATION, SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ECT.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4646173&pc=MTA | 2023826-2015-00402 | 2014-12-14 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-31 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20 AS OF (B)(6) 2015. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). IF ROTATION, SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ECT.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954612&pc=MTA | 2023826-2015-00965 | 2014-12-14 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4488989&pc=MTA | 2023826-2015-00110 | 2014-12-16 | Injury | STAAR SURGICAL COMPANY | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4488990&pc=MTA | 2023826-2015-00109 | 2014-12-16 | Injury | STAAR SURGICAL COMPANY | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) - NO CODE AVAILABLE (NEW INCISION). SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4974443&pc=MTA | 2023826-2015-00942 | 2014-12-16 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND EXCESSIVE VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4650220&pc=MTA | 2023826-2015-00404 | 2014-12-18 | Injury | STAAR SURGICAL COMPANY | 2015-04-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/24 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4): EVALUATION: METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY AND NO VISIBLE DAMAGE WAS OBSERVED. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4736017&pc=MTA | 2023826-2015-00481 | 2014-12-19 | Injury | STAAR SURGICAL COMPANY | 2015-04-29 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743956&pc=MTA | 2023826-2015-00503 | 2014-12-19 | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885640&pc=MTA | 2023826-2015-00723 | 2014-12-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885662&pc=MTA | 2023826-2015-00724 | 2014-12-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) /2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954761&pc=MTA | 2023826-2015-00959 | 2014-12-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084279&pc=MTA | 2023826-2015-01201 | 2014-12-19 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -08.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4816577&pc=MTA | 2023826-2015-00637 | 2014-12-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2006. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON 03/17/2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN BCVA WAS 20/16 ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY AND THERE WAS NO VISIBLE DAMAGE OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5011470&pc=MTA | 2023826-2015-01039 | 2014-12-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND EXCESSIVE VAULT WAS NOTED THE NEXT DAY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: -09.50. (B)(4). A LENS WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084256&pc=MTA | 2023826-2015-01199 | 2014-12-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266267&pc=MTA | 2023826-2015-01596 | 2014-12-27 | Injury | STAAR SURGICAL COMPANY | 2015-12-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -15.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. LOW VAULT WAS OBSERVED, THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -15.50. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4872765&pc=MTA | 2023826-2015-00719 | 2014-12-29 | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012297&pc=MTA | 2023826-2015-01016 | 2014-12-29 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089551&pc=MTA | 2023826-2015-01235 | 2014-12-29 | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/073 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4529080&pc=MTA | 2023826-2015-00172 | 2014-12-30 | Injury | STAAR SURGICAL COMPANY | 2015-02-19 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR SULCUS OR CILIARY SULCUS CYST ETC). STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUBCAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/OR REPLACE THE ICL INCREASES THE RISK FOR ANTERIOR SUBCAPSULAR CATARACT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, THE MOST PROBABLE CAUSE OF THE EVENT WAS LOW VAULTING SECONDARY TO A SHORT LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4529082&pc=MTA | 2023826-2015-00173 | 2014-12-30 | Injury | STAAR SURGICAL COMPANY | 2015-02-19 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR SULCUS OR CILIARY SULCUS CYST ETC). STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUBCAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/OR REPLACE THE ICL INCREASES THE RISK FOR ANTERIOR SUBCAPSULAR CATARACT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, THE MOST PROBABLE CAUSE OF THE EVENT WAS LOW VAULTING SECONDARY TO A SHORT LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975166&pc=MTA | 2023826-2015-00950 | 2014-12-31 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4). RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012077&pc=MTA | 2023826-2015-01044 | 2014-12-31 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULT WAS NOTED THE NEXT DAY. THE LENS WAS EXCHANGED FOR LONGER ONE AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. DIOPTER: -12.00. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4529096&pc=MTA | 2023826-2015-00171 | 2015-01-02 | Injury | STAAR SURGICAL COMPANY | 2015-02-19 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975033&pc=MTA | 2023826-2015-00968 | 2015-01-02 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5288060&pc=MTA | 2023826-2015-01616 | 2015-01-02 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S.(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4872755&pc=MTA | 2023826-2015-00729 | 2015-01-03 | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4872770&pc=MTA | 2023826-2015-00730 | 2015-01-03 | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2013. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4872778&pc=MTA | 2023826-2015-00731 | 2015-01-03 | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5168270&pc=MTA | 2023826-2015-01284 | 2015-01-05 | Injury | STAAR SURGICAL COMPANY | 2015-10-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS WITH -13.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, PATIENT'S UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -13.50. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE(S): EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885788&pc=MTA | 2023826-2015-00735 | 2015-01-06 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855058&pc=MTA | 2023826-2015-00689 | 2015-01-07 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016259&pc=MTA | 2023826-2015-01050 | 2015-01-10 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND LOW VAULTING WAS NOTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4596277&pc=MTA | 2023826-2015-00293 | 2015-01-12 | Injury | STAAR SURGICAL COMPANY | 2015-03-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA 20/20. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE OBSERVED. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842190&pc=MTA | 2023826-2015-00664 | 2015-01-12 | Injury | STAAR SURGICAL COMPANY | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12MM ICM120V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A CLEAR SURGICAL RESIDUE. NO VISIBLE DAMAGE WAS OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855565&pc=MTA | 2023826-2015-00683 | 2015-01-12 | Injury | STAAR SURGICAL COMPANY | 2015-06-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5082645&pc=MTA | 2023826-2015-01227 | 2015-01-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4, -13.0/+1.5/X100 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -13.0/+1.5/X100. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5163947&pc=MTA | 2023826-2015-01269 | 2015-01-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT AND WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/13. Manufacturer Narrative: DIOPTER: -13.00. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855103&pc=MTA | 2023826-2015-00695 | 2015-01-13 | Injury | STAAR SURGICAL COMPANY | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2015. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016020&pc=MTA | 2023826-2015-01047 | 2015-01-13 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST UCVA WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855081&pc=MTA | 2023826-2015-00684 | 2015-01-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084257&pc=MTA | 2023826-2015-01200 | 2015-01-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/06. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171632&pc=MTA | 2023826-2015-01380 | 2015-01-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016113&pc=MTA | 2023826-2015-01049 | 2015-01-14 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND LOW VAULTING WAS NOTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST UCVA WAS 20/17. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4618641&pc=MTA | 2023826-2015-00336 | 2015-01-15 | Injury | STAAR SURGICAL COMPANY | 2015-03-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) NEW INCISION. EXPLANTED. VAULTING. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE WAS DETERMINED. AS REPORTED ON THE MEDICAL REVIEW, EXCESSIVE VAULTING IS A CONSEQUENCE OF WRONG LENS USE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF WHITE TO WHITE MEASUREMENTS BASED ON TECHNIQUES, IMPROPER SULCUS TO SULCUS MEASUREMENT, AND PATIENT CONDITION, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS, IRREGULAR SULCUS OR CILIARY SULCUS CYST). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4667817&pc=MTA | 2023826-2015-00417 | 2015-01-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENGTH ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: - MEDICAL REVIEW. RESULTS:VISUAL INSPECTION OF THE RETURNED LENS SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSION: BASED ON COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, IT HAS BEEN DETERMINE THE IMPROPER VAULT IS LIKELY DUE TO INACCURATE WHITE TO WHITE MEASUREMENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4802612&pc=MTA | 2023826-2015-00616 | 2015-01-16 | Injury | STAAR SURGICAL COMPANY | 2015-05-28 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. IMPLANT DATE: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ACT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885510&pc=MTA | 2023826-2015-00756 | 2015-01-17 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. LOW VAULT AND LENS ROTATION WAS OBSERVED. THE LENS EXPLANTED ON (B)(6)2015 AND EXCHANGED FOR A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP UCVA WAS 20/16. Manufacturer Narrative: (B)(4): (EVALUATION ANTICIPATED BUT NOT YET DONE). (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786506&pc=MTA | 2023826-2015-00568 | 2015-01-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXPLANTED ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS DRY AND A HAPTIC WAS TORN. THE LENS WAS RE-HYDRATED AND BOTH THE LENGTH AND WIDTH WERE REMEASURE AND DETERMINED TO BE WIHIN ORIGINAL DESIGN SPECIFICATIONS. THEREFORE, IT CAN BE JUSTIFIED THE COMPLAINT WAS NOT PRODUCT RELATED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORK ORDER. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUB-CAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/OR REPLACE THE ICL INCREASE THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016394&pc=MTA | 2023826-2015-01031 | 2015-01-18 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4740584&pc=MTA | 2023826-2015-00487 | 2015-01-19 | Injury | STAAR SURGICAL COMPANY | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/25 AS OF (B)(6) 2015. REFERENCE MFR REPORT #2023826-2015-00488 FOR LEFT EYE (OS). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PATIENT WEIGHT: UNK. (B)(4). EVALUATION, METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH NO VISIBLE DAMAGE. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4740706&pc=MTA | 2023826-2015-00488 | 2015-01-19 | Injury | STAAR SURGICAL COMPANY | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN BCVA WAS 20/20 AS OF (B)(6) 2014. REFERENCE MFR REPORT #2023826-2015-00487 FOR PATIENT'S RIGHT EYE (OD). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). EVALUATION, METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH NO VISIBLE DAMAGE. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842137&pc=MTA | 2023826-2015-00678 | 2015-01-19 | Injury | STAAR SURGICAL COMPANY | 2015-06-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE (1) SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012315&pc=MTA | 2023826-2015-01026 | 2015-01-19 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4529079&pc=MTA | 2023826-2015-00170 | 2015-01-20 | Injury | STAAR SURGICAL COMPANY | 2015-02-19 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAS A LOW VAULT. THE PATIENT EXPERIENCED A SHALLOW ANTERIOR CHAMBER, A MANIFEST REFRACTION OF -2.5 AND TRANSITORY RESIDUAL REFRACTIVE ERROR. THE ICL REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE PATIENT WAS BEING MONITORED AND OVERTIME THE PATIENT RECOVERED. THE PROBLEM WAS RESOLVED AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). EVAL, METHOD - WORK ORDER SEARCH; RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855082&pc=MTA | 2023826-2015-00693 | 2015-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2015. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855102&pc=MTA | 2023826-2015-00694 | 2015-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016041&pc=MTA | 2023826-2015-01048 | 2015-01-21 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND EXCESSIVE VAULTING WAS NOTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5252018&pc=MTA | 2023826-2015-01552 | 2015-01-21 | Injury | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/037 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A LOW VAULT AND A REFRACTIVE SURPRISE. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4613544&pc=MTA | 2023826-2015-00326 | 2015-01-22 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON ENLARGED THE INCISION AND THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF REMAINING VISCO/FLUIDS. THE LENS WAS EXCHANGED FOR A SHORTER LENS BUT THE LENS WAS STILL A LITTLE LARGE. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084309&pc=MTA | 2023826-2015-01228 | 2015-01-22 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4, -14.0/+2.0/X095 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -14.0/+2.0/X095; DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885636&pc=MTA | 2023826-2015-00728 | 2015-01-24 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012307&pc=MTA | 2023826-2015-01023 | 2015-01-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012057&pc=MTA | 2023826-2015-01058 | 2015-01-25 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -11.50. DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5011366&pc=MTA | 2023826-2015-01059 | 2015-01-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -10.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: -10.50 EVENT PROBLEM CODES: PATIENT CODE: (B)(4)- NO CODE AVAILABLE (NEW INCISION), LABELED DEVICE CODE: (B)(4)- NO CODE AVAILABLE (LENS EXPLANTED), UNLABELED (B)(4)- NO CODE AVAILABLE (LOW VAULTING), UNLABELED DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): DEVICE NOT RETURNED. EVALUATION CODES: CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614953&pc=MTA | 2023826-2015-00360 | 2015-01-27 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). EVALUATION, METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614954&pc=MTA | 2023826-2015-00361 | 2015-01-27 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (NEW INCISION); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4620457&pc=MTA | 2023826-2015-00355 | 2015-01-27 | Injury | STAAR SURGICAL COMPANY | 2015-03-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS ON (B)(6) 2014 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS LOW, RESULTING IN AN EXPLANT OF THE LENS ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS EVALUATION ANTICIPATED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS EVALUATION ANTICIPATED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5017059&pc=MTA | 2023826-2015-01028 | 2015-01-27 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD : WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5384464&pc=MTA | 2023826-2016-00022 | 2015-01-27 | Injury | STAAR SURGICAL COMPANY | 2016-01-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -12.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2006. LOW VAULT WAS OBSERVED AND ON (B)(6) 2015, THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS. THIS RESOLVED THE PROBLEM. PATIENT'S LAST OFFICE VISIT ON (B)(6) 2015, BCVA WAS 20/20. SEE MFR. #2023826-2016-00021 FOR RIGHT EYE. Manufacturer Narrative: DIOPTER: -12.50. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842218&pc=MTA | 2023826-2015-00656 | 2015-01-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS (DATE UNKNOWN). LOW VAULT WAS NOTED, THE LENS WAS EXPLANTED (DATE UNKNOWN). EXCHANGE FOR ANOTHER LENS IS PLANNED. PATIENT'S LAST KNOWN BEST CORRECTED VISUAL ACUITY WAS 20/40 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. EVALUATION CONCLUSION: (NO CONCLUSION CAN BE DRAWN). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012302&pc=MTA | 2023826-2015-01019 | 2015-01-29 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4725291&pc=MTA | 2023826-2015-00430 | 2015-01-30 | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2014. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO EXCESSIVE VAULTING AND A REFRACTIVE SURPRISE/REFRACTIVE SURPRISE OVERTIME. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/13.3. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH AND DIOPTER POWER WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN LENS, WRONG LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016379&pc=MTA | 2023826-2015-01017 | 2015-01-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -07.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016387&pc=MTA | 2023826-2015-01018 | 2015-01-31 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -19.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016384&pc=MTA | 2023826-2015-01029 | 2015-02-01 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON 01/31/2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5043079&pc=MTA | 2023826-2015-01085 | 2015-02-02 | Injury | STAAR SURGICAL COMPANY | 2015-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -08.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGEL CLOSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UNCORRECTED VISUAL ACUITY - 20/19. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -08.00. (B)(4). CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553033&pc=MTA | 2023826-2016-00402 | 2015-02-02 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016390&pc=MTA | 2023826-2015-01030 | 2015-02-03 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5174745&pc=MTA | 2023826-2015-01387 | 2015-02-03 | Injury | STAAR SURGICAL COMPANY | 2015-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/33. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578350&pc=MTA | 2023826-2015-00265 | 2015-02-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT AGE AND WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). EVALUATION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE DETERMINED. AS REPORTED ON THE MEDICAL REVIEW, EXCESSIVE VAULTING IS A CONSEQUENCE OF WRONG LENS USE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF WHITE TO WHITE MEASUREMENTS BASED ON TECHNIQUES, IMPROPER SULCUS TO SULCUS MEASUREMENT, AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS, IRREGULAR SULCUS OR CILIARY SULCUS CYST). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842136&pc=MTA | 2023826-2015-00677 | 2015-02-05 | Injury | STAAR SURGICAL COMPANY | 2015-06-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE SURPRISE AND EXCESSIVE VAULT WAS NOTED. CUSTOMER IS CURRENTLY UNDER CONSULTATION. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: CORRECTED DATA/ADDITIONAL INFORMATION: CORRECTED TO ADVERSE EVENT. STATED EXCESSIVE VAULTING. SUPPLEMENTAL INFORMATION RECEIVED STATED NO EXCESSIVE VAULT OCCURRED. PREVIOUS "OTHER" SECTION: STATED INCORRECT DIOPTER. ACTUAL DIOPTER IS -6/+2.5/10. PREVIOUS EXP DATE IS INCORRECT. PREVIOUS MANUFACTURE DATE INCORRECT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5050682&pc=MTA | 2023826-2015-01156 | 2015-02-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167163&pc=MTA | 2023826-2015-01273 | 2015-02-07 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -16.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. A REFRACTIVE SURPRISE WAS NOTED ON 02/07/2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR ANOTHER SAME MODEL AND LENGTH LENS BUT DIFFERENT DIOPTER SIZE . THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -16.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5076237&pc=MTA | 2023826-2015-01129 | 2015-02-09 | Injury | STAAR SURGICAL COMPANY | 2015-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -13.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THIS EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5556337&pc=MTA | 2023826-2016-00407 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: ADDITIONAL DESCRIPTION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 AS OF 01/15/2016. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4669884&pc=MTA | 2023826-2015-00418 | 2015-02-13 | Injury | STAAR SURGICAL COMPANY | 2015-04-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF ICA (IRIDO-CORNEAL ANGLES), THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525083&pc=MTA | 2023826-2016-00337 | 2015-02-14 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -16.50/+2.00/X053 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2015, PATIENT'S LAST VISIT; UCVA WAS 20/22. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -16.50/+2.00/X053, DEVICE EVALUATED BY MANUFACTURER? NO - DEVICE NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5050331&pc=MTA | 2023826-2015-01148 | 2015-02-14 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UCVA WAS 20/16. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4669900&pc=MTA | 2023826-2015-00419 | 2015-02-15 | Injury | STAAR SURGICAL COMPANY | 2015-04-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) - (SECONDARY SURGERY); (NEW INCISION); (REDUCTION OF IRIDO-CORNEAL ANGLES). SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH A TORN HAPTIC. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4960341&pc=MTA | 2023826-2015-00987 | 2015-02-16 | Injury | STAAR SURGICAL COMPANY | 2015-08-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE SURGEON OBSERVED LOW VAULT ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119162&pc=MTA | 2023826-2015-01292 | 2015-02-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). CLAIM # (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5370553&pc=MTA | 2023826-2016-00020 | 2015-02-17 | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -10.00 (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5050431&pc=MTA | 2023826-2015-01149 | 2015-02-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012295&pc=MTA | 2023826-2015-01025 | 2015-02-19 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND THE PATIENT HAD A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012298&pc=MTA | 2023826-2015-01024 | 2015-02-19 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND THE PATIENT HAD A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161898&pc=MTA | 2023826-2015-01282 | 2015-02-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/16. SEE MFR #2023826-2015-01281 FOR LEFT EYE. Manufacturer Narrative: DIOPTER: -14.00. DEVICE EVALUATED BY MANUFACTURER. NO - (OTHER): LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4803648&pc=MTA | 2023826-2015-00618 | 2015-02-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-28 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYES (OD) AND PATIENT EXPERIENCED AN UNEXPECTED REFRACTIVE OUTCOME AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SMALLER SIZE AND THIS RESOLVED THE ISSUES.Event Description: ADDITIONAL INFORMATION RECEIVED - THE SURGEON NOTICED ROTATION AFTER IMPLANTATION. THE LENS WAS EXCHANGED FOR ANOTHER SAME SIZE LENS BUT DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: (B)(4). METHOD (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH, DIOPTER AND SPHERE WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. PER MEDICAL REVIEW - THE MOST LIKELY CAUSE OF THE RESIDUAL REFRACTION IS A COMBINATION OF SELECTING A LENS POWER THAT WAS OVER-CORRECTING THE REFRACTIVE ERROR AND THE OFF-AXIS POSITION OF THE TORIC LENS SECONDARY TO ROTATION OR MISALIGNMENT. WE CANNOT COMPLETELY RULE OUT LOW VAULT SECONDARY TO A SHORT LENS AS CONTRIBUTING FACTOR. CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN OR TILTED ICL, ETC. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION CODE 17: (EVALUATION CONCLUSION): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. CLAIM # 707642
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842223&pc=MTA | 2023826-2015-00659 | 2015-02-26 | Injury | STAAR SURGICAL COMPANY | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378616&pc=MTA | 2023826-2015-01673 | 2015-02-26 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM12.1 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4650451&pc=MTA | 2023826-2015-00411 | 2015-02-27 | Injury | STAAR SURGICAL COMPANY | 2015-04-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO: LENS WAS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: NO CONCLUSION CAN BE DRAWN. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975582&pc=MTA | 2023826-2015-00952 | 2015-02-27 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND LOW VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5011078&pc=MTA | 2023826-2015-01057 | 2015-02-27 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -11.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REMOVED ON THE SAME DAY DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ADDITIONAL INFORMATION: DIOPTER -11.00 (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. DEVICE NOT RETURNED. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THIS EVENT COULD NOT BE DETERMINED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5124743&pc=MTA | 2023826-2015-01293 | 2015-02-27 | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/X087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724252&pc=MTA | 2023826-2015-00451 | 2015-02-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-23 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LOW VAULT WAS NOTED AND THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). IF ROTATION, SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ECT.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842191&pc=MTA | 2023826-2015-00663 | 2015-03-02 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12MM ICM120V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN BEST CORRECTED VISUAL ACUITY WAS 20/20 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A CLEAR SURGICAL RESIDUE. NO VISIBLE DAMAGE WAS OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323565&pc=MTA | 2023826-2015-01730 | 2015-03-02 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5174736&pc=MTA | 2023826-2015-01381 | 2015-03-03 | Injury | STAAR SURGICAL COMPANY | 2015-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) /2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954577&pc=MTA | 2023826-2015-00966 | 2015-03-03 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323615&pc=MTA | 2023826-2015-01731 | 2015-03-03 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724365&pc=MTA | 2023826-2015-00453 | 2015-03-05 | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | VISIAN ICL IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). IF ROTATION, SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ECT.). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167030&pc=MTA | 2023826-2015-01373 | 2015-03-05 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.00/+3.0/X173 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012330&pc=MTA | 2023826-2015-01046 | 2015-03-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR LONGER ONE AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/23. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4973990&pc=MTA | 2023826-2015-00941 | 2015-03-09 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND LOW VAULT WAS OBSERVED THE NEXT DAY. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724363&pc=MTA | 2023826-2015-00454 | 2015-03-10 | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). EVALUATION CODES: METHOD: (OTHER) WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED IN LIQUID WITH NO VISIBLE DAMAGE. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287678&pc=MTA | 2023826-2015-01626 | 2015-03-10 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND SHALLOW CHAMBER WERE NOTED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY THE MANUFACTURER? NO, LENS WAS NOT RETURNED. (B)(4). CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4721369&pc=MTA | 2023826-2015-00466 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY | 2015-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855105&pc=MTA | 2023826-2015-00691 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2015. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE ICL LENS WAS EXPLANTED FROM THE RIGHT EYE (OD). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ICL LENS WAS EXPLANTED FROM THE RIGHT EYE (OD). VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171642&pc=MTA | 2023826-2015-01382 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.0/+3.5/X111 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5050510&pc=MTA | 2023826-2015-01151 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -03.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE LENS WAS EXPLANTED THE NEXT DAY DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724348&pc=MTA | 2023826-2015-00452 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS). LOW VAULT WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). EVALUATION, METHOD: (OTHER) WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE).OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5051428&pc=MTA | 2023826-2015-01153 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161901&pc=MTA | 2023826-2015-01281 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/20. SEE MFR #2023826-2015-01282 FOR RIGHT EYE. Manufacturer Narrative: DIOPTER: -14.50. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266060&pc=MTA | 2023826-2015-01602 | 2015-03-14 | Injury | STAAR SURGICAL COMPANY | 2015-12-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -14.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. LOW VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -14.50. (B)(4). CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5051340&pc=MTA | 2023826-2015-01152 | 2015-03-15 | Injury | STAAR SURGICAL COMPANY | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? LENS NOT RETURNED. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5051523&pc=MTA | 2023826-2015-01155 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5050969&pc=MTA | 2023826-2015-01150 | 2015-03-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE LENS WAS EXPLANTED THE NEXT DAY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806337&pc=MTA | 2023826-2015-00589 | 2015-03-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS - 68 (OTHER): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378474&pc=MTA | 2023826-2015-01666 | 2015-03-20 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND HAVE FOREIGN MATERIAL ON LENS SURFACE (HAIR). (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743957&pc=MTA | 2023826-2015-00504 | 2015-03-21 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4770082&pc=MTA | 2023826-2015-00541 | 2015-03-21 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS REMOVED AND NOT REPLACED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OS THE RETURNED LENS SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. THEREFORE, IT CAN BE JUSTIFIED THAT THE COMPLAINT WAS NOT PRODUCT RELATED. MEDICAL REVIEW - THERE IS NOT ENOUGH CLINICAL DATA TO DETERMINE A EXACT CAUSE OF THE EVENT, HOWEVER, WE CANNOT RULE OUT LOW VAULTING SECONDARY TO A SHORT LENS AS A CONTRIBUTING FACTOR. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5043434&pc=MTA | 2023826-2015-01087 | 2015-03-24 | Injury | STAAR SURGICAL COMPANY | 2015-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/+1.50/101 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND PATIENT HAD A SHALLOW CHAMBER. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/+1.50/101. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4791922&pc=MTA | 2023826-2015-00603 | 2015-03-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-22 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND EXCESSIVE VAULT WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(6). (B)(4). (OTHER): LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ACT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. THE LENS WAS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5051698&pc=MTA | 2023826-2015-01154 | 2015-03-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE LENS WAS EXPLANTED THE NEXT DAY DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855107&pc=MTA | 2023826-2015-00690 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5043693&pc=MTA | 2023826-2015-01088 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2015-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/+1.00/97 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND PATIENT HAD A SHALLOW CHAMBER. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/+1.00/97. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5565337&pc=MTA | 2023826-2016-00396 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2016-04-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -18.00/+1.00/X089 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED THE NEXT DAY. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S VISIT ON (B)(6) 2015, UCVA WAS 20/30. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -18.00/+1.00/X089 DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885658&pc=MTA | 2023826-2015-00732 | 2015-03-26 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5212116&pc=MTA | 2023826-2015-01440 | 2015-03-27 | Injury | STAAR SURGICAL COMPANY | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED AND THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.50. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743827&pc=MTA | 2023826-2015-00498 | 2015-03-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR REPORT# 2023826-2015-00498 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. LENS EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE OPTIC WAS TORN AND THE LENS WAS RETURNED DRY. THE WORK ORDER SEARCH SHOWED NO OTHER COMPLAINT RECORDED RELATED TO THE SAME WORK ORDER. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5042413&pc=MTA | 2023826-2015-01089 | 2015-04-01 | Injury | STAAR SURGICAL COMPANY | 2015-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND THE PATIENT HAD A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST -OP VISIT ON (B)(6) 2015, UNCORRECTED VISUAL ACUITY: 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -11.50. DEVICE EVALUATED BY MANUFACTURER? NO LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) : LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5042575&pc=MTA | 2023826-2015-01095 | 2015-04-01 | Injury | STAAR SURGICAL COMPANY | 2015-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND THE PATIENT HAD A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST -OP VISIT ON (B)(6) 2015, UNCORRECTED VISUAL ACUITY: 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -07.00. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643996&pc=MTA | 2023826-2016-00588 | 2015-04-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4, -21.0 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773366&pc=MTA | 2023826-2015-00554 | 2015-04-04 | Injury | STAAR SURGICAL COMPANY | 2015-05-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4940663&pc=MTA | 2023826-2015-00930 | 2015-04-06 | Injury | STAAR SURGICAL COMPANY | 2015-07-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE PATIENT HAD CATARACT SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - AN ANTERIOR OPACITY WAS NOTED ON (B)(6) 2015. THE PATIENT HAD CATARACT SURGERY AND AN IOL WAS IMPLANTED. Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, ICL WAS EXPLANTED TWO MONTHS POSTOPERATIVELY TO ADDRESS INADEQUATE VAULT. ACCORDING TO THE COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, AT THE SAME SECONDARY SURGERY DATE, CATARACT SURGERY WITH IOL IMPLANTATION WAS PERFORMED DUE TO ANTERIOR LENS OPACITY OF UNKNOWN CAUSE. PER SAME REPORT, UCVA WAS 20/30 +2 AND PATIENT WAS SCHEDULED FOR REPLACEMENT OF NATURAL LENS OF THE OTHER EYE. IT SHOULD BE NOTED THAT AT THE TIME OF THE INITIAL SURGERY THE PATIENT WAS (B)(6) AND PER DFU: "INDICATIONS :THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" THEREFORE, THERE IS NO SUFFICIENT DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. CONCLUSION : BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4) - (NEW INCISION); (SECONDARY SURGERY); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579343&pc=MTA | 2023826-2016-00478 | 2015-04-06 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+2.50/086 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S UNCORRECTED VISUAL ACUITY IS 20/25. DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. PERFORMED VISUAL INSPECTION AND FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5054920&pc=MTA | 2023826-2015-01161 | 2015-04-08 | Injury | STAAR SURGICAL COMPANY | 2015-09-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18/1.5/105 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ASSOCIATED WITH EXCESSIVE VAULTING WAS OBSERVED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE EYE IS QUIET AND DEEP AC NOW. PATIENT IS DOING WELL AND EXTREMELY HAPPY NOW. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086240&pc=MTA | 2023826-2015-01202 | 2015-04-08 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -19.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5050795&pc=MTA | 2023826-2015-01147 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5212748&pc=MTA | 2023826-2015-01437 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/22. SEE MFR. #2023826-2015-01438 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(6). DIOPTER: -09.00. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5212132&pc=MTA | 2023826-2015-01438 | 2015-04-10 | Injury | STAAR SURGICAL COMPANY | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/22. SEE MFR. #2023826-2015-01437 FOR LEFT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DIOPTER: -08.50. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5649019&pc=MTA | 2023826-2016-00601 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5720705&pc=MTA | 2023826-2016-00611 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4974919&pc=MTA | 2023826-2015-00946 | 2015-04-16 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND LOW VAULT WAS OBSERVED THE NEXT DAY. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4974976&pc=MTA | 2023826-2015-00947 | 2015-04-20 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND LOW VAULT WAS OBSERVED THE NEXT DAY. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086250&pc=MTA | 2023826-2015-01210 | 2015-04-20 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933038&pc=MTA | 2023826-2015-00911 | 2015-04-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. SEE MFR REPORT# 2023826-2015-00910 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. ANTICIPATED BUT NOT YET BEGUN. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4803384&pc=MTA | 2023826-2015-00617 | 2015-04-21 | Injury | STAAR SURGICAL COMPANY | 2015-05-28 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ASSOCIATED WITH EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (SECONDARY SURGERY); (NEW INCISION); SIGNIFICANT REDUCTION OF ICA. SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OS THE RETURNED LENS SHOWED THE LENS WAS RETURNED IN LIQUID WITH A TORN HAPTIC. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4835331&pc=MTA | 2023826-2015-00650 | 2015-04-21 | Injury | STAAR SURGICAL COMPANY | 2015-06-10 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND VAULT IS STILL LOW. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4836536&pc=MTA | 2023826-2015-00653 | 2015-04-22 | Injury | STAAR SURGICAL COMPANY | 2015-06-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ASSOCIATED WITH EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR SHORTER LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINTS WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885870&pc=MTA | 2023826-2015-00782 | 2015-04-23 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIE6CED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS -(OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5405799&pc=MTA | 2023826-2016-00046 | 2015-04-23 | Injury | STAAR SURGICAL COMPANY | 2016-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.50/+4.00/X080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015, FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR.#2023826-2016-00045 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-14.50/+4.00/X080 (B)(4). METHOD CODE(S): (EVALUATION METHOD): WORK ORDER SEARCH. RESULTS CODE(S): (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5107010&pc=MTA | 2023826-2015-01192 | 2015-04-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015.THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING AND THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2015, PATIENT'S LAST VISIT; UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: -9.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154925&pc=MTA | 2023826-2015-01349 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266238&pc=MTA | 2023826-2015-01595 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-12-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -13.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. LOW VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -13.00. (B)(4). CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5535802&pc=MTA | 2023826-2016-00319 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/+1.50/X088 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND SHALLOW CHAMBER WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. UCVA IS 20/24. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.00/+1.50/X088 (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5174752&pc=MTA | 2023826-2015-01386 | 2015-04-27 | Injury | STAAR SURGICAL COMPANY | 2015-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -17.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4930378&pc=MTA | 2023826-2015-00910 | 2015-04-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (ODS) ON (B)(6) 2013 AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST BCVA WAS 20/16. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, ANTICIPATED BUT NOT YET BEGUN. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: PATIENT'S LAST POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/16. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5042513&pc=MTA | 2023826-2015-01090 | 2015-04-28 | Injury | STAAR SURGICAL COMPANY | 2015-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND THE PATIENT HAD A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST -OP VISIT ON (B)(6) 2015, UNCORRECTED VISUAL ACUITY: 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -11.50. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919456&pc=MTA | 2023826-2015-00837 | 2015-04-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND EXCESSIVE VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. SEE MFR REPORT# 2023826-2015-00836 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086252&pc=MTA | 2023826-2015-01239 | 2015-04-29 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4813664&pc=MTA | 2023826-2015-00633 | 2015-05-02 | Injury | STAAR SURGICAL COMPANY | 2015-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): MEDICAL REVIEW, WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): PER MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND. CONCLUSION CODE: (EVALUATION CONCLUSION): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4813665&pc=MTA | 2023826-2015-00634 | 2015-05-02 | Injury | STAAR SURGICAL COMPANY | 2015-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON 05/02/2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084418&pc=MTA | 2023826-2015-01134 | 2015-05-02 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -8.0/+1.5/X086 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. LAST DOCTOR VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -8.0/+1.5/X086. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THIS EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5399897&pc=MTA | 2023826-2016-00081 | 2015-05-02 | Injury | STAAR SURGICAL COMPANY | 2016-01-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4997291&pc=MTA | 2023826-2015-01036 | 2015-05-03 | Injury | STAAR SURGICAL COMPANY | 2015-08-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM AND THE PATIENT'S LAST BCVA WAS 20/13.3. SEE MFR REPORT (B)(4) FOR THE EVENT ON THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: -10.50 (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4998074&pc=MTA | 2023826-2015-01035 | 2015-05-03 | Injury | STAAR SURGICAL COMPANY | 2015-08-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM AND THE PATIENT'S LAST BCVA WAS 20/28.5. SEE MFR REPORT #2023826-2015-01036 FOR THE EVENT ON THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4922990&pc=MTA | 2023826-2015-00823 | 2015-05-04 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047744&pc=MTA | 2023826-2015-01112 | 2015-05-04 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119167&pc=MTA | 2023826-2015-01297 | 2015-05-04 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5048142&pc=MTA | 2023826-2015-01086 | 2015-05-04 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -07.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: -07.50 DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4835834&pc=MTA | 2023826-2015-00652 | 2015-05-05 | Injury | STAAR SURGICAL COMPANY | 2015-06-10 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST BCVA WAS 20/25. SEE MFR REPORT #2023826-2015-00651 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4871739&pc=MTA | 2023826-2015-00741 | 2015-05-05 | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE VAULT IS STILL TOO LOW. SURGEON PLANS ON EXCHANGING THE 13.2 FOR A 13.7. SEE MFR REPORT #2023826-2015-00742 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(6). VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. PER MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525075&pc=MTA | 2023826-2016-00330 | 2015-05-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525077&pc=MTA | 2023826-2016-00331 | 2015-05-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/070DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218066&pc=MTA | 2023826-2015-01491 | 2015-05-06 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855084&pc=MTA | 2023826-2015-00699 | 2015-05-09 | Injury | STAAR SURGICAL COMPANY | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432677&pc=MTA | 2023826-2016-00147 | 2015-05-09 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408506&pc=MTA | 2023826-2016-00096 | 2015-05-10 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+0.50/X109 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENS WITH A SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DIOPTER:-12.00/+0.50/X109. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4949679&pc=MTA | 2023826-2015-00895 | 2015-05-12 | Injury | STAAR SURGICAL COMPANY | 2015-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218059&pc=MTA | 2023826-2015-01496 | 2015-05-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+0.5/174 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323365&pc=MTA | 2023826-2015-01732 | 2015-05-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323370&pc=MTA | 2023826-2015-01734 | 2015-05-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4872763&pc=MTA | 2023826-2015-00718 | 2015-05-16 | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378477&pc=MTA | 2023826-2015-01670 | 2015-05-16 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4872762&pc=MTA | 2023826-2015-00717 | 2015-05-17 | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4): EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4871948&pc=MTA | 2023826-2015-00742 | 2015-05-19 | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 . THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR REPORT #2023826-2015-00741 FOR THE OTHER INCIDENT INVOLVING THE SAME EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086245&pc=MTA | 2023826-2015-01204 | 2015-05-20 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -08.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154937&pc=MTA | 2023826-2015-01355 | 2015-05-20 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/X081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378611&pc=MTA | 2023826-2015-01664 | 2015-05-21 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408918&pc=MTA | 2023826-2016-00123 | 2015-05-23 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2M VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.00/+2.00/005 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167604&pc=MTA | 2023826-2015-01374 | 2015-05-25 | Injury | STAAR SURGICAL COMPANY | 2015-10-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. VAULTING. EXPLANTED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5386394&pc=MTA | 2023826-2016-00056 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY | 2016-01-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -16.5/+3.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975056&pc=MTA | 2023826-2015-00972 | 2015-05-27 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING AND THE LENS DISLOCATED/SUBLUXED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975057&pc=MTA | 2023826-2015-00973 | 2015-05-27 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING AND THE LENS DISLOCATED/SUBLUXED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885859&pc=MTA | 2023826-2015-00783 | 2015-05-28 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4886970&pc=MTA | 2023826-2015-00784 | 2015-05-28 | Injury | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5412037&pc=MTA | 2023826-2016-00044 | 2015-05-28 | Injury | STAAR SURGICAL COMPANY | 2016-02-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -16.50/+2.00/X085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULTING WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DIOPTER:-16.50/+2.00/X085. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4973714&pc=MTA | 2023826-2015-00937 | 2015-05-29 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS OBSERVED AND THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4930289&pc=MTA | 2023826-2015-00854 | 2015-06-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND EXCESSIVE VAULT WAS NOTED (ACD 2.86). THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(6). (OTHER): LENS NOT RETURNED. CLAIM # (B)(4): LENS NOT RETURNED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). LENS WAS THROWN AWAY. Manufacturer Narrative: THE DATE RECEIVED BY MANUFACTURER FOR THE PREVIOUS FOLLOW-UP WAS INCORRECT. THE CORRECT DATE IS (B)(4) 2017.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975030&pc=MTA | 2023826-2015-00970 | 2015-06-02 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING EXTREMELY WELL AND HAPPY. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4904126&pc=MTA | 2023826-2015-00811 | 2015-06-03 | Injury | STAAR SURGICAL COMPANY | 2015-07-09 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES ASSOCIATED WITH EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5224357&pc=MTA | 2023826-2015-01458 | 2015-06-04 | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.50/X086 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WITH LENS ROTATION WAS NOTED. REPOSITIONING OF THE LENS WAS PERFORMED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -17.50/+1.50/X086. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167027&pc=MTA | 2023826-2015-01375 | 2015-06-05 | Injury | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: CORRECTIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084662&pc=MTA | 2023826-2015-01131 | 2015-06-09 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -04.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. POST-OP, THE PATIENT IS STILL COMPLAINING OF REDUCED FIELD OF VISION, WITH TEMPORAL "SHADOWS" BILATERALLY. LAST DOCTOR VISIT ON (B)(6) 2015, UCVA WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -04.50. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5124741&pc=MTA | 2023826-2015-01301 | 2015-06-09 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5107002&pc=MTA | 2023826-2015-01189 | 2015-06-10 | Injury | STAAR SURGICAL COMPANY | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015.THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: (B)(4). CONCLUSION: REVIEW OF THE FILE INDICATES THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ACD BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER. EXCESSIVE VAULT, AND EXPLANT IS ADDITIONAL NOTES ASSOCIATED WITH THE RETURN. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -10.50. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5400115&pc=MTA | 2023826-2016-00075 | 2015-06-10 | Injury | STAAR SURGICAL COMPANY | 2016-01-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD(PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119169&pc=MTA | 2023826-2015-01302 | 2015-06-15 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933392&pc=MTA | 2023826-2015-00908 | 2015-06-17 | Injury | STAAR SURGICAL COMPANY | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPOSITIONED BUT WITH NO SUCCESS. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD (OTHER): WORK ORDER SEARCH RESULTS -(OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5223684&pc=MTA | 2023826-2015-01456 | 2015-06-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +4.00/+2.00/X123 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. A REFRACTIVE SURPRISE WAS IDENTIFIED AFTER IMPLANTATION AND THE ICL WAS IMMEDIATELY EXPLANTED. NO OTHER LENS WAS IMPLANTED. THE CALCULATION WAS INCORRECT. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING NO VISABLE DAMAGE TO LENS. A DIMENSIONAL AND FUNCTIONAL INSPECTION WAS PERFORMED, FINDING THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: +4.00/+2.00/X123. (B)(4). CONCLUSIONS CODE(S): (USER ERROR CAUSED OR CONTRIBUTED TO EVENT): BASED ON THE COMPLAINT HISTORY, IT WAS DETERMINED THAT THE ROOT CAUSE OF THIS EVENT WAS DUE TO ERROR IN CALCULATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5178917&pc=MTA | 2023826-2015-01323 | 2015-06-18 | Injury | STAAR SURGICAL COMPANY | 2015-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -9.0/+1.0/X018 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH AND SIMILAR DIOPTER LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -9.0/+1.0/X018. (B)(4). METHOD CODES: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODES: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5405766&pc=MTA | 2023826-2016-00049 | 2015-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR.#2023826-2016-00048 FOR RIGHT EYE. Manufacturer Narrative: DIOPTER:-11.50. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5405768&pc=MTA | 2023826-2016-00048 | 2015-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR. #2023826-2016-00049 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-15.00. (B)(4). METHOD CODE(S): (EVALUATION METHOD): WORK ORDER SEARCH. RESULTS CODE(S): (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323390&pc=MTA | 2023826-2015-01736 | 2015-06-24 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642732&pc=MTA | 2023826-2016-00650 | 2015-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4, -15.50/+3.0/X076 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENGTH LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2016-00649 FOR RIGHT EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. (B)(4). NO NEW INCISION WAS MADE. THIS CODE IS NOT APPLICABLE UNDER DEVICE CODE; IT SHOULD BE LISTED UNDER PATIENT CODE INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171630&pc=MTA | 2023826-2015-01379 | 2015-06-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642727&pc=MTA | 2023826-2016-00649 | 2015-06-24 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4, -16.00/+4.00/X089 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENGTH LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2016-00650 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON THE LENS SURFACE (NOT POSSIBLE TO SPECIFY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086242&pc=MTA | 2023826-2015-01203 | 2015-06-25 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5386360&pc=MTA | 2023826-2016-00054 | 2015-06-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218716&pc=MTA | 2023826-2015-01480 | 2015-06-26 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293163&pc=MTA | 2023826-2015-01651 | 2015-06-27 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -17.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. LOW VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A LONGER LENS. THIS RESOLVED THE PROBLEM. ON LAST VISIT, THE PATIENT'S UCVA WAS 20/17. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: DIOPTER: -17.50. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171639&pc=MTA | 2023826-2015-01389 | 2015-06-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533712&pc=MTA | 2023826-2016-00346 | 2015-06-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -09.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533715&pc=MTA | 2023826-2016-00347 | 2015-06-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -09.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5072035&pc=MTA | 2023826-2015-01124 | 2015-06-30 | Injury | STAAR SURGICAL COMPANY | 2015-09-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR #2023826-2015-01122 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5082424&pc=MTA | 2023826-2015-01122 | 2015-06-30 | Injury | STAAR SURGICAL COMPANY | 2015-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR #2023826-2015-01124 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.50. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167012&pc=MTA | 2023826-2015-01367 | 2015-07-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933376&pc=MTA | 2023826-2015-00901 | 2015-07-02 | Injury | STAAR SURGICAL COMPANY | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2006. THE LENS HAD AN INADEQUATE VAULT AND REFRACTIVE CHANGE OVER TIME. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933377&pc=MTA | 2023826-2015-00900 | 2015-07-02 | Injury | STAAR SURGICAL COMPANY | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON 09/18/2006. THE LENS HAD AN INADEQUATE VAULT AND REFRACTIVE CHANGE OVER TIME. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016310&pc=MTA | 2023826-2015-01063 | 2015-07-02 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. SUBSEQUENTLY, THE PATIENT HAD SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). CONCLUSION: PATIENT IS OVER THE AGE OF 45 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS OVER THE AGE OF 45 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -10.00. (B)(4)- EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WAS ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323407&pc=MTA | 2023826-2015-01738 | 2015-07-02 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5562577&pc=MTA | 2023826-2016-00406 | 2015-07-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323191&pc=MTA | 2023826-2015-01704 | 2015-07-06 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5371920&pc=MTA | 2023826-2016-00063 | 2015-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+4.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171706&pc=MTA | 2023826-2015-01287 | 2015-07-07 | Injury | STAAR SURGICAL COMPANY | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER -10.5/+1.0/X109 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A SMALLER LENS. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS . (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.5/+1.0/X109. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4940687&pc=MTA | 2023826-2015-00931 | 2015-07-08 | Injury | STAAR SURGICAL COMPANY | 2015-07-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO BEING TOO LARGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2015. THE ICL WAS NOTED TO BE LOO LARGE RIGHT AFTER THE LENS WAS IMPLANTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT IS DOING GREAT WITH NO SIGNS OR SYMPTOMS AFTER CORRECT SIZE LENS WAS IMPLANTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF AND MISSING. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ACCURATE DETERMINATION OF ANTERIOR SEGMENT DIMENSIONS IS KEY TO THE SAFETY OF IMPLANTED ICLS. WTW AND ACD MEASUREMENTS, THE RECOMMENDED METHOD FOR LENS SIZING PER THE FDA APPROVED DFU, WAS USED BY THE FACILITY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5224353&pc=MTA | 2023826-2015-01508 | 2015-07-08 | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -02.75 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION(UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5224356&pc=MTA | 2023826-2015-01509 | 2015-07-08 | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -02.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS -(EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525076&pc=MTA | 2023826-2016-00332 | 2015-07-08 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). ADDITIONAL INFORMATION: CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.92MM AT THE TIME OF IMPLANTATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5213887&pc=MTA | 2023826-2015-01413 | 2015-07-10 | Injury | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -08.50. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628359&pc=MTA | 2023826-2016-00557 | 2015-07-12 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: LENS DESCRIPTION IS -11.0/+1.0/093 (SPHERE/ CYLINDER/ AXIS). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/+1.0/093 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 WITH NO REPORTED PATIENT ISSUES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325786&pc=MTA | 2023826-2015-01715 | 2015-07-13 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/X113 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325934&pc=MTA | 2023826-2015-01716 | 2015-07-13 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+0.5/X083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5099291&pc=MTA | 2023826-2015-01173 | 2015-07-14 | Injury | STAAR SURGICAL COMPANY | 2015-09-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/+2.0/X040 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2015, EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. ON (B)(6) 2015. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SMALLER LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.50/+2.0/X040. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5234735&pc=MTA | 2023826-2015-01518 | 2015-07-14 | Injury | STAAR SURGICAL COMPANY | 2015-11-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS RETURNED BENT. THE CORRECT REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323406&pc=MTA | 2023826-2015-01737 | 2015-07-15 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5040393&pc=MTA | 2023826-2015-01109 | 2015-07-16 | Injury | STAAR SURGICAL COMPANY | 2015-08-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167018&pc=MTA | 2023826-2015-01365 | 2015-07-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(6). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5037266&pc=MTA | 2023826-2015-01106 | 2015-07-17 | Injury | STAAR SURGICAL COMPANY | 2015-08-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323382&pc=MTA | 2023826-2015-01735 | 2015-07-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IN TO THE PATIENT'S LEFTT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5792720&pc=MTA | 2023826-2016-00889 | 2015-07-18 | Injury | STAAR SURGICAL COMPANY | 2016-07-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -12.5/1.00/094 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL (BLOOD) WAS PRESENT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DOB-UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5697787&pc=MTA | 2023826-2016-00802 | 2015-07-20 | Injury | STAAR SURGICAL COMPANY | 2016-06-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5411888&pc=MTA | 2023826-2016-00047 | 2015-07-21 | Injury | STAAR SURGICAL COMPANY | 2016-02-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULTING WITHOUT LENS ROTATION WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-13.00. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE(S): (EVALUATION METHOD): WORK ORDER SEARCH RESULTS CODE(S): (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5023054&pc=MTA | 2023826-2015-01073 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2015-08-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -13.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER. ON (B)(6) 2015 UCVA: 20/25. SEE MFR# 2023826-2015-01072 FOR RIGHT EYE. Manufacturer Narrative: (B)(4) - EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5023223&pc=MTA | 2023826-2015-01072 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2015-08-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER. ON (B)(6) 2015 UCVA: 20/25. SEE MFR#2023826-2015-01073 FOR LEFT EYE. Manufacturer Narrative: (B)(4) - EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5362051&pc=MTA | 2023826-2016-00006 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2016-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.00/+2.50/X079 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2016-00007 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00/+2.50/X079. DEVICE EVALUATED BY MANUFACTURER? NO. DEVICE NOT RETURNED. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167015&pc=MTA | 2023826-2015-01368 | 2015-07-23 | Injury | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. VAULTING. EXPLANTED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5209473&pc=MTA | 2023826-2015-01436 | 2015-07-23 | Injury | STAAR SURGICAL COMPANY | 2015-11-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -15.00/+2.00/X142 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED AND THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/22. SEE MFR. #2023829-2015-01432 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(6). DIOPTER: -15.00/+2.00/X142. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167014&pc=MTA | 2023826-2015-01372 | 2015-07-24 | Injury | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/080 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655172&pc=MTA | 2023826-2016-00682 | 2015-07-24 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED AND THE LENS WAS REMOVED ON (B)(6) 2015. A LARGER LENS WAS IMPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210640&pc=MTA | 2023826-2015-01432 | 2015-07-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.50/+3.00/X005 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED AND THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/22. SEE MFR. #2023829-2015-01436 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(6). DIOPTER: -14.50/+3.00/X005. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579569&pc=MTA | 2023826-2016-00476 | 2015-07-26 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND THERE WAS A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218099&pc=MTA | 2023826-2015-01499 | 2015-07-28 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154918&pc=MTA | 2023826-2015-01347 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218094&pc=MTA | 2023826-2015-01497 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/114 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218820&pc=MTA | 2023826-2015-01501 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218971&pc=MTA | 2023826-2015-01492 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247873&pc=MTA | 2023826-2015-01558 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +05.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5362048&pc=MTA | 2023826-2016-00007 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.00/+4.00/X095 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2016-00006 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00/+4.00/X095 DEVICE EVALUATED BY MANUFACTURER? NO. DEVICE NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119108&pc=MTA | 2023826-2015-01303 | 2015-07-30 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND AN UNREACTIVE PUPIL (FIXED). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4): LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218992&pc=MTA | 2023826-2015-01493 | 2015-07-30 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638606&pc=MTA | 2023826-2016-00591 | 2015-07-30 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.5/127 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. PATIENT HAD KERATOCONUS AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH KERATOCONUS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826405&pc=MTA | 2023826-2016-00985 | 2015-07-30 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151468&pc=MTA | 2023826-2015-01242 | 2015-07-31 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/X093 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REMOVED ON THE SAME DAY DUE TO EXCESSIVE VAULTING AND THE PATIENT HAD SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA: 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/+2.0/X093. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) SHALLOW CHAMBER. VAULTING. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154933&pc=MTA | 2023826-2015-01353 | 2015-07-31 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643746&pc=MTA | 2023826-2016-00583 | 2015-07-31 | Injury | STAAR SURGICAL COMPANY | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.0MM ICM130V4, -11.0 DIOPTER IMPANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS REPORTED, THE LENS WAS REPLACED WITH A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154908&pc=MTA | 2023826-2015-01344 | 2015-08-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.5/+5.5/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579570&pc=MTA | 2023826-2016-00477 | 2015-08-03 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTED INDICATED THAT THE VAULT OF THE LENS WAS LOW. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED. (B)(4). (NEW INCISION) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: DIOPTER: -5.50. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323369&pc=MTA | 2023826-2015-01733 | 2015-08-04 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323636&pc=MTA | 2023826-2015-01728 | 2015-08-04 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS REPLACED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5042735&pc=MTA | 2023826-2015-01103 | 2015-08-06 | Injury | STAAR SURGICAL COMPANY | 2015-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -08.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -08.50. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5243470&pc=MTA | 2023826-2015-01533 | 2015-08-08 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.50/1.0/X101 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR SHORTER ICL WITH A SIMILAR DIOPTER SIZE. THE RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/16.6. SEE MFR. #2023826-2015-01532 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50/1.0/X101 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER SAME MODEL LENS WITH SIMILAR DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2015 WITH POST-OP UNCORRECTED VISUAL ACUITY (UCVA) 20/16.6. SEE MFR. #2023826-2015-01532 FOR RIGHT EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -07.50/1.0/X101. (B)(4). EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5243666&pc=MTA | 2023826-2015-01532 | 2015-08-08 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.50/1.50/X103 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR SHORTER ICL WITH A SIMILAR DIOPTER SIZE. THE RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/16.6. SEE MFR. #2023826-2015-01533 FOR LEFT EYE. Manufacturer Narrative: (B)(4). THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50/1.50/X103 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER SAME MODEL LENS WITH SIMILAR DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2015 WITH POST-OP UNCORRECTED VISUAL ACUITY (UCVA) 20/16.6. SEE MFR. #2023826-2015-01533 FOR LEFT EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -07.50/1.5/X103. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5215816&pc=MTA | 2023826-2015-01485 | 2015-08-09 | Injury | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146312&pc=MTA | 2023826-2015-01240 | 2015-08-12 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/X068 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 015 DUE TO EXCESSIVE VAULTING AND THE PATIENT HAD A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA: 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00/+2.0/X068. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378475&pc=MTA | 2023826-2015-01667 | 2015-08-12 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -09.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5386366&pc=MTA | 2023826-2016-00053 | 2015-08-12 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5215836&pc=MTA | 2023826-2015-01484 | 2015-08-13 | Injury | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323345&pc=MTA | 2023826-2015-01729 | 2015-08-14 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655196&pc=MTA | 2023826-2016-00690 | 2015-08-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -8.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WERE TWO RETURNED LENSES STUCK TOGETHER. THE OPTIC WAS TORN/SPLIT AND THE HAPTIC WAS TORN/BENT/BROKEN/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). REFERENCE MFG # 2023826-2016-00691 FOR LEFT EYE (OS). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655205&pc=MTA | 2023826-2016-00691 | 2015-08-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -7.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THE LENS WAS RETURNED IN PIECES AND COULD NOT BE EVALUATED. CORRECTED DATA: (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). REFERENCE MFG # 2023826-2016-00690 FOR RIGHT EYE (OD). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210046&pc=MTA | 2023826-2015-01426 | 2015-08-15 | Injury | STAAR SURGICAL COMPANY | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS -19.00 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AND THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01425 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -19.00. METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210060&pc=MTA | 2023826-2015-01425 | 2015-08-15 | Injury | STAAR SURGICAL COMPANY | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS -21.0 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AND THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. THE FACILITY DISCARDED THE ICL. SEE MFR. #2023826-2015-01426 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4): LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): THE ICL WAS DISCARDED BY THE FACILITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5409015&pc=MTA | 2023826-2016-00093 | 2015-08-15 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50/+2.00/X090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DIOPTER:-8.50/+2.00/X090. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5219011&pc=MTA | 2023826-2015-01503 | 2015-08-17 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE US. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5219058&pc=MTA | 2023826-2015-01504 | 2015-08-17 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.5/+2.0/107 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4) METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247885&pc=MTA | 2023826-2015-01556 | 2015-08-17 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/X011 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218087&pc=MTA | 2023826-2015-01494 | 2015-08-18 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) . THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5093135&pc=MTA | 2023826-2015-01171 | 2015-08-20 | Injury | STAAR SURGICAL COMPANY | 2015-09-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. SURGERY IS PENDING TO EXPLANT LENS AND EXCHANGE FOR SHORTER LENS. SEE MFR #2023826-2015-01172 FOR LEFT EYE. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS ON THE SAME DAY. THE EXCHANGE RESOLVED THE PROBLEM. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). "IRIDO CORNEAL ANGLES", NOT "INDO-CORNEAL ANGLES". DEVICE MANUFACTURE DATE- CORRECTED TO 02/13/2015. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DIOPTER: -18.00. EXPLANT DATE: NA. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE: EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5212839&pc=MTA | 2023826-2015-01414 | 2015-08-21 | Injury | STAAR SURGICAL COMPANY | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED AND THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -09.00. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5227461&pc=MTA | 2023826-2015-01424 | 2015-08-21 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -16.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. POST-OP VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -16.00. (B)(4). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323447&pc=MTA | 2023826-2015-01740 | 2015-08-23 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218985&pc=MTA | 2023826-2015-01415 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SHALLOW CHAMBER WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR SHORTER ICL WITH SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325833&pc=MTA | 2023826-2015-01710 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/X090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643977&pc=MTA | 2023826-2016-00582 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -2.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AND THE LENS WAS REMOVED ON (B)(6) 2015. A SMALLER LENS WAS IMPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323479&pc=MTA | 2023826-2015-01741 | 2015-08-26 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). METHOD CODE 3331 (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS CODE (EVALUATION RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5212788&pc=MTA | 2023826-2015-01422 | 2015-08-27 | Injury | STAAR SURGICAL COMPANY | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. SEE MFR. #2023826-2015-01423 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.50. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED.(B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5213886&pc=MTA | 2023826-2015-01423 | 2015-08-27 | Injury | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.50/X085 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. SEE MFR. #2023826-2015-01422 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.00/+1.50/X085. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293214&pc=MTA | 2023826-2015-01657 | 2015-08-27 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -13.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. LOW VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A LONGER LENS. THIS RESOLVED THE PROBLEM. LAST VISIT ON (B)(6) 2015, THE PATIENT'S UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE (B)(4). Manufacturer Narrative: DIOPTER: -13.50. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323421&pc=MTA | 2023826-2015-01739 | 2015-08-27 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). METHOD CODE (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS CODE (EVALUATION RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655175&pc=MTA | 2023826-2016-00683 | 2015-08-28 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM ICM115V4, -11.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655197&pc=MTA | 2023826-2016-00687 | 2015-08-28 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM ICM115V4, -12.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218184&pc=MTA | 2023826-2015-01418 | 2015-08-30 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SHALLOW CHAMBER WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A SHORTER ICL WITH SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.50. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5386326&pc=MTA | 2023826-2016-00052 | 2015-08-31 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869089&pc=MTA | 2023826-2016-01041 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/127 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRIED DISCOLORED MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378478&pc=MTA | 2023826-2015-01671 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308690&pc=MTA | 2023826-2015-01662 | 2015-09-03 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308710&pc=MTA | 2023826-2015-01663 | 2015-09-03 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/112 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655198&pc=MTA | 2023826-2016-00688 | 2015-09-03 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -14.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5288068&pc=MTA | 2023826-2015-01624 | 2015-09-04 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY THE MANUFACTURER? LENS WAS NOT RETURNED. (B)(4). METHOD CODE (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS CODE (PROCESS RESULT): A WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT. CONCLUSION CODE (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891382&pc=MTA | 2023826-2016-00679 | 2015-09-07 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. THERE WAS FOREIGN MATERIAL ON THE LENS SURFACE. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5312167&pc=MTA | 2023826-2015-01677 | 2015-09-09 | Injury | STAAR SURGICAL COMPANY | 2015-12-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325832&pc=MTA | 2023826-2015-01708 | 2015-09-10 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/X094 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432297&pc=MTA | 2023826-2016-00146 | 2015-09-15 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5208764&pc=MTA | 2023826-2015-01427 | 2015-09-16 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS -11.50 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AND THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S., BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.50. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287677&pc=MTA | 2023826-2015-01625 | 2015-09-16 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND SHALLOW CHAMBER WERE NOTED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY THE MANUFACTURER? NO, LENS WAS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (PROCESS RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293081&pc=MTA | 2023826-2015-01649 | 2015-09-17 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -09.5 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. LOW VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01650 FOR LEFT EYE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION-LENS RETURNED DRY IN LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH-ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DIOPTER: -09.5. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533618&pc=MTA | 2023826-2016-00349 | 2015-09-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577766&pc=MTA | 2023826-2016-00447 | 2015-09-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 -11.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2004. LOW VAULT WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS DID NOT RESOLVE THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/50. SEE MFR. #2023826-2016-00438 FOR CATARACT SURGERY AND IOL IMPLANTATION. Manufacturer Narrative: ADDITIONAL DATA: PMA510(K): NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED DRY AND BENT IN THE LENS CONTAINER WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: THE ICL WAS EXCHANGED FOR A LONGER LENS AND THIS RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: DIOPTER: -11.50. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5556363&pc=MTA | 2023826-2016-00404 | 2015-09-19 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 AS OF (B)(6) 2015. LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5212793&pc=MTA | 2023826-2015-01459 | 2015-09-21 | Injury | STAAR SURGICAL COMPANY | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.00/+4.00/X086 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -1.00/+4.00/X086. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578662&pc=MTA | 2023826-2015-00260 | 2015-09-22 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND THE LENS WAS EXPLANTED THE NEXT DAY DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST UCVA WAS 20/60. Manufacturer Narrative: (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW -ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628348&pc=MTA | 2023826-2016-00559 | 2015-09-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS REPORTED AND THE LENS WAS REMOVED ON (B)(6) 2015. A LARGER LENS WAS IMPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. CORRECTED DATA: -12/2.5/091 DIOPTER. DEVICE EVALUATION - LENS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628350&pc=MTA | 2023826-2016-00558 | 2015-09-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED AND THE LENS WAS REMOVED ON (B)(6)2015. A LARGER LENS WAS IMPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVENT DATE: UNK. PREVIOUS DATE IS NOT CONFIRMED DATE OF EVENT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218660&pc=MTA | 2023826-2015-01421 | 2015-09-23 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE ICL WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER ICL WITH SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5241793&pc=MTA | 2023826-2015-01529 | 2015-09-23 | Injury | STAAR SURGICAL COMPANY | 2015-11-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -06.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. POST-OP VISIT, UCVA WAS 20/20. Manufacturer Narrative: CORRECTED DATA: AGE AT TIME OF EVENT: PREVIOUS AGE (B)(6) TO CORRECTED AGE (B)(6); PREVIOUS AGE IS INCORRECT. (B)(4). (NEW INCISION) INCORRECT - NO NEW INCISION WAS MADE. (B)(4). LENS WAS EXCHANGED FOR A SHORTER LENS. DEVICE EVALUATION - PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. EVALUATION CONCLUSION CODE: PATIENT'S AGE WAS OUTSIDE RECOMMENDED RANGE OF IMPLANTATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -06.00. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325780&pc=MTA | 2023826-2015-01711 | 2015-09-23 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/X075 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5509260&pc=MTA | 2023826-2016-00280 | 2015-09-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5509266&pc=MTA | 2023826-2016-00279 | 2015-09-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655162&pc=MTA | 2023826-2016-00590 | 2015-09-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4, -10.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287698&pc=MTA | 2023826-2015-01635 | 2015-09-24 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/X080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325785&pc=MTA | 2023826-2015-01713 | 2015-09-24 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+0.5/X087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199929&pc=MTA | 2023826-2015-01360 | 2015-09-26 | Injury | STAAR SURGICAL COMPANY | 2015-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS LENGTH WITH A SIMILAR DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST VISIT ON (B)(6) 2015 - BCVA 20/13.3 AND UCVA 20/16.6. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -07.00. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323489&pc=MTA | 2023826-2015-01743 | 2015-09-26 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5234730&pc=MTA | 2023826-2015-01519 | 2015-09-28 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247874&pc=MTA | 2023826-2015-01557 | 2015-09-30 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.5/X119 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193763&pc=MTA | 2023826-2015-01445 | 2015-10-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE. THE SURGEON IS PLANNING ON EXPLANTING THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533623&pc=MTA | 2023826-2016-00348 | 2015-10-02 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5186014&pc=MTA | 2023826-2015-01361 | 2015-10-05 | Injury | STAAR SURGICAL COMPANY | 2015-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -16.00 DIOPTER IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SMALLER LENS. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULT CODE(S): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: DIOPTER: -16.00. DEVICE EVALUATED BY MANUFACTURER? NO -LENS DISCARDED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225033&pc=MTA | 2023826-2015-01464 | 2015-10-07 | Injury | STAAR SURGICAL COMPANY | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -11.00 DIOPER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE ICL WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01463 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. (B)(4). METHOD CODE: EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225035&pc=MTA | 2023826-2015-01463 | 2015-10-07 | Injury | STAAR SURGICAL COMPANY | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -11.00 DIOPER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE ICL WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01464 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432273&pc=MTA | 2023826-2016-00143 | 2015-10-07 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON 2015. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE, RESULTING IN THE EXPLANT OF THE LENS ON (B)(6) 2015 AND EXCHANGE FOR A LENS WITH A SMALLER VAULT. THE ISSUE WAS RESOLVED WITH THE IMPLANTATION OF THE NEW LENS. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432364&pc=MTA | 2023826-2016-00144 | 2015-10-07 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE, RESULTING IN THE EXPLANT OF THE LENS ON (B)(6) 2015 AND EXCHANGE FOR A LENS WITH A SMALLER VAULT. THE ISSUE WAS RESOLVED WITH THE IMPLANTATION OF THE NEW LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT DISTRIBUTED IN THE U.S. DEVICE EVALUATED BY THE MANUFACTURER? NO. DEVICE WAS NOT RETURNED. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5518855&pc=MTA | 2023826-2016-00295 | 2015-10-07 | Injury | STAAR SURGICAL COMPANY | 2016-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED A COUPLE OF HOURS LATER DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5312164&pc=MTA | 2023826-2015-01676 | 2015-10-08 | Injury | STAAR SURGICAL COMPANY | 2015-12-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572299&pc=MTA | 2023826-2016-00436 | 2015-10-10 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -20.0/+2.0/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN. CORRECTED DATA: PATIENT CODE: (NEW INCISION) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5649010&pc=MTA | 2023826-2016-00569 | 2015-10-10 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -18.5/+2.0/X087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408374&pc=MTA | 2023826-2016-00094 | 2015-10-13 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-06.50; EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE(S): (EVALUATION METHOD): WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225081&pc=MTA | 2023826-2015-01513 | 2015-10-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +10.0/+1.5/139 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408412&pc=MTA | 2023826-2016-00095 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DIOPTER:-18.00. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5375455&pc=MTA | 2023826-2016-00051 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533611&pc=MTA | 2023826-2016-00351 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) NEW INCISION. SECONDARY SURGERY. VAULTING. EXPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533612&pc=MTA | 2023826-2016-00352 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5242709&pc=MTA | 2023826-2015-01528 | 2015-10-15 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/-1.50/X090 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT AND LENS ROTATION WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER. PROBLEM WAS RESOLVES. POST VISIT ON (B)(6) 2015. UCVA WAS 20/20. Manufacturer Narrative: CORRECTED DATA: PREVIOUS MODEL VICMO13.2 IS INCORRECT, CORRECT DATA IS VTICMO13.2. (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -10.50/+1.50/X090 DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628366&pc=MTA | 2023826-2016-00560 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED AND THE LENS WAS REMOVED ON (B)(6) 2015. A LARGER LENS WAS IMPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). CONCLUSION: THE PATIENT'S ANTERIOR CHAMBER DEPTH IS LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT:UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5252473&pc=MTA | 2023826-2015-01554 | 2015-10-21 | Injury | STAAR SURGICAL COMPANY | 2015-11-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.0/+4.5/019 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247889&pc=MTA | 2023826-2015-01555 | 2015-10-22 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/X082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5480934&pc=MTA | 2023826-2016-00190 | 2015-10-22 | Injury | STAAR SURGICAL COMPANY | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 -08.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND ON (B)(6) 2015, THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. PATIENT'S LAST OFFICE VISIT ON (B)(6) 2015, UCVA: 20/20. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643966&pc=MTA | 2023826-2016-00584 | 2015-10-25 | Injury | STAAR SURGICAL COMPANY | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VICMO13.2, -14 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS REPORTED AND THE LENS WAS REMOVED ON (B)(6) 2015. A SMALLER LENS WAS IMPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PREVIOUS DATE OF EVENT IS UNCONFIRMED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643971&pc=MTA | 2023826-2016-00585 | 2015-10-25 | Injury | STAAR SURGICAL COMPANY | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -14.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS REPORTED AND THE LENS WAS REMOVED ON (B)(6) 2015. A SMALLER LENS WAS IMPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308728&pc=MTA | 2023826-2015-01661 | 2015-10-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -02.25 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE ANTERIOR CHAMBER WAS IRRIGATED AND AN INTRAOCULAR INJECTION WAS ADMINISTERED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5349511&pc=MTA | 2023826-2016-00003 | 2015-10-27 | Injury | STAAR SURGICAL COMPANY | 2016-01-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -7.50/+2.00/X107 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -7.50/+2.00/X107. DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): DEVICE NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909075&pc=MTA | 2023826-2016-01188 | 2015-10-27 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909099&pc=MTA | 2023826-2016-01189 | 2015-10-27 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LONGER LENS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5472107&pc=MTA | 2023826-2016-00097 | 2015-10-28 | Injury | STAAR SURGICAL COMPANY | 2016-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -12.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DIOPTER: -12.50. PT WEIGHT: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5472508&pc=MTA | 2023826-2016-00098 | 2015-10-28 | Injury | STAAR SURGICAL COMPANY | 2016-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A SMALLER LENGTH LENS WITH SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DIOPTER: -18.00. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602195&pc=MTA | 2023826-2016-00518 | 2015-10-28 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND LENS HAPTIC TORN/BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655176&pc=MTA | 2023826-2016-00686 | 2015-10-28 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4, -15.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293099&pc=MTA | 2023826-2015-01629 | 2015-10-30 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS OBSERVED AND THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM.SEE MFR. # 2023826-2015-01628 FOR LEFT EYE. Manufacturer Narrative: (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293222&pc=MTA | 2023826-2015-01628 | 2015-10-30 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -07.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS OBSERVED AND THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. # 2023826-2015-01629 FOR RIGHT EYE. Manufacturer Narrative: DIOPTER: -07.00. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5312177&pc=MTA | 2023826-2015-01679 | 2015-10-30 | Injury | STAAR SURGICAL COMPANY | 2015-12-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5312372&pc=MTA | 2023826-2015-01678 | 2015-10-30 | Injury | STAAR SURGICAL COMPANY | 2015-12-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TORN/BROKEN AND FIBERS ON LENS SURFACE. DIMENSION OF INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247870&pc=MTA | 2023826-2015-01545 | 2015-10-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -2.5/+2.5/X011 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A LOW VAULT AND REFRACTIVE SURPRISE. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5429557&pc=MTA | 2023826-2016-00130 | 2015-11-02 | Injury | STAAR SURGICAL COMPANY | 2016-02-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION PERFORMED FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5429386&pc=MTA | 2023826-2016-00129 | 2015-11-03 | Injury | STAAR SURGICAL COMPANY | 2016-02-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION CONDUCTED ON THE RETURNED LENS FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5445774&pc=MTA | 2023826-2016-00172 | 2015-11-03 | Injury | STAAR SURGICAL COMPANY | 2016-02-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS INTO AN INDIVIDUAL WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.91MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408763&pc=MTA | 2023826-2016-00090 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.00/X094 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. PATIENT'S LAST OFFICE VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: DIOPTER:-12.00/+2.50/X094. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655180&pc=MTA | 2023826-2016-00684 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM ICM115V4, -14.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LOW VAULT WAS REPORTED AND THE LENS WAS REMOVED ON (B)(6) 2015. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM ICM115V4, -14.5 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LOW VAULT WAS REPORTED AND THE LENS WAS REMOVED ON (B)(6) 2015. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5482380&pc=MTA | 2023826-2016-00188 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 1.50/+3.00/X046 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015 EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES AND A UNREACTIVE PUPIL. ON (B)(6) 2015, THE LENS WAS EXPLANTED AND THIS TIME NOT REPLACED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). METHOD - WORK ORDER SEARCH RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5296541&pc=MTA | 2023826-2015-01631 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.50/+3.0/X095 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015. THE EXCHANGED LENS, A SHORTER LENS WAS IMPLANTED ON (B)(6) 2015. Manufacturer Narrative: AGE AT TIME OF EVENT IS INCORRECT. DATE OF BIRTH PROVIDED. (B)(4). THE PROBLEM WAS RESOLVED. DEVICE EVALUATION : PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS CLEAR, SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. HOWEVER, THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.50/+3.00/X095. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5482069&pc=MTA | 2023826-2016-00189 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 +2.00/+0.50/X071 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND ON (B)(6) 2015, THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER SPHERICAL LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/ SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (NOT POSSIBLE TO SPECIFY) ON LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5505912&pc=MTA | 2023826-2016-00275 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/071 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5349510&pc=MTA | 2023826-2016-00002 | 2015-11-10 | Injury | STAAR SURGICAL COMPANY | 2016-01-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -8.50/+2.0/X094 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT IS DOING WELL AND HAPPY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -8.50/+2.00/X094, DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5384639&pc=MTA | 2023826-2016-00019 | 2015-11-10 | Injury | STAAR SURGICAL COMPANY | 2016-01-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -10.00. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5496061&pc=MTA | 2023826-2016-00261 | 2015-11-11 | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +06.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5649011&pc=MTA | 2023826-2016-00568 | 2015-11-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -19.5/+2.5/X088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5506283&pc=MTA | 2023826-2016-00274 | 2015-11-12 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE CASE/VIAL; VISUAL INSPECTION FOUND UNSPECIFIED FOREIGN MATERIAL ON SURFACE BUT NO VISIBLE DAMAGE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686811&pc=MTA | 2023826-2016-00214 | 2015-11-15 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (LIQUID). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5429384&pc=MTA | 2023826-2016-00126 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-02-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5429545&pc=MTA | 2023826-2016-00127 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-02-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). PROCESS EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643932&pc=MTA | 2023826-2016-00581 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 -11.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655188&pc=MTA | 2023826-2016-00689 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -14.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN, BROKEN AND DEFORMED. CORRECTED DATA: (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.8MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5349526&pc=MTA | 2023826-2016-00004 | 2015-11-19 | Injury | STAAR SURGICAL COMPANY | 2016-01-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 -13.00/+2.50X094 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.0/+2.50/X094. DEVICE EVALUATED BY MANUFACTURER: DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5399902&pc=MTA | 2023826-2016-00074 | 2015-11-23 | Injury | STAAR SURGICAL COMPANY | 2016-01-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5665326&pc=MTA | 2023826-2016-00739 | 2015-11-23 | Injury | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN, BROKEN AND BENT. THERE WAS DRY, DISCOLORED MATERIAL PRESENT ON THE LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870251&pc=MTA | 2023826-2016-01055 | 2015-11-23 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +2.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +2.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25 AND UNCORRECTED VISUAL ACUITY WAS 20/32. THE LENS WAS EXCHANGED FOR A LONGER AND DIFFERENT DIOPTER LENS ON (B)(6) 2016. METHOD: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5370126&pc=MTA | 2023826-2016-00018 | 2015-11-26 | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -08.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. LAST OFFICE VISIT ON (B)(6) 2015 BCVA: 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE OPTIC TORN. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - DEVICE NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432303&pc=MTA | 2023826-2016-00148 | 2015-11-26 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816259&pc=MTA | 2023826-2016-00969 | 2015-11-26 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+3.5/175 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED INFORMATION: (B)(4) SHOULD BE INCLUDED. (B)(4) CODE NOT APPLICABLE, SHOULD BE UNDER PATIENT CODE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE (B)(4) BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676333&pc=MTA | 2023826-2016-00712 | 2015-11-27 | Injury | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5480832&pc=MTA | 2023826-2016-00191 | 2015-11-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 -7.50/+2.50/X118 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. ON (B)(6) 20105, EXCESSIVE VAULTING WAS NOTED WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. ON (B)(6) 2015, THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR. REPORT #2023826-2016-00192 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -7.50/+2.50/X118, (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5481076&pc=MTA | 2023826-2016-00192 | 2015-11-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 -4.0/+4.00/X031 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. ON (B)(6) 20105, EXCESSIVE VAULTING WAS NOTED WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. ON (B)(6) 2015, THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR. REPORT #2023826-2016-00191 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -4.0/+4.0/X031. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5505898&pc=MTA | 2023826-2016-00268 | 2015-11-30 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +08.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE VAULTING IS GOOD AND THE PATIENT IS HAPPY. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES (UNABLE TO EVALUATE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PROCESS EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628372&pc=MTA | 2023826-2016-00551 | 2015-12-01 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.5 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATIONS - LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686817&pc=MTA | 2023826-2016-00186 | 2015-12-01 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -8 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SEE MFR REPORT #2023826-2016-00187 FOR LEFT EYE (OS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5411886&pc=MTA | 2023826-2016-00087 | 2015-12-02 | Injury | STAAR SURGICAL COMPANY | 2016-02-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULTING WITHOUT LENS ROTATION WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. PATIENT'S LAST OFFICE VISIT ON (B)(6) 2015, UCVA :20/20. Manufacturer Narrative: DIOPTER:-08.50. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5681975&pc=MTA | 2023826-2016-00571 | 2015-12-03 | Injury | STAAR SURGICAL COMPANY | 2016-05-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -8.00 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST- CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873690&pc=MTA | 2023826-2016-00986 | 2015-12-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5446118&pc=MTA | 2023826-2016-00171 | 2015-12-04 | Injury | STAAR SURGICAL COMPANY | 2016-02-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/130 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577550&pc=MTA | 2023826-2016-00448 | 2015-12-04 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM TICM130V4 -14.00/+3.50/X106 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. LOW VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED WITH A DIFFFERENT SIZE LENS ON (B)(6) 2016. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. CONCLUSION - AS PER DFU FOR THIS LENS MODEL, TICLS ARE CONTRAINDICATED FOR PATIENTS OVER 45 YEARS OLD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/+3.50/X106. DEVICE NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575868&pc=MTA | 2023826-2016-00462 | 2015-12-05 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.80MM AT THE TIME OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579298&pc=MTA | 2023826-2016-00473 | 2015-12-06 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE WITH SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: DIOPTER: -10.0. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593294&pc=MTA | 2023826-2016-00483 | 2015-12-09 | Injury | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -1.5/+5.5/096 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: PREVIOUS REPORT SOURCE, (B)(6), WAS INCORRECT. CORRECT REPORT SOURCE IS (B)(6). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5505910&pc=MTA | 2023826-2016-00270 | 2015-12-11 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577775&pc=MTA | 2023826-2016-00445 | 2015-12-11 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -08.50/+2.0/X080 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WITH SHALLOW CHAMBER WAS NOTED. THE ICL WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -08.50/+2.00/X080. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628371&pc=MTA | 2023826-2016-00556 | 2015-12-11 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536813&pc=MTA | 2023826-2016-00312 | 2015-12-12 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 -8.50 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 IN TO THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS LOW. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. SEE MFR REPORT #2023826-2016-00311 FOR RIGHT EYE (OD). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -8.50. EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A WORK ORDER SEARCH FOUND TWO SIMILAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536914&pc=MTA | 2023826-2016-00311 | 2015-12-12 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 -8.50 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS LOW. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. SEE MFR REPORT #2023826-2016-00312 FOR LEFT EYE (OS). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.98MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -8.50. (B)(4). METHOD CODE (EVALUATION METHOD): WORK ORDER SEARCH RESULTS CODE (EVALUATION RESULT): A WORK ORDER SEARCH FOUND TWO SIMILAR COMPLAINTS. CONCLUSION CODE (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5429379&pc=MTA | 2023826-2016-00124 | 2015-12-13 | Injury | STAAR SURGICAL COMPANY | 2016-02-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5506350&pc=MTA | 2023826-2016-00216 | 2015-12-13 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 -08.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. ON (B)(6) 2015, EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. SEE MFR. # 2023826-2016-00215 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -08.00 DEVICE EVALUATED BY MANUFACTURER? NO - DEVICE NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WAS ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830684&pc=MTA | 2023826-2016-00574 | 2015-12-15 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/100 DIOPTER, IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5495680&pc=MTA | 2023826-2016-00259 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/034 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ANTERIOR CHAMBER IS DEEP AND THE PATIENT IS HAPPY. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575693&pc=MTA | 2023826-2016-00399 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575694&pc=MTA | 2023826-2016-00400 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/SPLIT AND HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628357&pc=MTA | 2023826-2016-00554 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -9.50 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON "(B)(6) 2016" AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT:UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628369&pc=MTA | 2023826-2016-00555 | 2015-12-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -10.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.86MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579681&pc=MTA | 2023826-2016-00449 | 2015-12-21 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. LOW VAULT WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016 FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, BCVA WAS 20/20. SEE MFR #2023826-2016-00452 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOT. (B)(4). Manufacturer Narrative: DIOPTER: -10.00. (B)(4). DEVICE HAS NOT BEEN RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5580036&pc=MTA | 2023826-2016-00452 | 2015-12-21 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. LOW VAULT WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016 FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. SEE MFR #2023826-2016-00449 FOR RIGHT EYE. Manufacturer Narrative: THE LENS IS NOT MARKETED IN THE U.S. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC AND HAPTIC WERE TORN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DIOPTER: -10.00. (B)(4). DEVICE HAS NOT BEEN RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536785&pc=MTA | 2023826-2016-00317 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 -10.50 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 IN TO THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS LOW. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2016 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.50. (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536790&pc=MTA | 2023826-2016-00316 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 -8.50 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 INTO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS LOW. THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2016 AND THIS RESOLVED THE PROBLEM. SEE MFR REPORT #2023826-2016-00317 FOR LEFT EYE (OS). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -8.50. PT WEIGHT: UNK. (B)(4). METHOD: (EVALUATION METHOD): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540229&pc=MTA | 2023826-2016-00365 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 -09.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. LOW VAULT WAS NOTED. THE LENS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -09.50. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638603&pc=MTA | 2023826-2016-00597 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/X119 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686816&pc=MTA | 2023826-2016-00187 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -8 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN BCVA WAS 20/20 AND UCVA WAS 20/20 AS OF (B)(6) 2015. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: PATIENT CODE: (B)(4) (NEW INCISION) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THE DESCRIPTION INCORRECTED STATED THAT EXCESSIVE VAULT WAS REPORTED. LOW VAULT WAS REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SEE MFR REPORT #2023826-2016-00186 FOR RIGHT EYE (OD).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688978&pc=MTA | 2023826-2016-00185 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -10.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A SMALLER LENS ON 01/07/2016 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/20 ON (B)(6) 2016. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. SEE MFG. REPORT #2023826-2016-00184 FOR LEFT EYE (OS). (B)(4). LENS NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688980&pc=MTA | 2023826-2016-00184 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -10.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS REPORTED, THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/20 ON (B)(6) 2016. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT:UNK. (B)(4). SEE MFG. REPORT #2023826-2016-00185 FOR RIGHT EYE (OD). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5958044&pc=MTA | 2023826-2016-01273 | 2015-12-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0 MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING A TORN/SPLIT OPTIC, HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE (IMPOSSIBLE TO IDENTIFY). (B)(4). ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694841&pc=MTA | 2023826-2016-00251 | 2015-12-24 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5878874&pc=MTA | 2023826-2016-01068 | 2015-12-25 | Injury | STAAR SURGICAL COMPANY | 2016-08-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -12.00/1.0/075 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5827714&pc=MTA | 2023826-2016-00503 | 2015-12-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4) Manufacturer Narrative: CORRECT IMPLANTED DATE IS (B)(6) 2015. DATE RECEIVED BY MFR: 03/11/2016 IS NOT THE CORRECT DATE THAT THE PRODUCT EVALUATION WAS RECEIVED. THE CORRECT DATE IS 05/09/2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602194&pc=MTA | 2023826-2016-00521 | 2015-12-29 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676346&pc=MTA | 2023826-2016-00711 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.0/+1.5/125 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830651&pc=MTA | 2023826-2016-00577 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/024 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.9MM AT THE TIME OF IMPLANTATION. CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628412&pc=MTA | 2023826-2016-00553 | 2015-12-31 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WAS REPORTED, THE LENS WAS REPLACED WITH A LARGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTION: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668389&pc=MTA | 2023826-2016-00723 | 2016-01-01 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. LOW VAULT WITH NO ROTATION WAS NOTED ON (B)(6) 2016. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/30. Manufacturer Narrative: CONCLUSION CODE: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ACD OF 2.89MM IS BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919708&pc=MTA | 2023826-2016-01194 | 2016-01-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, BUT FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919715&pc=MTA | 2023826-2016-01195 | 2016-01-07 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602191&pc=MTA | 2023826-2016-00520 | 2016-01-09 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. A LONGER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EVENT DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540154&pc=MTA | 2023826-2016-00363 | 2016-01-11 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 -14.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016 AND EXCESSIVE VAULTING WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016; UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540214&pc=MTA | 2023826-2016-00364 | 2016-01-11 | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 -13.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016 AND EXCESSIVE VAULTING WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016; UCVA WAS 20/20. SEE MFR. #2023826-2016-00363 FOR RIGHT EYE. Manufacturer Narrative: ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50. DEVICE EVALUATED BY MANUFACTURER: NO, DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676394&pc=MTA | 2023826-2016-00715 | 2016-01-11 | Injury | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+4.0/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT AND THAT THERE WAS DRIED, DISCOLORED MATERIAL ON THE LENS SURFACE. CORRECTED DATA: CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.97MM. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830418&pc=MTA | 2023826-2016-00616 | 2016-01-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536868&pc=MTA | 2023826-2016-00315 | 2016-01-12 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 -9.0 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2016 IN TO THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND THE CHAMBER WAS SHALLOW. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2016 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION, METHOD: WORK ORDER SEARCH. EVALUATION, RESULT: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: UNABLE TO CONFIRM COMPLAINT - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5565079&pc=MTA | 2023826-2016-00394 | 2016-01-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -12.00/+2.00/X049 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED ON (B)(6) 2016. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.00/2.00/X049. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575866&pc=MTA | 2023826-2016-00459 | 2016-01-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VICMOS ARE CONTRAINDICATED FOR PATIENTS UNDER AGE OF 21 YEARS OLD. DEVICE EVALUATION-LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT, AND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5621835&pc=MTA | 2023826-2016-00533 | 2016-01-14 | Injury | STAAR SURGICAL COMPANY | 2016-04-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50/-0.5/075 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN, OPTIC TORN/SPLIT AND PRESENCE OF FOREIGN MATERIAL (SPOT/BLOCK) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5621848&pc=MTA | 2023826-2016-00534 | 2016-01-14 | Injury | STAAR SURGICAL COMPANY | 2016-04-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00/+1.0/080 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536849&pc=MTA | 2023826-2016-00309 | 2016-01-15 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.5/+2.5/90 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2016 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2016 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.5/+2.5/90 (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596434&pc=MTA | 2023826-2016-00488 | 2016-01-15 | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919731&pc=MTA | 2023826-2016-01197 | 2016-01-15 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN AND HAVING FOREIGN MATERIAL ON LENS SURFACE (NOT POSSIBLE TO SPECIFY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5648608&pc=MTA | 2023826-2016-00667 | 2016-01-17 | Injury | STAAR SURGICAL COMPANY | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+4.0/077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC AND THE OPTIC WERE BOTH BENT. CORRECTED DATA: CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.82MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540046&pc=MTA | 2023826-2016-00367 | 2016-01-19 | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 -7.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. LOW VAULT WAS NOTED. THE LENS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT, BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -07.50. METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575686&pc=MTA | 2023826-2016-00434 | 2016-01-19 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575869&pc=MTA | 2023826-2016-00435 | 2016-01-19 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533732&pc=MTA | 2023826-2016-00350 | 2016-01-19 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5900788&pc=MTA | 2023826-2016-00620 | 2016-01-22 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15/+2/090 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND UNREACTIVE PUPIL. THE PATIENT'S INCISION WAS ENLARGED AND THE ANTERIOR CHAMBER WAS IRRIGATED OF REMAINING VISCO FLUIDS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525105&pc=MTA | 2023826-2016-00335 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO 12.1 -7.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. LOW VAULT WAS NOTED. THE LENS WAS EXPLANTED ON 01/30/2016 AND EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, PATIENT'S LAST VISIT; BCVA WAS 20/25. SEE MFR. #2023826-2016-00334 FOR LEFT EYE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -7.50. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525106&pc=MTA | 2023826-2016-00334 | 2016-01-24 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO 12.1 -9.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. LOW VAULT WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, PATIENT'S LAST VISIT; BCVA WAS 20/25. SEE MFR. #2023826-2016-00335 FOR RIGHT EYE. Manufacturer Narrative: CORRECTED DATA: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY ON THE PRODUCT IN THE LENS CASE/VIAL. THERE WAS ALSO CLEAR RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -9.00. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831903&pc=MTA | 2023826-2016-00929 | 2016-01-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). CLAIM # (B)(4). NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575697&pc=MTA | 2023826-2016-00460 | 2016-01-25 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.8MM. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579678&pc=MTA | 2023826-2016-00471 | 2016-01-25 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS ON (B)(6) 2016 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS LOW. THE LENS WAS EXCHANGED FOR A LARGER LENS THE SAME DAY, BUT THE VAULT DID NOT IMPROVE AND A FOLLOW UP IS NEEDED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PREVIOUS DEVICE CODE WAS INCORRECT; (SECONDARY SURGERY) WAS A PATIENT CODE INSTEAD OF A DEVICE CODE. (B)(4). Manufacturer Narrative: DIOPTER: -11.0. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525072&pc=MTA | 2023826-2016-00339 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -1.50/+2.5X095 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, PATIENT'S LAST VISIT; UCVA WAS 20/20. Manufacturer Narrative: CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0 MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0 MM. CORRECTED DATA: (B)(4). CONCLUSION: CONCLUSION WAS FOUND AND UNABLE TO CONFIRM COMPLAINT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -1.50/+2.50/X095. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540226&pc=MTA | 2023826-2016-00361 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 -15.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. LOW VAULT WAS NOTED ON (B)(6) 2016. THE LENS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/40. Manufacturer Narrative: LENS WAS RETURNED DRY, IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -15.50 (B)(4). METHOD CODE(S): LENS WORK ORDER SEARCH. RESULTS CODE(S): A LENS WORK ORDER SEARCH REVEALED THERE WAS ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODES: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5903748&pc=MTA | 2023826-2016-01139 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY | 2016-08-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5 MM TICM125V4 IMPLANTABLE COLLAMER LENS, -20.0/+3.0/79 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919723&pc=MTA | 2023826-2016-01198 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831914&pc=MTA | 2023826-2016-00934 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.00/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL PRESENT ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919725&pc=MTA | 2023826-2016-01200 | 2016-01-28 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (DRIED DISCOLORED MATERIAL) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS MODEL IS CONTRAINDICATED IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM AND IN A PATIENT UNDER THE AGE OF 21. THE PATIENT WAS (B)(6) (UNDER 21) AND HAD AN ACD OF 2.80MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525100&pc=MTA | 2023826-2016-00338 | 2016-01-28 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -08.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2014. LOW VAULT WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, PATIENT'S LAST VISIT; UCVA WAS 20/20. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, AND THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: DIOPTER: -8.50, (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5464649&pc=MTA | 2023826-2016-00201 | 2016-01-29 | Injury | STAAR SURGICAL COMPANY | 2016-02-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+0.5/135 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION CODE: (B)(4)-PER DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIAL TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD LESS THAN 3.0MM AT THE TIME OF IMPLANTATION. PER DFU FOR THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED FOR PATIENTS ABOVE THE AGE OF 45 YEARS OLD. THIS PATIENT WAS OVER 45 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579728&pc=MTA | 2023826-2016-00479 | 2016-01-29 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/ SPLIT, HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE - SPOT, HAIR. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955120&pc=MTA | 2023826-2016-01268 | 2016-01-29 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHTT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (IMPOSSIBLE TO SPECIFY) ON LENS SURFACE. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5720695&pc=MTA | 2023826-2016-00614 | 2016-01-31 | Injury | STAAR SURGICAL COMPANY | 2016-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN WITH HAIR ON LENS SURFACE AND A BLACK SPOT ON OPTIC. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5720698&pc=MTA | 2023826-2016-00613 | 2016-01-31 | Injury | STAAR SURGICAL COMPANY | 2016-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN AND UNSPECIFIED FOREIGN MATERIAL ON LENS SURFACE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525102&pc=MTA | 2023826-2016-00336 | 2016-02-01 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 -4.00/+4.00/X087 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. LOW VAULT WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, PATIENT'S LAST VISIT; UCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.83MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -4.00/+4.00/X087. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5542874&pc=MTA | 2023826-2016-00362 | 2016-02-01 | Injury | STAAR SURGICAL COMPANY | 2016-04-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -11.50/+1.00/X091 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016 AND EXCESSIVE VAULTING WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: DIOPTER: -11.50/+1.00/X091. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5621903&pc=MTA | 2023826-2016-00532 | 2016-02-01 | Injury | STAAR SURGICAL COMPANY | 2016-04-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638252&pc=MTA | 2023826-2016-00627 | 2016-02-02 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7, -12.00/+3.00/083 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/25. Manufacturer Narrative: CORRECTION TO DATE OF EVENT: (B)(6) 2016. CORRECTION TO DESCRIBE EVENT OR PROBLEM: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00/+3.00/083 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016, BCVA WAS 20/25. CORRECTION TO IMPLANT DATE: (B)(6) 2016. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON THE LENS SURFACE (SPOT). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593287&pc=MTA | 2023826-2016-00511 | 2016-02-03 | Injury | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND NARROW ANGLE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642751&pc=MTA | 2023826-2016-00599 | 2016-02-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: CONCLUSION - AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.89MM. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5672071&pc=MTA | 2023826-2016-00645 | 2016-02-05 | Injury | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+2.0/077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ANTERIOR ENDOTHELIUM CHAMBER DEPTH BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS HAIR ON HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572289&pc=MTA | 2023826-2016-00433 | 2016-02-09 | Injury | STAAR SURGICAL COMPANY | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5664128&pc=MTA | 2023826-2016-00695 | 2016-02-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS BUT THERE WAS FOREIGN MATERIAL ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577767&pc=MTA | 2023826-2016-00446 | 2016-02-10 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -16.00/+3.50/X022 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE ICL WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. PATIENT CONDITION ON LAST VISIT, STERILE UVEITIS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: DIOPTER: -16.00/+3.50/X022. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638263&pc=MTA | 2023826-2016-00629 | 2016-02-12 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO 13.7, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS WITH SIMILAR DIOPTER SIZE ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/20. SEE MFR #2023826-2016-00628 FOR RIGHT EYE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. FOREIGN MATERIAL (UNABLE TO SPECIFY) WAS FOUND ON THE LENS SURFACE. (B)(4). CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT WAS (B)(6) YEARS OLD AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638258&pc=MTA | 2023826-2016-00628 | 2016-02-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO 13.7, -11.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS WITH SIMILAR DIOPTER SIZE ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/20. SEE MFR #2023826-2016-00629 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. FOREIGN MATERIAL (UNABLE TO SPECIFY) WAS FOUND ON THE LENS SURFACE. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT WAS 20 YEARS OLD AT THE TIME OF IMPLANTATION. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5487763&pc=MTA | 2023826-2016-00237 | 2016-02-16 | Injury | STAAR SURGICAL COMPANY | 2016-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND NARROW ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2016. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED WHEN THE LENS WAS EXCHANGED THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW, AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC FOLDED OVER. A PIECE OF THE OTHER HAPTIC WAS TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WERE PARTICLES STUCK ON THE LENS SURFACE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULT: (OPERATIONAL PROBLEM): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5488111&pc=MTA | 2023826-2016-00238 | 2016-02-16 | Injury | STAAR SURGICAL COMPANY | 2016-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND NARROW ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2016. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED WHEN THE LENS WAS EXCHANGED THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW, AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. EVALUATION RESULT) A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WERE PARTICLES STUCK ON THE LENS SURFACE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638614&pc=MTA | 2023826-2016-00596 | 2016-02-16 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +2.0/+4.5/X095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE. CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. USER FACILITY IS THE INCORRECT REPORT SOURCE, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5472399&pc=MTA | 2023826-2016-00222 | 2016-02-17 | Injury | STAAR SURGICAL COMPANY | 2016-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +6.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT AGE AND WEIGHT: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5621632&pc=MTA | 2023826-2016-00539 | 2016-02-19 | Injury | STAAR SURGICAL COMPANY | 2016-04-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -6.50/+1.50/X133 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. LOW VAULTING WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS (B)(6) 2016. THIS RESOLVED THE PROBLEM. THE PATIENT'S LAST VISIT ON (B)(6) 2016, BCVA:20/25. THE ICL WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: CORRECTED DATA: CORRECT DIOPTER OF THE LENS IS -6.50/+1.50/X113. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE WAS DISCARDED BY THE FACILITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5621841&pc=MTA | 2023826-2016-00531 | 2016-02-22 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DATE OF BIRTH - UNK. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5890954&pc=MTA | 2023826-2016-00618 | 2016-02-24 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938269&pc=MTA | 2023826-2016-01246 | 2016-02-24 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON EVACUATED REMAINING VISCO/FLUIDS FROM THE ANTERIOR CHAMBER. AN INTRAOCULAR INJECTION (MEDICATION) WAS ADMINISTERED. AN ADDITIONAL SUTURE WAS REQUIRED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (DRIED, DISCOLORED MATERIAL). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540148&pc=MTA | 2023826-2016-00377 | 2016-02-25 | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5/+3.5/098 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS HAD AN EXCESSIVE VAULT AND THE PUPIL WAS UNREACTIVE (FIXED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED. THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572300&pc=MTA | 2023826-2016-00326 | 2016-02-27 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5673301&pc=MTA | 2023826-2016-00636 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724404&pc=MTA | 2023826-2016-00789 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -9.00/+1.0/64 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5691581&pc=MTA | 2023826-2016-00783 | 2016-03-01 | Injury | STAAR SURGICAL COMPANY | 2016-06-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891352&pc=MTA | 2023826-2016-00677 | 2016-03-02 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5884618&pc=MTA | 2023826-2016-00619 | 2016-03-03 | Injury | STAAR SURGICAL COMPANY | 2016-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6337664&pc=MTA | 2023826-2016-01294 | 2016-03-04 | Injury | STAAR SURGICAL COMPANY | 2017-02-16 | MTA | ICL (INTRAOCULAR COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO 12.6 IMPLANTABLE COLLAMER LENS -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). REPORT RESUBMITTED PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549238&pc=MTA | 2023826-2016-00369 | 2016-03-05 | Injury | STAAR SURGICAL COMPANY | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -14.00/1.50/X171 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES. ICL REMAINS IMPLANTED. SURGERY IS PENDING TO EXPLANT AND REPLACE WITH ANOTHER ICL. Manufacturer Narrative: PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: DIOPTER: -14.00/1.50/X171. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596423&pc=MTA | 2023826-2016-00499 | 2016-03-07 | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -08.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. LOW VAULT WITH NO ROTATION WAS OBSERVED. THE ICL WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS WITH A SIMILAR DIOPTER. LAST VISIT UCVA: 20/20. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926673&pc=MTA | 2023826-2016-00678 | 2016-03-07 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5797038&pc=MTA | 2023826-2016-00920 | 2016-03-10 | Injury | STAAR SURGICAL COMPANY | 2016-07-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, BUT FOUND DRIED DISCOLORED MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5799978&pc=MTA | 2023826-2016-00926 | 2016-03-12 | Injury | STAAR SURGICAL COMPANY | 2016-07-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/098 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/ SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (SPOT) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5649012&pc=MTA | 2023826-2016-00600 | 2016-03-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5799970&pc=MTA | 2023826-2016-00925 | 2016-03-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE - SPOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5720693&pc=MTA | 2023826-2016-00612 | 2016-03-14 | Injury | STAAR SURGICAL COMPANY | 2016-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND LIQUID AND HAIR ON LENS SURFACE CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5890953&pc=MTA | 2023826-2016-00622 | 2016-03-14 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THERE WAS NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891374&pc=MTA | 2023826-2016-00676 | 2016-03-14 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). CONCLUSION CODE: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ANTERIOR ENDOTHELIUM CHAMBER DEPTH BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUEDEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676420&pc=MTA | 2023826-2016-00716 | 2016-03-15 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -19.0/+1.0/012 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) IS 20/20. Manufacturer Narrative: CORRECTED DATA: AGE AT TIME OF EVENT: 37 TO 40 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. A RESIDUE WAS PRESENT ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5629473&pc=MTA | 2023826-2016-00540 | 2016-03-16 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -3.00/+3.50/X129 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH A KERATOCONUS EYE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5890928&pc=MTA | 2023826-2016-00621 | 2016-03-16 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917486&pc=MTA | 2023826-2016-01179 | 2016-03-16 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5 MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN, AND UNSPECIFIED FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: WEIGHT: UNK. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668561&pc=MTA | 2023826-2016-00598 | 2016-03-17 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN AND HAVE FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: THREE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676358&pc=MTA | 2023826-2016-00736 | 2016-03-17 | Injury | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. CORRECTED DATA: CONCLUSION - AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.88MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642723&pc=MTA | 2023826-2016-00647 | 2016-03-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -18.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. OFFICE VISIT ON (B)(6) 2016, UCVA: 20/50. SEE MFR. #2023826-2016-00648 FOR LEFT EYE. Manufacturer Narrative: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS 2.82MM WHICH IS BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS NOT POSSIBLE TO SPECIFY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642724&pc=MTA | 2023826-2016-00648 | 2016-03-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -18.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. OFFICE VISIT ON (B)(6) 2016, UCVA: 20/50. SEE MFR. #2023826-2016-0064 FOR RIGHT EYE. Manufacturer Narrative: CORRECTED DATA: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ANTERIOR ENDOTHELIUM CHAMBER DEPTH OF 2.931MM IS BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CONTAINER. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOTS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593285&pc=MTA | 2023826-2016-00510 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK) AND SHALLOW ANTERIOR CHAMBER. A YAG WAS PERFORMED ON THE PERIPHERAL IRIDECTOMIES AND PATIENT WAS GIVEN MEDICATION BUT IT WAS UNSUCCESSFUL, THE LENS WAS EXPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5651122&pc=MTA | 2023826-2016-00662 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/072 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: DEVICE MANUFACTURE DATE 02/05/2016 IS INCORRECT, CORRECT DATA IS 01/20/2016. (B)(4) - CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH WAS BELOW 3.0MM FOR VICL/VTICL. ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, THE HAPTIC TORN/BROKEN/BENT/DEFORMED, AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5940525&pc=MTA | 2023826-2016-01251 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/098 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBERS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5940533&pc=MTA | 2023826-2016-01252 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT FIBERS WERE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655199&pc=MTA | 2023826-2016-00703 | 2016-03-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, WITH FOREIGN MATERIAL ON PRODUCT IN THE LENS CASE/VIAL. (B)(4) NEW INCISION IS NOT CORRECT AND SHOULD BE DELETED. A SECOND SURGERY WAS PERFORMED, NOT A NEW INCISION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5659939&pc=MTA | 2023826-2016-00666 | 2016-03-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE HAPTIC TORN. PREVIOUS PATIENT CODE (B)(4) (INCISION OPENED) IS NOT CORRECT. SECONDARY SURGERY WAS ALREADY REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5659940&pc=MTA | 2023826-2016-00665 | 2016-03-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ON (B)(6) 2016 IT WAS CONFIRMED THAT CUSTOMER DISCARDED EXPLANTED LENS. LENS UNABLE TO BE EVALUATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5673334&pc=MTA | 2023826-2016-00634 | 2016-03-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE, BUT IS NOT POSSIBLE TO BE SPECIFIED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5673188&pc=MTA | 2023826-2016-00742 | 2016-03-23 | Injury | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND SCRATCHES ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. AGE - UNK. WEIGHT: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5673220&pc=MTA | 2023826-2016-00646 | 2016-03-23 | Injury | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +5.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID, WITH DRIED, DISCOLORED MATERIAL ON THE LENS SURFACE. PRODUCT WAS RETURNED IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE HAPTIC BENT. CORRECTED DATA: PREVIOUS PATIENT CODE, "INCISION OPENED" IS NOT CORRECT. LENS WAS EXPLANTED AND EXCHANGED IN A SECOND SURGERY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. AGE - UNK, DATE OF BIRTH - UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817531&pc=MTA | 2023826-2016-00917 | 2016-03-23 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/073 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20.Event Description: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE (B)(4). Manufacturer Narrative: UCVA, INSTEAD OF BCVA, WAS REPORTED. PREVIOUS DEVICE CODE (B)(4) WAS INCORRECT; (B)(4) (SECONDARY SURGERY) WAS A PATIENT CODE INSTEAD OF A DEVICE CODE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5810831&pc=MTA | 2023826-2016-00943 | 2016-03-29 | Injury | STAAR SURGICAL COMPANY | 2016-07-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/1.0. Manufacturer Narrative: PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS LIQUID. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5810849&pc=MTA | 2023826-2016-00944 | 2016-03-29 | Injury | STAAR SURGICAL COMPANY | 2016-07-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/1.2. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/ SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED, FOREIGN MATERIAL (LIQUID) AND A SPOT ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632874&pc=MTA | 2023826-2016-00562 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO INCORRECT SIZE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PATIENTS CORRECT INITIALS (B)(6). AGE (B)(6). DATE OF BIRTH IS (B)(6). SEX: FEMALE. DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND SCRATCHES ON ONE HAPTIC. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655217&pc=MTA | 2023826-2016-00705 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/083 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WHO IS LESS THAN 21 YEARS OF AGE IS CONTRAINDICATED. THIS PATIENT IS (B)(6). CORRECTED INFORMATION: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5696628&pc=MTA | 2023826-2016-00673 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL PRESENT ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5696632&pc=MTA | 2023826-2016-00674 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND PRESENCE OF FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5756525&pc=MTA | 2023826-2016-00791 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -8.00 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATIONS: THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN/BENT AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (HAIR). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CONCLUSION COD: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.93MM AT THE TIME OF IMPLANTATION. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH BEST CORRECTED VISUAL ACUITY (BCVA) 20/20. RESULT CODE: PREVIOUS CODE NOT APPLICABLE, DIMENSIONAL INSPECTION FOUND LENS MEASUREMENTS TO BE WITHIN SPECIFICATIONS. CONCLUSION CODE: PREVIOUS CODE NOT APPLICABLE, DIMENSIONAL INSPECTION FOUND LENS MEASUREMENTS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5756722&pc=MTA | 2023826-2016-00790 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -10.5/+0.5/062 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE PATIENTS POST OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: DIMENSIONAL INSPECTION PERFORMED FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926678&pc=MTA | 2023826-2016-00680 | 2016-03-31 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, BUT THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED HAIR. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867095&pc=MTA | 2023826-2016-01017 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5884733&pc=MTA | 2023826-2016-01121 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/89 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT' POST-OP UNCORRECTED VISUAL ACUITY IS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED INFORMATION: "USER FACILITY" TO "DISTRIBUTOR". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655368&pc=MTA | 2023826-2016-00721 | 2016-04-03 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -07.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS OBSERVED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SMALLER LENGTH LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. LAST VISIT ON (B)(6) 2016, UCVA: 20/32. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH-NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE LENS PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5662053&pc=MTA | 2023826-2016-00719 | 2016-04-04 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/+1.50/X079 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/20. Manufacturer Narrative: (B)(4). ADDITIONAL DATA: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIAL TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD BELOW 3.0MM AT THE TIME OF IMPLANTATION. PER DFU FOR THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. THIS PATIENT WAS BELOW AGE 21 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH SOME LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE ON IT. THE LIQUID IN VIAL WAS THE SALINE SOLUTION DRYING UP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5918359&pc=MTA | 2023826-2016-01183 | 2016-04-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/87 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL WAS OBSERVED TO BE ON THE LENS (DRIED MATERIAL). UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655206&pc=MTA | 2023826-2016-00704 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+0.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938270&pc=MTA | 2023826-2016-01245 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+4.0/120 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DRIED, DISCOLORED MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688631&pc=MTA | 2023826-2016-00751 | 2016-04-12 | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: (B)(4). ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5741917&pc=MTA | 2023826-2016-00841 | 2016-04-12 | Injury | STAAR SURGICAL COMPANY | 2016-06-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.50/+2.5/123 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED. FOREIGN MATERIAL IN THE FORM OF SPOTS AND FIBERS WERE PRESENT ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5752003&pc=MTA | 2023826-2016-00795 | 2016-04-12 | Injury | STAAR SURGICAL COMPANY | 2016-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE AND HAPTIC SCRATCH. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5788409&pc=MTA | 2023826-2016-00899 | 2016-04-12 | Injury | STAAR SURGICAL COMPANY | 2016-07-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL - SPOT PRESENT ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655329&pc=MTA | 2023826-2016-00706 | 2016-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THAT THE LENS WAS WITHIN SPECIFICATIONS. METHOD CODE: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5659955&pc=MTA | 2023826-2016-00693 | 2016-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/20. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL AND THAT LENS WAS RETURNED BENT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5662096&pc=MTA | 2023826-2016-00718 | 2016-04-14 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -10.50/+2.50/X083 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/16. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE EXPLANT DATE, (B)(6) 2016, IS INCORRECT. THE CORRECT EXPLANT IS (B)(6) 2016. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND A PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5900800&pc=MTA | 2023826-2016-01138 | 2016-04-16 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO 12.1 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(6). (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND LENS OPTIC TORN/SPLIT, LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE NOT POSSIBLE TO SPECIFY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688639&pc=MTA | 2023826-2016-00752 | 2016-04-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688643&pc=MTA | 2023826-2016-00753 | 2016-04-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN ALONG WITH HAVING DRIED MATERIAL ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5884738&pc=MTA | 2023826-2016-00890 | 2016-04-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5685146&pc=MTA | 2023826-2016-00749 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: CONCLUSION: ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923612&pc=MTA | 2023826-2016-01209 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL PRESENT ON THE LENS SURFACE. CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5685207&pc=MTA | 2023826-2016-00754 | 2016-04-20 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT WAS (B)(6) YEARS OLD AT THE TIME OF THE EVENT. DEVICE CODE: (EXPLANTED) - THIS CODE IS NOT APPLICABLE UNDER DEVICE CODE; IT SHOULD BE LISTED UNDER PATIENT CODE INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5752027&pc=MTA | 2023826-2016-00794 | 2016-04-21 | Injury | STAAR SURGICAL COMPANY | 2016-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -17.00/3.5/85 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668547&pc=MTA | 2023826-2016-00710 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. LOW VAULT WITH NO ROTATION WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680342&pc=MTA | 2023826-2016-00746 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE LENS WAS TOO LONG. THE SURGEON NOTED THE EYE WAS BULGING AND THE IRIS WAS BEING PUSHED FORWARD, SO DECIDED TO REMOVE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS, WHICH RESOLVED THE PROBLEM. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: RESULT - (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870294&pc=MTA | 2023826-2016-01052 | 2016-04-25 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ACD BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5691561&pc=MTA | 2023826-2016-00782 | 2016-04-27 | Injury | STAAR SURGICAL COMPANY | 2016-06-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE LENS IS PLANNED TO EXCHANGE WITH ANOTHER LENS. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016. THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FIBER ON LENS SURFACE. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724446&pc=MTA | 2023826-2016-00785 | 2016-04-28 | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/72 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724461&pc=MTA | 2023826-2016-00784 | 2016-04-28 | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5903664&pc=MTA | 2023826-2016-01137 | 2016-04-30 | Injury | STAAR SURGICAL COMPANY | 2016-08-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SHALLOW CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND FOREIGN MATERIAL (HAIR) BUT NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4) DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724383&pc=MTA | 2023826-2016-00788 | 2016-05-04 | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/106 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923608&pc=MTA | 2023826-2016-01205 | 2016-05-07 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -17.5/+3.5/074 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SHALLOW CHAMBER. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION-LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN AND DEFORMED, AND FIBERS ON LENS SURFACES. DIMENSIONAL INSPECTION CONCLUDE THAT THE LENS IS IN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902495&pc=MTA | 2023826-2016-01147 | 2016-05-21 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.5/+2.0/175 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: THE REPORTER INDICATED THAT THE LENS HAS SINCE BEEN EXCHANGED. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH PREEXISTING KERATOCONUS EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919746&pc=MTA | 2023826-2016-01204 | 2016-05-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED.. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: AGE AT TIME OF EVENT: PREVIOUS AGE (B)(6) TO CORRECTED AGE (B)(6) - PREVIOUS AGE IS INCORRECT. ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 ON (B)(6) 2016. (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE OPTIC WAS TORN/SPLIT AND THE HAPTIC WAS TORN/BROKEN. THERE WAS FOREIGN MATERIAL ON THE LENS SURFACE WHICH WAS UNABLE TO BE SPECIFIED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873739&pc=MTA | 2023826-2016-01074 | 2016-05-24 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. (B)(4). Manufacturer Narrative: DATE OF EVENT: UNK TO (B)(6) 2016. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(6). PREVIOUS AGE IS INCORRECT. DATE OF EVENT: (B)(6) 2015 TO UNK - PREVIOUS DATE OF EVENT IS INCORRECT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5736184&pc=MTA | 2023826-2016-00832 | 2016-05-25 | Injury | STAAR SURGICAL COMPANY | 2016-06-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO OVERSIZED (HIGH) VAULT. THE LENS WAS EXCHANGED FOR A 13.2MM ICL WITH THE SAME DIOPTER. Manufacturer Narrative: WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938118&pc=MTA | 2023826-2016-01237 | 2016-05-27 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -2.25 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5957333&pc=MTA | 2023826-2016-01264 | 2016-05-27 | Injury | STAAR SURGICAL COMPANY | 2016-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+5.5/025 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND IRIS-ENDOTHELIUM TOUCH. THE SURGEON WILL PERFORM EXPLORATION OF RETROPUPILLARY SPACE TO TRY TO LIBERATE POSSIBLE POSTERIOUS PERIPHERIC SYNECHIAE. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5780931&pc=MTA | 2023826-2016-00884 | 2016-05-30 | Injury | STAAR SURGICAL COMPANY | 2016-07-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5783098&pc=MTA | 2023826-2016-00885 | 2016-05-30 | Injury | STAAR SURGICAL COMPANY | 2016-07-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870287&pc=MTA | 2023826-2016-01051 | 2016-05-30 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923618&pc=MTA | 2023826-2016-01220 | 2016-05-30 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) IN (B)(6)1999. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5850908&pc=MTA | 2023826-2016-01023 | 2016-05-31 | Injury | STAAR SURGICAL COMPANY | 2016-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A SHORTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5866938&pc=MTA | 2023826-2016-01012 | 2016-05-31 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM513.2 IMPLANTABLE COLLAMER LENS, -7.00 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION-LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5918178&pc=MTA | 2023826-2016-01180 | 2016-05-31 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO 12.1 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION (THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND UNSPECIFIED FOREIGN MATERIAL ON SURFACE AND HAPTIC TORN). (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5827728&pc=MTA | 2023826-2016-00946 | 2016-06-01 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN / BROKEN / BENT / DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870503&pc=MTA | 2023826-2016-01059 | 2016-06-01 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND WILL BE EXPLANTED. THE PUPIL WAS LARGE AND SLUGGISH. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SMALLER LENS. AFTER THIS SECONDARY SURGICAL INTERVENTION, PUPIL WAS MID-DILATED AND NON-REACTIVE. DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. PATIENT CODE - (B)(4) SECONDARY INTERVENTION, EXCHANGED LENS. MID-DILATED AND NON-REACTIVE. WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE CORRECT REPORT SOURCE COUNTRY IS THE (B)(4) . (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5877567&pc=MTA | 2023826-2016-01075 | 2016-06-01 | Injury | STAAR SURGICAL COMPANY | 2016-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+4.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/23. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.875MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5900906&pc=MTA | 2023826-2016-00893 | 2016-06-03 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -5.5/1/142 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAD NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5836860&pc=MTA | 2023826-2016-01004 | 2016-06-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/097 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867041&pc=MTA | 2023826-2016-01014 | 2016-06-05 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+3/112 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20. IMPLANTED (B)(6) 2016 . DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL/RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5960051&pc=MTA | 2023826-2016-01167 | 2016-06-05 | Injury | STAAR SURGICAL COMPANY | 2016-09-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+3/80 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870300&pc=MTA | 2023826-2016-01053 | 2016-06-06 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/097 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND EXPLANTED WITHIN THE SAME DAY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PREVIOUS CLAIM # (B)(4) WAS INCORRECT; THE CORRECT CLAIM # (B)(4). (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5751220&pc=MTA | 2023826-2016-00866 | 2016-06-07 | Injury | STAAR SURGICAL COMPANY | 2016-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND NOTED AN EXCESSIVE VAULT IMMEDIATELY AFTER THE SURGERY. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD ON THE SAME DAY WHEN THE PATIENT CAME BACK TO THE FACILITY. THE LENS WAS EXCHANGED FOR A 12.6MM ICL AND THE ADDITIONAL PERIPHERAL IRIDECTOMY (PI) WAS PERFORMED. THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS NOT PRODUCT-RELATED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5886814&pc=MTA | 2023826-2016-01126 | 2016-06-07 | Injury | STAAR SURGICAL COMPANY | 2016-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED WITH THE EXCHANGED LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5814342&pc=MTA | 2023826-2016-00874 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2016-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES IS REPORTED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5814350&pc=MTA | 2023826-2016-00873 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2016-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, THE HAPTIC TORN/BROKEN, AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS IN PATIENTS OVER THE AGE OF 45 YEARS OLD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873723&pc=MTA | 2023826-2016-01072 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). CONCLUSION CODE: 24-PER DFU FOR THIS LENS MODEL, VTCMOS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. THIS PATIENT WAS BELOW 21 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5947591&pc=MTA | 2023826-2016-01151 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2016-09-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +2.0/+2.5/85 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: CORRECTED DATA: USER SOURCE-PREVIOUSLY STATED USER FACILITY. SHOULD HAVE STATED DISTRIBUTOR INSTEAD. (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951518&pc=MTA | 2023826-2016-01256 | 2016-06-10 | Injury | STAAR SURGICAL COMPANY | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY IS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816255&pc=MTA | 2023826-2016-00966 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. Manufacturer Narrative: CONCLUSION: OFF-LABEL USE: PATIENT YOUNGER THAN AGE 21 AND ACD BELOW 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5835216&pc=MTA | 2023826-2016-01003 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923150&pc=MTA | 2023826-2016-01215 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS, -7.00/+1.0/095 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN AND BENT AND FOREIGN MATERIAL AND SCRATCHES ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS MODEL IS CONTRAINDICATED IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH(ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM. THE PATIENT HAD AN ACD OF 2.986MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873877&pc=MTA | 2023826-2016-01064 | 2016-06-15 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS, AN ADDITIONAL SUTURE WAS USED AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: REPORT SOURCE: "DISTRIBUTOR" REPLACES "USER FACILITY". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5905516&pc=MTA | 2023826-2016-01173 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/108 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16.5. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FIBER AND SMALL PARTICULATES ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909200&pc=MTA | 2023826-2016-01172 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16.5. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FIBER AND SMALL PARTICULATES ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831907&pc=MTA | 2023826-2016-00922 | 2016-06-22 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. PATIENT'S ACD IS 2.8 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5845333&pc=MTA | 2023826-2016-01020 | 2016-06-22 | Injury | STAAR SURGICAL COMPANY | 2016-08-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/86 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT DEVELOPED A REFRACTIVE SURPRISE AND THE LENS REMAINS IN THE PATIENTS EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938059&pc=MTA | 2023826-2016-01236 | 2016-06-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY IS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, OBSERVING NO VISIBLE DAMAGE TO THE LENS. CONCLUSION CODE: PER DFU, THIS LENS MODEL IS CONTRAINDICATED IN THE PATIENT'S AGE BEING UNDER 21 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938261&pc=MTA | 2023826-2016-01244 | 2016-06-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/051 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5835313&pc=MTA | 2023826-2016-00979 | 2016-06-23 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: (B)(4). CONCLUSION CODE: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ACD BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5877566&pc=MTA | 2023826-2016-01076 | 2016-06-23 | Injury | STAAR SURGICAL COMPANY | 2016-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831910&pc=MTA | 2023826-2016-00994 | 2016-06-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CONCLUSION: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.8 MM AT THE TIME OF IMPLANTATION. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5835224&pc=MTA | 2023826-2016-00951 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5884703&pc=MTA | 2023826-2016-01097 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM513.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5884721&pc=MTA | 2023826-2016-01098 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM513.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5903890&pc=MTA | 2023826-2016-01150 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-08-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. DEVICE EVALUATION FOUND THAT THE LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869125&pc=MTA | 2023826-2016-01038 | 2016-06-25 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT; VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926656&pc=MTA | 2023826-2016-01218 | 2016-06-25 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831905&pc=MTA | 2023826-2016-00947 | 2016-06-27 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENTS UNDER THE AGE OF 21 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). CLAIM # (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831924&pc=MTA | 2023826-2016-00949 | 2016-06-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5./X094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (HAIR) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834829&pc=MTA | 2023826-2016-00933 | 2016-06-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT; VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870296&pc=MTA | 2023826-2016-01054 | 2016-06-29 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED, BUT THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951519&pc=MTA | 2023826-2016-01259 | 2016-06-29 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.0/+4.5/068 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. THIS PATIENT WAS UNDER 21 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5971272&pc=MTA | 2023826-2016-01292 | 2016-07-01 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM 13.2 IMPLANTABLE COLLAMER LENS, -12.5/+4.5/110 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN THE LENS CASE/VIAL AND IN LIQUID. VISUAL INSPECTION FOUND THE LENS TO HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870077&pc=MTA | 2023826-2016-01013 | 2016-07-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/99 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL WAS OBSERVED TO BE ON THE LENS (SPOT). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED DUE TO THE AGE OF THE PATIENT BEING HIGHER THAN WHAT WAS INDICATED IN THE DFU. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5879317&pc=MTA | 2023826-2016-01067 | 2016-07-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7/+1/38 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN, AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4) SHOULD BE INCLUDED. SHOULD NOT BE LISTED AS DEVICE CODE. (B)(4) SHOULD BE INCLUDED IN PATIENT CODE, NOT DEVICE CODE. CONCLUSION CODE: LENS RETURNED AND EVALUATED, CAN'T CONFIRM COMPLAINT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867030&pc=MTA | 2023826-2016-01016 | 2016-07-05 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND PRESENCE OF FOREIGN MATERIAL (SPOT) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909289&pc=MTA | 2023826-2016-01175 | 2016-07-05 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870491&pc=MTA | 2023826-2016-01057 | 2016-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: ADDITIONAL DATA: THE LENS EXCHANGE RESOLVED THE PROBLEM. THE VAULT VALUE WENT FROM 90 MICROMETERS TO 350 MICROMETERS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870496&pc=MTA | 2023826-2016-01058 | 2016-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: THE VAULT VALUE WAS 130 UM. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH VAULT VALUE 270 UM. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.88MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902506&pc=MTA | 2023826-2016-01164 | 2016-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/081 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902509&pc=MTA | 2023826-2016-01165 | 2016-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/091 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5840815&pc=MTA | 2023826-2016-01008 | 2016-07-10 | Injury | STAAR SURGICAL COMPANY | 2016-08-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/24. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.95MM AT THE TIME OF IMPLANTATION. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5840816&pc=MTA | 2023826-2016-01007 | 2016-07-10 | Injury | STAAR SURGICAL COMPANY | 2016-08-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/24. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.86MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5886826&pc=MTA | 2023826-2016-01127 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2016-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON 07/11/2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VTICMOS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. THIS PATIENT WAS UNDER 21 AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926657&pc=MTA | 2023826-2016-01223 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873728&pc=MTA | 2023826-2016-01073 | 2016-07-13 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5905508&pc=MTA | 2023826-2016-01171 | 2016-07-14 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817545&pc=MTA | 2023826-2016-00928 | 2016-07-15 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/076 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN, OPTIC TORN/SPLIT AND PRESENCE OF FOREIGN MATERIAL (SPOT) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909072&pc=MTA | 2023826-2016-01177 | 2016-07-16 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5886866&pc=MTA | 2023826-2016-01128 | 2016-07-17 | Injury | STAAR SURGICAL COMPANY | 2016-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934951&pc=MTA | 2023826-2016-01242 | 2016-07-17 | Injury | STAAR SURGICAL COMPANY | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY AND BENT, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5898663&pc=MTA | 2023826-2016-01146 | 2016-07-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: THE PATIENT IS UNDER THE AGE OF 21 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955099&pc=MTA | 2023826-2016-01263 | 2016-07-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT AGE BELOW 21 OR OVER 45 YEARS. PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5850918&pc=MTA | 2023826-2016-01025 | 2016-07-20 | Injury | STAAR SURGICAL COMPANY | 2016-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN MATERIAL ON THE LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5898664&pc=MTA | 2023826-2016-01145 | 2016-07-21 | Injury | STAAR SURGICAL COMPANY | 2016-08-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DRIED MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902549&pc=MTA | 2023826-2016-01166 | 2016-07-21 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS), THE LENS MISSING A PIECE OF HAPTIC AND THE LENS WAS CAUGHT WITH VIAL CAP. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926650&pc=MTA | 2023826-2016-01216 | 2016-07-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE PRESENCE OF FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938106&pc=MTA | 2023826-2016-01235 | 2016-07-24 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE LENS. VISUAL INSPECTION FOUND THAT THE OPTIC WAS TORN AND THE HAPTIC WAS TORN AND BENT. FOREIGN MATERIAL (FIBERS) WAS PRESENT ON THE LENS SURFACE. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5850886&pc=MTA | 2023826-2016-01024 | 2016-07-26 | Injury | STAAR SURGICAL COMPANY | 2016-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5884722&pc=MTA | 2023826-2016-01123 | 2016-07-26 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE REPORTER INDICATED THE LENS WAS INSERTED AND REMOVED DURING THE SAME SURGERY. THERE WAS NO REPORT OF ANY PATIENT INJURY. THE PATIENT'S CURRENT CONDITION ON (B)(6) 2016 WAS 20/20 +2 UNCORRECTED WITH PHOTOPHOBIA. THE LENS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5931257&pc=MTA | 2023826-2016-01168 | 2016-07-26 | Injury | STAAR SURGICAL COMPANY | 2016-09-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENTS UNDER THE AGE OF 21 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5940518&pc=MTA | 2023826-2016-01250 | 2016-07-26 | Injury | STAAR SURGICAL COMPANY | 2016-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). DEVICE EVALUATION: VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS IS INCORRECT. CORRECT DATA IS VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS CLEAR RESIDUE. REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT AGE BELOW 21 OR OVER 45 YEARS, ANTERIOR ENDOTHELIUM CHAMBER DEPTH BELOW 3.0MM FOR VICL/VTICL. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5975342&pc=MTA | 2023826-2016-01315 | 2016-07-27 | Injury | STAAR SURGICAL COMPANY | 2016-09-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC TORN/BROKEN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869107&pc=MTA | 2023826-2016-01034 | 2016-07-28 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). CONCLUSION: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENTS UNDER THE AGE OF 21 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869115&pc=MTA | 2023826-2016-01035 | 2016-07-28 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE PRESENCE OF FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934948&pc=MTA | 2023826-2016-01243 | 2016-07-28 | Injury | STAAR SURGICAL COMPANY | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENTS UNDER THE AGE OF 21 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869130&pc=MTA | 2023826-2016-01037 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/084 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A EXCESSIVE VAULT AND EXPLANTED WITHIN THE SAME DAY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951511&pc=MTA | 2023826-2016-01254 | 2016-08-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/076 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. A LONGER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955107&pc=MTA | 2023826-2016-01257 | 2016-08-05 | Injury | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951513&pc=MTA | 2023826-2016-01255 | 2016-08-12 | Injury | STAAR SURGICAL COMPANY | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND HAPTIC DEFORMED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5535804&pc=MTA | 2023826-2016-00318 | 2016-10-07 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/+2.00/X098 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND SHALLOW CHAMBER WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015; UCVA WAS 20/13. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.00/+2.00/X098. (B)(4). A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6091600&pc=MTA | 2023826-2016-01604 | 2016-10-12 | Injury | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE OPTIC WAS TORN AND THE HAPTIC WAS TORN AND BENT. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. THERE WERE FIBERS ON THE LENS SURFACE. CORRECTED DATA: OFF-LABEL USE INDICATED AS PATIENT IS PREGNANT AND ACD IS 2.88MM, BOTH OF WHICH ARE CONTRAINDICATIONS TO THE USE OF THIS LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643989&pc=MTA | 2023826-2016-00586 | 2016-11-16 | Injury | STAAR SURGICAL COMPANY | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 -8.00 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS REPORTED, THE LENS WAS EXCHANGED WITH A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5810842&pc=MTA | 2023826-2016-00919 | 2016-11-23 | Injury | STAAR SURGICAL COMPANY | 2016-07-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PATIENT' POST-OP UNCORRECTED VISUAL ACUITY IS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4488993&pc=MTA | 2023826-2015-00114 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON 03/27/2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4509427&pc=MTA | 2023826-2015-00118 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE DEVICE WAS DISCARDED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED AND THE SURGERY WENT WELL. THAT NIGHT, THE PATIENT WAS HAVING PROBLEMS AND RETURNED TO SURGERY. THE LENS WAS EXPLANTED AND WAS EXCHANGED WITH A SHORTER LENS. THE PATIENT IS DOING WELL. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. INITIAL REPORTER: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). THE MEDICAL DEVICE WAS UNLIKELY THE CAUSE OF THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. CLAIM # 427509 (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4509436&pc=MTA | 2023826-2015-00117 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE (18MMHG), NARROWING OF THE ANGLE AND WAS EXCHANGED WITH A SHORTER LENS. Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. AS REPORTED ON THE MEDICAL REVIEW, THE MEDICAL DEVICE WAS UNLIKELY THE CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: (B)(4) - INTRAOCULAR PRESSURE RISE. (B)(4). Manufacturer Narrative: EVALUATION METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). THE MEDICAL DEVICE WAS UNLIKELY THE CAUSE OF THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578987&pc=MTA | 2023826-2015-00225 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE.: UNK (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629122&pc=MTA | 2023826-2015-00384 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAS AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT THE LENS AND EXCHANGE FOR A SHORTER LENS. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. Manufacturer Narrative: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4754468&pc=MTA | 2023826-2015-00506 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. PER MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4754658&pc=MTA | 2023826-2015-00508 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE ICL WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DATE OF EVENT: UNK. (B)(4). EVALUATION, METHOD - WORK ORDER SEARCH. RESULTS : A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE NOT AVAILABLE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4759147&pc=MTA | 2023826-2015-00520 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, THE LENS WAS REPLACED FOR A SHORTER LENS ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (EXPLANTED); (VAULTING). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE LENS SHOWED IT WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ACT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4816832&pc=MTA | 2023826-2015-00551 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS WAS TOO BIG AND THE PRESSURE SPIKED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - THE INTRAOCULAR PRESSURE SPIKE COULD BE THE CONSEQUENCE OF EXCESSIVE VAULTING. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4835809&pc=MTA | 2023826-2015-00651 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-10 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST BCVA WAS 20/20. SEE MFR REPORT #2023826-2015-00652 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. DEVICE EVALUATED BY THE MANUFACTURER? LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4836882&pc=MTA | 2023826-2015-00629 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855059&pc=MTA | 2023826-2015-00682 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4870160&pc=MTA | 2023826-2015-00745 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY INDICATED THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OPERATIVE BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT IS DOING WELL. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. THE LENS SIZING ERROR MOST LIKELY CAUSED THE EXCESSIVE LENS VAULTING. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - NARROW ANGLE AND SHALLOW ANTERIOR CHAMBER COULD BE CONSEQUENCES OF EXCESSIVE LENS VAULTING. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4870162&pc=MTA | 2023826-2015-00746 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY INDICATED THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OPERATIVE BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT IS DOING WELL. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. THE LENS SIZING ERROR MOST LIKELY CAUSED THE EXCESSIVE LENS VAULTING. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - NARROW ANGLE AND SHALLOW ANTERIOR CHAMBER COULD BE CONSEQUENCES OF EXCESSIVE LENS VAULTING. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4872373&pc=MTA | 2023826-2015-00743 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 AND EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4872425&pc=MTA | 2023826-2015-00744 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 AND EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). (OTHER): LENS NOT RETURNED. EVALUATION CODES: METHOD (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917239&pc=MTA | 2023826-2015-00833 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: (B)(4). NEW INCISION. SECONDARY SURGERY. EXPLANTED. VAULTING. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH PIECES TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF REDDISH RESIDUE ON THE LENS SURFACE. CONCLUSIONS: UNABLE TO CONFIRM COMPLAINT, BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). EVALUATION RESULT: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: EVALUATION RESULT: PER MEDICAL REVIEW - ACCURATE DETERMINATION OF ANTERIOR SEGMENT DIMENSIONS IS KEY TO MAINTENANCE OF ADEQUATE VAULT. THE METHOD OF ICL SIZING FOR THIS PATIENT WAS BASED ON ACD AND WTW RECOMMENDATIONS PER DFU. THE REPORT PROVIDES NO SPECIFIC INFORMATION ON THE DEGREE OF VAULTING, THE CAUSE OF EXCESSIVE VAULT, OR THE CLINICAL EVIDENCE FOR EXCESSIVE VAULTING. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917260&pc=MTA | 2023826-2015-00832 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS WAS TOO BIG. THE LENS WAS REPLACED WITH A SMALLER DIOPTER LENS. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS DRY, WITH DRIED SURGICAL RESIDUE ON THE LENS SURFACE. WAS UNABLE TO DETECT ANY DAMAGE TO THE LENS DUE TO THE RETURNED CONDITION. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS : BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. AS REPORTED ON THE MEDICAL REVIEW, THE OVER-SIZED LENS MOST LIKELY CAUSED THE EXCESSIVE VAULTING. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917262&pc=MTA | 2023826-2015-00831 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS WAS TOO BIG. THE LENS WAS REPLACED WITH A SMALLER DIOPTER LENS. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). NEW INCISION. SECONDARY SURGERY. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS STUCK IN THE LENS VIAL. THE LENS WAS DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. UNABLE TO DETECT ANY DAMAGE TO THE LENS DUE TO THE RETURNED CONDITION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION, WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. AS REPORTED ON THE MEDICAL REVIEW, THE OVER-SIZED LENS MOST LIKELY CAUSED THE EXCESSIVE VAULTING. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4920268&pc=MTA | 2023826-2015-00844 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2015. AT ONE DAY POST-OPERATIVE, THE PERIPHERAL IRIDOTOMIES WERE FOUND TO BE INSUFFICIENT. THE SURGEON PERFORMED A SECOND SURGERY TO OPEN THE IRIDOTOMIES AND THE IRIDOTOMIES WERE FULLY PATENT. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND HYPEROPIC OUTCOME. THE PATIENT WAS NOT HAPPY WITH HIS REFRACTIVE VISION. THE ICL WAS EXCHANGED FOR A SHORTER LENS (THIS LENS WAS ALSO EXPLANTED A COUPLE OF HOURS LATER - SEE MFR REPORT # 2023826-2015-00845). THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: (B)(4) NEW INCISION. SECONDARY SURGERY. REFRACTIVE SURPRISE. NEW IRIDECTOMY. EXPLANTED. VAULTING. DEVICE EVALUATED BY MANUFACTURER? NO: LENS DISCARDED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS DISCARDED. Manufacturer Narrative: BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923015&pc=MTA | 2023826-2015-00825 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE LENS HAS AN EXCESSIVE VAULT AND THERE IS A REFRACTIVE SURPRISE. THE LENS WILL BE EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933186&pc=MTA | 2023826-2015-00922 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, WORK ORDER SEARCH, AND THE PRODUCT EVALUATION, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO ROOT CAUSE WAS DETERMINED. PER MEDICAL REVIEW: REPORTEDLY, ICL (13.7 MM) WAS EXPLANTED/EXCHANGED FOR A SHORTER LENGTH ICL (13.2 MM) ONE WEEK POSTOPERATIVELY TO ADDRESS EXCESSIVE VAULTING AND SUBSEQUENT SHALLOWING OF ANTERIOR CHAMBER. NO OTHER COMPLICATIONS SUCH AS ELEVATED IOP WERE OBSERVED. ACCORDING TO (B)(4) THE CAUSE OF THE EVENT WAS UNKNOWN. THE SAME REPORT STATED THAT BCVA UPON EXCHANGE WAS 20/20. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933192&pc=MTA | 2023826-2015-00923 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE EXCESSIVE VAULT OF THE IMPLANTED LENS. ADDITIONALLY, MAJORITY OF THE LENS FROM WO# (B)(4) WERE IMPLANTED WITH NO ISSUES. PER MEDICAL REVIEW: REPORTEDLY, MICL (13.2 MM) WAS EXPLANTED/EXCHANGED FOR A SHORTER LENGTH MICL (12.6 MM) TWO WEEKS POSTOPERATIVELY, TO ADDRESS LENS OVER- VAULTING AND SUBSEQUENT PUPILLARY BLOCK AND ELEVATED IOP. UPON LENS EXCHANGE THE PATIENT WAS DOING VERY WELL AND HEALTHY (VA: 20/20).THE CCQS REPORT STATED THAT THE EVENT WAS NOT PRODUCT RELATED AND THAT THE MOST LIKELY CAUSE WAS PATIENT ANATOMIC ISSUE THAT LED TO PRE-OP MISCALCULATION DUE TO INCONSISTENT WTW MEASUREMENT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954993&pc=MTA | 2023826-2015-00906 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975008&pc=MTA | 2023826-2015-00967 |  | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975014&pc=MTA | 2023826-2015-00964 |  | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5033673&pc=MTA | 2023826-2015-01137 |  | Injury | STAAR SURGICAL COMPANY | 2015-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS FOR THIS SERIAL #(B)(4), REMAINS IMPLANTED AND WAS REPORTED AS BEING EXPLANTED IN ERROR. Manufacturer Narrative: LENS NOT RETURNED. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO: LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086248&pc=MTA | 2023826-2015-01206 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -08.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089549&pc=MTA | 2023826-2015-01234 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -19.0/2.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089554&pc=MTA | 2023826-2015-01253 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND THE PATIENT WAS NOT HAPPY WITH THE OUTCOME. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PER MEDICAL REVIEW - ACCURATE DETERMINATION OF ANTERIOR SEGMENT DIMENSIONS IS KEY TO MAINTENANCE OF ADEQUATE VAULT. THE METHOD OF ICL SIZING FOR THIS PATIENT WAS BASED ON ACD AND WTW RECOMMENDATIONS PER DFU. THE REPORT PROVIDES NO SPECIFIC INFORMATION ON THE DEGREE OF VAULTING, THE CAUSE OF EXCESSIVE VAULT, OR THE CLINICAL EVIDENCE FOR EXCESSIVE VAULTING. BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089559&pc=MTA | 2023826-2015-01236 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -23.0/+2.5/071 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5158109&pc=MTA | 2023826-2015-01363 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: PER MEDICAL REVIEW - REPORTEDLY, ICL(13.2 MM) WAS EXPLANTED/EXCHANGED FOR A SHORTER LENGTH ICL (12.6 MM) EIGHT MONTHS POSTOPERATIVELY TO ADDRESS EXCESSIVE VAULTING AND NARROWING OF THE ANGLE. NO INCREASE IN IOP WAS REPORTED AND BCVA WAS 20/20 PRIOR TO THE EXPLANTATION. ACCORDING TO THE CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, DATED ON (B)(6) 2015, THE EVENT WAS NOT DEVICE RELATED. IT WAS ATTRIBUTED TO PATIENT FACTOR ("BORDERLINE WHITE TO WHITE MEASUREMENT ") AND SUBSEQUENT SURGEON`S DECISION TO IMPLANT LONGER (13.2 MM) ICL. NO REPORTED POSTOPERATIVE SEQUELAE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, THE CAUSE OF THIS COMPLAINT WAS PATIENT FACTOR. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE SURGERY FACILITY REPORTED, THE EVENT WAS NOT DEVICE RELATED AND WAS ATTRIBUTED TO PATIENT FACTOR (BORDERLINE WHITE TO WHITE MEASUREMENT) AND SURGEON DECISION TO IMPLANT LONGER LENS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD CODE: - WORK ORDER SEARCH. RESULTS CODE: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167000&pc=MTA | 2023826-2015-01366 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -18.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167047&pc=MTA | 2023826-2015-01377 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171646&pc=MTA | 2023826-2015-01383 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/X080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5174769&pc=MTA | 2023826-2015-01394 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE PATIENT WAS NOT HAPPY WITH THE RESULTS AND THE LENS WAS EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO OVERVAULT/EXCESSIVE VAULTING AND NARROWING OF THE ANGLE. THE PATIENT EXPERIENCED GLARE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP UNCORRECTED VA WAS 20/20. Manufacturer Narrative: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO ROOT CAUSE WAS DETERMINED. PER MEDICAL REVIEW: REPORTEDLY, MICL (13.2MM;-9D) WAS EXPLANTED/EXCHANGED, TWO MONTHS POSTOPERATIVELY, FOR A SAME DIOPTER BUT SHORTER LENS LENGTH (12.6 MM;-9D) TO ADDRESS LENS OVERVAULT AND SUBSEQUENT NARROWING OF THE ANGLE AND PATIENT SUBJECTIVE DISTURBANCES (GLARE). THE IOP WAS NOT ELEVATED. POST-OP UCVA WAS 20/20. ACCORDING TO (B)(4), THE EXACT CAUSE OF THE EVENT WAS UNKNOWN BUT WAS NOT PRODUCT RELATED. THE SAME FORM ALSO STATED THAT EVER SINCE THIS HAPPENED THEY REFINED VAULT NOMOGRAM TO AVOID THIS TYPE OF EVENT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULTS (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, VISUAL INSPECTION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: REPORTEDLY, ICL (13.2 MM;-9 D) WAS EXPLANTED/EXCHANGED, ALMOST TWO MONTHS POSTOPERATIVELY, FOR ANOTHER ICL (UNKNOWN SIZE; UNKNOWN DIOPTER) TO ADDRESS PATIENT DISSATISFACTION WITH VA. NO REPORTED PROBLEM WITH THE LENS POSITION PRIOR TO THE EXPLANTATION. NO REPORTED POSTOPERATIVE SEQUELAE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193778&pc=MTA | 2023826-2015-01399 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS EXPLANTED DUE TO NARROWING OF THE ANGLE AND DECREASED PERIPHERAL VISION. A POSSIBLE CAUSE OF THE EVENT WAS A SIZING ISSUE. THE VAULT IS NOW NORMAL AND UCVA IS 20/20. Manufacturer Narrative: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE MEDICAL REVIEW REPORTS DEVICE CAUSALITY. THE QUESTIONNAIRE INDICATED THE CAUSE OF THE EVENT AS NEEDED OPTIMIZED SIZING. THE PRODUCT WAS NOT RETURNED TO STAAR FOR A PRODUCT EVALUATION. PER MEDICAL REVIEW: REPORTEDLY, MICL (13. 7 MM) WAS EXPLANTED/EXCHANGED, FIVE WEEKS POSTOPERATIVELY FOR A SHORTER LENGTH ICL (13.2MM) TO ADDRESS HIGH VAULT AND SUBSEQUENT NARROWING OF THE ANGLE AND SHALLOWING OF ANTERIOR CHAMBER. THE PATIENT WAS EXPERIENCING DECREASE OF PERIPHERAL VISION. ACCORDING TO THE SAME REPORT CAUSE OF THE EVENT WAS ATTRIBUTED TO THE DEVICE (SIZING). UPON LENS EXCHANGE UCVA WAS 20/20 AND THE LENS WAS PROPERLY POSITIONED WITH NORMAL VAULT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW - REPORTEDLY ICL (13.7 MM) WAS EXPLANTED/EXCHANGED , 5 WEEKS POSTOPERATIVELY, FOR A SHORTER LENGTH ICL (13.2MM) TO ADDRESS "EXCESSIVE VAULTING AND SHALLOWING OF THE ANTERIOR CHAMBER". NO REPORTED LOSS OF BCVA PRIOR TO THE EXCHANGE. NO REPORTED POSTOPERATIVE SEQUELAE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER - NO - (OTHER): LENS DISCARDED. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS DISCARDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193785&pc=MTA | 2023826-2015-01398 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO NARROWING OF THE ANGLE AND THE PATIENT HAD DECREASED PERIPHERAL VISION. THE REPORTER INDICATED A POSSIBLE CAUSE FOR THE EVENT WAS THE LENS DIAMETER TO WHITE-TO-WHITE RATIO WAS NOT CORRECT. THE PATIENT IS DOING VERY WELL, WITH NORMAL VAULT AND 20/20 UCVA OU. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS (EVALUATION RESULT): ): BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE REPORTER INDICATED A POSSIBLE CAUSE FOR THE EVENT WAS THE LENS DIAMETER TO WHITE-TO-WHITE RATIO WAS NOT CORRECT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. THE PRODUCT WAS NOT RETURNED TO STAAR FOR A PRODUCT EVALUATION. UPON LENS EXCHANGE WITH A SHORTER LENS, THE PATIENT WAS REPORTED WELL, NORMAL VAULT AND UCVA 20/20. PER MEDICAL REVIEW: REPORTEDLY, MICL (13. 7 MM) WAS EXPLANTED/EXCHANGED, FIVE WEEKS POSTOPERATIVELY, FOR A SHORTER LENGTH ICL (13.2MM) TO ADDRESS HIGH VAULT AND SUBSEQUENT NARROWING OF THE ANGLE AND SHALLOWING OF ANTERIOR CHAMBER. THE PATIENT WAS EXPERIENCING DECREASE OF PERIPHERAL VISION. ACCORDING TO THE SAME REPORT THE CAUSE OF THE EVENT WAS ATTRIBUTED TO THE DEVICE SINCE LENS DIAMETER TO WHITE-TO-WHITE RATIO WAS NOT CORRECT. UPON LENS EXCHANGE THE PATIENT WAS DOING VERY WELL (UCVA 20/20) AND THE LENS WAS PROPERLY POSITIONED WITH NORMAL VAULT. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD CODE (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS CODE (EVALUATION RESULT): PER MEDICAL REVIEW - REPORTEDLY ICL( 13.7 MM) WAS EXPLANTED/EXCHANGED , 5 WEEKS POSTOPERATIVELY, FOR A SHORTER LENGTH ICL TO ADDRESS "EXCESSIVE VAULTING AND SHALLOWING OF THE ANTERIOR CHAMBER". NO REPORTED LOSS OF BCVA PRIOR TO THE EXCHANGE. NO REPORTED POSTOPERATIVE SEQUELAE. CONCLUSION CODE (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS DISCARDED. METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5215723&pc=MTA | 2023826-2015-01483 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS,09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5215754&pc=MTA | 2023826-2015-01486 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5215773&pc=MTA | 2023826-2015-01487 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5217139&pc=MTA | 2023826-2015-01472 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2015. THE PERIPHERAL IRIDOTOMIES WERE REDONE TO OPEN THE ANGLE BUT THE CHAMBER WAS SHALLOW. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: RESULTS (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, VISUAL INSPECTION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. PER MEDICAL REVIEW - REPORTEDLY, ICL (13.7 MM) WAS EXPLANTED/EXCHANGED FOR A SHORTER LENGTH ICL (13.2 MM), SIX DAYS POSTOPERATIVELY, TO ADDRESS EXCESSIVE VAULTING AND SHALLOWING OF ANTERIOR CHAMBER. PRIOR TO THE EXCHANGE THE SURGEON PERFORMED SECONDARY IRIDOTOMY TO OPEN THE ANGLE. FOLLOWING SECONDARY YAG, ANTERIOR CHAMBER WAS STILL SHALLOW NASALLY SO THE SURGEON DECIDED TO EXCHANGE THE LENS. VA FOUR HOURS AFTER EXCHANGE WAS 20/25. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218718&pc=MTA | 2023826-2015-01479 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). (PROCESS EVALUATION): WORK ORDER SEARCH (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218799&pc=MTA | 2023826-2015-01489 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5224354&pc=MTA | 2023826-2015-01506 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+3.50/X073 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5224355&pc=MTA | 2023826-2015-01507 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+1.0/101 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325831&pc=MTA | 2023826-2015-01705 |  | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5371776&pc=MTA | 2023826-2016-00064 |  | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS RE-POSITIONED ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5471957&pc=MTA | 2023826-2016-00210 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO VAULTING AND WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (OPERATIONAL PROBLEM): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE OR NOT DETECT A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE RESULTED IN THE EXCESSIVE VAULTING OF THE LENS. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: PATIENT CODE: (B)(4) METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5472397&pc=MTA | 2023826-2016-00209 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO VAULTING AND WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE OR NOT DETECT A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE RESULTED IN THE EXCESSIVE VAULTING OF THE LENS. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5495698&pc=MTA | 2023826-2016-00262 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5505907&pc=MTA | 2023826-2016-00271 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5511738&pc=MTA | 2023826-2016-00239 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO TOO MUCH VAULT (EXCESSIVE VAULTING). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: (OPERATIONAL PROBLEM): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE OR NOT DETECT A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE RESULTED IN THE EXCESSIVE VAULTING OF THE LENS. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5518777&pc=MTA | 2023826-2016-00296 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.00/+3.0/170 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISUAL DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553051&pc=MTA | 2023826-2016-00412 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S EYE. THE LENS HAD A HIGH VAULT AND WAS TILTED ON ONE SIDE. THE PATIENT HAD NORMAL VISION AND PRESSURE BUT THE EVENT HAD INDUCED SOME PUPIL SIZE DIFFERENCE. THE REPORTER INDICATED AT A POST-OP VISIT, THE VAULT HAD GONE DOWN A BIT, BUT WAS STILL HIGH. THE VISION WAS 20/20 AND NORMAL PRESSURE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5581304&pc=MTA | 2023826-2016-00481 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON EXPLANTED A MICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, ON 04/04/2016. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS , -11.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO INADEQUATE VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. THE LENS WAS RETURNED IN LIQUID; (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, PRODUCT EVALUATION AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: EXPIRATION DATE - UNK. (B)(4). (B)(4) LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602211&pc=MTA | 2023826-2016-00517 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -21.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602215&pc=MTA | 2023826-2016-00516 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -22.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: ONLY SECONDARY SURGERY FOR LENS EXCHANGE WAS REPORTED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN AND BENT (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638607&pc=MTA | 2023826-2016-00592 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE (NOT SPECIFIED). (B)(4). CORRECTED DATA: REPORT SOURCE: INITIAL REPORT IS INCORRECT, SHOULD BE DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638610&pc=MTA | 2023826-2016-00593 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBER). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5659933&pc=MTA | 2023826-2016-00692 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4). CONCLUSION: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENTS UNDER THE AGE OF 21 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5665325&pc=MTA | 2023826-2016-00733 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5667073&pc=MTA | 2023826-2016-00575 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A LONGER LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN AND BENT. (B)(4). CORRECTED DATA: (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5673195&pc=MTA | 2023826-2016-00635 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.50/+1.0/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5707621&pc=MTA | 2023826-2016-00675 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724421&pc=MTA | 2023826-2016-00787 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -11.0/+20/093 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION:THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -11.0/+20/093 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724456&pc=MTA | 2023826-2016-00786 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM 125V4 IMPLANTABLE COLLAMER LENS, -22.00/+4.5/081 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, TICMS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EVENT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5736183&pc=MTA | 2023826-2016-00833 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7, -10.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO VAULT WAS TOO HIGH. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED - THE REPORTER STATED THE LENS WAS DOWNSIZED TO A MORE APPROPRIATE VAULT AND THERE WAS NO DEFECT, INJURY OR ANY OTHER ISSUE. DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: AGE/DATE OF BIRTH: UNK. WEIGHT: UNK. DATE OF EVENT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5743540&pc=MTA | 2023826-2016-00847 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S EYE AND THE LENS HAD A HIGH VAULT, WITH MECHANICAL ANISOCORIA POST ICL IMPLANTATION. THE PATIENT WAS NOT HAPPY AND THE LENS WILL BE EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5780973&pc=MTA | 2023826-2016-00887 |  | Injury | STAAR SURGICAL COMPANY | 2016-07-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) IS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PREVIOUSLY, IT WAS REPORTED THAT VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE. THIS IS INCORRECT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN AND BENT AND FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817537&pc=MTA | 2023826-2016-00945 |  | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT, HAPTIC DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830576&pc=MTA | 2023826-2016-00617 |  | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830637&pc=MTA | 2023826-2016-00578 |  | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED. CORRECTED DATA: (B)(4). THIS IS A RE-SUBMISSION TO CORRECT MFR# FOR SUPPLEMENTAL #2 IN WHICH THE MFR NUMBER WAS INCORRECT. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5866933&pc=MTA | 2023826-2016-01043 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO A SIZE ISSUE. ANOTHER ICL LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC AND ONE HAPTIC TORN, WITH A PIECE OF HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867033&pc=MTA | 2023826-2016-01015 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/62 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (SPOT) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870485&pc=MTA | 2023826-2016-01044 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO SIZING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902492&pc=MTA | 2023826-2016-01159 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO INCREASED VAULT AND ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). RESULTS CODE(S): (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF THE LENS OPTIC AND ONE HAPTIC TORN OFF. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909285&pc=MTA | 2023826-2016-01174 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/101 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, WITH FIBER AND SMALL PARTICULATES ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919713&pc=MTA | 2023826-2016-01193 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO LOW VAULTING. A LONGER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND THE PRESENCE OF FOREIGN MATERIAL (SPOTS AND HAIR) ON THE LENS SURFACE. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926675&pc=MTA | 2023826-2016-01219 |  | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629123&pc=MTA | 2023826-2015-00379 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. AGE - AND WEIGHT: UNK. EVENT DATE: UNK. IMPLANT DATE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629142&pc=MTA | 2023826-2015-00378 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. AGE AND WEIGHT: UNK. EVENT DATE: UNK. EXPLANT DATE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4736044&pc=MTA | 2023826-2015-00479 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-04-29 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE (B)(4) BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786493&pc=MTA | 2023826-2015-00565 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-05-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2014, AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/20 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) - (SECONDARY SURGERY); (NEW INCISION); (EXPLANTED); (VAULTING). (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED IN LIQUID AND NO VISIBLE DAMAGE WAS OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855083&pc=MTA | 2023826-2015-00697 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2014. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4949721&pc=MTA | 2023826-2015-00955 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954563&pc=MTA | 2023826-2015-00962 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954994&pc=MTA | 2023826-2015-00958 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5033672&pc=MTA | 2023826-2015-01138 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE PATIENT HAD A REACTION AND THE LENS WAS THE WRONG DIOPTER. THERE WAS NO PATIENT INJURY. A DIFFERENT TYPE LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN AND CRACKED. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. BASED ON THE COMPLAINT HISTORY, EVALUATION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD: (PROCESS EVALUATION): - WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084246&pc=MTA | 2023826-2015-01196 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). CLAIM # (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084252&pc=MTA | 2023826-2015-01198 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -18.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084260&pc=MTA | 2023826-2015-01207 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5189897&pc=MTA | 2023826-2015-01406 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-10-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, -3.5/+3.0/X078 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5189899&pc=MTA | 2023826-2015-01407 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-10-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, 10.5/+3.0/X051 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225077&pc=MTA | 2023826-2015-01512 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +9.5/+1.0/022 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5648761&pc=MTA | 2023826-2016-00567 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.0/+3.5/086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6021746&pc=MTA | 2023826-2016-01443 | 2013-05-31 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/012 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2013. THE LENS HAD BEEN REPOSITIONED TWICE WITH RE-ROTATION DUE TO LOW VAULT. THE LENS WAS REMOVED ON (B)(6) 2016, AND WILL BE EXCHANGED FOR A LENS OF LONGER LENGTH. Manufacturer Narrative: EVENT DATE: (B)(6) 2013. LENS REMOVED (B)(6) 2016; CORRECTED - LENS WAS EXCHANGED ON (B)(6) 2016 AND REPLACED WITH THE SAME SIZE LENS. PROBLEM WAS RESOLVED. LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. (PREVIOUS)-(B)(6) 2016, (CORRECTED)-(B)(6) 2016. (PREVIOUS)-NO, (CURRENT)-YES. (PREVIOUS)-LENS NOT RETURNED, (CURRENT)-LENS RETURNED BUT NOT YET EVALUATED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404638&pc=MTA | 2023826-2017-00396 | 2013-08-27 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.0/+6.0/99 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT EXPERIENCED LOW VAULT AND LENS ROTATION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN SMALL CENTRIFUGE VIAL AND THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. THERE IS DEBRIS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059288&pc=MTA | 2023826-2016-01529 | 2014-02-07 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS - -15.5/+3.0/86 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041201&pc=MTA | 2023826-2016-01091 | 2015-01-30 | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+4.50/081 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND ROTATION OF THE LENS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND DRIED DISCOLORATED RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041198&pc=MTA | 2023826-2016-01092 | 2015-02-08 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/085 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND ROTATION OF THE LENS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND MULTIPLE DRIED, DISCOLORED PARTICULATES ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6112596&pc=MTA | 2023826-2016-01681 | 2015-07-07 | Injury | STAAR SURGICAL COMPANY | 2016-11-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING WITH ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND FIBERS ON LENS SURFACE, BUT NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6509080&pc=MTA | 2023826-2017-00565 | 2015-10-06 | Injury | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2, DIOPTER -17.0/+1.5/093 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: IN THE INITIAL MDR, THE LENS IS INCORRECTLY DESCRIBED AS A VTICM13.2. THE LENS IS A VTICMO13.2. IN THE INITIAL MDR, THE LENS IS INCORRECTLY DESCRIBED AS A VTICM13.2. THE LENS IS A VTICMO13.2. (B)(4). Manufacturer Narrative: CORRECTED DATA: IN FOLLOW-UP #1, THE DATE IS NOT LISTED. IT IS (B)(6) 2017. (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6028659&pc=MTA | 2023826-2016-01458 | 2016-01-22 | Injury | STAAR SURGICAL COMPANY | 2016-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO HIGH VAULTING AND TILTING OF THE LENS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024720&pc=MTA | 2023826-2016-01457 | 2016-02-15 | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/089 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND ROTATION OF THE LENS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6036774&pc=MTA | 2023826-2016-01472 | 2016-02-15 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/094 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND ROTATION OF THE LENS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.98MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5984997&pc=MTA | 2023826-2016-01324 | 2016-05-07 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.5/+3.5/72 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND FOREIGN MATERIAL ON SURFACE AND SPOTTY OPTIC (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5987394&pc=MTA | 2023826-2016-01339 | 2016-05-20 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS - -7.50/+2.0/060 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULT AND LENS ROTATION "NOT" ASSOCIATED WITH LOW VAULT. THE LENS WAS REPOSITIONED IN THE EYE. THE LENS REMAINS IN THE EYE. THE SURGEON PLANS TO EXPLANT THE LENS. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: LENS EXCHANGED WAS PERFORMED, PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783462&pc=MTA | 2023826-2018-01213 | 2016-06-07 | Injury | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+4.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2 IMPLANTABLE COLLAMER LENS -14.5/+4.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. REPORTEDLY, LENS REPOSITIONING WAS PREVIOUSLY ATTEMPTED BUT DID NOT REACH DESIRED OUTCOME. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE EXCHANGE FOR THE LONGER LENS RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS WITH FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6253455&pc=MTA | 2023826-2017-00078 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.00/+1.5/088 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: THERE IS LENS ROTATION ASSOCIATED WITH LOW VAULT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6216619&pc=MTA | 2023826-2016-01852 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-12-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -4.50/+6.0/003 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REPOSITIONED. ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND ROTATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. CORRECTIONAL DATA: PREVIOUS EXPLANT DATE NOT APPLICABLE, ON 12/05/2016 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND ROTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4), (LENS REPOSITIONING, LENS EXPLANT, SECONDARY SURGERY). (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6620720&pc=MTA | 2023826-2017-00936 | 2016-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -12.50/+1.0/080 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS DISLOCATION/SUBLUXATION. THE PROBLEM HAS BEEN RESOLVED. Manufacturer Narrative: THE FOLLOWING INFORMATION WAS NOT INCLUDED IN THE INITIAL REPORT: THE PATIENT ALSO EXPERIENCED ENLARGEMENT OF INCISION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6197149&pc=MTA | 2023826-2016-01791 | 2016-08-17 | Injury | STAAR SURGICAL COMPANY | 2016-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INPSECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE BUT HAS FIBERS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: SECONDARY SURGERY. LENS EXCHANGE. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191114&pc=MTA | 2023826-2016-01786 | 2016-08-30 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THIS WAS PERFORMED ON (B)(6) 2016. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS ON (B)(6) 2016 DUE TO LOW VAULT AND LENS ROTATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191103&pc=MTA | 2023826-2016-01789 | 2016-08-31 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THIS WAS PERFORMED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/91 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULTING AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. Manufacturer Narrative: THE DIOPTER WAS PREVIOUSLY INCOMPLETE. THE LENS EXCHANGE RESOLVED THE PROBLEM. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6392833&pc=MTA | 2023826-2017-00358 | 2016-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-03-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.50/133 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULT, AND LENS DISLOCATION/SUBLUXATION. THE LENS WAS REPOSITIONED BUT IT ROTATED AGAIN AFTER REPOSITIONING. ON (B)(6) 2017 THE LENS WAS THEN EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET BEEN EVALUATED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION-LENS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH NARROW ANTERIOR CHAMBER ANGLES (I.E. LESS THAN GRADE III AS DETERMINED BY GONIOSCOPIC EXAM). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531949&pc=MTA | 2023826-2017-00731 | 2016-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/074 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR LONGER LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6124382&pc=MTA | 2023826-2016-01671 | 2016-09-28 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/093 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6407222&pc=MTA | 2023826-2017-00405 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/99 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT EXPERIENCED LOW VAULT AND LENS ROTATION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A SMALL PIECE OF HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404194&pc=MTA | 2023826-2017-00387 | 2016-12-02 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.50/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED VERY LOW VAULT, AND LENS ROTATION 60 DEGREES CW. ON (B)(6) 2017 THE LENS WAS REPOSITIONED BUT IT ROTATED AGAIN AFTER REPOSITIONING. SURGEON PLANS TO EXCHANGE LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR THE UNITS WITHIN THE SAME LOT. OFF LABEL USE (OVER 45 YEARS OF AGE AT THE TIME OF SURGERY, ACD<3.00MM) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6253838&pc=MTA | 2023826-2017-00082 | 2016-12-05 | Injury | STAAR SURGICAL COMPANY | 2017-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT REPORTING LOW VAULT AND ROTATION OF THE LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: CONCLUSION: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.82MM AT THE TIME OF IMPLANTATION. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589767&pc=MTA | 2023826-2017-00877 | 2016-12-19 | Injury | STAAR SURGICAL COMPANY AG | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6539264&pc=MTA | 2023826-2017-00748 | 2017-02-02 | Injury | STAAR SURGICAL COMPANY | 2017-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/094 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442669&pc=MTA | 2023826-2017-00522 | 2017-02-15 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, AND LENS ROTATION. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR LONGER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: METHOD CODE: CODE IS NOT APPLICABLE, THERE WAS INSPECTION OF OBJECTS OR SURFACES REQUIRING MAGNIFICATION USING AN OPTICAL MICROSCOPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN SMALL VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE, MISSING PIECE OF HAPTIC. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6513105&pc=MTA | 2023826-2017-00685 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-04-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, DIOPTER -8.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LARGER LENS, SAME DIOPTER DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR SURGICAL RESIDUE AND DEBRIS). (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). DEVICE NOT YET RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6513129&pc=MTA | 2023826-2017-00686 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-04-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1, DIOPTER -7.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LARGER LENS, SIMILAR DIOPTER DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE DEVICE WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS). (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE DEVICE WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589757&pc=MTA | 2023826-2017-00875 | 2017-02-22 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR RESIDUE ON LENS. CONCLUSION: PATIENT'S ACD IS BELOW 3.0 MM AND AS PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRADICTED FOR PATIENTS WITH ACD BELOW 3.0 MM. CLAIM# (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR RESIDUE ON LENS. CONCLUSION: PATIENT'S ACD IS BELOW 3.0 MM AND AS PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRADICTED FOR PATIENTS WITH ACD BELOW 3.0 MM. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6491230&pc=MTA | 2023826-2017-00638 | 2017-03-06 | Injury | STAAR SURGICAL COMPANY | 2017-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM12.0, DIOPTER -15.0/1.5/072 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR ANOTHER LENS, DIFFERENT MODEL AND DIOPTER, ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6491432&pc=MTA | 2023826-2017-00637 | 2017-03-06 | Injury | STAAR SURGICAL COMPANY AG | 2017-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM12.0, DIOPTER -13.5/1.5/097 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR ANOTHER LENS, DIFFERENT MODEL AND DIOPTER, ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. CONCLUSION: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6486114&pc=MTA | 2023826-2017-00615 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, AND LENS ROTATION. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR LONGER LENS, AND THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED AND DEBRIS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6624611&pc=MTA | 2023826-2017-00950 | 2017-03-21 | Injury | STAAR SURGICAL COMPANY | 2017-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589766&pc=MTA | 2023826-2017-00876 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.50/X119 DIOPTER, IN THE PATIENT'S RIGHT +EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING FROM THE LENS AND CLEAR RESIDUE ON THE LENS SURFACE. CONCLUSION: OFF_LABEL USE- AS PER DFU, THE IMPLANTATION OF LENS MODEL VTICMO13.2 IS CONTRAINDICATED FOR PATIENTS OF AGE LESS THAN 21 YEARS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6655602&pc=MTA | 2023826-2017-01021 | 2017-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-06-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULTING AND LENS ROTATION ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN THE SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND RED SURGICAL RESIDUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6655646&pc=MTA | 2023826-2017-01022 | 2017-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-06-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULTING AND LENS ROTATION ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6624727&pc=MTA | 2023826-2017-00953 | 2017-05-02 | Injury | STAAR SURGICAL COMPANY | 2017-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/081 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED, THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE), AND THE LENS HAVING A CURVED SHAPE. CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670131&pc=MTA | 2023826-2017-01033 | 2017-05-22 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, AND LENS ROTATION. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6682228&pc=MTA | 2023826-2017-01091 | 2017-05-22 | Injury | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, DIOPTER -11.5 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: NOTE: THIS COMPLAINT IS A DUPLICATE. REFERENCE MDR#2023826-2017-01033. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6723380&pc=MTA | 2023826-2017-01138 | 2017-05-25 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION AND WAS EXCHANGED FOR A LONGER LENS OF SIMILAR DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2017, THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: LENS EXCHANGE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6791947&pc=MTA | 2023826-2017-01219 | 2017-06-18 | Injury | STAAR SURGICAL COMPANY | 2017-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICMO13.2, -8.50/1.5/57 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT AND LENS ROTATION WAS NOTICED. THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE VAULT WAS STILL TOO LOW AND THE LENS ROTATED AGAIN AFTER REPOSITIONING. THE LENS IS PLANNED TO BE EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER" IN FOLLOW-UP # 1 SHOULD BE (B)(6) 2017. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LONGER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF FIBERS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993880&pc=MTA | 2023826-2017-01662 | 2017-06-27 | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -10.5/+2.5/079 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS REPOSITIONED. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT, LENS ROTATION, AND LENS DISLOCATION/SUBLUXATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS RED AND CLEAR RESIDUE AND DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6796062&pc=MTA | 2023826-2017-01307 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/101 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND WAS REPOSITIONED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS -16.5/+3.5/101 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND ROTATION AND WAS REPOSITIONED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6796078&pc=MTA | 2023826-2017-01306 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+3.00/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND WAS REPOSITIONED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS HAD LOW VAULT, LENS ROTATION BY 90 DEGREES AND WAS REPOSITIONED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254846&pc=MTA | 2023826-2019-00063 | 2017-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+5.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE CAUSE OF THE EVENT IS LISTED AS "WRONG SIZE." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. METHOD CODE ADDED CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6769228&pc=MTA | 2023826-2017-01211 | 2017-07-14 | Injury | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND MISALIGNMENT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY WITH IMPLANT. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7131207&pc=MTA | 2023826-2017-01902 | 2017-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -2.0/+3.5/96 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL OBSERVATION FOUND THE LENS WITH NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS SURFACE. CORRECTIONS: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -2.0/+3.5/96 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND LENS ROTATION. THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978096&pc=MTA | 2023826-2017-01616 | 2017-07-21 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4, DIOPTER -15.5/+3.0/081 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE), AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7033554&pc=MTA | 2023826-2017-01727 | 2017-08-24 | Injury | STAAR SURGICAL COMPANY | 2017-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -9.5/+2.5/090 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6)2017, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE AND PARTICLES ON THE LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6959958&pc=MTA | 2023826-2017-01594 | 2017-09-06 | Injury | STAAR SURGICAL COMPANY | 2017-10-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOSITURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4). SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR LONGER LENS (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050902&pc=MTA | 2023826-2017-01716 | 2017-10-25 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -18.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND LENS DISLOCATION/SUBLUXATION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7198057&pc=MTA | 2023826-2018-00017 | 2017-10-31 | Injury | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.50/+1.5/059 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR LONGER LENS THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7149243&pc=MTA | 2023826-2017-01957 | 2017-11-04 | Injury | STAAR SURGICAL COMPANY | 2017-12-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.0/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER IN INITIAL MDR CHANGED FROM 08-DEC-2017 TO 07-DEC-2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183434&pc=MTA | 2023826-2018-00055 | 2017-11-09 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -11.0/+1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF DEBRIS AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183457&pc=MTA | 2023826-2018-00054 | 2017-11-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -8.0/+2.0/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND THICK CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7184025&pc=MTA | 2023826-2018-00056 | 2017-11-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -8.50/+2.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR AND WHITE RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201661&pc=MTA | 2023826-2018-00092 | 2017-12-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICH12.1 IMPLANTABLE COLLAMER LENS, -1.50/+4.00/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTES LOW VAULT, LENS ROTATION, LENS DISLOCATION/SUBLUXATION. REPORTEDLY, THE LENS WAS ALSO REPOSITIONED. ADDITIONAL COMMENTS STATE, "PRIMARY ICL IS TOO SHORT AND TOO MOBILE IN THE SULCUS. NEEDS TO BE REPLACED WITH A 12.6 MM MODEL." CURRENTLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182196&pc=MTA | 2023826-2018-01868 | 2017-12-28 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING, WITH LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639928&pc=MTA | 2023826-2018-00959 | 2018-01-13 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.0/+2.5/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT AND THE OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456981&pc=MTA | 2023826-2018-00589 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -13.0/+3.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS HAS A CURVED SHAPE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571606&pc=MTA | 2023826-2018-00842 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. VAULT WAS RESOLVED. ROTATION WAS NOT RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456982&pc=MTA | 2023826-2018-00591 | 2018-02-21 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -16.0/+3.0/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465301&pc=MTA | 2023826-2018-00606 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481617&pc=MTA | 2023826-2018-00663 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.0/+2.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT . THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: SHOULD HAVE BEEN MENTIONED IN INTIAL MDR. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND LENS ROTATION. (B)(4) Manufacturer Narrative: PATIENT CODE: (B)(4), SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7495916&pc=MTA | 2023826-2018-00715 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+5.5/171 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A LARGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE TORN PIECE MISSING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594259&pc=MTA | 2023826-2018-00889 | 2018-04-04 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.0/+2.0/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. REPORTEDLY, THE LENS WAS ALSO REPOSITIONED PRIOR TO LENS EXCHANGE. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853006&pc=MTA | 2023826-2018-01331 | 2018-04-06 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765090&pc=MTA | 2023826-2018-01168 | 2018-04-30 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765118&pc=MTA | 2023826-2018-01166 | 2018-05-24 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND A HOLE IN THE OPTIC. (B)(4). Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED TO "16-OCT-2018" IN THE SUPPLEMENTAL MDR#2 . CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704644&pc=MTA | 2023826-2018-00987 | 2018-05-26 | Injury | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704666&pc=MTA | 2023826-2018-01051 | 2018-05-26 | Injury | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7885165&pc=MTA | 2023826-2018-01398 | 2018-05-26 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FIBERS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783456&pc=MTA | 2023826-2018-01211 | 2018-06-07 | Injury | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+4.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/4.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. IT WAS REPORTED THAT THERE WAS LENS ROTATION DUE TO LOW VAULT. LENS REPOSITIONING WAS PERFORMED BUT DID NOT REACH DESIRED OUTCOME. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7997736&pc=MTA | 2023826-2018-01560 | 2018-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740028&pc=MTA | 2023826-2018-01142 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6-METHOD CODE ADDED. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765097&pc=MTA | 2023826-2018-01167 | 2018-06-27 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824411&pc=MTA | 2023826-2018-01261 | 2018-07-10 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -2.50/+3.5/136 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD LOW VAULTING. THE LENS HAD SHIFTED DOWN TO LOWER MARGIN OF PUPIL AND THE CONJUCTIVA WAS OVERHANGING FOR ALMOST 1MM ALL AROUND CORNEA. THE LENS REMAINS IMPLANTED. THE REPORTER INDICTED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LOW VAULT OF THE LENS WAS FIRST NOTICED ON THE 1ST POST-OP DAY AND THE DECENTRATION WAS FOUND ON THE 6TH POST-OP DAY. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7973764&pc=MTA | 2023826-2018-01474 | 2018-07-26 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, EXCESSIVE VAULT, AND ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2018. THE LENS WAS REMOVED AND A SHORTER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870353&pc=MTA | 2023826-2018-01386 | 2018-08-02 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870364&pc=MTA | 2023826-2018-01383 | 2018-08-02 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870361&pc=MTA | 2023826-2018-01381 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS REPOSITIONED. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE IS LISTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7972154&pc=MTA | 2023826-2018-01526 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870358&pc=MTA | 2023826-2018-01385 | 2018-08-16 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE IS LISTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129212&pc=MTA | 2023826-2018-01729 | 2018-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND (B)(6) 2018, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE:(B)(4). ADDITIONAL DATA: DEVICE EVALUATED BY MFR-DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN, AND TORN PIECE IS MISSING. CLAIM #(B)(4). Manufacturer Narrative: UNK. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7890063&pc=MTA | 2023826-2018-01414 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON DECIDED TO IMPLANT AN 13.2 MM VTICMO13.2, -10.0/+2.5/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS LENGTH WHICH WAS TOO SHORT CAUSING LOW VAULT. IT WAS HIS FAULT NOT TO SELECT THE 13.7MM LENS LENGTH WHICH IS NOW IN THE PATIENTS EYE (AFTER EXCHANGE ON ((B)(6) 2018) AND SHOWS A GOOD VAULTING OF ~400¿M. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED, LENS HAS A CURVED SHAPE, AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON DECIDED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10/2.5/62 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON 08/07/2018. THE LENS WAS EXCHANGED ON 08/30/2018 DUE TO LENS LENGTH BEING TOO SHORT AND CAUSING LOW VAULT. IT WAS REPORTED "IT WAS HIS FAULT NOT TO SELECT THE 13.7MM LENS LENGTH WHICH IS NOW IN THE EYE (AFTER EXCHANGE) AND SHOWS A GOOD VAULTING OF ~400M." THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906081&pc=MTA | 2023826-2018-01345 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.05/+2.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.5/+2.5/72 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8092652&pc=MTA | 2023826-2018-01714 | 2018-09-01 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATAION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WASRESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194497&pc=MTA | 2023826-2018-01902 | 2018-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8059983&pc=MTA | 2023826-2018-01621 | 2018-09-18 | Injury | STAAR SURGICAL COMPANY | 2018-11-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.50/2.0/089, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8054766&pc=MTA | 2023826-2018-01625 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-11-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.50/2.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194048&pc=MTA | 2023826-2018-01900 | 2018-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8102860&pc=MTA | 2023826-2018-01727 | 2018-10-19 | Injury | STAAR SURGICAL COMPANY | 2018-11-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO LOW VAULT (SUBLUXATION). IT WAS REPORTED EVENT WAS DUE TO PATIENT RELATED FACTOR. THERE ARE NO PLANS TO IMPLANT ANOTHER ICL. Manufacturer Narrative: (B)(4). LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8284935&pc=MTA | 2023826-2019-00037 | 2018-11-06 | Injury | STAAR SURGICAL COMPANY | 2019-01-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) /2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: DATE SHOULD BE CORRECTED TO 02-JAN-2019 IN THE ORIGINAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409328&pc=MTA | 2023826-2019-00390 | 2018-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT AND LENS ROTATION. THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS. REPORTEDLY, THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CORRECTED DATA: DATE OF EVENT: CORRECTED FROM "(B)(6) 2018" TO "(B)(6) 2018" FOR INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354487&pc=MTA | 2023826-2019-00293 | 2018-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 12.1MM, VTICM5_12.1, -11.5/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, INA MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: CORRECTED DATA EVENT: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 12.1MM, VTICM5_12.1, -11.5/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON 06 JUL 2018. ON 29 DEC 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH LENS ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. DEVICE CODE:(B)(4) - ADDED TO INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9284956&pc=MTA | 2023826-2019-02043 | 2019-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM, VICMO12.6, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE: 2923 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8932875&pc=MTA | 2023826-2019-01549 | 2019-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTICS BENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8937586&pc=MTA | 2023826-2019-01555 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT WAS REMOVED FROM THE OPERATING ROOM AND RETURNED 2 HOURS LATER TO CHECK THE VAULT. THE LENS WAS FOUND TO HAVE LOW VAULTING, WITH LENS ROTATION, AND THE LENS WAS EXPLANTED THE SAME DAY. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS WHITE-TO-WHITE BORDER. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4741318&pc=MTA | 2023826-2015-00167 |  | Injury | STAAR SURGICAL COMPANY | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 OF -10.5/+1.6/100 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS ROTATION ASSOCIATED WITH A LOW VAULT. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). RESULTS: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. THIS IS A RE-SUBMISSION FOR SUPPLEMENTAL MDR #2 WITH CORRECT MFR NUMBER PER FDA REQUEST. SUBMITTED IN ERROR WITH INCORRECT MFR NUMBER USED WAS (2023826-2016-00167) THIS IS NOW CORRECTED TO MFR # 2023826-2015-00167. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS IS A RE-SUBMISSION OF THE INITIAL MDR ELECTRONICALLY THROUGH WEB TRADER PER FDA REQUEST. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6519134&pc=MTA | 2023826-2017-00687 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND LENS ROTATION. LARGE CYSTS WERE PRESENT IN THE EYE, INTERFERING WITH FOOTPLATES POSITIONING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6791745&pc=MTA | 2023826-2017-01316 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-14 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/10.5/5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6)2017 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7046770&pc=MTA | 2023826-2017-01746 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND THE LENS HAD TILTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/19. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. PATIENT CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152460&pc=MTA | 2023826-2017-01964 |  | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -16.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. A LONGER LENS WAS THEN IMPLANTED (DATE UNKNOWN). THE SURGEON STATES THAT THE "LENS SHIFTED TO THE TEMPORAL SIDE AND THE CENTRAL HOLE SHIFTED TO THE IRIS EDGE." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. CORRECTED INFORMATION-THE PROBLEM WAS NOT RESOLVED WITH THE LENS EXCHANGE. SEE MFR REPORT# 2023826-2018-00813 (CLAIM# (B)(4)) FOR CASE CONCLUSION. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: DATE OF EVENT-UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152555&pc=MTA | 2023826-2017-01963 |  | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE SURGEON STATES THAT THE "LENS SHIFTED TO TEMPORAL SIDE AND THE CENTRAL HOLE SHIFTED TO THE IRIS EDGE." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557851&pc=MTA | 2023826-2018-00813 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -16.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. REPORTEDLY, IT WAS REPLACED WITH AN IOL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: (B)(4) Manufacturer Narrative: "PATIENT HAS NO INCIDENCE OF CATARACT BUT HE REFUSED TO IMPLANT DIAMETER 13.7". CLAIM # (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9079765&pc=MTA | 2023826-2019-01661 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -7.00 DIOPTER LENS INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. THE SURGEON REPORTED EXCESSIVE VAULT AND PERFORMED LENS ROTATION TO 90 DEGREE ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. THE PATIENT IS COMFORTABLE AND THE REMAINING REFRACTION WAS CORRECTED WITH LASER SURGERY. Manufacturer Narrative: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LASER SURGERY SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE 2923 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. RESULT CODE 213 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553052&pc=MTA | 2023826-2016-00388 | 2014-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.00/+4.50/X 89 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2014. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED (B)(6) 2014. THERE WAS ENLARGEMENT OF OR ADDITION OF A NEW PERIPHERAL IRIDECTOMY. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT IS (B)(6). (B)(4) Manufacturer Narrative: DIOPTER: -12.00/+4.50/X089. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5734100&pc=MTA | 2023826-2016-00831 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2016-06-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT THE LENS WAS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS WITH ANY PREVIOUS CORNEAL/REFRACTIVE SURGERY AND OVER 45 YEARS OF AGE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552477&pc=MTA | 2023826-2019-00657 | 2019-03-07 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT, IRIS STRETCHED, SHALLOW AC PERIPHERALLY, RESIDUAL REFRACTION INCREASED TO -0.75 CYL, AND THAT THE LENS MOVED 1MM FROM THE RETINA. THE LENS CURRENTLY REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5989111&pc=MTA | 2023826-2016-01352 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND HAPTIC TORN AND SPOTTY OPTIC. (B)(4): (UNINTENDED MOVEMENT) WAS INCORRECT; ROTATION DUE TO LOW VAULTING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923601&pc=MTA | 2023826-2016-01206 | 2016-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION (THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND UNSPECIFIED FOREIGN MATERIAL ON SURFACE AND NO VISIBLE DAMAGE). (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6428721&pc=MTA | 2023826-2017-00463 | 2015-11-24 | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THERE WAS A LENS TEAR DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6031741&pc=MTA | 2023826-2016-01133 | 2016-03-12 | Injury | STAAR SURGICAL COMPANY | 2016-10-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/92 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT DEVELOPED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED IN THE PATIENTS EYE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). MODEL# - CORRECTED MODEL# VTICMO13.2. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684956&pc=MTA | 2023826-2017-01053 | 2017-04-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -8.5/+0.5/072 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE PIGMENT DISPERSION. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE PATIENT EXPERIENCED PIGMENT DISPERSION. CORRECTED INFORMATION RECEIVED FROM THE REPORTER STATES THAT THE PATIENT EXPERIENCED HIGH VAULT, WHICH THEN CAUSED THE PIGMENT DISPERSION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6737203&pc=MTA | 2023826-2017-01132 |  | Injury | STAAR SURGICAL COMPANY | 2017-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -16 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS THEN EXPLANTED. Manufacturer Narrative: DEVICE MANUFACTURE DATE: IT IS CORRECTED TO 24-MAY-2016 FROM 30-NOV-2018. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -16 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS THEN EXPLANTED DUE TO EXCESSIVE VAULTING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5534893&pc=MTA | 2023826-2016-00258 | 2014-04-19 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4, -17.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2011. DECREASED VISION OVER TIME WAS NOTED (B)(6) 2014. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015; UCVA WAS 20/16. SEE MFR. #2023826-2016-00257 FOR LEFT EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. DECREASED VISION OVER TIME WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PRODUCT IS NOT MANUFACTURED IN THE U.S. AND NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: DIOPTER: -17.00, (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5642753&pc=MTA | 2023826-2016-00594 | 2016-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+1.50/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+1.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS ROTATION, AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5890898&pc=MTA | 2023826-2016-00892 | 2016-05-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT SEEING DOUBLE IMAGES. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE THE OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN, AND FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199517&pc=MTA | 2023826-2015-01358 | 2015-10-04 | Injury | STAAR SURGICAL COMPANY | 2015-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Insufficient Information | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER -11.5 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. LOW VAULT WITH REFRACTIVE SURPRISE WAS OBSERVED FIVE YEARS AFTER ICL IMPLANTATION. SURGERY IS PENDING TO EXPLANT AND EXCHANGE ICL. SEE MFR. #2023826-2015-01359 FOR LEFT EYE(OS). Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED WITH VICMO13.2 ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE PRESENCE OF FOREIGN MATERIAL (HAIR) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.50. DEVICE EVALUATED BY MANUFACTURER: NO LENS REMAINS IMPLANTED. (B)(4). NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287715&pc=MTA | 2023826-2015-01633 | 2015-07-29 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Missing Value Reason | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SHALLOW ANTERIOR CHAMBER (PERIPHERAL AND CENTRAL) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834807&pc=MTA | 2023826-2016-00998 | 2016-03-22 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Torn Material; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/006 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND LENS TEAR. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/ SPLIT, HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL PRESENT ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5830560&pc=MTA | 2023826-2016-00992 | 2016-04-04 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Torn Material; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THAT THE LENS WAS TORN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6332495&pc=MTA | 2023826-2017-00148 | 2016-12-20 | Injury | STAAR SURGICAL COMPANY | 2017-02-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Unintended Movement | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.50/+6/008 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THERE WAS A PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. AFTER THE CATARACT SURGERY, THE PATIENT GOT HIS BCVA PRE OP. ON PATIENT'S LAST VISIT ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: REVIEW OF THE FILE INDICATES THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS BECAUSE PATIENT HAS KERATOCONUS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199519&pc=MTA | 2023826-2015-01359 | 2015-10-07 | Injury | STAAR SURGICAL COMPANY | 2015-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Insufficient Information | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER -11.0 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. LOW VAULT WITH REFRACTIVE SURPRISE WAS OBSERVED FIVE YEARS AFTER ICL IMPLANTATION. SURGERY IS PENDING TO EXPLANT AND EXCHANGE ICL. SEE MFR. #2023826-2015-01358 FOR RIGHT EYE(OD). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR 13.2 MM VICMO13.2, IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER ON (B)(6) 2015. THIS EXCHANGE WITH LONGER LENS RESOLVED THE PROBLEM. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4816518&pc=MTA | 2023826-2015-00639 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-06-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Malposition of Device | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT AND ROTATION OF THE LENS WERE NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. REFERENCE MFR. REPORT # 2023826-2015-00638 FOR PATIENT'S RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH NO VISIBLE DAMAGE OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549342&pc=MTA | 2023826-2016-00386 | 2015-02-04 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Missing Value Reason | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -11.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2013. LOW VAULT WAS NOTED. ON (B)(6) 2015 THE LENS WAS EXPLANTED AND EXCHANGED FOR ANOTHER, SAME MODEL AND DIOPTER SIZE. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -11.50. PT WEIGHT: UNK. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552486&pc=MTA | 2023826-2019-00662 | 2019-03-21 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT, IRIS STRETCHED, SHALLOW AC PERIPHERALLY, RESIDUAL REFRACTION OF -0.75 SPH -0.25 CYL X 110. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7617151&pc=MTA | 2023826-2018-00926 | 2018-05-21 | Injury | STAAR SURGICAL COMPANY | 2018-06-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO LOW VAULT. REPORTEDLY, THIS IS DUE TO A PATIENT-RELATED FACTOR. THE REPORTER STATES THAT THE SULCUS IS MUCH BIGGER THAN THE MEASURED WHITE-TO-WHITE VALUE AND IS TOO BIG FOR A LONGER 13.7MM LENS. ON (B)(6) 2018 AN ARTIFLEX LENS WAS IMPLANTED INTO THE PATIENT'S EYE. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913423&pc=MTA | 2023826-2020-00722 | 2019-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. AFTER THE EXCHANGE, THE VAULT DIDN'T IMPROVE TOO MUCH AND THE SURGEON CONSIDERED IT MAYBE CONNECTED TO THE SUSPENSORY LIGAMENT. CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. THE PATIENT IS UNDER CONTINUOUS OBSERVATION. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8983237&pc=MTA | 2023826-2019-01692 | 2019-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE CODE 2682 SHOULD HAVE BEN INCLUDED IN INITIAL MDR. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9515913&pc=MTA | 2023826-2019-02447 | 2019-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9560313&pc=MTA | 2023826-2020-00056 | 2019-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. IN THE REPORTERS OPINION PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS WAS CURVED IN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED THAT ON (B)(6) 2019 THE LENS WAS EXCHANGED DURING THE SAME SURGERY WITH THE REPLACEMENT LENS. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9578497&pc=MTA | 2023826-2020-00089 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9578500&pc=MTA | 2023826-2020-00091 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025651&pc=MTA | 2023826-2020-00998 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -12.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9158760&pc=MTA | 2023826-2019-01846 | 2019-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: "ADDITIONAL TREATMENT REPORTED AS "EXPLANT REPOSITION." SHOULD BE ADDED TO THE INITIAL MDR. H6- DEVICE CODE (B)(6)- SECONDARY SURGICAL INTERVENTION; REPOSITIONING SHOULD BE ADDED TO THE INITIAL MDR. H6- DEVICE CODE (B)(6)- PATIENT DEVICE INCOMPATIBILITY SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# 714671. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9578492&pc=MTA | 2023826-2020-00086 | 2017-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS COVERED IN CLEAR DRY RESIDUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652082&pc=MTA | 2023826-2016-00652 | 2016-03-28 | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Positioning Problem; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM VTICMO12, -17.00/+2.00/X080 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS ROTATED 70 DEGREES AND WAS REPOSITIONED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS WITH A SIMILAR DIOPTER ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (DRY DISCOLORED MATERIAL). (B)(4). CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER THE AGE 45 YEARS OLD. CORRECTED DATA: "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -17.00/+2.00/X080 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS ROTATED 70 DEGREES AND WAS REPOSITIONED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER ON (B)(6) 2016. THIS RESOLVED THE PROBLEM." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668415&pc=MTA | 2023826-2016-00713 | 2014-08-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Positioning Problem; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4, -19.50/1.50/X009 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2014. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. A REPOSITIONING OF THE LENS WAS PERFORMED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/33. SEE MFR. #2023826-2016-00714 FOR LEFT EYE. Manufacturer Narrative: (B)(4). CHANGE "INDO-CORNEAL ANGLES" TO "IRIDO-CORNEAL ANGLES". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668391&pc=MTA | 2023826-2016-00714 | 2014-08-26 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Positioning Problem; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4, -17.00/+1.50/X005 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2014. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED. A REPOSITIONING OF THE LENS WAS PERFORMED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. SEE MFR. #2023826-2016-00713 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5706645&pc=MTA | 2023826-2016-00254 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-06-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Torn Material; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/115 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694858&pc=MTA | 2023826-2016-00250 | 2015-12-24 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Torn Material; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600891&pc=MTA | 2023826-2015-00314 | 2014-10-21 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS ROTATION, ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED. THE ICL WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DATE OF BIRTH - UNK. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4613563&pc=MTA | 2023826-2015-00329 | 2014-12-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) - SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING); UNINTENDED MOVEMENT. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4744632&pc=MTA | 2023826-2015-00494 | 2015-03-17 | Injury | STAAR SURGICAL COMPANY | 2015-05-02 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED, AS WELL AS LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW AND DEVICE HISTORY REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH PEN MARKS. NO VISIBLE DAMAGE WAS OBSERVED. MEDICAL REVIEW - THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST. SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS AND HAPTIC POSITION, ECT). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842194&pc=MTA | 2023826-2015-00679 | 2015-05-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS ROTATED AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SMALLER SIZE ON (B)(6) 2015 AND THIS RESOLVED THE ISSUE. SEE MFR# 2023826-2015-00680 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT INFORMATION AND WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090070&pc=MTA | 2023826-2018-01701 | 2018-10-11 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT. REFERENCE MFR REPORT 2023826-2018-01698 FOR PREVIOUS IMPLANTED LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE REPORTER STATES THE LENS MOVED IN THE EYE, "THE LENS OPTICAL ZONE MIGRATION DOWN." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795453&pc=MTA | 2023826-2019-01232 | 2019-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | P030016 | 1 | Event Description: A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXPLANTED, THE SURGEON REPORTING LOW VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND THAT THE LENS WAS IMPLANTED INVERTED. ON (B)(6) 2019 A LONGER LENS WAS IMPLANTED AND THE PROBLEM IS NOTED AS RESOLVED. SURGEON STATES THE PATIENT IS "DOING WELL." THE CAUSE OF THE EVENT IS REPORTED AS DEVICE, STATES THE LENS "FLIP[PPED] DURING INSERTION." Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4772745&pc=MTA | 2023826-2015-00543 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-05-14 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN BOTH OF THE PATIENT'S EYES AND NOTICED ROTATION OF BOTH LENSES AFTER IMPLANTATION. ICL VAULT WAS NOTICED TO BE LOW IN BOTH EYES AND THE ROTATION WAS ASSOCIATED WITH UNEXPECTED REFRACTIVE OUTCOME. THE SURGEON JUDGED THE LENS TO BE OFF-AXIS BY 27 DEGREES CW IIN THE RIGHT EYE AND 20 DEGREES CW IN THE LEFT EYE. BOTH LENSES REMAIN IMPLANTED. THE SURGEON IS PLANNING TO EXCHANGE BOTH LENSES. IT IS UNKNOWN IF THE LENSES WILL BE LONGER WITH A DIFFERENT POWER OR ONLY LONGER DUE TO INCREASE VAULTING. SEE MFR# 2023826-2015-00542 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. EVALUATION METHOD: LENS WORK ORDER SEARCH & DEVICE HISTORY REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE DEVICE HISTORY REVIEW DETERMINED THAT NOTHING IN THE MANUFACTURING OF THE LENS COULD HAVE CAUSED THE REPORTED EVENT. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND DEVICE HISTORY REVIEW, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY REVIEW & MEDICAL REVIEW. RESULTS: BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING OF THE LENS COULD HAVE CAUSED THE REPORTED EVENT. MEDICAL REVIEW - ROTATION WAS ASSOCIATED WITH UNEXPECTED REFRACTIVE OUTCOME. POST-OP REFRACTION WAS SLIGHTLY MYOPIC (SEQ OF -0.25 D IN THE RIGHT EYE AND -0.75 IN THE LEFT EYE) THE SURGEON JUDGED THE LENS TO BE OFF AXIS BY 27 DEGREES CW IN THE RIGHT EYE AND BY 20 DEGREES CW IN LEFT EYE. THE SURGEON IS PLANNING TO EXCHANGE THESE LENSES WITH NEW ONES (IT IS UNKNOWN IF THE NEW LENSES WILL BE LONGER AND WITH DIFFERENT POWER OR ONLY LONGER TO INCREASE VAULTING). LOW VAULT AND ROTATION MAY BE DUE TO SHORT LENSES. THE CAUSE OF THE RESIDUAL REFRACTION MAY BE MULTIFACTORIAL, INCLUDING OFF-AXIS POSITION SECONDARY TO ROTATION, INACCURATE PREOPERATIVE MEASUREMENTS, SURGICALLY INDUCED CHANGES, TILTING, ETC. CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION,ECT). ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, REVIEW HISTORY AND MEDICAL REVIEWS, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806956&pc=MTA | 2023826-2015-00585 | 2014-04-11 | Injury | STAAR SURGICAL COMPANY | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD - WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED IN LIQUID AND NO VISIBLE DAMAGE WAS OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORKORDER. MEDICAL REVIEW: ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5291346&pc=MTA | 2023826-2015-01612 | 2014-08-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -22.5/+6.0/93 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/17. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012305&pc=MTA | 2023826-2015-01020 | 2014-12-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/089 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO INADEQUATE VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016333&pc=MTA | 2023826-2015-01051 | 2015-01-13 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS WITH -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LENS ROTATION ASSOCIATED WITH LOW VAULT WAS OBSERVED. THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UCVA WAS 20/15. Manufacturer Narrative: (B)(4). METHOD: LENS WORK ORDER SEARCH . EVALUATION: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4633022&pc=MTA | 2023826-2015-00388 | 2015-01-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-25 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM ICM11.5 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012 AND LOW VAULT WAS NOTED. ROTATION OF THE LENS WAS ALSO NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015. Manufacturer Narrative: (B)(4) - (SECONDARY SURGERY); (NEW INCISION); SIZE INCORRECT FOR PATIENT; (VAULTING); (EXPLANTED; UNINTENDED MOVEMENT OF THE DEVICE. DEVICE EVALUATED BY THE MANUFACTURER? NO. (LENS EVALUATION ANTICIPATED). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. A DEVICE HISTORY REVIEW WAS PERFORMED AND DETERMINED NOTHING IN THE MANUFACTURING OF THE LENS WAS THE ROOT CAUSE OF THE EVENT . MEDICAL REVIEW-CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN (B)(4) OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC, POSITION,ECT.). BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, DEVICE HISTORY AND MEDICAL REVIEW, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4841450&pc=MTA | 2023826-2015-00675 | 2015-02-02 | Injury | STAAR SURGICAL COMPANY | 2015-06-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS ROTATED AND WAS REPOSITIONED. THE SURGEON PLANS ON EXCHANGING THIS LENS FOR A LONGER LENGTH DUE TO LOW VAULT. THE PATIENT'S LAST BCVA ON (B)(6) 2015 WAS 20/25. Manufacturer Narrative: (B)(4) (NEW INCISION), (SECONDARY SURGERY), (REPOSITIONING OF LENS). (B)(4) - SIZE INCORRECT FOR PATIENT - (LENS EXPLANTED), (VAULTING), UNINTENDED MOVEMENT. LENS REMAINS IMPLANTED EVALUATION METHOD - WORK ORDER SEARCH. RESULTS; A LENS WORK ORDER SEARCH WAS PERFORMED AND THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT INFORMATION AND WORK ORDER SEARCH, WE ARE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4) Manufacturer Narrative: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN (B)(4) OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC, POSITION ECT.). ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). THE MOST PROBABLE CAUSE OF THIS EVENT WAS LOW VAULTING SECONDARY TO A SHORT LENS. BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. THE LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842195&pc=MTA | 2023826-2015-00680 | 2015-02-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS ROTATED AND LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SMALLER SIZE ON (B)(6) 2015 AND THIS RESOLVED THE ISSUE. SEE MFR # 2023826-2015-00679 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT INFORMATION AND WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4816519&pc=MTA | 2023826-2015-00638 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-06-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT AND ROTATION OF THE LENS WERE NOTED, THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. REFERENCE MFR. REPORT # 2023826-2015-00639 FOR PATIENT'S LEFT EYE. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH A BROKEN HAPTIC, HOWEVER, THERE WAS NO VISIBLE DAMAGE OBSERVED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITON, ETC.). CONCLUSION : BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4921703&pc=MTA | 2023826-2015-00819 | 2015-06-03 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5386383&pc=MTA | 2023826-2016-00055 | 2015-06-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -22.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225078&pc=MTA | 2023826-2015-01514 | 2015-07-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/X102 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL WITH PERFECT VAULTING OF THE LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225082&pc=MTA | 2023826-2015-01516 | 2015-08-05 | Injury | STAAR SURGICAL COMPANY | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/X071 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225083&pc=MTA | 2023826-2015-01515 | 2015-08-19 | Injury | STAAR SURGICAL COMPANY | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/X131 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972077&pc=MTA | 2023826-2016-01314 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+2.0/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT, WITH LENS ROTATION AND A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/25. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5904275&pc=MTA | 2023826-2016-00888 | 2015-10-31 | Injury | STAAR SURGICAL COMPANY | 2016-08-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, 12.5/+2.0/88 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS REPOSITIONED, THEN EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(6). DEVICE MANUFACTURE DATE - CORRECTED FROM: 06/30/2018; TO: 07/29/2015. DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC BENT AND FOREIGN MATERIAL ON THE LENS SURFACE THAT WAS NOT POSSIBLE TO SPECIFY. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5453090&pc=MTA | 2023826-2016-00174 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2016-02-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+2.0/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ROTATION AND REFRACTIVE CHANGE OVERTIME . THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919722&pc=MTA | 2023826-2016-01199 | 2016-01-28 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS DISLOCATION/SUBLUXATION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND FOREIGN MATERIAL (DRIED DISCOLORED MATERIAL) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS MODEL IS CONTRAINDICATED IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH(ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM AND IN A PATIENT UNDER THE AGE OF 21. THE PATIENT WAS 19 (UNDER 21) AND HAD AN ACD OF 2.80MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668354&pc=MTA | 2023826-2016-00699 | 2016-02-01 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/+3.50/X141 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. LOW VAULT WITH LENS ROTATION WAS NOTED. THE PHYSICIAN TRIED TO REPOSITION THE LENS, BUT THE LENS MOVED AGAIN ON (B)(6) 2016. THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5665324&pc=MTA | 2023826-2016-00740 | 2016-03-10 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TICM125V4 IMPLANTABLE COLLAMER LENS, -17.5/+4.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS HAD DISLOCATED/SUBLUXED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5733031&pc=MTA | 2023826-2016-00830 | 2016-03-23 | Injury | STAAR SURGICAL COMPANY | 2016-06-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, LENS ROTATION AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED BENT AND DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN WITH FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694863&pc=MTA | 2023826-2016-00755 | 2016-04-25 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+4.5/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT, ROTATION, AND LENS DISLOCATION/SUBLUXATION. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5850883&pc=MTA | 2023826-2016-01028 | 2016-05-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+4.5/084 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. CORRECTED STATEMENT-THE PROBLEM WAS NOT RESOLVED WITH THE LENS EXCHANGE. SEE MFR REPORT# 2023826-2018-01213 FOR CASE CONTINUATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5850885&pc=MTA | 2023826-2016-01027 | 2016-05-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+4.5/102 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC TORN/BROKEN/BENT/DEFORMED. CLAIM #(B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. CORRECTED INFORMATION-THE PROBLEM WAS NOT RESOLVED WITH THE LENS EXCHANGE. SEE MFR REPORT# 2023826-2018-01211 FOR CASE CONTINUATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873703&pc=MTA | 2023826-2016-01063 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -1.5/+6.0./X080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS FOUND ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572539&pc=MTA | 2023826-2016-00411 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -18.0/+6.0/X141 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED INFORMATION: THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED ON (B)(6) 2015 FOR A LONGER LENS AND THE PROBLEMS WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5648953&pc=MTA | 2023826-2016-00502 | 2016-02-15 | Injury | STAAR SURGICAL COMPANY | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Missing Value Reason | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6-8.0/+1.50/X171 IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED LOW VAULT, LENS ROTATION AND GLARE AND HALOS. THE LENS WAS EXPLANTED ON (B)(6) 2016. A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER 45 YEARS OF AGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5583365&pc=MTA | 2023826-2016-00395 | 2006-03-29 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 -08.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2006. A REFRACTIVE CHANGE OVERTIME WAS OBSERVED ON (B)(6) 2016. PLANS TO EXCHANGE LENS IS PENDING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CONCLUSIONS CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT INFORMATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4555180&pc=MTA | 2023826-2015-00201 | 2014-12-25 | Injury | STAAR SURGICAL COMPANY | 2015-02-27 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS HAD A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6340150&pc=MTA | 2023826-2017-00240 | 2015-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-02-17 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.0/059 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD). THIS WAS PERFORMED ON (B)(6) 2015. ON (B)(6) 2015, THE LENS WAS THEN EXPLANTED. Manufacturer Narrative: LENS WAS EXCHANGED FOR THE SAME SIZE AND MODEL, SIMILAR DIOPTER LENS ON (B)(6) 2015. AS OF THE DATE OF SUPPLEMENTAL MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. (B)(4). METHOD CODE: WORK ORDER SEARCH PERFORMED. NO SIMILAR COMPLAINTS WERE FOUND FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5106749&pc=MTA | 2023826-2015-01183 | 2015-08-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/X102 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE THE LENS REMAINS IMPLANTED. A SUPPLEMENTAL WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS CODE(S): (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INVESTIGATION PERFORMED, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. EVALUATION CONCLUSION CODE(S): - BASED ON THE COMPLAINT HISTORY, DEVICE HISTORY RECORD REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5908631&pc=MTA | 2023826-2016-00370 | 2015-12-03 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, UNKNOWN DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT DEVELOPED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: DIOPTER OF THE LENS IS -10/+2,5/70. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7558981&pc=MTA | 2023826-2018-00835 | 2018-03-12 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), DURING THE SAME SURGERY DUE TO A CAPSULE TEAR. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF LENS OPTIC AND ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS REVIEWED FOR SHARP/ROUGH EDGES AND NO EVIDENCE OF SHARP/ROUGH EDGES WERE FOUND. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631809&pc=MTA | 2023826-2019-00845 | 2019-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | P030016 | 1 | Event Description: REPORTEDLY, A 13.2MM, MICL 13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS "TOO LONG". THE LENS WAS REMOVED AND SCHEDULED EXCHANGE OF A SHORTER LENGTH LENS. NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT PROBLEM (E.G. MALFUNCTION) IS NOT APPLICABLE IN ORIGINAL MDR; CORRECTED TO ADVERSE EVENT BECAUSE LENS WAS EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: MALFUNCTION REPORTED IN ORIGINAL MDR IS NOT APPLICABLE; CORRECTED TO SERIOUS INJURY. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY TAPED TO CARTRIDGE CASE. VISUAL INSPECTION: LENS HAPTIC IS BENT WITH RESIDUE. NO VISIBLE DAMAGE TO SFC-45 CARTRIDGE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819379&pc=MTA | 2023826-2020-00448 | 2020-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2MM, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED, A REPLACEMENT LENS (SAME MODEL,LENGTH AND DIOPTER) WAS LATER IMPLANTED. THE REPORTER STATED " THE SURGEON THINKS THAT THE PROBLEM IS CAUSED BY LENS ERROR. HE EXCHANGED WITH THE SAME LENS AND PROBLEM HAS RESOLVED". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL DEBRIS IN AMICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC ORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4636965&pc=MTA | 2023826-2015-00394 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Code Available | UNK | 1 | Event Description: THE PATIENT REPORTED THROUGH (B)(6), SHE HAD AN ICL IMPLANTABLE COLLAMER LENS IMPLANTED IN HER LEFT EYE (OS) 4 WEEKS AGO AND HER VISION WAS GREAT (-19.5). SHE WILL HAVE PRK SURGERY IN THE NEXT COUPLE OF WEEKS. THE PATIENT'S CONCERN WAS HER LEFT EYE WAS ACHY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE PATIENT REPORTED HAVING A HEADACHE AFTER SURGERY. THE PATIENT HAD A HEAVY FEELING OVER THE EYE AND A PULLING SENSATION. THE PATIENT REPORTED PRESSURE, SWELLING AND NUMBNESS AROUND THE EYE. THE PATIENT'S SURGEON HAS INDICATED THE LENS LOOKS TEXTBOOK AND THE EYE IS FINE. Manufacturer Narrative: (B)(4). LENS IMPLANTED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4767218&pc=MTA | 2023826-2015-00539 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Code Available | UNK | 1 | Event Description: THE PATIENT REPORTED HAD IRIDOTOMIES DONE AND THE GLARE IS REALLY BAD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GLARE COULD BE CAUSED BY MALPOSITION OR OVERSIZE OF AN IRIDOTOMY PERFORMED BEFORE AN ICL LENS IMPLANTATION. BUT MANY FACTORS SUCH AS LENS DISLOCATION, ANGLE CLOSURE GLAUCOMA, CORNEA EDEMA OR DYSTROPHY AND CATARACT ARE ASSOCIATED WITH GLARE. GIVEN THE REPORTED LIMITED INFORMATION, THE CAUSE OF THE EVENT IS UNKNOWN. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6542763&pc=MTA | 2023826-2017-00765 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND EXPLANTED AN MICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S EYE. THE REPORTER INDICATED THIS WAS THE 1ST OF 2 LENSES USED FOR THIS PATIENT - SEE MFR REPORT # 2023826-2017-00754. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577108&pc=MTA | 2023826-2017-00850 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN. THE REPORTER INDICATED THE LENS WAS EXPLANTED AND THE PATIENT WAS GOING TO HAVE AN ARTIFICIAL IRIS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT DATE: UNK. EXPIRATION DATE - UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. MFR DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7509058&pc=MTA | 2023826-2018-00691 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | P030016 | 1 | Event Description: A 13.2MM, MICL13.2,-4.0 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS RETURNED UNDAMAGED. NOTE ON THE LENS PACKAGING BOX INDICATED "(OS)OPENED, IMPLANTED, EXCHANGED". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: DATE OF THIS REPORT -IT IS CORRECTED TO 12-APRIL-2018 IN THE INITIAL MDR. RECEIVED BY MFR -IT IS CORRECTED TO 12-APRIL-2018 IN THE INITIAL MDR. RETURNED TO MANUFACTURER : 30-MAR-2018. (B)(4). Manufacturer Narrative: (B)(4). LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720893&pc=MTA | 2023826-2018-01099 |  | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | NA | 1 | Event Description: A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+3.0/017 (SPHERE/CYLINDER/AXIS), WAS REPORTED AS BEING TAKEN OUT (REMOVED) FROM THE PATIENTS LEFT EYE (OS) AND REPLACED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS REMOVED AND REPLACED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7789908&pc=MTA | 2023826-2018-01236 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2 MM MICL13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION- LENS WAS RECEIVED DRY, IN A SMALL ENVELOPE. VISUAL INSPECTION FOUND A TORN HAPTIC AND RESIDUE ON THE LENS SURFACE, TORN HAPTIC PIECE IS MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8278110&pc=MTA | 2023826-2019-00178 |  | Injury | STAAR SURGICAL COMPANY | 2019-01-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | N/A | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT CANCELLED THE SURGERY AFTER THE EXPLANTED OD LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA THIS IS A DUPLICATE FILE OF MDR # 2023826-2019-00177. PLEASE REFER TO MDR # 2023826-2019-00177 FOR FURTHER DETAILS. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8425520&pc=MTA | 2023826-2019-00415 |  | Injury | Unknown Manufacturer | 2019-03-15 | MTA | UNK | Insufficient Information | No Code Available | N/A | 1 | Event Description: THE REPORTER STATED " AN UNSPECIFIED INCIDENT IN ONE EYE OR BOTH EYES OF PATIENT "GIGLIA". THE LENSES HAD BEEN IMPLANTED AT THIS CLINIC IN (B)(6)/(B)(6) 2017. PROBABLY AT LEAST ONE LENS WAS REMOVED AT THE EYE CLINIC OF (B)(6) SHORTLY AFTER THE IMPLANTATION. UNFORTUNATELY WE HAVE NO MORE INFORMATION YET. WE WILL KEEP YOU UPDATED". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK, NOT MTA, NO SERIAL NUMBER REPORTED. MANUFACTURER NAME, CITY, AND STATE: UNK, NO SERIAL NUMBER REPORTED. EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: (B)(6)/(B)(6) 2017. EXPLANT DATE: UNK. MANUFACTURING SITE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8425690&pc=MTA | 2023826-2019-00416 |  | Injury | Unknown Manufacturer | 2019-03-15 | MTA | UNK | Insufficient Information | No Code Available | N/A | 1 | Event Description: THE REPORTER STATED " AN UNSPECIFIED INCIDENT IN ONE EYE OR BOTH EYES OF PATIENT "GIGLIA". THE LENSES HAD BEEN IMPLANTED AT THIS CLINIC IN NOVEMBER/DECEMBER 2017. PROBABLY AT LEAST ONE LENS WAS REMOVED AT THE EYE CLINIC OF THE UNIVERSITY TUEBINGEN SHORTLY AFTER THE IMPLANTATION. UNFORTUNATELY WE HAVE NO MORE INFORMATION YET. WE WILL KEEP YOU UPDATED". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (PRODUCT CODE): UNK NOT MTA, NO SERIAL NUMBER REPORTED (ESUBMITTER REQUIRES AN ENTRY FOR THIS SECTION, CANNOT BE UNK OR LEFT BLANK) UNK, NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (IMPLANT DATE): NOVEMBER/DECEMBER 2017 UNK. (MANUFACTURING SITE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671671&pc=MTA | 2023826-2019-01022 | 2019-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Material Discolored | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXPLANTED. THE SURGEON REPORTS OPACITY AND YELLOWING OF THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND THE LENS WAS DISCOLORED (YELLOW). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8553235&pc=MTA | 2023826-2019-00462 | 2019-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT AND SURGEON ARE HAPPY. CAUSE OF EVENT WAS REPORTED AS USER ERROR, REFRACTION WAS CAUSED THROUGH A WRONG CALCULATION MEASUREMENT AND MISCALCULATION THROUGH THE DOCTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED DUE TO A REPORTED REFRACTIVE SURPRISE. USER FACILITY CHECKED IN ERROR. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. LENGTH MEASUREMENT AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : WORK ORDER SEARCH WAS PERFORMED AND NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8697864&pc=MTA | 2023826-2019-01083 | 2019-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT IS HAPPY. CAUSE OF EVENT WAS REPORTED AS USER ERROR, REFRACTION WAS CAUSED THROUGH A WRONG CALCULATION MEASUREMENT AND MISCALCULATION THROUGH THE DOCTOR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9335809&pc=MTA | 2023826-2019-02187 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_13.7, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BLOOD-LIKE SUBSTANCE; DEBRIS (THREADISH). DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073825&pc=MTA | 2023826-2018-01676 | 2018-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2018. IT WAS REPORTED THAT THIS EYE HAS LENS COLOBOMA AND THE PUPIL WAS MID-DILATED AT LAST VISIT. Manufacturer Narrative: CHANGED TO ADVERSE EVENT. EVENT - THE PATIENT DID NOT EXPERIENCE REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401679&pc=MTA | 2023826-2020-01703 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME LENGTH TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSION AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9323812&pc=MTA | 2023826-2019-02145 | 2019-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A TORIC LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 OF -16.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH SPHERICAL LENS OF -15.00 DIOPTER ON (B)(6) 2019. THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH STATEMENT IN PREVIOUS MDR SHOULD BE CORRECTED TO "ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND." (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149327&pc=MTA | 2023826-2018-01804 | 2018-11-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+2.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE "PATIENT IS DISSATISFIED WITH THE MYOPIA REMAINING". THE LENS REMAINS IMPLANTED. SEE MFR REPORT# 2023826-2018-01801 FOR PREVIOUSLY IMPLANTED LENS. Manufacturer Narrative: DEVICE CODE 1401: PREVIOUS CODE IN INITIAL MDR NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9272512&pc=MTA | 2023826-2019-02017 | 2019-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) AND BLURRED VISION WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY "AT EXAMINATION AFTER 1.5 HOUR FROM THE SURGERY, FLARE IN THE FRONT CHAMBER AND FIBRIN AROUND ICL HOLE. 0.1% FLUMETHOLON WAS PRESCRIBED. ONE DAY LATER, MASS OF FIBRIN WAS OBSERVED, PRESCRIPTION CHANGED TO RINDERON" AND THE EYE IS "GETTING BETTER NOW." PATIENT IS CONTINUOUSLY MONITORED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (MEDICATION: FLUMETHOLON, RINDERON). PATIENT CODE 3191: NO CODE AVAILABLE (MASS OF FIBRIN). DEVICE CODE 1494: OFF-LABEL USE (PATIENT'S AGE IS OVER 45 YEARS OLD AT TIME OF IMPLANT). DEVICE CODE: 2993 SHOULD BE CORRECTED TO 1401. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9272513&pc=MTA | 2023826-2019-02018 | 2019-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) AND BLURRED VISION WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY "AT EXAMINATION AFTER 1.5 HOUR FROM THE SURGERY, FLARE IN THE FRONT CHAMBER AND FIBRIN AROUND ICL HOLE. LENS REMAINS IMPLANTED. 0.1% FLUMETHOLON WAS PRESCRIBED. ONE DAY LATER, MASS OF FIBRIN WAS OBSERVED, PRESCRIPTION CHANGED TO RINDERON" AND THE EYE IS "GETTING BETTER NOW." PATIENT IS CONTINUOUSLY MONITORED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (MEDICATION: FLUMETHOLON, RINDERON). PATIENT CODE 3191: NO CODE AVAILABLE (MASS OF FIBRIN). DEVICE CODE 1494: OFF-LABEL USE (PATIENT'S AGE IS OVER 45 YEARS OLD AT TIME OF IMPLANT). DEVICE CODE: 2993 SHOULD BE CORRECTED TO 1401. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720997&pc=MTA | 2023826-2018-01107 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8392893&pc=MTA | 2023826-2019-00362 | 2019-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS THE PATIENT WAS "DISSATISFIED WITH THE QUALITY OF VISION AFTER ICL IMPLANT." THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. ON (B)(6) 2019 THE LENS WAS EXPLANTED. THE PATIENT IS "OKAY" AFTER EXPLANT. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8392898&pc=MTA | 2023826-2019-00359 | 2019-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS THE PATIENT WAS "DISSATISFIED WITH THE QUALITY OF VISION AFTER ICL IMPLANT." THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. ON (B)(6) 2019 THE LENS WAS EXPLANTED. THE PATIENT IS "OKAY" AFTER EXPLANT. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552451&pc=MTA | 2023826-2019-00650 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ANISOPIA, AS THE ICL IMPLANTED IN THE FELLOW EYE HAD TO BE REMOVED (REFERENCE MFR# 2023826-2019-00655). THE CAUSE IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372520&pc=MTA | 2023826-2019-00324 | 2018-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372523&pc=MTA | 2023826-2019-00325 | 2018-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATED BY MFR : DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824522&pc=MTA | 2023826-2020-00354 | 2018-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652058&pc=MTA | 2023826-2019-00919 | 2018-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237923&pc=MTA | 2023826-2019-00080 | 2018-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/121 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE ON LENS SURFACE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CORRECTED DATA DEVICE CODE- (B)(4). ADDED TO INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652120&pc=MTA | 2023826-2019-00918 | 2018-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TEAR. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918167&pc=MTA | 2023826-2019-01526 | 2018-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND TEARS ON THE LENS HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237944&pc=MTA | 2023826-2019-00081 | 2018-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CORRECTED DATA DEVICE CODE- (B)(4). COMPLAINT-(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9515900&pc=MTA | 2023826-2019-02312 | 2018-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISISBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645474&pc=MTA | 2023826-2019-00862 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8296803&pc=MTA | 2023826-2019-00138 | 2018-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.0/086 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR IN LENS HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843597&pc=MTA | 2023826-2019-01335 | 2018-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8834101&pc=MTA | 2023826-2019-01345 | 2018-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICMO12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS DRY WITH CLEAR SURGICAL RESIDUE ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC IN TWO LOCATIONS AND RESIDUE ON THE LENS. DEVICE CODE 1494: OFF LABEL USE: ACD<3.0MM; AGE OF PATIENT < 21 YEARS AT THE TIME OF IMPLANTATION. PATIENT CODE 3191: NO CODE AVAILABLE - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. PATIENT ALSO EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES PRIOR TO THE EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813861&pc=MTA | 2023826-2019-01282 | 2018-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813860&pc=MTA | 2023826-2019-01283 | 2018-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT/DEFORMED AND RESIDUE ON THE LENS. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8401460&pc=MTA | 2023826-2019-00375 | 2018-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICMO13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966513&pc=MTA | 2023826-2019-01643 | 2018-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836129&pc=MTA | 2023826-2019-01251 | 2018-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372534&pc=MTA | 2023826-2019-00327 | 2018-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372536&pc=MTA | 2023826-2019-00331 | 2018-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER- CORRECTED FROM "26-MAR-2019" TO "20-MAR-2019" FOR MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564628&pc=MTA | 2023826-2019-00681 | 2018-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. (B)(4): PATIENT'S ACD WAS LESS THAN 3.0MM CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8805175&pc=MTA | 2023826-2019-01256 | 2018-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8805170&pc=MTA | 2023826-2019-01255 | 2018-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTICS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957638&pc=MTA | 2023826-2020-00792 | 2018-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1, -05.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CORRECTED DATA: B3-DATE OF EVENT IN INITIAL MDR SHOULD BE CORRECTED TO UNK. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307842&pc=MTA | 2023826-2019-00224 | 2018-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/DEFORMED; WITH RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: YES;DEVICE EVALUATED BY MANUFACTURER VISUAL INSPECTION FOUND HAPTIC TORN STATEMENT IN PREVIOUS MDR IS NOT APPLICABLE. VISUAL INSPECTION FOUND HAPTIC DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8421242&pc=MTA | 2023826-2019-00348 | 2018-10-31 | Injury | STAAR SURGICAL COMPANY | 2019-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.5/094 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652153&pc=MTA | 2023826-2019-00944 | 2018-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DEVICE CODE 1494: PATIENT WAS OUTSIDE THE INDICATED AGE RANGE ON THE DIRECTED INDICATIONS FOR USE AT THE TIME OF IMPLANTATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772270&pc=MTA | 2023826-2019-01135 | 2018-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317884&pc=MTA | 2023826-2019-02137 | 2018-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8805182&pc=MTA | 2023826-2019-01257 | 2018-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A TEAR ON THE LENS HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: "THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND PROBLEM WAS RESOLVED." SHOULD BE CORRECTED TO THE FOLLOWING: "THE LENS WAS EXCHANGED ON (B)(6) 2018 FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED." DEVICE CODE 1494: OFF LABEL USE (ACD<3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429451&pc=MTA | 2023826-2020-01789 | 2018-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9331858&pc=MTA | 2023826-2019-02109 | 2018-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H6: DEVICE CODE 1494: OFF-LABEL USE- ACD <3.00MM SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9180047&pc=MTA | 2023826-2019-01938 | 2018-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8289267&pc=MTA | 2023826-2019-00133 | 2018-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND RESIDUE ON THE LENS. CORRECTED DATA: CONCLUSION CODE 4315: CODE SHOULD BE ADDED TO ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203023&pc=MTA | 2023826-2019-01919 | 2018-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF LABEL USE (AGE < 21 YEARS OLD). DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813864&pc=MTA | 2023826-2019-01287 | 2018-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. REFERENCE MFR REPORT # 2023826-2019-01288 FOR CLAIM CONTINUATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: "DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS" SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: INJECTOR MODEL-MSI-PF; LOT#-UNK SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9361665&pc=MTA | 2023826-2019-02153 | 2019-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: DATE OF THIS REPORT: (B)(6) 2019 SHOULD BE CORRECTED TO 04-NOV-2019 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9361668&pc=MTA | 2023826-2019-02152 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER: (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412428&pc=MTA | 2023826-2019-02289 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9838139&pc=MTA | 2023826-2020-00438 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE DID NOT RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED DRY INSIDE A CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND FIBERS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552445&pc=MTA | 2023826-2019-00651 | 2019-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837787&pc=MTA | 2023826-2020-00523 | 2019-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9838381&pc=MTA | 2023826-2020-00437 | 2019-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8851472&pc=MTA | 2023826-2019-01272 | 2019-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO 12.1, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: 1494: OFF-LABEL USE; ACD= 2.92MM WHICH IS BELOW THE APPROVED INDICATIONS FOR USE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9221382&pc=MTA | 2023826-2019-01971 | 2019-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401658&pc=MTA | 2023826-2020-01690 | 2019-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -05.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9205119&pc=MTA | 2023826-2019-01934 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9416995&pc=MTA | 2023826-2019-02210 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND DEBRIS ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957622&pc=MTA | 2023826-2020-00787 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957624&pc=MTA | 2023826-2020-00786 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT CONDITION ON (B)(6) 2019 "IN GOOD CONDITION WITHOUT OTHER COMPLICATIONS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664750&pc=MTA | 2023826-2019-00991 | 2019-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DEVICE CODE 1494: OFF LABEL USE: ACD<3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8392895&pc=MTA | 2023826-2019-00360 | 2019-01-29 | Injury | STAAR SURGICAL COMPANY | 2019-03-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/077 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS A PATIENT-RELATED FACTOR. STATUS OF THE EYE IS REPORTED AS "ANTERIOR CHAMBER NARROWING." Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403837&pc=MTA | 2023826-2019-00333 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODE- 4111 ADDED TO INITIAL MDR. PREVIOUS METHOD CODE NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403839&pc=MTA | 2023826-2019-00334 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -08.00, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497495&pc=MTA | 2023826-2019-00574 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8421273&pc=MTA | 2023826-2019-00349 | 2019-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/097 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660740&pc=MTA | 2023826-2019-00957 | 2019-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DEVICE CODE 1494: OFF-LABEL USE (ACD<3.0MM). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918901&pc=MTA | 2023826-2019-01529 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICH12.6, 6.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS; RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645635&pc=MTA | 2023826-2019-00899 | 2019-02-19 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652102&pc=MTA | 2023826-2019-00921 | 2019-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774156&pc=MTA | 2023826-2019-01190 | 2019-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, PIGMENT DISPERSION, AND IRIS ATROPHY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645470&pc=MTA | 2023826-2019-00901 | 2019-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631516&pc=MTA | 2023826-2019-00847 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894306&pc=MTA | 2023826-2019-01439 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894310&pc=MTA | 2023826-2019-01440 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9226076&pc=MTA | 2023826-2019-02024 | 2019-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM, VICMO 12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE: ACD < 3.00MM SHOULD BE ADDED IN THE INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETUNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717755&pc=MTA | 2023826-2020-00254 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9967625&pc=MTA | 2023826-2020-00913 | 2019-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645468&pc=MTA | 2023826-2019-00860 | 2019-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518148&pc=MTA | 2023826-2019-00559 | 2019-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518165&pc=MTA | 2023826-2019-00562 | 2019-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. CLEAR SURGICAL RESIDUE WAS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC TORN WITH RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518167&pc=MTA | 2023826-2019-00564 | 2019-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518168&pc=MTA | 2023826-2019-00563 | 2019-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824648&pc=MTA | 2023826-2020-00494 | 2019-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824660&pc=MTA | 2023826-2020-00495 | 2019-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373328&pc=MTA | 2023826-2019-02225 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER THE SURGEON DOES NOTE THE PATIENT'S CILIARY BODY HAS "POOR SHAPE." Manufacturer Narrative: B4 - CORRECTED TO 04-NOV-2019 IN INITIAL MDR. G4 - CORRECTED TO 04-NOV-2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: SEX, WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). DEVICE CODE 1494: OFF-LABEL USE (PRE-EXISTING CONDITION OF PRIMARY OPEN ANGLE OR NARROW ANGLE GLAUCOMA). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983360&pc=MTA | 2023826-2020-00937 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652105&pc=MTA | 2023826-2019-00922 | 2019-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664344&pc=MTA | 2023826-2019-00981 | 2019-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO 13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664349&pc=MTA | 2023826-2019-00982 | 2019-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO 13.2, -03.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966521&pc=MTA | 2023826-2019-01638 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9216441&pc=MTA | 2023826-2019-01985 | 2019-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICMO 12.6, 13.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATED BY MFR:"HAPTIC TORN AND BENT" SHOULD BE CORRECTED TO "HAPTIC TORN AND BROKEN" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 1494: OFF-LABEL: ACD < 3.00MM SHOULD BE ADDED IN THE INITIAL MDR. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9216486&pc=MTA | 2023826-2019-01986 | 2019-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICMO 12.6, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 1494: OFF- LABEL : ACD < 3.00MM SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8957681&pc=MTA | 2023826-2019-01537 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION CORRECTED TO: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795477&pc=MTA | 2023826-2019-01246 | 2019-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957643&pc=MTA | 2023826-2020-00790 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -05.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429446&pc=MTA | 2023826-2020-01796 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA. B5- THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8663869&pc=MTA | 2023826-2019-00989 | 2019-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE CODE: 1494 OFF-LABEL USE SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8917676&pc=MTA | 2023826-2019-01528 | 2019-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9225977&pc=MTA | 2023826-2019-02025 | 2019-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICMO 12.6, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO ENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: PATIENT CODE 1494- OFF-LABEL USE: ACD < 3.00MM SHOULD BE ADDED IN THE INTIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9397392&pc=MTA | 2023826-2019-02257 | 2019-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: B4 - CORRECTED TO (B)(6)-2019 IN INITIAL MDR. G4 - CORRECTED TO (B)(6)2019 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). DEVICE CODE 1494, OFF-LABEL USE, ACD <3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8967214&pc=MTA | 2023826-2019-01595 | 2019-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE INCISION WAS ENLARGED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON ENLARGED THE INCISION TO 3.0MM TO REMOVE THE INITIAL LENS AND NO SUTURES WERE REQUIRED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CONCLUSION CODE: 24 SHOULD BE REMOVED FROM SUPPLEMENTAL #1 MEDWATCH REPORT AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8686767&pc=MTA | 2023826-2019-01049 | 2019-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -05.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO LOW VAULT. REPORTEDLY, "THE FIRST DAY AFTER SURGERY, THERE WAS ALMOST NO VAULT. THE SECOND DAY WAS ABOUT 100UM. AFTER COMMUNICATION WITH SURGEON, THE PATIENT DECIDED TO GIVE UP THE SURGERY." Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125528&pc=MTA | 2023826-2019-01774 | 2019-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION:LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS, AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988461&pc=MTA | 2023826-2020-00958 | 2019-05-03 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664335&pc=MTA | 2023826-2019-00938 | 2019-05-05 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913412&pc=MTA | 2023826-2020-00728 | 2019-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#(B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8764088&pc=MTA | 2023826-2019-01106 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -10.50 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664343&pc=MTA | 2023826-2019-00939 | 2019-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910972&pc=MTA | 2023826-2020-00684 | 2019-05-10 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON 01-MAY-2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8746388&pc=MTA | 2023826-2019-01157 | 2019-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISIBLE INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. DEVICE CODE: 1494 - OFF LABEL USE (ACD < 3.0MM). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8751290&pc=MTA | 2023826-2019-01170 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9176974&pc=MTA | 2023826-2019-01906 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532329&pc=MTA | 2023826-2019-02460 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774165&pc=MTA | 2023826-2019-01196 | 2019-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795455&pc=MTA | 2023826-2019-01238 | 2019-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10049167&pc=MTA | 2023826-2020-01067 | 2019-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISIAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10049174&pc=MTA | 2023826-2020-01066 | 2019-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306681&pc=MTA | 2023826-2019-01998 | 2019-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8967223&pc=MTA | 2023826-2019-01599 | 2019-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9433793&pc=MTA | 2023826-2019-02323 | 2019-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA 510(K): THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317886&pc=MTA | 2023826-2019-02141 | 2019-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8957680&pc=MTA | 2023826-2019-01538 | 2019-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISABLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412464&pc=MTA | 2023826-2019-02278 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUTION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401613&pc=MTA | 2023826-2020-01681 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION:LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401615&pc=MTA | 2023826-2020-01682 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454412&pc=MTA | 2023826-2019-02376 | 2019-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- DEVICE CODE 1494: ACD <3.00MM SHOULD BE ADDED INTO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # 715029
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8931071&pc=MTA | 2023826-2019-01557 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -8.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. PER THE REPORTER, "IOP IS RETAINED NORMAL AFTER EXCHANGE OF THE LENS TO 12.6 DIAMETER." Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9447915&pc=MTA | 2023826-2019-02371 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8971743&pc=MTA | 2023826-2019-01608 | 2019-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. UDI CORRECTED TO N/A IN INITIAL MDR. DEVICE CODE 1494 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. METHOD CODE 4111 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9173468&pc=MTA | 2023826-2019-01889 | 2019-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND TEARS ON THE LENS HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186703&pc=MTA | 2023826-2019-01888 | 2019-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317871&pc=MTA | 2023826-2019-02140 | 2019-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND "INACCURATE DATA MEASUREMENT." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913426&pc=MTA | 2023826-2020-00723 | 2019-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373326&pc=MTA | 2023826-2019-02226 | 2019-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B4- "03-NOV-2019" SHOULD BE CORRECTED TO "04-NOV-2019" IN THE INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8851483&pc=MTA | 2023826-2019-01351 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880969&pc=MTA | 2023826-2020-00658 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8937552&pc=MTA | 2023826-2019-01552 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9139938&pc=MTA | 2023826-2019-01822 | 2019-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869320&pc=MTA | 2023826-2020-00624 | 2019-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -2.75 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869328&pc=MTA | 2023826-2020-00623 | 2019-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -2.75 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983371&pc=MTA | 2023826-2020-00931 | 2019-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412442&pc=MTA | 2023826-2019-02285 | 2019-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10049217&pc=MTA | 2023826-2020-01064 | 2019-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9369707&pc=MTA | 2023826-2019-02250 | 2019-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- DEVICE CODE 1494: OFF-LABEL USE: AGE < 21 YEARS OLD SHOULD BE ADDED TO THE INITIAL MDR. (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9433775&pc=MTA | 2023826-2019-02325 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910955&pc=MTA | 2023826-2020-00685 | 2019-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559335&pc=MTA | 2023826-2020-00043 | 2019-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559337&pc=MTA | 2023826-2020-00044 | 2019-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889408&pc=MTA | 2023826-2020-00606 | 2019-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936255&pc=MTA | 2023826-2020-00767 | 2019-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY STATED THAT THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. ADDITIONAL INFORMATION-THE LENS WAS EXCHANGED WITH A 13.2MM LENS AND THE PROBLEM WAS NOT RESOLVED. REFERENCE MFR REPORT# 2023826-2020-01004 FOR FINAL EXCHANGED LENS AND PROBLEM RESOLUTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587402&pc=MTA | 2023826-2020-00111 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869319&pc=MTA | 2023826-2020-00615 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9031822&pc=MTA | 2023826-2019-01664 | 2019-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: "THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019" SHOULD BE CORRECTED TO "THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587410&pc=MTA | 2023826-2020-00114 | 2019-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427886&pc=MTA | 2023826-2019-02320 | 2019-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9597950&pc=MTA | 2023826-2020-00125 | 2019-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9175898&pc=MTA | 2023826-2019-01831 | 2019-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE OPTIC AND HAPTIC WERE TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9447906&pc=MTA | 2023826-2019-02366 | 2019-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- DEVICE CODE 1494: OFF- LABEL USE- ACD< 3.00MM; AGE < 21 SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUA INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412463&pc=MTA | 2023826-2019-02288 | 2019-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -17.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8983236&pc=MTA | 2023826-2019-01689 | 2019-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE SURGEON EXPLANTED THE LENS BECAUSE IT IS "TOO BIG." REPORTEDLY, NO PATIENT INJURY OCCURRED AND NO REPLACEMENT LENS WAS IMPLANTED. Manufacturer Narrative: RETURNED TO MANUFACTURER DATE CORRECTED TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A STERILE POUCH WRAPPED IN GAUZE. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE LENS WAS RETURNED CURVED IN A COIL. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335341&pc=MTA | 2023826-2020-01538 | 2019-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936294&pc=MTA | 2023826-2020-00758 | 2019-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9417015&pc=MTA | 2023826-2019-02205 | 2019-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9219462&pc=MTA | 2023826-2019-01988 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICMO 13.2, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISBLE DAMAGE TO LENS. CORRECTED DATA: PREVIOUSLLY STATED VICMO 13.2 IS CORRECTED TO VICMO 12.1. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9199822&pc=MTA | 2023826-2019-01929 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308835&pc=MTA | 2023826-2020-01522 | 2019-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889427&pc=MTA | 2023826-2020-00604 | 2019-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9187334&pc=MTA | 2023826-2019-01864 | 2019-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: -UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308787&pc=MTA | 2023826-2020-01517 | 2019-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608208&pc=MTA | 2023826-2020-00130 | 2019-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398523&pc=MTA | 2023826-2020-01748 | 2019-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9639477&pc=MTA | 2023826-2020-00189 | 2019-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION. LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND A YELLOWISH APPERANCE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WIHTIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9899346&pc=MTA | 2023826-2020-00649 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9977660&pc=MTA | 2023826-2020-00920 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION. H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837546&pc=MTA | 2023826-2020-00550 | 2019-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9354900&pc=MTA | 2023826-2019-02248 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE REPORTER STATED THAT THE PROBLEM RESOLVED AFTER THE LENS EXPLANT. THE DOCTOR IMPLANTED A SHORTER LENGTH LENS (12.6MM). PATIENTS CURRENT BCVA 20/20, UCVA 20/20. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 1494: OFF-LABEL USE (AGE < 21 YEARS OLD) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373321&pc=MTA | 2023826-2019-02223 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936284&pc=MTA | 2023826-2020-00762 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837892&pc=MTA | 2023826-2020-00524 | 2019-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936260&pc=MTA | 2023826-2020-00763 | 2019-10-10 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163048&pc=MTA | 2023826-2020-01261 | 2019-10-10 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9416990&pc=MTA | 2023826-2019-02202 | 2019-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -08.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. PER THE REPORTER ON (B)(6) 2019, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9923401&pc=MTA | 2023826-2020-00830 | 2019-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559330&pc=MTA | 2023826-2020-00036 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON THE SAME DATE IN A SECOND SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889424&pc=MTA | 2023826-2020-00601 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512606&pc=MTA | 2023826-2019-02431 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512592&pc=MTA | 2023826-2019-02432 | 2019-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936274&pc=MTA | 2023826-2020-00748 | 2019-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -12.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE/DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H1- "MALFUNCTION" SHOULD BE CORRECTED TO "SERIOUS INJURY" IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586857&pc=MTA | 2023826-2020-00014 | 2019-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889407&pc=MTA | 2023826-2020-00605 | 2019-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9364601&pc=MTA | 2023826-2019-02157 | 2019-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512637&pc=MTA | 2023826-2019-02439 | 2019-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOSITURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9607848&pc=MTA | 2023826-2020-00168 | 2019-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9979184&pc=MTA | 2023826-2020-00924 | 2019-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9924748&pc=MTA | 2023826-2020-00839 | 2019-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310503&pc=MTA | 2023826-2020-01508 | 2019-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936239&pc=MTA | 2023826-2020-00770 | 2019-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION:LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532321&pc=MTA | 2023826-2019-02468 | 2019-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532327&pc=MTA | 2023826-2019-02461 | 2019-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753271&pc=MTA | 2023826-2020-00370 | 2019-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936308&pc=MTA | 2023826-2020-00746 | 2019-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936245&pc=MTA | 2023826-2020-00755 | 2019-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189457&pc=MTA | 2023826-2020-01287 | 2019-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512600&pc=MTA | 2023826-2019-02435 | 2019-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512619&pc=MTA | 2023826-2019-02436 | 2019-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL DEBRIS AND MOISTURE ON THE LENS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10: "MOISTURE ON THE LENS" SHOULD BE REMOVED FROM THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10125277&pc=MTA | 2023826-2020-01213 | 2019-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL WIT MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10139148&pc=MTA | 2023826-2020-01230 | 2019-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION-LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED INFORMATION: B5-THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE DID NOT RESOLVED THE PROBLEM. REFERENCE MFR 2023826-2020-01231 FOR REPLACEMENT LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9922460&pc=MTA | 2023826-2020-00828 | 2019-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401633&pc=MTA | 2023826-2020-01689 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC TORN WITH THE HAPTIC TORN, BENT AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9595579&pc=MTA | 2023826-2020-00081 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO LOW VAULT AND SUSPENSORY LIGAMENT BEING VERY LOOSE. CAUSE OF THE EVENT IS REPORTED AS PATIENT'S SUSPENSORY LIGAMENT BEING VERY LOOSE. REPORTEDLY, "THE PATIENT'S SUSPENSORY LIGAMENT IS VERY LOOSE, SO SHE DECIDED TO REMOVE THE LENS." Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9595587&pc=MTA | 2023826-2020-00082 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2019 DUE TO LOW VAULT AND "SUSPENSORY LIGAMENT BEING VERY LOOSE". THIS DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT'S SUSPENSORY LIGAMENT BEING VERY LOOSE. REPORTEDLY, "THE VAULT WASN'T IDEAL AFTER EXCHANGE SURGERY, THE PATIENT DECIDED TO REMOVE THE LENS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586868&pc=MTA | 2023826-2020-00015 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587407&pc=MTA | 2023826-2020-00112 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -12.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. REFERENCE MFR REPORT# 2023826-2020-00113 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824584&pc=MTA | 2023826-2020-00490 | 2019-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION WAS NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936242&pc=MTA | 2023826-2020-00761 | 2019-12-10 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9987307&pc=MTA | 2023826-2020-00941 | 2019-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6,-7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9939165&pc=MTA | 2023826-2020-00841 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403850&pc=MTA | 2023826-2020-01741 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587403&pc=MTA | 2023826-2020-00113 | 2019-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -12.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. REFERENCE MFR REPORT# 2023826-2020-00115 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FIBERS ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -12.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. REFERENCE MFR REPORT# 2023826-2020-00115 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# : (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9595588&pc=MTA | 2023826-2020-00083 | 2019-12-16 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO LOW VAULT AND "SUSPENSORY LIGAMENT BEING VERY LOOSE". CAUSE OF THE EVENT IS REPORTED AS PATIENT'S SUSPENSORY LIGAMENT BEING VERY LOOSE. REPORTEDLY, "THE VAULT WASN'T IDEAL AFTER EXCHANGE SURGERY, THE PATIENT DECIDED TO REMOVE THE LENS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9589379&pc=MTA | 2023826-2020-00103 | 2019-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (VICL IS INDICATED FOR USE IN PHAKIC EYE TREATMENT FOR CORRECTION/REDUCTION OF MYOPIA IN ADULTS RANGING FROM -0.5 TO -20.0 D. PATIENT HAD -21.) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9589390&pc=MTA | 2023826-2020-00099 | 2019-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2020-00102 FOR THE EXCHANGED LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (VICL IS INDICATED FOR USE IN PHAKIC EYE TREATMENT FOR THE CORRECTION/REDUCTION OF MYOPIA IN ADULTS RANGING FROM -0.5D TO -20.0D. PATIENT WAS -24.0D. SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9898242&pc=MTA | 2023826-2020-00644 | 2019-12-17 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189096&pc=MTA | 2023826-2020-01378 | 2019-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1,-9.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORCED AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DAT: B4-DATE OF THIS REPORT IN INITIAL MDR SHOULD BE CORRECTED TO 05-JUN-2020. G4- DATE RECEIVED BY MANUFACTURER IN INITIAL MDR SHOULD BE CORRECTED TO 05-JUN-2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093695&pc=MTA | 2023826-2020-01074 | 2019-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889413&pc=MTA | 2023826-2020-00592 | 2019-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10114924&pc=MTA | 2023826-2020-01137 | 2019-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, PATIENT DOES NOT PLANT TO IMPLANT ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10114928&pc=MTA | 2023826-2020-01136 | 2019-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, PATIENT DOES NOT PLANT TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869322&pc=MTA | 2023826-2020-00612 | 2019-12-26 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025644&pc=MTA | 2023826-2020-01004 | 2019-12-27 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957584&pc=MTA | 2023826-2020-00777 | 2020-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954142&pc=MTA | 2023826-2020-00886 | 2020-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9668836&pc=MTA | 2023826-2020-00234 | 2020-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE SURGEON EXCHANGED THE LENS WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A ZIPLOCK BAG. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10001067&pc=MTA | 2023826-2020-00991 | 2020-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1,-11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188380&pc=MTA | 2023826-2020-01363 | 2020-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928330&pc=MTA | 2023826-2020-00705 | 2020-01-16 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICMO 13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403906&pc=MTA | 2023826-2020-01743 | 2020-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214850&pc=MTA | 2023826-2020-01357 | 2020-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965182&pc=MTA | 2023826-2020-00901 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965193&pc=MTA | 2023826-2020-00902 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN, WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075073&pc=MTA | 2023826-2020-01095 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332954&pc=MTA | 2023826-2020-01572 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965197&pc=MTA | 2023826-2020-00903 | 2020-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163049&pc=MTA | 2023826-2020-01262 | 2020-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197374&pc=MTA | 2023826-2020-01336 | 2020-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -4.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880994&pc=MTA | 2023826-2020-00673 | 2020-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332476&pc=MTA | 2023826-2020-01591 | 2020-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332866&pc=MTA | 2023826-2020-01592 | 2020-03-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353755&pc=MTA | 2023826-2020-01638 | 2020-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: D9 - RETURNED TO MANUFACTURER DATE CORRECTED TO 12-AUG-2020 IN SUPPLEMENTAL MDR#1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222173&pc=MTA | 2023826-2020-01416 | 2020-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214961&pc=MTA | 2023826-2020-01351 | 2020-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10044526&pc=MTA | 2023826-2020-01052 | 2020-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6,-9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICOCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457351&pc=MTA | 2023826-2020-01869 | 2020-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445843&pc=MTA | 2023826-2020-01814 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTINO FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451103&pc=MTA | 2023826-2020-01856 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451121&pc=MTA | 2023826-2020-01855 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457334&pc=MTA | 2023826-2020-01893 | 2020-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457357&pc=MTA | 2023826-2020-01892 | 2020-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457361&pc=MTA | 2023826-2020-01876 | 2020-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403856&pc=MTA | 2023826-2020-01742 | 2020-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440592&pc=MTA | 2023826-2020-01809 | 2020-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10333805&pc=MTA | 2023826-2020-01599 | 2020-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-DEVICE EVALUATION: LENS RETURNED IN A MIRCO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10496539&pc=MTA | 2023826-2020-01965 | 2020-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10434161&pc=MTA | 2023826-2020-01804 | 2020-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -3.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10434209&pc=MTA | 2023826-2020-01805 | 2020-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140598&pc=MTA | 2023826-2020-01206 | 2020-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491996&pc=MTA | 2023826-2020-01967 | 2020-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K-THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10112203&pc=MTA | 2023826-2020-01152 | 2020-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094020&pc=MTA | 2023826-2020-01134 | 2020-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426571&pc=MTA | 2023826-2020-01761 | 2020-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140604&pc=MTA | 2023826-2020-01196 | 2020-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332957&pc=MTA | 2023826-2020-01567 | 2020-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361988&pc=MTA | 2023826-2020-01625 | 2020-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-DEVICE EVALUATION: LENS RETURNED IN A MIRCO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10418323&pc=MTA | 2023826-2020-01783 | 2020-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. D4 - EXPIRATION DATE: 31/05/2022 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10233520&pc=MTA | 2023826-2020-01440 | 2020-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188329&pc=MTA | 2023826-2020-01362 | 2020-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285735&pc=MTA | 2023826-2020-01483 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429121&pc=MTA | 2023826-2020-01844 | 2020-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184835&pc=MTA | 2023826-2020-01321 | 2020-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476019&pc=MTA | 2023826-2020-01941 | 2020-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476098&pc=MTA | 2023826-2020-01939 | 2020-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445862&pc=MTA | 2023826-2020-01811 | 2020-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND WAS DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457342&pc=MTA | 2023826-2020-01884 | 2020-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIPTER -11.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). A SHORTER LENS WAS IMPLANTED AT A LATER DATED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10224722&pc=MTA | 2023826-2020-01392 | 2020-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335299&pc=MTA | 2023826-2020-01557 | 2020-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429125&pc=MTA | 2023826-2020-01837 | 2020-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429136&pc=MTA | 2023826-2020-01838 | 2020-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10270662&pc=MTA | 2023826-2020-01470 | 2020-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332984&pc=MTA | 2023826-2020-01578 | 2020-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT, AND STRETCHED OUT WITH DEBRIS ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10355337&pc=MTA | 2023826-2020-01650 | 2020-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332959&pc=MTA | 2023826-2020-01570 | 2020-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354273&pc=MTA | 2023826-2020-01617 | 2020-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS SIZE WAS LARGE AND THE IRIDOCORNEAL ANGLES WERE REDUCED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. IT WAS DIFFICULT TO DETERMINE THE LENS SIZE FOR THIS PATIENT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401657&pc=MTA | 2023826-2020-01698 | 2020-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429086&pc=MTA | 2023826-2020-01834 | 2020-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -3.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429145&pc=MTA | 2023826-2020-01833 | 2020-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -3.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10279936&pc=MTA | 2023826-2020-01456 | 2020-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354233&pc=MTA | 2023826-2020-01614 | 2020-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353753&pc=MTA | 2023826-2020-01636 | 2020-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354272&pc=MTA | 2023826-2020-01612 | 2020-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10355336&pc=MTA | 2023826-2020-01651 | 2020-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354238&pc=MTA | 2023826-2020-01615 | 2020-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH DEBRIS ON THE LENS AND THE LENS WAS DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403863&pc=MTA | 2023826-2020-01734 | 2020-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403858&pc=MTA | 2023826-2020-01733 | 2020-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401664&pc=MTA | 2023826-2020-01692 | 2020-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392437&pc=MTA | 2023826-2020-01673 | 2020-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ANF RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401655&pc=MTA | 2023826-2020-01694 | 2020-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429077&pc=MTA | 2023826-2020-01832 | 2020-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K-THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429098&pc=MTA | 2023826-2020-01831 | 2020-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K-THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483349&pc=MTA | 2023826-2020-01955 | 2020-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9244517&pc=MTA | 2023826-2019-01975 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -02.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2017. PER REPORTER ON (B)(6) 2019 THE LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: DATE OF EVENT: CORRECTED TO (B)(6) 2019. DEVICE CODE 1494: OFF LABEL USE, ACD < 3.0MM. DEVICE CODE 2993 NOT APPLICABLE. EVENT: THE LENS WAS IMPLANTED ON (B)(6) 2019. LOW VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT WERE ALSO REPORTED. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE WAS REPORTED AS DUE TO THE DEVICE. DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE 3191: NO CODE AVAILABLE, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT CLAIM# (B)(4). Manufacturer Narrative: B5 - A LASER PI WAS PERFORMED 1 WEEK PRIOR TO IMPLANTATION. LOW VAULT UPDATED TO EXCESSIVE VAULT. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH - ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307849&pc=MTA | 2023826-2019-00215 | 2018-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+4.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335338&pc=MTA | 2023826-2020-01540 | 2020-01-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. WITHIN THE SAME SURGERY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587397&pc=MTA | 2023826-2020-00117 | 2019-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE PATIENT REFUSED A 2ND IMPLANT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH CLEAR SURGICAL RESIDUE/DEBRIS IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISISBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS NOT WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8629823&pc=MTA | 2023826-2019-00893 | 2019-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO OBSERVATION OF LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED ON THE SAME DAY IN A SECOND SURGERY. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS IMPLANTED AFTER THE LENS EXPIRATION DATE. REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND RESIDUE AND DEBRIS ON LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9398431&pc=MTA | 2023826-2019-02258 | 2019-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -16.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8978123&pc=MTA | 2023826-2019-01666 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND MYDRIASIS WAS REPORTER. THE REPORTER STATED THAT THE MYDRIASIS WAS NOTICED ON (B)(6) 2019. BRIMONIDINE DROPLETS WERE USED TWICE A DAY TO TREAT PATIENT. PATIENTS CURRENT BCVA IS 20/40, INTRAOCULAR PRESSURE IS STABLE, BUT THERE IS VAULT AND NARROW ANGLES. LENS REMAINS IMPLANTED, DOCTOR PLANS TO IMPLANT A 12.6MM LENGTH LENS. Manufacturer Narrative: ADDITIONAL DATA: EVENT: THE LENS WAS REMOVED ON (B)(6) 2019 AND REPLACED WITH A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. ON PATIENTS LAST VISIT ON (B)(6) 2019, THE PATIENTS BCVA WAS 20/30 AND UCVA WAS REPORTED TO BE 20/60. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: OTHER RELEVANT HISTORY: IRITIS, SYNECHIA, PIGMENT DISPERSION AND PSEUDOEXFOLIATION IN INITIAL MDR IS NOT APPLICABLE. DEVICE CODE: 2682 IN INITIAL MDR CORRECTED TO 1583. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8978124&pc=MTA | 2023826-2019-01667 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -17.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND MYDRIASIS WAS REPORTER. THE REPORTER STATED THAT THE MYDRIASIS WAS NOTICED ON (B)(6) 2019. BRIMONIDINE DROPLETS WERE USED TWICE A DAY TO TREAT PATIENT. PATIENTS CURRENT BCVA IS 20/40, INTRAOCULAR PRESSURE US STABLE, BUT THERE IS VAULT AND NARROW ANGLES. LENS REMAINS IMPLANTED, DOCTOR PLANS TO IMPLANT A 12.6MM LENGTH LENS. Manufacturer Narrative: THE LENS WAS REMOVED ON (B)(6) 2019 AND REPLACED WITH A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. ON PATIENTS LAST VISIT ON (B)(6) 2019, THE PATIENTS BCVA WAS 20/40 AND UCVA WAS REPORTED TO BE 20/70. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426602&pc=MTA | 2023826-2020-01779 | 2020-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem; Patient Device Interaction Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -07.50 DIOPTER IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020 DUE TO "THE PATIENT REQUEST TO CHANGE THE LENS TO LOWER POWER BECAUSE PATIENT FEELS EYESIGHT IS TOO HIGH. THE EYESIGHT COUD BE JUST ON TARGET." THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE: 3001 - SHOULD BE REMOVED AS IT'S NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426557&pc=MTA | 2023826-2020-01778 | 2020-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient Device Interaction Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020 DUE TO "THE PATIENT REQUEST TO CHANGE THE LENS TO LOWER POWER BECAUSE PATIENT FEELS EYESIGHT IS TOO HIGH. THE EYESIGHT COUD BE JUST ON TARGET." THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9399134&pc=MTA | 2023826-2019-02299 | 2019-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED. IT WAS REPORTED THAT THE PATIENT REQUESTED THE LENS REMOVAL. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4- "03-NOV-2019" SHOULD BE CORRECTED TO "04-NOV-2019" IN THE INITIAL MDR. G4- "03-NOV-2019" SHOULD BE CORRECTED TO "04-NOV-2019" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9855328&pc=MTA | 2023826-2020-00537 | 2019-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. LENS DISLOCATION/SUBLUXATION WAS OBSERVED. ON (B)(6) 2019, A "SLIT-LAMP MICROSCOPY OF THE EYE SHOWED A CLEAR CORNEA AND PARTIAL DISLOCATION OF THE LENS INTO THE ANTERIOR CHAMBER, WITH PUPILLARY CAPTURE", THE LENS WAS REPOSITIONED, AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, "THE PATIENT TOOK DIAZEPINE AND BUPROPION FOR DEPRESSION, WHICH CAUSED THE PUPIL TO OPEN WIDE." Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (PUPIL OVALISATION). H6 - DEVICE CODE: 2923 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (SAFETY AND EFFECTIVENESS OF THE LENS HAS NOT BEEN ESTABLISHED IN PATIENTS WITH HISTORY OF PREVIOUS OCULAR SURGERY INCLUDING REFRACTIVE CORNEAL SURGERY). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020" IN SUPPLEMENTAL MEDWATCH REPORTER #2. H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #2. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512594&pc=MTA | 2023826-2019-02430 | 2019-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS SUBLUXATION BY MARFAN SYNDROME. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512609&pc=MTA | 2023826-2019-02429 | 2019-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS SUBLUXATION BY MARFAN SYNDROME. REFERENCE MFR REPORT 2023826-2019-02427 FOR INITIAL LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION-LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5092392&pc=MTA | 2023826-2015-01170 | 2015-08-24 | Injury | STAAR SURGICAL COMPANY | 2015-09-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Optical Decentration | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/+3.0/X086 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2015 A REFRACTIVE SURPRISE WAS NOTED. THE LENS REMAINS IMPLANTED. A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED THEN ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.0/3.0/X086. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINT WAS FOUND. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254839&pc=MTA | 2023826-2019-00096 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 | MTA | IMPLANTABLE COLLAMER LENS | Optical Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.0/+6.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE SURGEON REPORTS THAT, AT 2 MONTH POST-OP, THE PATIENT DOES NOT SEE WELL. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320739&pc=MTA | 2023826-2018-00326 | 2017-12-08 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -18.0, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO ZONULAR ABNORMALITIES. REPORTEDLY, AFTER IMPLANT THE SURGEON FELT THE ZONULE WAS WEAK. ONE DAY POST-OP THE SURGEON DECIDED TO REMOVE THE LENS AS HE FELT IT MIGHT FALL INTO THE VITREOUS CAVITY. THERE IS NO PLAN FOR ADDITIONAL SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: THE SURGEON REMOVED THE ICL LENS AND NOT HER CRYSTALLINE LENS. THE SURGEON ALSO SAID THAT HE WAS NOT SURE THAT THE PATIENT HAD ANY OCULAR HISTORY BUT HE FELT HER ZONULE WAS TOO WEAK WHEN HE IMPLANTED THE ICL. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9282350&pc=MTA | 2023826-2019-02035 | 2019-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS BECAUSE THE PATIENT WANTED LOW CORRECTION. THIS EXCHANGE RESOLVED THE PROBLEM. IT WAS REPORTED THAT VISUAL ACUITY AFTER IMPLANTING THE INITIAL LENS WAS 20/10 AND AFTER EXCHANGING THE LENS IT WAS 20/16. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: B4- "07-OCT-2019" SHOULD BE CORRECTED TO "06-OCT-2019" IN THE INITIAL MDR. D11- INJECTOR MODEL: "MSI-TF" SHOULD BE CORRECTED TO "MSI-PF" IN THE INITIAL MDR. G4- "07-OCT-2019" SHOULD BE CORRECTED TO "06-OCT-2019" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913464&pc=MTA | 2023826-2020-00742 | 2019-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO SUSPENSORY LIGAMENT RELAXATION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT IS RECOVERING WELL AFTER THE OPERATION. CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS WAS EXCHANGED FOR A LONGER LENS AND THE PATIENTS CURRENT CONDITION IS GOOD. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6048751&pc=MTA | 2023826-2016-01119 | 2016-01-28 | Injury | STAAR SURGICAL COMPANY | 2016-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO THE LENS BEING UPSIDE-DOWN. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6036811&pc=MTA | 2023826-2016-01465 | 2016-02-03 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO THE LENS BEING UPSIDE-DOWN. THE LENS WAS EXCHANGED FOR ANOTHER LENS OF THE SAME MODEL AND LENGTH, AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS IN PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6432887&pc=MTA | 2023826-2017-00492 | 2017-03-01 | Injury | STAAR SURGICAL COMPANY | 2017-03-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, ON (B)(6) 2017. THE LENS WAS FOUND TO BE UPSIDE DOWN IN THE PATIENT'S EYE AND WAS EXPLANTED ON (B)(6) 2017, WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6752274&pc=MTA | 2023826-2017-01140 | 2017-06-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. A SIMILAR SECOND LENS WAS IMPLANTED SUCCESSFULLY ON (B)(6) 2017. PATIENT IS DOING WELL. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND WHITE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114231&pc=MTA | 2023826-2017-01837 | 2017-10-25 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. DURING INSERTION, THE LENS INVERTED AND ROTATED. THE SURGEON HAD TO ENLARGE THE INCISION TO EXPLANT THE LENS. THE LENS WAS INTRAOPERATIVELY EXCHANGED WITH A BACK-UP LENS AND IT WAS SUCCESSFUL. THE PATIENT IS CURRENTLY FINE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A CARTRIDGE CASE. VISUAL INSPECTION FOUND WHITE AND CLEAR RESIDUE ON LENS SURFACE AND A PIECE OF HAPTIC MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7177089&pc=MTA | 2023826-2018-00010 | 2017-10-26 | Injury | STAAR SURGICAL COMPANY | 2018-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO THE CONCAVE FACE OF THE IMPLANTED LENS WAS UPWARDS. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7156005&pc=MTA | 2023826-2017-01973 | 2017-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS | Positioning Failure | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2017. ON THE SAME DATE (UNKNOWN IF THE SAME SURGERY), THE LENS WAS EXPLANTED. AS THE SURGEON ATTEMPTED TO RE-IMPLANT THE LENS, HE DISCOVERED THE LENS HAD BEEN DAMAGED DURING EXPLANT. A BACK-UP LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S.(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197902&pc=MTA | 2023826-2018-00075 | 2017-12-18 | Injury | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -7.0/+2/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT, A TEAR THROUGH THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A, 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, OF -7.0/+2/088 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. MODEL #: IT IS CORRECTED FROM "VTICM5_12.6" TO "VTICM5_13.2". SERIAL #: IT IS CORRECTED FROM "(B)(4)" TO "(B)(4)". Manufacturer Narrative: METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335043&pc=MTA | 2023826-2018-00339 | 2018-01-06 | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM5_12.1, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. AFTER THE SURGERY THE SURGEON FELT SOMETHING WRONG BECAUSE "THE VISUAL ACUITY OUTCOME WAS POOR." ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO THE LENS BEING UPSIDE DOWN. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. SECONDARY SURGERY, LENS EXCHANGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7328923&pc=MTA | 2023826-2018-00355 |  | Injury | Unknown Manufacturer | 2018-03-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Code Available |  | 1 | Event Description: A FACEBOOK POST FROM A PATIENT STATED THAT THE SURGEON IMPLANTED A ICL BACK TO FRONT IN HIS RIGHT EYE (OD). THE LENS HAD TO BE EXTRACTED 4 WEEKS LATER DROPPING ENDO CELL COUNT SO THEY COULDNT TURN IT AND NOW THE PATIENT IS LEFT IMBALANCED. HE REPORTED THAT THEY ARE UNABLE TO IMPLANT AGAIN DUE TO DAMAGE TO ENDO CELL WHICH CANT REPLENISH. ATTEMPTED TO CONTACT THE PATIENT BUT THE PATIENT DELETED HIS REVIEW FROM THE VISIAN FACEBOOK PAGE (THIS CAN ONLY BE DONE BY THE PATIENT). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7051357&pc=MTA | 2023826-2017-01784 | 2017-10-26 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure; Inadequacy of Device Shape and/or Size | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50/+1.00/102 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/100. Manufacturer Narrative: THE REPORTER INDICATED THE ICL LENS (CRYSTAL) WAS IMPLANTED UPSIDE-DOWN UPON INJECTION. THE ICL LENS WAS TORN WHILE BEING REMOVED FROM THE PATIENT'S EYE ON (B)(4) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC DEFORMED AND MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4). SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5583323&pc=MTA | 2023826-2016-00450 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-04-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Problem; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 -15.00/+05.00/X089 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2008. THERE WAS A LENS ROTATION AND LENS REPOSITIONING. THE PATIENT EXPERIENCED A SECOND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. LENS REMAINS IMPLANTED. PLANNING TO EXCHANGE LENS. SEE MFR. #2023826-2016-00451 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE (B)(4). BUT NOT MARKETED IN THE U.S. DIOPTER: -15.00/+05.00/X089. DEVICE EVALUATED BY MANUFACTURER? NO. DEVICE REMAINS IMPLANTED. CONCLUSIONS: BASED ON THE COMPLAINT INFORMATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4962422&pc=MTA | 2023826-2015-01006 | 2015-06-22 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-08-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT (OD) EYE UPSIDE DOWN. THE LENS WAS BROKEN WHILE REMOVING FROM THE EYE. NO PATIENT INJURY WAS REPORTED. Manufacturer Narrative: WHILE THE LENS IS BEING EJECTED INTO THE ANTERIOR CHAMBER THE SURGEON MUST CHECK FOR LENS LANDMARKS TO ENSURE PROPER LENS ORIENTATION. IF THIS IS NOT DONE, THE LENS MAY END-UP BEING IMPLANTED UPSIDE-DOWN. ONCE THE SURGEON REALIZES OF THE IMPROPER ORIENTATION, THE LENS MUST BE REMOVED TO AVOID COMPLICATIONS AND A NEW LENS SHOULD BE IMPLANTED. THIS IS NORMALLY DONE USING THE SAME INCISION. THE MOST PROBABLE CAUSE OF THIS EVENT WAS USER'S ERROR WHILE LOADING/INJECTING THE LENS IN THE ANTERIOR CHAMBER. THE VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5309168&pc=MTA | 2023826-2015-01692 | 2015-12-03 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED AND REMOVED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, DUE TO THE HAPTIC WAS NOTICED TO BE RIPPED UPON INSERTION. THE LENS WAS REMOVED IN SMALL PIECES AND THERE WAS NO PATIENT INJURY. THE SURGEON REPORTED THE LENS WAS NOT NOTICED TO BE DAMAGED WHILE LOADING, BUT FELT THE LENS HAD A SMALL DEFECT WHEN RECEIVED AND THEN WAS DAMAGED AS IT WENT THROUGH THE INJECTOR. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE/LOSS OF MECHANICAL INTEGRITY TO THE HAPTIC OF THE IMPLANTED LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED THE DAMAGE. PER MEDICAL REVIEW - REPORTEDLY, INTRAOPERATIVE MICL(12.6 MM) REMOVAL/ EXCHANGE WAS PERFORMED TO ADDRESS HAPTIC DAMAGE NOTED UPON INSERTION. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ACCORDING TO SURGEON`S EMAIL, DATED ON 12/14/15 , HE DIDN`T NOTICE DAMAGE DURING LOADING SO HE ATTRIBUTED EVENT TO SMALL DEFECT AS RECEIVED THAT RESULTED IN HAPTIC TEAR WHILE HE WAS PUSHING IT THROUGH INJECTOR. HE CONFIRMED THAT THERE WAS NO PATIENT INJURY DURING REMOVAL. BACKUP LENS WAS SUCCESSFULLY IMPLANTED. NO REPORTED POSTOPERATIVE SEQUELAE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5822523&pc=MTA | 2023826-2016-00980 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS INSERTING A MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE. PART OF THE LENS REMAINED IN THE CARTRIDGE AND THE OTHER PART WAS IN THE PATIENT'S EYE. THE INCISION WAS ENLARGED TO REMOVE THE LENS FROM THE EYE WITHIN THE SAME SURGERY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. THE REPORTER WAS UNSURE ABOUT THE CAUSE OF THE EVENT BUT THERE WAS A TENSION DURING INSERTION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). RESULTS: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS TORN INTO TWO PIECES. BOTH HAPTICS WERE TORN, WITH PIECES MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059286&pc=MTA | 2023826-2016-01526 | 2016-09-02 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS - -7.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A TORN LENS. THE SURGEON IMPLANTED A NEW LENS ON (B)(6) 2016. THE ISSUE IS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5918361&pc=MTA | 2023826-2016-01184 | 2016-02-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND THE LENS WAS NOTED TO BE DAMAGED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A DIFFERENT MODEL AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(6). (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN, AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS DRIED, DISCOLORED MATERIAL. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592591&pc=MTA | 2023826-2015-00283 | 2014-10-20 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | VISIAN TICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ROTATION OF THE LENS WAS NOTED ON (B)(6)2014 AND THE LENS WAS REPOSITIONED. THE LENS WAS REMOVED AND REPLACED AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA - 20/20. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH, DEVICE HISTORY RECORD RESULTS -(OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER (B)(4). Manufacturer Narrative: EVALUATION METHOD: MEDICAL REVIEW. RESULTS: - VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION ECT.). THERE IS NOT ENOUGH INFORMATION TO DETERMINE THE EXACT ROOT CAUSE, HOWEVER, IT SEEMS THE ICL WAS SHORT FOR THIS EYE. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789589&pc=MTA | 2023826-2015-00580 | 2015-02-03 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/SIZE LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. EXPLANTED. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. PER MEDICAL REVIEW - WHILE THE LENS IS BEING EJECTED INTO THE ANTERIOR CHAMBER THE SURGEON MUST CHECK FOR LENS LANDMARKS TO ENSURE PROPER LENS ORIENTATION. IF THIS IS NOT DONE, THE LENS MAY END-UP BEING IMPLANTED UPSIDE-DOWN. ONCE THE SURGEON REALIZES OF THE IMPROPER ORIENTATION, THE LENS MUST BE REMOVED TO AVOID COMPLICATIONS AND A NEW LENS SHOULD BE IMPLANTED. THIS IS NORMALLY DONE USING THE SAME INCISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, THE MOST PROBABLE CAUSE OF THIS EVENT WAS USER'S ERROR WHILE LOADING/INJECTING THE LENS IN THE ANTERIOR CHAMBER. (B)(4). Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. LENS IMPLANTED UPSIDE DOWN, EXCHANGE RESOLVED THE PROBLEM. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4930179&pc=MTA | 2023826-2015-00914 | 2015-02-27 | Injury | STAAR SURGICAL COMPANY | 2015-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE LENS ROTATED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THIS RESOLVED THE PROBLEM. SEE MFR REPORT# 2023826-2015-00915 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4903566&pc=MTA | 2023826-2015-00808 | 2015-05-20 | Injury | STAAR SURGICAL COMPANY | 2015-07-09 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND FIVE DAYS POST-OP THE LENS HAD ROTATED 90 DEGREES. THE LENS WAS REPOSITIONED AND THREE DAYS LATER IT HAD ROTATED AGAIN. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND NOT REPLACED. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS -(OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7943035&pc=MTA | 2023826-2018-01476 | 2018-09-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM MICL13.2, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS REMAINED FOLDED AND WHEN IT UNFOLDED, THE HAPTIC WERE IN THE AC ANGLE (LENS FLIPPED), THE SURGEON WAS UNABLE TO VIEW THE ORIENTATION DOT OF LENS. ON (B)(6) 2018 THE LENS WAS REMOVED AND A SAME TYPE MODEL AND LENGTH LENS WAS IMPLANTED. PATIENT IS DOING VERY WELL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, EXPLANT. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149323&pc=MTA | 2023826-2018-01808 | 2018-10-30 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS | Unintended Movement | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6)2018. ON (B)(6) 2018 THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS DUE TO THE LENS WAS IMPLANTED "REVERSE SIDE." THE LENS SURFACE THAT SHOULD HAVE BEEN ON THE IRIS SIDE WAS INSERTED SO THAT IT WAS FACING THE CAPSULE; "FRONT AND BACK SIDE ARE REVERSED." THE CAUSE IS REPORTED AS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870926&pc=MTA | 2023826-2019-01402 | 2019-07-09 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | P030016 | 1 | Event Description: A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2019. REPORTEDLY, THE LENS FLIPPED DURING INSERTION. IT WAS REMOVED AND REPLACED WITH ANOTHER ICL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10032665&pc=MTA | 2023826-2020-01025 | 2019-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND EXCHANGED WITH THE SAME SIZE LENS ON (B)(6) 2019 DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR AND DEVICE. PER THE REPORTER ON (B)(6) 2020, LENS BEING IMPLANTED UPSIDE DOWN IS USER ERROR AND ONLY ON ORIENTATION HOLE IS DEVICE ERROR." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9757928&pc=MTA | 2023826-2020-00380 | 2020-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME LENGTH TORIC LENS DUE TO A LOW VAULT THAT WAS CAUSED BY THE LENS BEING IMPLANTED UPSIDE DOWN. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED " THE DOCTOR CHOSE A TORIC LENS WHEN IMPLANTING THE SECOND LENS BECAUSE HE WAS CONCERNED ABOUT INDUCED ASTIGMATISM THAT MIGHT HAVE OCCURRED DURING FIRST SURGERY." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10112136&pc=MTA | 2023826-2020-01148 | 2020-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426599&pc=MTA | 2023826-2020-01772 | 2020-07-20 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 13.2MM MICL13.2, -3.5 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. REPORTEDLY, SURGEON "DID NOT RECOGNIZE THAT LENS WAS IMPLANTED UPSIDE DOWN AT SURGERY, BUT IT WAS EVIDENT AT THE ONE-WEEK POST-OP VISIT." Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4757685&pc=MTA | 2023826-2015-00530 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Code Available | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS) ON (B)(6) 1999. THE PATIENT REPORTED SHE FEELS AND SEES THE LENS MOVING AROUND. THE LENS REMAINS IMPLANTED. THE PATIENT WAS A PARTICIPANT IN THE ICL MYOPIA CLINICAL STUDY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4)- UNINTENDED MOVEMENT. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4800434&pc=MTA | 2023826-2015-00615 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-27 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) AND THERE WAS AN INCIDENT OF REFRACTIVE SURPRISE. IT IS SUSPECTED THE LENS WAS ROTATED BUT IT HAS NOT YET BEEN CONFIRMED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. EVENT DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS REMAINS IMPLANTED. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8853047&pc=MTA | 2023826-2019-01382 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENTS IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) WITH THE AIM OF CORRECTING MYOPIA OR MYOPIC ASTIGMATISM. PERI-OPERATIVELY, THE FOLLOWING COMPLICATIONS WERE OBSERVED - ONE LENS BEGAN TO DEVELOP IN A REVERSED POSITION IN THE ANTERIOR CHAMBER, AND IT WAS NECESSARY TO EXTEND THE TEMPORAL INCISION, REMOVE THE LENS AND RE-IMPLANT IT INTO THE CORRECT POSITION, WITH SUBSEQUENT SUTURE OF THE WOUND. THE FURTHER COURSE WAS UNCOMPLICATED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: COMMON DEVICE NAME CORRECTED TO PHAKIC INTRAOCULAR LENS & PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: MTA & QCB (COHORT INCLUDED SEVERAL LENS MODEL) DEVICE INFORMATION CORRECTED TO ICM V4, VICMO, TICM V4, VTICM, VTICMO (COHORT INCLUDED SEVERAL LENS MODELS). DATE RECEIVED BY MANUFACTURER: SHOULD HAVE INCLUDED LITERATURE IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4838979&pc=MTA | 2023826-2015-00672 | 2014-09-09 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THE LENS ROTATED AND REPOSITIONED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. THE LENS WAS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4930113&pc=MTA | 2023826-2015-00915 | 2015-02-27 | Injury | STAAR SURGICAL COMPANY | 2015-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND THE LENS ROTATED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THIS RESOLVED THE PROBLEM. SEE MFR REPORT# 2023826-2015-00914 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704647&pc=MTA | 2023826-2015-00436 | 2015-03-18 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS. THE SURGERY WENT WELL AND THE PATIENT WAS MOVED FROM THE OPERATING ROOM TO BE MONITORED. THE PATIENT WAS BROUGHT BACK SHORTLY AFTER TO CHECK THE INTRAOCULAR PRESSURE AND THE SURGEON NOTED THE ICL WAS IMPLANTED UPSIDE DOWN. THE INCISION WAS RE-OPENED AND THE LENS WAS EXPLANTED WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE PATIENT IS DOING WELL AND THE VISUAL ACUITY IS 20/15. SUTURES WERE NOT REQUIRED. Manufacturer Narrative: (B)(4) - METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. THE INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER WERE NOT RETURNED FOR EVALUATION. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GIVEN THE REPORTED INFORMATION, THE EVENT WAS SURGICAL FACTOR RELATED. THE MEDICAL DEVICE IS NOT THE ROOT CAUSE OF THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, THE ROOT CAUSE OF THE EVENT WAS SURGICAL FACTOR RELATED. THE MEDICAL DEVICE IS NOT THE ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4940667&pc=MTA | 2023826-2015-00897 | 2015-06-11 | Injury | STAAR SURGICAL COMPANY | 2015-07-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON 06/11/2015 DUE TO LENS ROTATION. THE LENS WAS EXCHANGED FOR A SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5741984&pc=MTA | 2023826-2016-00842 | 2015-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-06-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.50/+5.5/071 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND ROTATION, AND WAS REPOSITIONED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. PERFORMED VISUAL INSPECTION AND FOUND THE HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4625042&pc=MTA | 2023826-2015-00363 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE ICL WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL FACTORS MAY CONTRIBUTE TO ROTATION OF AN ICL (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4759154&pc=MTA | 2023826-2015-00521 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS NOTED, THE LENS WAS EXCHANGED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) -(SECONDARY SURGERY); (NEW INCISION); UNINTENDED MOVEMENT (ROTATION); (EXPLANTED). EVALUATION METHOD: (OTHER) WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5478575&pc=MTA | 2023826-2016-00229 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016. THE INCISIONS WERE ENLARGED TO REMOVE AND SUTURES WERE REQUIRED. Manufacturer Narrative: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE OR NOT DETECT A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE CAUSED INVERSION, FRONT TO BACK, OF THE LENS DURING INSERTION. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5957291&pc=MTA | 2023826-2016-01289 |  | Injury | STAAR SURGICAL COMPANY | 2016-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. SOON AFTER, THE SURGEON REALIZED THE LENS WAS IMPLANTED UPSIDE DOWN. THE LENS WAS EXPLANTED ON (B)(6) 2016 WITH NO PATIENT INJURY AND WAS ACCIDENTLY DISCARDED. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENTS POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS DISCARDED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356128&pc=MTA | 2023826-2018-00433 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS | Use of Device Problem | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+5.0/131 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED AS THE SURGEON IMPLANTED THE LENS INTO THE WRONG EYE-IT SHOULD HAVE BEEN IMPLANTED IN THE OTHER EYE. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356134&pc=MTA | 2023826-2018-00436 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS | Use of Device Problem | No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A DIFFERENT LENS DUE TO THE SURGEON IMPLANTING THE INCORRECT LENS INTO THE EYE. THE SURGEON MIXED UP THE LENSES FOR THE OS AND OD EYES. THE PROBLEM WAS RESOLVED. REFERENCE MFR REPORT# 2023826-2018-00283 (CLAIM# (B)(4)) FOR PREVIOUS PART OF THIS CASE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660094&pc=MTA | 2023826-2018-00999 | 2018-05-17 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS | Use of Device Problem | No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/059 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS DUE TO THE INCORRECT LENS BEING IMPLANTED INTO THE EYE (OD LENS IMPLANTED INTO OS EYE). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS MARKED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10404449&pc=MTA | 2023826-2020-01705 | 2020-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL 12.6, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED. ENLARGEMENT OF PI AND YAG PERFORMED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS "MAYBE ASSOCIATED WITH BPI TO PERIPHERAL?" REPORTEDLY, "APPEARS WNL/NO SYMPTOMS.SIGNIFIGANT INCREASE IN IRIDO-COREANL ANGLES. PATIENT CURRENTLY ON PILO% TICL." Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED A SHALLOW ANTERIOR CHAMBER. SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SHALLOW ANTERIOR CHAMBER. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10418205&pc=MTA | 2023826-2020-01784 | 2020-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - HEALTH EFFECT CLINICAL CODE: 4581 (SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES). SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8041444&pc=MTA | 2023826-2018-01590 | 2018-07-21 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available; Missing Value Reason | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.00/+2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2018 THE LENS WAS REPOSITIONED BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549244&pc=MTA | 2023826-2016-00393 | 2016-01-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available; Missing Value Reason | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -11.50/2.00/X113 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES WAS NOTED (B)(6) 2016. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. AC IS QUIET AND DEEP NOW. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: DIOPTER: -011.50/+2.00/X113. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401662&pc=MTA | 2023826-2020-01699 | 2019-12-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; No Apparent Adverse Event | No Code Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -09.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS DISLOCATION/SUBLUXATION WAS OBSERVED. LENS WAS REMOVED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: B1- ADVERSE EVENT CORRECTED TO PRODUCT PROBLEM. B3-DATE OF EVENT CORRECTED TO (B)(6) 2019. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 OF -9.00 DIOPTER IMPLANTED COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY WITHIN INITIAL SURGERY DUE TO LENS DISLOCATION / SUBLUXATION. A REPLACEMENT LENS WAS LATER IMPLANTED ON (B)(6) 2019 AND THIS RESOLVED THE PROBLEM. D6B- DATE OF REMOVAL CORRECTED TO (B)(6) 2019. H1- SERIOUS INJURY CORRECTED TO MALFUNCTION. H6-IMPACT CODE 3191 CORRECTED TO 2199. INVESTIGATION CODE 4110 SHOULD HAVE BEEN REPORTED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: 716936.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457358&pc=MTA | 2023826-2020-01868 | 2018-11-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATE: H6- PATIENT CODE 3191 SHOULD BE CORRECTED TO HEALTH EFFECT CLINICAL CODE 4582 IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332992&pc=MTA | 2023826-2020-01582 | 2020-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2020 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: H6 - PATIENT CODE: 3191 LENS EXCHANGE SHOULD BE REMOVED. LENS WAS REPLACED INTRAOPERATIVELY AND SHOULD BE CORRECTED TO CODE 4582. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440446&pc=MTA | 2023826-2020-01895 | 2019-04-03 04:00 | Injury | STAAR SUGICAL COMPANY | 2020-08-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available; Toxic Anterior Segment Syndrome (TASS) | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.5 DIOPTER, INTRAOCULAR LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. TREATMENT PERFORMED AND INTRAOCULAR PRESSURE VALUE OF 39.3MMHG WAS REPORTED. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. IT WAS REPORTED THAT THE PROBLEM RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. IT WAS REPORTED THAT TREATMENT WAS PERFORMED AND BACTERIAL CULTURE CAME BACK NEGATIVE." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7467152&pc=MTA | 2023826-2018-00603 | 2018-03-14 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Activation Failure | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, AND WAS ATTEMPTING TO TUCK THE HAPTICS UNDER THE IRIS. THE SURGEON WAS ABLE TO TUCK ONE HAPTIC BUT THE PUPIL CONSTRICTED AND SURGEON WAS UNABLE TO TUCK THE OTHER HAPTIC. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY. THIS WAS THE 1ST OF 2 LENSES USED ON THIS PATIENT. THE 2ND LENS, SAME MODEL, WAS USED ON (B)(6) 2018 BUT THE SURGEON WAS UNABLE TO TUCK THE HAPTIC DUE TO PUPIL CONSTRICTED - SEE MFR REPORT # 2023826-2018-00602. THE SURGEON DID NOT IMPLANT A 3RD LENS. THE REPORTER INDICATED THE CAUSE OF THIS EVENT WAS PATIENT RELATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6573851&pc=MTA | 2023826-2017-00847 | 2017-04-19 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, THE PATIENT MOVED THE EYE AND THE LENS FLIPPED. THE LENS WAS REMOVED AND REPLACED WITH THE BACKUP LENS WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS NOT PRODUCT RELATED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION OF THE LENS FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7155973&pc=MTA | 2023826-2017-01985 | 2017-11-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A CQ2015A COLLAMER THREE PIECE LENS,+21.5 DIOPTER, WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY BECAUSE THE SURGEON NOTICED A PIGMENT ON THE LENS AND THE LENS LOOKED THICKER THAN NORMAL. THE REPORTER STATED THAT THE OR TECHNICIAN INSPECTED THE LENS PRIOR TO LOADING BUT DID NOT NOTICE ANYTHING UNUSUAL WITH THE LENS. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: SHOULD HAVE BEEN IN INITIAL MDR. INJECTOR MODEL: MSI-PM AND CARTRIDGE MODEL: CQ. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TORN IN TWO PIECES, HAPITCS SEPARATED FROM LENS AND A TEAR IN THE OPTIC. CLAIM #: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7591522&pc=MTA | 2023826-2018-00683 | 2018-02-19 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -11.5/+1.5/98 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED DURING THE SAME SURGERY BECAUSE THE PATIENT MOVED DURING THE PROCEDURE AND THE LENS TURNED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR IN A HAPTIC AND RESIDUE ON THE LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7676120&pc=MTA | 2023826-2018-01026 | 2018-05-29 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), IT TORE. THE LENS WAS REMOVED INTRAOPERATIVELY ON (B)(6) 2018. DURING THE SAME OPERATION, AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THERE WAS NO PATIENT INJURY AND THE CAUSE IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7770618&pc=MTA | 2023826-2018-01179 | 2018-06-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS IMPLANTING A VTICMO12.6, -13.5/2.00/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS AND THE PLUNGER OVERSHOT THE LENS AND IT BECAME STUCK. IT WAS REPORTED THAT THE LENS TORE DURING REMOVAL AND THERE WAS NO PATIENT INJURY. DELIVERY SYSTEM MODEL AND LOT NUMBERS ARE NOT AVAILABLE. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED.Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THE REPORTER STATES THAT DURING IMPLANT, THE PLUNGER OVERSHOT THE LENS AND GOT STUCK. IT TORE DURING REMOVAL. THE LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2018. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047801&pc=MTA | 2023826-2015-01119 | 2015-01-22 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO USE A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, BUT THE SURGEON NOTED A BLACK SPOT (FOREIGN BODY ADHESION) ON THE LENS. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS FOREIGN BODY ADHESION. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (SPOT ON LENS). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5114844&pc=MTA | 2023826-2015-01191 | 2015-02-06 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS INSERTED UPSIDE DOWN. THE LENS WAS REMOVED AND EXCHANGED WITH ANOTHER LENS, SAME MODEL AND SIZE. NO ADVERSE EVENT WAS REPORTED. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.50. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975055&pc=MTA | 2023826-2015-00971 | 2015-06-05 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS AND THE LENS WAS DAMAGED. THERE WAS NO REPORTED PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING VERY WELL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM CODES: PATIENT CODE: (B)(4)- NO CONSEQUENCE OR IMPACT TO PATIENT, LABELED. DEVICE CODE: (B)(4)- NO CODE AVAILABLE (LENS DAMAGE), LABELED DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS NOT RETURNED. EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). OTHER TEXT : LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975029&pc=MTA | 2023826-2015-00969 | 2015-06-16 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS AND THE LENS WAS DAMAGED WHILE INJECTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5072622&pc=MTA | 2023826-2015-01193 | 2015-07-30 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE FACILITY RECEIVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, AND NOTED A "STAIN" ON THE OUTSIDE OF THE LENS VIAL. THE SURGEON DECIDED TO IMPLANT THE LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: AFTER A REVIEW OF THE DHR, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. HOWEVER, THE RETURNED PRODUCT DID NOT REVEAL ANY STAIN ON THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND SCRATCHES ON THE OUTSIDE OF THE VIAL AND DRIED RESIDUE INSIDE THE VIAL. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: - WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5068907&pc=MTA | 2023826-2015-01102 | 2015-08-10 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6, -13.00/+2.50/X121 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD). LENS GOT DAMAGED WHILE INJECTING INTO THE EYE. THE LENS WAS REMOVED AND NO ADVERSE EVENT WAS REPORTED. A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5102508&pc=MTA | 2023826-2015-01178 | 2015-08-26 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/X092 DIOPTER IN THE PATIENT'S RIGHT EYE (OD). THE LENS FLIPPED UPSIDE DOWN DURING INSERTION INTO THE EYE. THE LENS WAS ACCIDENTALLY TORN WHILE IT WAS BEING REMOVED FROM THE EYE. NO ADVERSE EVENT WAS REPORTED. NO OTHER LENS WAS IMPLANTED AT THE TIME. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -15.5/+1.5/X092. (B)(4). METHOD CODE: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION : CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248932&pc=MTA | 2023826-2015-01473 | 2015-10-20 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER, FROM THE VIAL AND NOTED, PRIOR TO LOADING, SOME SORT OF "MARK/FLAW" ON THE LENS, POSSIBLY A PUNCTURE/TEAR. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THEY FELT THE LENS WAS RECEIVED THAT WAY. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS THAT SUGGESTS A CONTRIBUTORY FACTOR IN THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308643&pc=MTA | 2023826-2015-01699 | 2015-11-18 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON NOTICED A BLACK SPECK IN THE FLUID INSIDE THE LENS VIAL OF A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER. THE LENS WAS NOT LOADED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. THE LENS VIAL WAS REVIEWED BY QUALITY ENGINEER, HOWEVER, NO LIQUID WAS FOUND AND THE LENS WAS YELLOWISH, HENCE NO OBSERVATION WAS MADE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN A SMALL AMOUNT OF LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5309194&pc=MTA | 2023826-2015-01694 | 2015-12-10 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, FROM THE LENS VIAL AND NOTED A "MARK" OR SOMETHING SIMILAR ON THE LENS SURFACE. THE LENS WAS NOT USED OR LOADED AND WAS PLACED BACK INTO THE VIAL. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE TO THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. NO OBSERVABLE MARKS WERE FOUND ON THE RETURNED LENS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323634&pc=MTA | 2023826-2015-01727 | 2015-12-10 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON EXAMINED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, UNDER THE MICROSCOPE AND NOTED ONE OF THE HAPTIC FOOTPLATES WAS FOUND TO BE SMALLER THAN THE OTHER FOOTPLATE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS OPTIC AND THE OTHER HAPTIC WERE TORN. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE TO THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5381805&pc=MTA | 2023826-2016-00069 | 2015-12-31 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+3.0/092 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2015. THE SURGERY WAS UNEVENTFUL. AFTER THE SURGERY, THEY NOTED MICROBIAL GROWTH/DEPOSITS ON THE EXTERIOR OF THE GLASS BOTTLE, ESPECIALLY ON THE RIM AND NECK. THE LENS REMAINS IMPLANTED AND THE PATIENT IS DOING WELL. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: REVIEW OF THE COMPLAINT BY THE MANUFACTURING TEAM CONCLUDED THE "MICROBIAL GROWTH-LIKE DEPOSIT" ON THE EXTERIOR OF THE VIAL, IS MOST LIKELY THE RESIDUAL BALANCED SALT SOLUTION (BSS) ON THE VIAL. AS PART OF THE PRODUCT PROCESSING STEPS, IT IS POSSIBLE THAT BSS COULD BE DEPOSITED ON THE VIAL(S) AND IF THE BSS DRIES UP AND THE VIAL SURFACE IS NOT THOROUGHLY WIPED DOWN, AS PER THE PROCEDURE, THIS COULD CAUSE "MICROBIAL GROWTH-LIKE DEPOSIT" ON THE VIAL SURFACE. HENCE, RE-TRAINING OF THE CONCERNED OPERATORS WAS COMPLETED BY MANUFACTURING AS A CORRECTIVE ACTION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536715&pc=MTA | 2023826-2016-00308 | 2016-01-05 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS INSERTED INTO THE EYE UPSIDE DOWN. DURING THE SAME PROCEDURE, THE LENS WAS REMOVED AND EXCHANGED FOR A SAME MODEL AND DIOPTER LENS. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/32. NO ADVERSE EVENT WAS REPORTED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5409054&pc=MTA | 2023826-2016-00131 | 2016-01-28 | Malfunction | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON REMOVED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, FROM THE LENS VIAL, TO PREPARE FOR LOADING, AND NOTED THE LENS LOOKED STRANGE. THE SURGEON STATED ONE HAPTIC LOOKED DEFORMED. THE SURGEON DECIDED NOT TO USE THE LENS AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE LENS WAS NOT LOADED AND THERE WAS NO PATIENT CONTACT. THE REPORTER STATED THEY FELT THE LENS WAS RECEIVED THAT WAY AND WAS DEFECTIVE. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: LENS NOT RETURNED. DEVICE: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632896&pc=MTA | 2023826-2016-00566 | 2016-02-09 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, WAS NOT IMPLANTED DUE TO A BAD LOAD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5487774&pc=MTA | 2023826-2016-00182 | 2016-02-10 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, FROM THE LENS VIAL AND WAS PREPARING TO LOAD THE LENS. THE SURGEON NOTED ONE OF THE HAPTICS LOOKED CRIMPED AND DECIDED NOT TO USE THE LENS. THE LENS WAS NOT LOADED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THE SURGEON FELT THE LENS WAS RECEIVED THAT WAY AND WAS DEFECTIVE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5488107&pc=MTA | 2023826-2016-00236 | 2016-02-16 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, BUT THE LENS DID NOT FEEL RIGHT. THE LENS WAS NOT FULLY INSERTED INTO THE PATIENT'S EYE AND WAS REMOVED WITH NO PATIENT INJURY, NO INCISION ENLARGEMENT OR SUTURES. THE BACKUP LENS WAS USED. THE REPORTER STATED THERE WAS NO PROBLEM WITH THE LENS OR INJECTION SYSTEM. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF BOTH HAPTICS TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DARK RESIDUE ON THE LENS SURFACE. (NO FAILURE DETECTED): THERE WAS NOT ENOUGH DETAIL PROVIDED BY THE COMPLAINANT TO PERFORM A ROOT CAUSE ANALYSIS. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE OR NOT DETECT A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION DURING LOADING AND INSERTION, EXCEPT THE SUBJECTIVE EXPERIENCE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5511778&pc=MTA | 2023826-2016-00226 | 2016-02-26 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS FOLDING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, FOR LOADING, AND NOTED A STAIN IN THE LENS HAPTIC. THE LENS WAS NOT IMPLANTED AND THERE WAS NO REPORT OF ANY PATIENT INJURY. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: THE FOREIGN RESIDUE/WHITE STAIN ON THE LENS HAPTIC WAS LIKELY TO BE INTRODUCED BY THE MANUFACTURING RELATED ACTIVITIES AT STAAR. IT IS LIKELY, THE RESIDUE FROM THE TOOL USED FOR INSPECTION WAS TRANSFERRED ONTO THE LENS HAPTIC SURFACE. NECESSARY CORRECTIVE ACTIONS WERE TAKEN TO UPDATE ASSOCIATED PROCEDURES TO IMPLEMENT THE USE OF THE REPLACEMENT TOOL. REVIEW OF THE RISK DOCUMENTATION FOR THIS LENS MODEL FOUND THAT THIS TYPE OF FAILURE MODE HAS BEEN EVALUATED AND APPROPRIATE MITIGATION STEPS HAVE BEEN IMPLEMENTED TO ENSURE THIS MODEL IS DESIGNED AND MANUFACTURED SUCH THAT THE LENS MEETS ESTABLISHED SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A FOREIGN MATERIAL/WHITE STAIN ON THE LENS SURFACE. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE INSPECTION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655428&pc=MTA | 2023826-2016-00725 | 2016-04-14 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON HAD DIFFICULTY LOADING A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: (B)(4). METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A SMALL TEAR ON ONE HAPTIC. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951508&pc=MTA | 2023826-2016-01281 | 2016-08-15 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, AND THE LENS EDGE WAS NOTED TO BE IRREGULAR. THERE WAS PATIENT CONTACT BUT NO INJURY. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS NOT INSERTED AND REMOVED. THE SURGEON DID NOT LIKE THE EDGE APPEARANCE AFTER LOADING INTO THE DELIVERY INSTRUMENT, DURING PROCEDURE. (B)(4). Manufacturer Narrative: (B)(4) LENS EDGE IRREGULAR. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5464650&pc=MTA | 2023826-2016-00208 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-02-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED AND REMOVED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, DUE TO THE LENS WOULD NOT UNFOLD PROPERLY. THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVALUATION RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. BOTH HAPTICS WERE FOLDED AND THERE WAS DRIED DARK RESIDUE ON THE LENS. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT); BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. Manufacturer Narrative: (B)(4). UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5648755&pc=MTA | 2023826-2016-00640 | 2016-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Bent | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON LOADED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, AND NOTED AFTER PULLING THE LENS DOWN THE CARTRIDGE BARREL, A PIECE OF THE LENS WAS PROTRUDING OUT OF THE CARTRIDGE. HE WASN'T SURE IF HE COULD INJECT THE LENS CORRECTLY INTO THE EYE, SO HE EJECTED THE LENS ONTO THE STERILE FIELD, WITH THE PLAN OF RELOADING THE LENS. AFTER EXAMINING THE LENS, HE NOTED ONE HAPTIC WAS SLIGHTLY BENT, SO HE DECIDED NOT TO USE THIS LENS. THERE WAS NO PATIENT CONTACT AND THE BACKUP WAS IMPLANTED WITH NO PROBLEM. THE PATIENT IS DOING VERY WELL. THE SURGEON STATED THE CAUSE OF THE EVENT WAS UNKNOWN BUT OCCASIONALLY WHEN HE IS PULLING THE LENS DOWN THE CARTRIDGE BARREL, HE FINDS IT HARD TO RELEASE THE LENS WITH THE FORCEPS. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. THE LENS WAS RETURNED IN LIQUID. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5088802&pc=MTA | 2023826-2015-01126 | 2015-05-14 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -12.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE AND WAS REMOVED. NO ADVERSE EVENT WAS REPORTED. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.00. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937354&pc=MTA | 2023826-2015-00926 | 2015-06-18 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS BROKE. THE LENS REMAINS IMPLANTED BUT WILL BE REMOVED. THE LENS WAS THE EXCHANGE LENS, SEE MFR REPORT (B)(4) FOR THE FIRST LENS.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #: (B)(4). LENS NOT RETURNED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089071&pc=MTA | 2023826-2015-01125 | 2015-06-25 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE AND WAS REMOVED. NO ADVERSE EVENT WAS REPORTED. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -08.50. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536899&pc=MTA | 2023826-2016-00359 | 2015-08-26 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND EXCHANGED FOR ANOTHER SAME MODEL AND DIOPTER SIZE. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2015, UCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -07.50. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5242700&pc=MTA | 2023826-2015-01527 | 2015-10-21 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND EXCHANGED FOR A ANOTHER SIMILAR MODEL AND DIOPTER SIZE. NO ADVERSE EVENTS WERE REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5244753&pc=MTA | 2023826-2015-01538 | 2015-10-23 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM VTICMO13.2, -13.50/+2.50/X055 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND NO OTHER LENS WAS IMPLANTED. NO ADVERSE EVENTS REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50/+2.50/X055. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S) (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5507754&pc=MTA | 2023826-2016-00213 | 2015-11-26 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON PARTIALLY INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE ON (B)(6) 2015. ON INSERTION, A TEAR DEVELOPED IN ONE OF THE FOOTPLATES. THE LENS WAS NOT FULLY IMPLANTED. THE LENS WAS REMOVED AT THE TIME BUT NOT REPLACED. ON (B)(6) 2016 A ICL WAS IMPLANTED. Manufacturer Narrative: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. DEVICE NOT RETURNED. (B)(4). CLAIM # (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643929&pc=MTA | 2023826-2016-00580 | 2015-10-29 | Injury | STAAR SURGICAL COMPANY | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 -15.50 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS REPORTEDLY TORE DURING INJECTION/DELIVERY IN TO THE EYE AND WAS REMOVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. CONCLUSION CODE - AS PER DFU OF THIS LENS MODEL, IMPLANTATION OF THIS LENS IN A PERSON WHO IS BELOW 21 YEARS OF AGE IS CONTRAINDICATED. THIS PATIENT WAS (B)(6) AT THE TIME OF EVENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5505902&pc=MTA | 2023826-2016-00272 | 2016-01-07 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, WAS NOTED TO BE BROKEN, BEFORE BEING LOADED. THERE WAS NO PATIENT CONTACT. THE REPORTER INDICATED THE LENS WAS RECEIVED BROKEN. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5440036&pc=MTA | 2023826-2016-00151 | 2014-12-30 | Malfunction | STAAR SURGICAL COMPANY | 2016-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO REPORT OF ANY PATIENT INJURY. Manufacturer Narrative: CORRECTED DATA: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH . RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287704&pc=MTA | 2023826-2015-01636 | 2015-07-20 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, AND NOTED THE LENS HAD TORN/BROKEN. THE LENS WAS REMOVED WITH NO REPORTED INJURY AND WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167639&pc=MTA | 2023826-2015-01370 | 2015-07-24 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-10-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -18.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THERE WAS NO REPORTED PATIENT INJURY. Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -18.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO REPORTED PATIENT INJURY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: PREVIOUS (B)(4) ARE INCORRECT. THERE IS NO BREAK OR TORN MATERIAL ASSOCIATED WITH A MATERIAL INTEGRITY ISSUE. CORRECT DEVICE CODE IS (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218815&pc=MTA | 2023826-2015-01502 | 2015-08-15 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+0.5/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE LENS WAS REMOVED AND THERE WAS NO REPORTED PATIENT INJURY. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: THIS CASE IS AN ADVERSE EVENT NOT A PRODUCT PROBLEM. DEVICE EVALUATION - VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593296&pc=MTA | 2023826-2016-00482 | 2016-03-01 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO REPORT OF ANY PATIENT INJURY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4957931&pc=MTA | 2023826-2015-00980 | 2014-03-14 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-07-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE. THERE WAS NO PATIENT INJURY REPORTED AND ANOTHER SAME MODEL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD - LENS WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5481740&pc=MTA | 2023826-2016-00167 | 2014-12-14 | Malfunction | STAAR SURGICAL COMPANY AG | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -9.5/2.5/086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED ON (B)(6) 2014 AND WAS EXCHANGED FOR ANOTHER SAME MODEL/DIFFERENT DIOPTER LENS. THE EXCHANGED LENS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN. CORRECTED DATA: (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: CORRECTED DATA: PLEASE DISREGARD PREVIOUS MDR SUPPLEMENTAL #2. SUBMITTED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743972&pc=MTA | 2023826-2015-00505 | 2015-03-11 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE. THE REPORTER INDICATED WHEN THE LENS WAS BEING EJECTED, THE LENS MAY HAVE BEEN CAUGHT BY THE FOAM TIP PLUNGER AND TORE. THE SURGERY WAS CANCELLED BUT ANOTHER SAME MODEL LENS WAS IMPLANTED SUCCESSFULLY THE NEXT DAY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVALUATION: RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4760623&pc=MTA | 2023826-2015-00517 | 2015-03-30 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-05-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THERE WAS NO APPARENT INJURY REPORTED AND ANOTHER SAME MODEL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4754639&pc=MTA | 2023826-2015-00507 | 2015-04-02 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING INJECTION. THE LENS WAS REMOVED WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789587&pc=MTA | 2023826-2015-00575 | 2015-04-02 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) BUT THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED INJURY AND WAS EXCHANGED FOR ANOTHER SAME SIZE LENS, WHICH RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4958120&pc=MTA | 2023826-2015-00986 | 2015-04-13 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-07-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE. THERE WAS NO PATIENT INJURY REPORTED AND ANOTHER SAME MODEL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD (OTHER) - LENS WORK ORDER SEARCH. RESULTS (OTHER):A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4957967&pc=MTA | 2023826-2015-00984 | 2015-05-04 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-07-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE. THERE WAS NO PATIENT INJURY REPORTED AND ANOTHER SAME MODEL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4958069&pc=MTA | 2023826-2015-00985 | 2015-05-04 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-07-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE. THERE WAS NO PATIENT INJURY REPORTED AND ANOTHER SAME MODEL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937372&pc=MTA | 2023826-2015-00907 | 2015-06-11 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED INJURY. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2015 BUT IT ALSO TORE - SEE MFR REPORT (B)(4) FOR THE EXCHANGED LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4949683&pc=MTA | 2023826-2015-00896 | 2015-07-01 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A LONGER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5518776&pc=MTA | 2023826-2016-00294 | 2015-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+4.0/117 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO REPORTED INJURY AND THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171634&pc=MTA | 2023826-2015-01402 | 2015-08-10 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS FROM THE LENS BOX AND NOTED THE LENS WAS TORN/BROKEN. THE LENS WAS NOT USED AND THERE WAS NO REPORTED PATIENT INJURY. THE REPORTER INDICATED THE LENS WAS RECEIVED BROKEN. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119124&pc=MTA | 2023826-2015-01304 | 2015-08-20 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/X084 DIOPTER, AND THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218084&pc=MTA | 2023826-2015-01495 | 2015-08-27 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO REPORTED PATIENT INJURY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5185745&pc=MTA | 2023826-2015-01403 | 2015-09-19 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.5/X140 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193757&pc=MTA | 2023826-2015-01443 | 2015-09-28 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED PATIENT INJURY AND WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5166980&pc=MTA | 2023826-2015-01376 | 2015-05-20 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Crack | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/096 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS WAS NOTED TO HAVE A CRACK UPON INSERTION. THERE WAS NO REPORTED PATIENT INJURY AND THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5478497&pc=MTA | 2023826-2016-00231 | 2016-02-04 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Defective Device | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE TECHNICIAN NOTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, APPEARED TO BE DEFECTIVE. THE LENS WAS IN THE VIAL AND WAS NOT USED. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632889&pc=MTA | 2023826-2016-00565 | 2016-04-07 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Defective Device | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON NOTED A DEFECT IN THE LEADING HAPTC OF A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5055258&pc=MTA | 2023826-2015-01094 | 2015-02-21 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contamination with Chemical or Other Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO USED A 13.2MM VICMO13.2, -12.00 DIOPTER IMPLANTABLE COLLAMER LENS. FOREIGN BODY MATERIAL (DUST) WAS NOTED ON THE LENS. THE LENS WAS NOT IMPLANTED. BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5335403&pc=MTA | 2023826-2015-01695 | 2015-11-17 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contamination with Chemical or Other Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/+3.00/X093 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THERE WAS A FIBRE ON THE ICL AFTER INSERTING THE LENS INTO THE EYE. THE LENS WAS EXPLANTED AND EXCHANGED FOR ANOTHER SIMILAR LENS AND SIZE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -17.00/+3.00/X093. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571622&pc=MTA | 2023826-2018-00841 | 2018-03-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT BEFORE THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+2.0/087 (SPHERE/CYLINDR/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD), HE NOTICED A FOREIGN BODY ON THE LENS. REPORTEDLY, A THREAD-SHAPED FOREIGN BODY WAS FOUND ON THE LENS DURING LOADING. THE LENS DID NOT CONTACT THE EYE. THIS OCCURRED ON (B)(6) 2018. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN WITH TORN PIECE MISSING AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. IT SHOULD BE NOTED THAT NO FOREIGN BODY WAS NOTED ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5755949&pc=MTA | 2023826-2016-00869 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contamination with Chemical or Other Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON NOTICED DEBRIS ON A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, DURING LOADING PROCESS. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL CONTACT OR INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4). Manufacturer Narrative: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. "THERE WAS NO PATIENT CONTACT. NO ADDITIONAL INFORMATION WAS PROVIDED." (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197906&pc=MTA | 2023826-2018-00069 | 2017-02-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Damaged by Another Device | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -17.0/+1.5/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DURING IMPLANTATION, THE CARTRIDGE CRACKED DURING INJECTION, CAUSING THE LENS TO TEAR. THE LENS WAS EXPLANTED DURING THE SAME SURGERY AND THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED WITH A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6583369&pc=MTA | 2023826-2017-00865 | 2017-05-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS ROTATED WHEN IT WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS REMOVED AND THERE WAS NO PATIENT INJURY. NO SUTURES USED AND NO ENLARGEMENT OF THE INCISION. Manufacturer Narrative: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND "TWO LENSES RETURNED IN A LENS CASE WRAPPED WITH BUBBLE WRAP AND PACKING SLIP." VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6500835&pc=MTA | 2023826-2017-00663 | 2015-09-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18/+4.5/112 DIOPTER, WAS NOTED TO BE TORN/ BROKEN BEFORE INJECTION INTO THE PATIENT'S EYE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. FOREIGN MATERIAL (HAIR) WAS FOUND ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6499917&pc=MTA | 2023826-2017-00656 | 2016-03-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+0.5/142 DIOPTER, WAS NOTED TO BE TORN/ BROKEN BEFORE INJECTION INTO THE PATIENT'S EYE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. ANOTHER TORRIC LENS OF THE SAME MODEL AND SIMILAR DIOPTER WAS USED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6415472&pc=MTA | 2023826-2017-00418 | 2016-03-16 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, TORE DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF HAPTIC TORN OFF AND MISSING. THE LENS OPTIC WAS SCRATCHED. THE LENS WAS RETURNED IN LIQUID. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6507935&pc=MTA | 2023826-2017-00658 | 2016-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD), BUT THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND FIBER ON LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6291276&pc=MTA | 2023826-2017-00202 | 2016-12-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL13.2, -14.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED. THE SURGEON NOTICED THE LENS WAS DAMAGED AFTER IT WAS IMPLANTED. THE LENS WAS REMOVED AND THE BACK UP LENS WAS IMPLANTED WITH NO ISSUES. THE REPORTED STATED THERE WAS NO PATIENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A LENS CASE. VISUAL INSPECTION FOUND THE HAPTIC TORN, PIECES OF THE HAPTIC MISSING AND DRY RESIDUE ALL OVER LENS. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR CLAIM WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393828&pc=MTA | 2023826-2017-00333 | 2017-01-10 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, TORE DURING LOADING. THERE WAS PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED IT WAS NOT DUE TO USER ERROR. THE REPORTER WAS UNABLE TO PROVIDE FURTHER INFORMATION. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ONE HAPTIC. WORK ORDER SEARCH: ONE SIMILAR CLAIM WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6344330&pc=MTA | 2023826-2017-00249 | 2017-01-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-20 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.N Manufacturer Narrative: WORK ORDER SEARCH: TWO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6432876&pc=MTA | 2023826-2017-00342 | 2017-02-15 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, TORE DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6473686&pc=MTA | 2023826-2017-00606 | 2017-02-28 | Injury | STAAR SURGICAL COMPANY | 2017-04-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS TORE PRIOR TO INJECTION INTO THE EYE. REPORTER STATED THAT THERE WAS RESISTANCE WHEN PULLING THE LENS INTO THE TIP OF THE CARTRIDGE. SUGGESTED THE POSSIBILITY THAT THERE MAY NOT HAVE BEEN ENOUGH BSS TO ACTIVATE THE COATING OF THE CARTRIDGE OR THAT THERE WAS NOT ENOUGH VISCOELASTIC USED, CREATING "DRY SPOTS." THE LENS TORE DUE TO TOO MUCH FORCE APPLIED WHILE PULLING. AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6675642&pc=MTA | 2023826-2017-00985 | 2017-04-13 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, AND THE LENS TORE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6586333&pc=MTA | 2023826-2017-00882 | 2017-05-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON WAS LOADING A MICL13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS AND NOTICED THAT THE LENS WAS TORN. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEMS. Manufacturer Narrative: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID; IN A LENS VIAL. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC. WORK ORDER SEARCH: NO SIMILAR CLAIM TYPE WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6583414&pc=MTA | 2023826-2017-00866 | 2017-05-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL12.6, -16.0 DIOPTER, TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. THE REPORTER WAS NOT ABLE TO DETERMINE IF THE EVENT WAS CAUSED DUE TO USER ERROR. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND "TWO LENSES RETURNED IN A LENS CASE WRAPPED WITH THE PACKING SLIP AND BUBBLE WRAP". VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC. NO SIMILAR CLAIM TYPE WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6679332&pc=MTA | 2023826-2017-01065 | 2017-05-16 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, IN THE PATIENTS RIGHT EYE WITHOUT DIFFICULTY AND NOTICED THERE WAS A TEAR ON THE TRAILING LEFT HAPTIC. THE LENS WAS REMOVED AND THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES. THE REPORTER STATED THAT THERE WAS NO PATIENT INJURY AND THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6685149&pc=MTA | 2023826-2017-01103 | 2017-05-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON HAD TWO 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENSES, -7.0 DIOPTER, THAT TORE DURING FOLDING. THE PRIMARY LENS HAD PATIENT CONTACT AND THE BACKUP LENS DID NOT. THE REPORTER STATED THAT THE CAUSE OF THE EVENT COULD HAVE BEEN DUE TO A LOADING ISSUE. THERE WAS NO PATIENT INJURY AND THE SURGEON DECIDED TO IMPLANT THE BACKUP LENS OF THE OPPOSITE EYE TO COMPLETE THE SURGERY. THE REPORTER WAS NOT SURE WHICH LENS WAS THE PRIMARY LENS THAT HAD PATIENT CONTACT BECAUSE THE LABEL WAS REMOVED WITH THE WRAP. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID ,IN LENS VIAL. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC. ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6655146&pc=MTA | 2023826-2017-00979 | 2017-05-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS,-13.5 DIOPTER, AND NOTED THE LENS WAS TORN. THE LENS WAS CUT OUT AND REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY, NO INCISION ENLARGEMENT, AND NO SUTURES WERE REQUIRED. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM AND THE PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6675648&pc=MTA | 2023826-2017-00986 | 2017-06-01 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, FROM THE VIAL. THE SURGEON WAS PREPARING TO LOAD THE LENS, AND NOTED A PUNCTURE/TEAR ON THE LENS. THE SURGEON DECIDED NOT TO USE THE LENS AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THERE WAS NO PATIENT CONTACT AND IT WAS UNKNOWN WHEN THE LENS WAS DAMAGED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7001416&pc=MTA | 2023826-2017-01678 | 2017-08-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017 AND THE LENS TORE/BROKE. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY. ANOTHER LENS WAS IMPLANTED 2 WEEKS LATER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS EXCHANGE RESOLVED THE PROBLEM AND THE EXACT DATE OF THE LENS EXCHANGE WAS UNKNOWN. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC WAS MISSING, HAPTIC TORN, AND CLEAR RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION CODE:OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7232833&pc=MTA | 2023826-2018-00155 | 2017-10-24 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), AND THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY. ANOTHER SAME MODEL, DIFFERENT DIOPTER LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN BUT THE LENS MAY HAVE BEEN SQUEEZED IN THE INJECTOR. Manufacturer Narrative: CORRECTED DATA: DATE OF THIS REPORT: IN THE INITIAL MDR, 05-JAN-2018 IS CORRECTED TO 04-JAN-2018. DATE RECEIVED BY MANUFACTURER: IN THE INITIAL MDR, 05-JAN-2018 IS CORRECTED TO 04-JAN-2018. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114230&pc=MTA | 2023826-2017-01864 | 2017-11-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A MICL13.7, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. PRIOR TO INJECTION, THE SURGEON NOTICED THAT THE LENS HAD A DENT. HE SAID THAT HE WAS UNSURE IF THE LENS CAME DENTED OR IF IT OCCURRED DURING LOADING. HE DID NOT IMPLANT THE LENS AND THERE WAS NO PATIENT CONTACT OR INJURY. HE USED A BACK-UP LENS AND IT WAS SUCCESSFUL. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL AND "SCRATCH INDENT ON LENS WHILE LOADING" WRIITEN ON BOX. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7617155&pc=MTA | 2023826-2018-00922 | 2017-11-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), IT WAS DISCOVERED THAT IT HAD TORN DURING INJECTION. THE LENS WAS REMOVED INTRAOPERATIVELY ON (B)(6) 2017. ON (B)(6) 2017 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE CAUSE OF THIS EVENT IS REPORTED AS UNKNOWN AND NO PATIENT INJURY IS INDICATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197908&pc=MTA | 2023826-2018-00073 | 2017-11-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED DURING THE SAME SURGERY DUE TO THE LENS BEING FOUND TORN AFTER INJECTION INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGE FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE LENS MIGHT HAVE TORN WHEN IT WAS PULLED WITH THE FORCEP. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336334&pc=MTA | 2023826-2018-00394 | 2018-02-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM MICL12.6,IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, WOULD NOT LOAD PROPERLY/LOADING ISSUE. THERE WAS NO PATIENT CONTACT WITH THE PRODUCT AND THE BACK UP LENS WAS USED WITH NO PROBLEM. THE REPORTER STATED THAT THE CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF OPTIC AND HAPTIC MISSING. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7499554&pc=MTA | 2023826-2018-00705 | 2018-02-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. DURING INJECTION, THE LENS TORN. THE LENS WAS INTRAOPERATIVELY EXCHANGED FOR ANOTHER LENS AND IT SUCCESSFUL. THERE WAS NO PATIENT INJURIES AND THE PATIENT IS IN GOOD CONDITION. THE CAUSE OF THE TEAR IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CORRECTED DATA: IT IS CORRECTED TO "(B)(6) 2018" IN THE INITIAL MDR. EXPIRATION DATE: IT IS CORRECTED TO "31-OCT-2020" IN THE INITIAL MDR. IT IS CORRECTED TO "(B)(6) 2018" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4)-SECONDARY SURGERY, LENS EXCHANGE. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7495137&pc=MTA | 2023826-2018-00671 | 2018-02-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/081 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH PRESENCE OF DEBRIS ON THE LENS SURFACE. CORRECTED DATA: DATE OF THIS REPORT: IT IS CORRECTED TO "(B)(6) 2018" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED TO "14-APR-2018" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7495932&pc=MTA | 2023826-2018-00716 | 2018-03-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-08 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. INTRAOPERATIVELY, THE LENS WAS EXPLANTED. NO PATIENT INJURY IS REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7434154&pc=MTA | 2023826-2018-00553 | 2018-03-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+1.5/108 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) AND THE SURGEON NOTED THE LENS WAS TORN. THE LENS WAS CUT TO REMOVE FROM THE EYE DURING THE SAME SURGERY. THERE WAS NO PATIENT INJURY. THE TORN PIECE OF THE LENS REMAINED IN THE CARTRIDGE. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL AND HAD RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC, HAPTIC TORN AND THE LENS RETURNED IN THREE PIECES. THERE WAS PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7434227&pc=MTA | 2023826-2018-00554 | 2018-03-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/088 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) AND THE SURGEON NOTED THE LENS WAS TORN. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/FLUIDS. THE LENS WAS REMOVED FROM THE EYE DURING THE SAME SURGERY. THERE WAS NO PATIENT INJURY. THE TORN PIECE OF THE LENS REMAINED IN THE CARTRIDGE. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS TORN IN THREE PIECES AND A PIECE OF HAPTIC MISSING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571590&pc=MTA | 2023826-2018-00848 | 2018-04-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), HE NOTICED IT HAD TORN. THE LENS WAS REMOVED INTRAOPERATIVELY ON (B)(6) 2018. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE REPORTER STATES THAT POSSIBLY THE LENS WAS TORN DURING LOADING "DUE TO THE POOR STATUS OF ICL FORCEPS." THE TEAR WAS NOT DISCOVERED UNTIL AFTER THE LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE TORN PIECE MISSING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7539102&pc=MTA | 2023826-2018-00767 | 2018-04-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS NOTICED THAT AFTER IMPLANTATION, THE LENS HAD A SCRATCH. THE SURGEON INTRAOPERATIVELY EXPLANTED THE LENS WITH A FORCEP INSTRUMENT AND IMPLANTED ANOTHER STAAR LENS. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM NUMBER (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7607748&pc=MTA | 2023826-2018-00914 | 2018-05-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/082 (SPHERE/CYLINDER/AXIS), FROM THE LENS VIAL AND NOTICED, WHILE LOADING, THE LENS WAS TORN. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL LENS WAS IMPLANTED. THE SURGEON FELT THE LENS WAS DEFECTIVE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7629848&pc=MTA | 2023826-2018-00947 | 2018-05-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.5/+4.5/099 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY , WITH NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN; HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7603593&pc=MTA | 2023826-2018-00906 | 2018-05-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON NOTICED THAT A 13.2 MM MICL13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS TORN DURING INSERTION. THE LENS WAS REMOVED AND THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES. THERE WAS NO PATIENT INJURY, NO INCISION ENLARGEMENT, AND NO SUTURE. THE CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE TORN PIECE MISSING. CORRECTED DATA: DATE RECEIVED BY MANUFACTURER: INITIAL MDR SHOULD BE 18 MAY 2018. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7636106&pc=MTA | 2023826-2018-00968 | 2018-05-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7665042&pc=MTA | 2023826-2018-01014 | 2018-05-31 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY. ON (B)(6) 2018 AND ALTERNATE LENS WAS IMPLANTED. REPORTEDLY, NO PATIENT INJURY OCCURRED. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. METHOD CODE CORRECTED: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6111712&pc=MTA | 2023826-2016-01653 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, TORE DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THERE WAS NO ENOUGH LENS IN THE FORCEPS AND THE LENS WAS TORN. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6415278&pc=MTA | 2023826-2017-00429 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON WAS LOADING A MICL12.6, -12.0 DIOPTER LENS AND NOTICED THE LENS WAS DAMAGED. HE DID NOT KNOW IF THE DAMAGE OCCURRED DURING THE LOADING PROCESS OR WHAT CAUSED THE DAMAGE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER WAS UNABLE TO PROVIDE FURTHER INFORMATION. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND MISSING PIECES OF THE HAPTIC. NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218046&pc=MTA | 2023826-2015-01411 | 2015-07-02 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND EXCHANGED FOR ANOTHER ICL, SAME MODEL. NO ADVERSE EVENT WAS REPORTED. LAST POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: (B)(4). ADDITIONAL DATA: CONCLUSION CODE: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENT. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21 - 45 YEARS OF AGE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668407&pc=MTA | 2023826-2016-00717 | 2016-04-18 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -08.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS TORE DURING INJECTION/INSERTION INTO THE EYE. DURING THE SAME PROCESURE, THE LENS WAS REMOVED AND EXCHANGED FOR ANOTHER LENS, SAME MODEL AND DIOPTER SIZE. NO PATIENT INJURY REPORTED. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THERE IS NO BREAK OR TORN MATERIAL ASSOCIATED WITH A MATERIAL INTEGRITY ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5216008&pc=MTA | 2023826-2015-01412 | 2015-07-03 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AN EXCHANGED FOR ANOTHER ICL, SAME MODEL AND DIOPTER SIZE. NO ADVERSE EVENT WAS REPORTED. LAST POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE OPTIC WAS TORN/SPLIT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.50. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323172&pc=MTA | 2023826-2015-01722 | 2015-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Advance; Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER BUT THE LENS BECAME STUCK IN THE CARTRIDGE AND TORE. THE REPORTER STATED IT WAS UNKNOWN IF THERE WAS ANY PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. MEDICAL REVIEW - REPORTEDLY, MICL (13.2 MM) WAS STUCK IN THE INJECTOR AND WAS SUBSEQUENTLY TORN OFF DURING INSERTION REQUIRING INTRAOPERATIVE EXCHANGE. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ACCORDING TO REPORT , BY AN OFFICE MANAGER, DATED ON (B)(6) 2015 THERE WAS NO CONFIRMATION IF THE LENS HAD A PATIENT CONTACT. NO REPORTED POSTOPERATIVE SEQUELAE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND TWO LENSES STUCK TOGETHER IN ONE VIAL. THE LENSES WERE RETURNED DRY AND COULD NOT BE EVALUATED DUE TO THE RETURNED CONDITION. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW AND THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575684&pc=MTA | 2023826-2016-00418 | 2016-03-14 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Fold, Unfold or Collapse | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, AND THE LENS WOULD NOT UNFOLD IN THE PATIENT'S EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO APPARENT INJURY AND ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PT AGE AND WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680583&pc=MTA | 2023826-2016-00760 | 2016-04-14 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Fold, Unfold or Collapse | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS WOULD NOT UNFOLD CORRECTLY. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5693629&pc=MTA | 2023826-2016-00796 | 2016-05-02 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Fold, Unfold or Collapse | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, AND THE LENS WOULD NOT UNFOLD IN THE PATIENT'S EYE. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULTS (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5983980&pc=MTA | 2023826-2016-01328 | 2016-07-21 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Fold, Unfold or Collapse | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON HAD PROBLEMS FOLDING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, AND THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6277723&pc=MTA | 2023826-2017-00160 | 2017-01-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-25 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Fold, Unfold or Collapse | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, WOULD NOT FOLD VERY WELL IN THE CARTRIDGE DURING THE LOADING PROCEDURE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THAT CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248830&pc=MTA | 2023826-2015-01574 | 2015-10-27 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Insert; Positioning Problem | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, BUT THE LENS UNFOLDED AND WAS ONLY PARTIALLY INSERTED. THE LENS WAS REMOVED WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. REPORTEDLY, MICL WAS ONLY PARTIALLY INSERTED IN THE EYE WHEN THE SURGEON OBSERVED UNEXPECTED UNFOLDING AND DECIDED TO REMOVE/EXCHANGE THE LENS. NO REPORT OF INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ANOTHER LENS WAS SUCCESSFULLY IMPLANTED. ACCORDING TO THE REPORT BY A NURSE, DATED ON (B)(6) 2015, THE CAUSE OF THE EVENT WAS UNKNOWN. THE SAME REPORT STATED THAT THERE WAS NO PATIENT INJURY DURING REMOVAL OF A DAMAGED LENS. NO REPORTED POSTOPERATIVE SEQUELAE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PT AGE AND WEIGHT: UNK. METHOD - 3331 (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE CARTRIDGE AND INJECTOR WERE RETURNED WITH NO VISIBLE DAMAGE TO EITHER DEVICE. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5072589&pc=MTA | 2023826-2015-01213 | 2015-08-11 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Open or Close | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON LOADED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE SFC-45 FP CARTRIDGE, BUT WAS UNABLE TO CLOSE THE CARTRIDGE. THE SURGEON FELT THE LENS WAS TOO BIG FOR THE CARTRIDGE BUT WAS NOT SURE IF THE EVENT WAS DUE TO THE LENS OR CARTRIDGE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. SEE MFR. 2023826-2015-01219 FOR THE CARTRIDGE. Manufacturer Narrative: EVALUATION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, CARTRIDGE AND FOAM TIP PLUNGER LOT SEARCHES AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS AND INJECTION SYSTEM NOT RETURNED. METHOD: WORK ORDER SEARCH, CARTRIDGE LOT NUMBER SEARCH, FOAM TIP PLUNGER LOT NUMBER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. A CARTRIDGE LOT NUMBER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. A FOAM TIP PLUNGER LOT NUMBER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY, LENS WORK ORDER SEARCH, CARTRIDGE AND FOAM TIP LOT NUMBER SEARCHES, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5125127&pc=MTA | 2023826-2015-01320 | 2015-08-31 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult To Position | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, BUT WHEN THE SURGEON ATTEMPTED TO TUCK THE HAPTICS, THE SURGEON NOTED THE HAPTICS WERE COVERING THE PERIPHERAL IRIDOTOMIES. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE LENS WAS TORN UPON REMOVAL. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: PER MEDICAL REVIEW - MALPOSITION OF THE HAPTICS COULD BE DUE TO PATIENT ANATOMY, PATHOLOGY SUCH AS CILIARY CYSTS, FAULT IN LENS MANUFACTURE, OR DUE TO SURGEON HANDLING. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119132&pc=MTA | 2023826-2015-01259 | 2015-09-03 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Failure to Unfold or Unwrap | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER BUT THE LENS WOULD NOT UNFOLD IN THE PATIENT'S EYE, THE LENS WAS ON ITS SIDE AND LOOKED LIKE A "TACO". THE LENS WAS REMOVED WITH NO PATIENT INJURY, NO ENLARGED INCISION OR SUTURES AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS DID NOT UNFOLD. THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF UNFOLDING WAS DUE TO IMPROPER LENS LOADING, OR SURGEON HANDLING EITHER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795473&pc=MTA | 2023826-2019-01231 | 2019-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Failure to Unfold or Unwrap | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2019. THE SURGEON REPORTS THE LENS "WOULDN'T LAY FLAT-LOOKED LIKE THERE WAS A FOLD IN IT." INTRAOPERATIVELY, THE SURGEON REMOVED THE LENS AND SUCCESSFULLY IMPLANTED A BACK-UP LENS. AFTER THE INITIAL LENS WAS REMOVED, NOTHING OBVIOUS WAS NOTED ON THE LENS. Manufacturer Narrative: "PHAKIC INTRAOCULAR LENS, PRODUCT CODE: MTA" SHOULD BE CORRECTED TO "PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: SURGERY OCCURRED IN THE LEFT (OS) EYE. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNKNOWN. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146293&pc=MTA | 2023826-2015-01248 | 2015-03-30 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Fracture | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND EXCHANGED FOR ANOTHER LENS, SAME MODEL AND DIOPTER LENS. NO INJURY TO THE PATIENT REPORTED. POST-OP VISIT ON (B)(6) 2015 - UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154928&pc=MTA | 2023826-2015-01266 | 2015-09-16 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Fracture | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -11.0/+5.0/X145 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS FOUND TO HAVE A LATERAL CRACK DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED. NO OTHER LENS WAS IMPLANTED. NO INJURY TO THE PATIENT REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT. IS NOT. MARKETED IN THE U.S. DIOPTER: -11.0/+5.0/X145. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146288&pc=MTA | 2023826-2015-01243 | 2015-08-07 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Fracture; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE AND WAS IMPLANTED UPSIDE DOWN. THE LENS WAS REMOVED AND EXCHANGED FOR ANOTHER LENS, SAME MODEL AND DIOPTER LENS. NO INJURY TO THE PATIENT REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -09.50. DEVICE EVALUATED BY MANUFACTURER? NO.: LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6385264&pc=MTA | 2023826-2017-00334 | 2017-02-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL13.2, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN DUE TO A LOADING ERROR. THE LENS WAS REMOVED DURING THE INITIAL SURGERY AND THE BACK UP LENS WAS IMPLANTED. THE REPORTER STATED THERE WERE NO ADVERSE EVENTS WITH THIS CASE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: NO SIMILAR CLAIM WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983358&pc=MTA | 2023826-2020-00934 | 2020-03-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -15.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS REMOVED AND EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116158&pc=MTA | 2023826-2020-01165 | 2020-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. INTRAOPERATIVELY, THE LENS WAS REMOVED AND EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116192&pc=MTA | 2023826-2020-01164 | 2020-05-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS REMOVED AND EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4702455&pc=MTA | 2023826-2015-00446 | 2015-02-23 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. LOW VAULT WAS NOTED ON (B)(6) 2015 AND AN ICL EXCHANGE FOR A LONGER LENS IS NEEDED. PATIENT'S LAST KNOWN BCVA WAS 20/20 AS OF (B)(6) 2015. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) - NO CONSEQUENCES OR IMPACT TO PATIENT; SIZE INCORRECT FOR PATIENT; (IMPLANTED); (VAULTING). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047799&pc=MTA | 2023826-2015-01056 | 2015-08-05 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, AND NOTED THE LENS WAS TOO LONG FOR THE PATIENT'S EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY AND A SHORTER LENS WAS IMPLANTED. THE PATIENT IS DOING WELL AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PER MEDICAL REVIEW - ACCURATE DETERMINATION OF ANTERIOR SEGMENT DIMENSIONS IS KEY TO THE SAFETY OF IMPLANTED ICLS. THE METHOD OF ICL SIZING FOR THIS PATIENT IS NOT KNOWN. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW, EVALUATION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6424882&pc=MTA | 2023826-2017-00399 | 2017-02-22 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A MICL13.7 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, AND SAW THAT IT WAS TOO BIG. THE LENS WAS REMOVED DURING THE SAME SURGERY AND A SMALLER LENS WAS IMPLANTED. THERE WAS NO INJURY TO THE PATIENT, NO ENLARGEMENT OF THE INCISION AND NO SUTURES WERE NEEDED. Manufacturer Narrative: LENS WAS RETURNED IN A SMALL AMOUNT OF LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6727165&pc=MTA | 2023826-2017-01214 | 2017-06-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, WAS IMPLANTED AND REMOVED. THE REPORTER INDICATED THAT THE LENS WAS REMOVED DUE TO LENGTH, VAULT TOO HIGH, WILL NEED TO GO WITH A 13.2MM LENGTH LENS, THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872460&pc=MTA | 2023826-2017-01446 | 2017-08-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT WAS NOTICED. THE PATIENT IS FINE AND THERE IS NO PLAN FOR FUTURE SURGERY. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: UNK. UNK. UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021347&pc=MTA | 2023826-2017-01664 | 2017-10-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ABOUT A MINUTE AFTER IMPLANTATION, THE SURGEON NOTICED EXCESSIVING VAULTING. INTRAOPERATIVELY, THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED WITH THE PATIENT CURRENTLY BEING FINE AND SEEING 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169879&pc=MTA | 2023826-2018-00052 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS REMOVED DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7239350&pc=MTA | 2023826-2018-00162 | 2017-12-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTED STATED THAT THE SURGEON IMPLANTED A MICL13.7, -10. DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED DURING THE SAME SURGERY DUE TO THE LENS SIZE . THERE WAS NO PATIENT INJURY. IT WAS REPORTED THAT THE SURGEON HAS BEEN REQUESTING A UBM ON EACH PATIENT WITH SULCUS TO SULCUS MEASUREMENT RECORDED. THE WHITE TO WHITE READING FROM THE UBM HAS BEEN RUNNING LARGER THAN THEIR WHITE TO WHITE READING FROM THEIR DIGITIAL CALIPER. THE REPORTER INDICATED THAT THE ICL CHOSEN WAS NOT A MISTAKE ON BEHALF OF THE ONLINE CALCULATOR JUST A CONFLICT IN UBM AND DITIGIAL CALIPER MEASUREMENTS AND IN THE FUTURE THEY WOULD LIKE TO ORDER THE ICL SIZE FOR BOTH MEASUREMENTS TO ENSURE MOVING FORWARD WITH ANY LAST MINUTE CHANGES THAT MAY NEED TO BE MADE. Manufacturer Narrative: CORRECTED DATA: IN THE INITIAL MDR, "UNK" IS NOT APPLICABLE THE CORRECT DATE IS (B)(6) 2017. IN THE INITIAL MDR, "UNK" IS NOT APPLICABLE THE CORRECT DATE IS (B)(6) 2018. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY,CLEAR RESIDUE ON THE LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594228&pc=MTA | 2023826-2018-00901 | 2018-05-14 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS ON (B)(6) 2018. THE LENS WAS INTRAOPERATIVELY EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. AS THE LENS WAS EXPLANTED, PIECES OF IT BROKE OFF. ALL THE PIECES WERE COLLECTED. THERE WAS NO PATIENT INJURY AND THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: PRIOR ENTRY OF 'MALE' IS INCORRECT, PATIENT IS 'FEMALE'. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE. VISUAL INSPECTION FOUND HAPTIC TORN, TORN HAPTIC PIECE MISSING. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7799160&pc=MTA | 2023826-2018-01245 | 2018-07-19 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM MICL12.6, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO LOW VAULT THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED DRY, IN VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT PATIENT EXPERIENCED EXCESSIVE VAULT AND DID NOT EXPERIENCE LOW VAULT AS INITIALLY REPORTED. (B)(4) Manufacturer Narrative: CORRECTED DATA: TYPE OF REPORTABLE EVENT IS MALFUNCTION WHICH IS CORRECTED FROM SERIOUS INJURY IN THE PREVIOUS MDRS. CLAIM #: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356063&pc=MTA | 2023826-2019-00281 | 2019-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING A LOW VAULT, BUT AT 3 DAYS POST-OPERATIVE, THE VAULT HAD IMPROVED AND THE PATIENT WAS BEING MONITORED. THE PATIENT HAD NO COMPLAINTS AND WAS DOING VERY WELL. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356308&pc=MTA | 2023826-2019-00282 | 2019-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING A LOW VAULT, BUT AT 3 DAYS POST-OPERATIVE, THE VAULT HAD IMPROVED AND THE PATIENT WAS BEING MONITORED. THE PATIENT HAD NO COMPLAINTS AND WAS DOING VERY WELL. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7294290&pc=MTA | 2023826-2018-00268 |  | Malfunction | Unknown Manufacturer | 2018-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | NA | 1 | Event Description: INCREASED VAULT WITH NO SYMPTOMS WAS REPORTED. EVERYTHING IS FINE WITH PATIENT, DOCTOR IS NOT WORRIED ABOUT THE EYE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248528&pc=MTA | 2023826-2015-01575 | 2015-11-03 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, BUT THE SURGEON NOTED THE LENS WAS TOO LONG FOR THE PATIENT'S EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY AND A SHORTER LENS WAS IMPLANTED. Manufacturer Narrative: RESULTS (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE LENS HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, MICL (12.6MM) WAS EXPLANTED/EXCHANGED INTRAOPERATIVELY FOR A SHORTER SIZE MICL TO ADDRESS SURGEON`S OBSERVATION UPON INSERTION THAT LENS WAS TOO LONG FOR THE PATIENT`S EYE. NO REPORTED PROBLEM WITH THE LENS OR INJECTOR SYSTEM. INCISION WAS NOT ENLARGED AND NO OTHER TISSUE DAMAGE WAS REPORTED. ACCORDING TO THE REPORT, BY A NURSE, DATED ON (B)(6) 2015 THE CAUSE OF THE EVENT WAS UNKNOWN AND THERE WAS NO PATIENT INJURY. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5720699&pc=MTA | 2023826-2016-00615 | 2015-11-23 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/60. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.93MM. CLAIM # (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4) .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5395120&pc=MTA | 2023826-2016-00107 | 2016-01-05 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, AND THE SURGEON NOTED THE VAULT OF THE LENS WAS HIGH. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY OR ANY COMPLICATIONS. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: (B)(4). RESULTS: (OPERATIONAL PROBLEM): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DEFORMATION OF THE LENS, OR RELEASE OF NONCONFORMING LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE CAUSED FAILURE TO THE LENS. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT , A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5905512&pc=MTA | 2023826-2016-01160 | 2016-07-25 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, BUT THE SURGEON FELT THE LENS WAS NOT LOADED CORRECTLY, SO THE LENS WAS REMOVED. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9873930&pc=MTA | 2023826-2020-00636 | 2019-04-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Mechanical Problem | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ACCIDENTLY IMPLANTED A 12.6MM, VICMO12.6, -7.00 DIOPTER IMPLANTABLE COLLAMER LENS UPSIDE DOWN. THE LENS WAS IMPLANTED, REMOVED AND EXCHANGED ON (B)(6) 2020 WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THAT A LENS TEAR WAS NOTICED BY THE SURGEON WHEN SHE REMOVED THE LENS FROM THE PATIENTS EYE AND THAT THE LENS WAS CONTAMINATED WHEN THEY REMOVED IT " THE LENS WAS REMOVED AND CONTACT TO THE NON-ASEPTIC AREAS." Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233854&pc=MTA | 2023826-2019-00083 | 2017-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, WITHIN THE SAME SURGERY, DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254849&pc=MTA | 2023826-2019-00095 | 2018-03-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS INTRA-OPERATIVELY. THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. METHOD CODE ADDED, RESULT CODE UPDATED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7728853&pc=MTA | 2023826-2018-01139 | 2018-03-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE EXCHANGE LENS DID NOT ENTIRELY RESOLVE THE RESIDUAL PROBLEM. UPON POST-OP EXAMINATION, THE RESIDUE IS STILL VISIBLE ON THE PERIPHERAL OF THE OPTIC , BUT NOT VISUALLY SIGNIFICANT FOR THE PATIENT. A NEW CARTRIDGE WAS USED TO IMPLANT THIS EXCHANGE, LOT NUMBER IS UNKNOWN. IT WAS REPORTED THAT THE SAME FOAM TIP PLUNGER THAT WAS USED FOR THE INITIAL LENS IMPLANT WAS USED FOR THIS LENS. LENS WAS LEFT IN PLACE AND PATIENT'S REFRACTION IS AS INTENDED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721255&pc=MTA | 2023826-2018-01082 | 2018-04-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE REMOVING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/104 (SPHERE/CYLINDER/AXIS) FROM THE VIAL, THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE LENS WAS RETURNED IN PIECES AND WAS UNABLE TO BE EVALUATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853017&pc=MTA | 2023826-2018-01338 | 2018-04-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), IT TORE. THE LENS WAS REMOVED INTRAOPERATIVELY ON (B)(6) 2018. DURING THE SAME OPERATION, AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THERE WAS NO PATIENT INJURY AND THE CAUSE IS LISTED AS UNKNOWN. Manufacturer Narrative: RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740020&pc=MTA | 2023826-2018-00998 | 2018-06-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S LEFT EYE (OS), THE LENS BROKE. THE LENS WAS IMPLANTED ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY AND REPORTEDLY, THE "PATIENT IS FINE." Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237928&pc=MTA | 2023826-2019-00084 | 2018-06-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, WITHIN THE SAME SURGERY, DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC TORN OFF. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721254&pc=MTA | 2023826-2018-01076 | 2018-06-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), IT TORE. THE LENS WAS REMOVED INTRAOPERATIVELY ON (B)(6) 2018. REPORTEDLY, ON THE SAME DATE, IN A 2ND OPERATION, AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THERE WAS NO PATIENT INJURY AND THE SURGEON STATES THE "LENS SEEMED TO BE TORN DURING THE LOADING [PROCESS] DUE TO CARELESSNESS." Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CORRECTED DATA: SURGEON IMPLANTED LENS INTO THE PATIENT'S "RIGHT EYE (OD)"; STATEMENT OF "PATIENT'S LEFT EYE (OS)" IS INCORRECT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7749557&pc=MTA | 2023826-2018-01158 | 2018-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -14.5/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND THE PATIENT WAS IN GOOD CONDITION AFTER SURGERY. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704799&pc=MTA | 2023826-2018-01053 | 2018-06-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/092 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8133980&pc=MTA | 2023826-2018-01786 | 2018-06-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON OPENED THE BLISTER OF VICH13.2 , + 8.00 DIOPTER, IMPLANTABLE LENS AND HE FOUND THAT THE LENS HAD A HOLE IN THE MIDDLE. THE SURGEON DID NOT KNOW THAT IT WAS NECESSARY TO PERFORM A PREOPERATIVE IRIDECTOMY FOR THIS LENS MODEL. HE CONTACTED THE TECHNICAL DEPARTMENT WHO EXPLAINED THE SURGICAL PROCEDURE FOR THIS TYPE OF LENS. THE SURGEON HAD NOT BEEN PROPERLY TRAINED. THE PROCEDURE HAD TO BE POSTPONED AND ANOTHER IMPLANT WAS PLACED. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DEVICE HISTORY RECORD REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAS BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. ADDITIONAL INFORMATION WAS RECEIVED ON 20 DEC 2018: THE SAME SURGEON THAT REPORTED THIS CASE PERFORMED THE IMPLANTATION OF THE REPLACEMENT LENS. CLAIM#: (B)(4). Manufacturer Narrative: METHOD CODE: "3331" NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. H6-DEVICE HISTORY RECORD (DHR) REVIEW:" DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE". NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716697&pc=MTA | 2023826-2018-01096 | 2018-06-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) . AFTER IMPLANTING THE LENS THE SURGEON SAW DEFECT. THE LENS WAS REMOVED AND REPLACED WITH BACK UP LENS WITHIN SAME SURGERY, NO PATIENT INJURY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION:LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH:NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7770789&pc=MTA | 2023826-2018-01178 | 2018-06-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS IMPLANTING A 12.6, VICM5_12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS AND THE PLUNGER OVERSHOT THE LENS AND IT BECAME STUCK. IT WAS REPORTED THAT THE LENS TORE SURING REMOVAL AND THER WAS NO PATIENT INJURY. MODEL AND LOT NUMBERS FOR DELIVERY SYSTEM ARE NOT AVAILABLE. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PREVIOUS STATEMENT: ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. CORRECTED STATEMENT: ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140923&pc=MTA | 2023826-2020-01223 | 2018-07-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A 12.6MM, VICMO12.6, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2018 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824050&pc=MTA | 2023826-2018-01283 | 2018-07-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/089 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED INTRAOPERATIVELY, THE SURGEON NOTING LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THIS OCCURRED ON (B)(6) 2018. REPORTEDLY, THE LENS WAS TORN DURING LOADING INTO THE CARTRIDGE. 'PATIENT INJURY' IS INDICATED AS NO. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7787039&pc=MTA | 2023826-2018-01216 | 2018-07-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT 13.2MM MICL13.2, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS FLIPPED UPSIDE DOWN IN THE EYE WHEN IT CAME OUT OF THE INJECTOR. THE LENS WAS REMOVED AND THE BACKUP LENS WAS IMPLANTED WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY AND THE CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN SPECIMEN CUP. VISUAL INSPECTION FOUND RESIDUE ON LENS, TORN HAPTIC, TORN HAPTIC PIECE MISSING. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824082&pc=MTA | 2023826-2018-01275 | 2018-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THE LENS DID TOUCH THE EYE BUT WAS NOT IMPLANTED. "NO PATIENT INJURY" IS REPORTED. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824089&pc=MTA | 2023826-2018-01285 | 2018-07-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2018. WITHIN THE SAME SURGERY, AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE REPORTER STATES THE SURGEON THOUGHT THE LENS MAY HAVE BEEN TORN DURING LOADING INTO THE CARTRIDGE BUT WAS NOT DISCOVERED UNTIL IMPLANT. PATIENT INJURY IS REPORTED AS "NO." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS WAS TORN IN TWO PIECES WHICH ARE STUCK TOGETHER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884853&pc=MTA | 2023826-2018-01315 | 2018-07-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO 13.7, -09.50 DIOPTER, INTRAOCULAR LENS TORE DURING INJECTION. THE LENS WAS REMOVED AND EXCHANGED DURING INITIAL SURGERY WITH NO PATIENT INJURY. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT/DEFORMED, RESIDUE ON THE LENS SURFACE, AND PIECES OF THE HAPTIC/OPTIC WERE MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255291&pc=MTA | 2023826-2019-00105 | 2018-07-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE EXCHANGED LENS RESOLVED THE PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140988&pc=MTA | 2023826-2020-01225 | 2018-08-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON 16 AUG 2018 A 12.6MM, VICMO12.6, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2018 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7898391&pc=MTA | 2023826-2018-01432 | 2018-08-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL13.2, -14.5 DIOPTER, IMPLANTABLE COLLAMER LENS HAD A TORN DURING INJECTION/DELIVERY INTO THE EYE. LENS WAS RIMPLANTED AND REMOVED ON (B)(6) 2018 WITHIN THE SAME SURGERY. THERE WAS NO PAITENT INJURY. REPORTER STATED THE CARTIDGE WAS AT FAULT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8408110&pc=MTA | 2023826-2019-00413 | 2018-08-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -15.0/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. PATIENT CONTACT BUT NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2018 A REPLACEMENT LENS WAS IMPLANTED AND THE REPORTER STATED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8171724&pc=MTA | 2023826-2018-01872 | 2018-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT WITH THE DEVICE BUT NO PATIENT INJURY. ON (B)(6) 2018 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8616249&pc=MTA | 2023826-2019-00817 | 2018-09-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED DURING INITIAL SURGERY AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. THERE WAS PATIENT CONTACT WITH THE DEVICE BUT NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC HAD TEARS AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7927074&pc=MTA | 2023826-2018-01468 | 2018-09-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, INTO THE PATIENT'S EYE ON (B)(6) 2018. INTRAOPERATIVELY, THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS. THE REASON FOR THE EXCHANGE IS THAT IT "WAS NOT POSITIONED CORRECTLY." "NO PATIENT INJURY" IS REPORTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372285&pc=MTA | 2023826-2019-00303 | 2018-09-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, - 12.0/+2.0/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: IN PREVIOUS MDR FOLLOW-UP TYPE WAS FOR DEVICE EVALUATION. DEVICE EVALUATED BY MANUFACTURER: YES; DEVICE EVALUATED BY MANUFACTURER. RESULT CODE 114: OPERATIONAL PROBLEM IDENTIFIED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7947554&pc=MTA | 2023826-2018-01487 | 2018-09-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-09 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.1, -13.50/1.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION. THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS LATER IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. CORRECTED DATA: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 -13.5/+1.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY; REPORTER STATES NO PATIENT INJURY. ON (B)(6) 2018 A REPLACEMENT LENS WAS SUCCESSFULLY IMPLANTED, RESOLVING THE PROBLEM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699730&pc=MTA | 2023826-2019-01052 | 2018-09-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VICMO 12.6, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, ON (B)(6) 2018. A REPLACEMENT LENS WAS IMPLANTED DURING INITIAL SURGERY AND PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8065575&pc=MTA | 2023826-2018-01623 | 2018-09-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -18.00/1.5/017 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING LOADING IN THE CARTRIDGE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE SURGERY WAS ABORTED AND A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2018. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED, LOADING AND SURGERY WENT WELL AND THE PATIENT IS DOING WELL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN A ND THE LENS WAS CURVED IN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8299125&pc=MTA | 2023826-2019-00211 | 2018-09-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -6.50/2.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THE TORN LENS WAS REMOVED ON ORIGINAL SURGERY DATE OF (B)(6) 2018 WITH NO INJURY TO THE PATIENT. A REPLACEMENT LENS WAS ORDERED ON THE SAME DAY OF THE INCIDENT AND IMPLANTED ON (B)(6) 2018 WITH NO INJURY TO THE PATIENT AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THAT THE CAUSE OF THIS EVENT WAS USER ERROR " LENS OUT OF AXIAL POSITION TURNED LENS AND LENS TORE". Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8028760&pc=MTA | 2023826-2018-01549 | 2018-10-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING INSERTION INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED FROM EYE DURING INITIAL SURGERY ON (B)(6) 2018 WITH NO PATIENT INJURY. BACKUP LENS WAS IMPLANTED AND RESOLVED PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND DEBRIS ON THE LENS AND ONLY HALF OF THE LENS RETURNED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8081339&pc=MTA | 2023826-2018-01703 | 2018-10-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, THE PATIENT MOVED AND THE LENS BENT UPWARDS. THE LENS WAS PARTIALLY INSERTED AND WAS REMOVED DURING THE SAME SURGERY, WITH NO PATIENT INJURY. THE SURGEON WAS NOT SURE IF THE LENS WAS DAMAGED, SO DECIDED TO IMPLANT THE BACKUP LENS WITH NO PROBLEM. THE SURGEON STATED THE CAUSE OF THE EVENT WAS PATIENT RELATED. Manufacturer Narrative: THE LENS WAS RETURNED IN DRY, IN A CARTRIDGE CASE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113536&pc=MTA | 2023826-2018-01757 | 2018-10-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/073 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED WITH ANOTHER SAME MODEL/DIOPTER LENS ON (B)(6) 2018. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT; PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8092546&pc=MTA | 2023826-2018-01686 | 2018-10-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF THE LENS OPTIC AND ONE HAPTIC TORN AND BROKEN. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8186162&pc=MTA | 2023826-2018-01878 | 2018-11-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 13.2 MM, MICL13.2, -3.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IN THE PROCESS OF BEING IMPLANTED, BUT IT WOULD NOT UNFOLD CORRECTLY OUT OF THE CARTRIDGE WHEN IT WAS PARTIALLY IMPLANTED IN THE EYE. THE LENS WAS THEN REMOVED IN ORDER TO RELOAD IT INTO THE CARTRIDGE AGAIN, WHEN IT WAS NOTED THAT IT WAS EXPIRED. THE BACKUP LENS WAS THEN OPENED AND IMPLANTED DURING THE SAME SURGERY. PER REPORTER THE BACK UP LENS UNFOLDED PERFECTLY. CAUSE OF EVENT WAS UNKNOWN. THE REPORTED STATED THAT THERE WAS NO INJURY TO THE PATIENT, CURRENTLY 20/20 AND HAPPY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. FOREIGN MATERIAL WAS OBSERVED ON THE LENS SURFACE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8806864&pc=MTA | 2023826-2019-01201 | 2018-11-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VICMO 12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENTS LEFT EYE (OS). THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED DURING INITIAL SURGERY. A NEW LENS WAS IMPLANTED ON (B)(6) 2019. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS / RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2018. (CORRECTED FROM (B)(6) 2019 IN ORIGINAL MDR). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8203049&pc=MTA | 2023826-2018-01827 | 2018-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS REMAIND FOLDED IN PATIENTS EYE. THE SURGEON REMOVED AND EXCHANGED THE LENS DURING INITIAL SURGERY ON (B)(6) 2018. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE". REPORTEDLY, THERE WAS NO PATIENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS REMAINED FOLDED IN THE PATIENT'S EYE. THE SURGEON REMOVED AND EXCHANGED THE LENS DURING THE INITIAL SURGERY ON (B)(6) 2018. THERE WAS NO PATIENT INJURY AND THE PROBLEM WAS RESOLVED. IN THE REPORTER'S OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255229&pc=MTA | 2023826-2019-00104 | 2018-11-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THEY FELT THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. POST-OP PATIENT CONDITION WAS RESIDUAL ASTIGMATISM OF 1.25 D. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298832&pc=MTA | 2023826-2019-00202 | 2018-11-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, DURING THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255268&pc=MTA | 2023826-2019-00085 | 2018-12-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE EXCHANGE RESOLVED THE PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF LENS OPTIC AND HAPTIC TORN OFF. THE CARTRIDGE WAS RETURNED WITH NO VISIBLE DAMAGE TO THE DEVICE. THERE WAS RESIDUE ON THE LENS AND CARTRIDGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8259211&pc=MTA | 2023826-2019-00119 | 2018-12-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY. IT WAS REPORTED THAT THERE WAS PATIENT CONTACT BUT NO INJURY. Manufacturer Narrative: DEVICE MANUFACTURE DATE CORRECTED TO: 03-MAY-2016. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298708&pc=MTA | 2023826-2019-00189 | 2018-12-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, - 12.0/+1.5/085 (SPHERE/CYLINDER/AXIS) DURING THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPIC AND HAPTIC TORN. PRESENCE OF RESIDUE WAS FOUND ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8333965&pc=MTA | 2023826-2019-00236 | 2018-12-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE, THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. REPORTEDLY, THE SURGEON NOTICED THE LENS DAMAGE DURING THE BEGINNING OF INSERTION. LENS WAS PARTIALLY IMPLANTED THEN REMOVED. ALTERNATE LENS IMPLANTED ALL WITHIN THE SAME SURGERY. "NO PATIENT INJURY" IS NOTED AND THE PATIENT IS NOW "DOING GREAT." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8336924&pc=MTA | 2023826-2019-00239 | 2019-01-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON NOTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, WAS DEFECTED WHEN IT WAS BEING LOADED. THE LENS WAS NOT INSERTED BECAUSE IT DID NOT LOOK RIGHT. THERE WAS NO PATIENT INJURY AND ANOTHER SAME MODEL LENS WAS IMPLANTED. THE SURGEON INDICATED THE LENS WAS DEFORMED AND WAS RECEIVED THAT WAY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS PARTIALLY INSERTED AND THE SURGEON NOTICED THE DEFECT AND DID NOT CONTINUE TO INSERT THE LENS. THE LENS WAS PULLED BACK OUT AND THE BACKUP LENS WAS IMPLANTED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTICS TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8382689&pc=MTA | 2023826-2019-00342 | 2019-01-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT WHILE INJECTING A 12.6MM, VTICMO12.6, -16/2.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(60 2019 THE LENS TORE. THE LENS WAS REMOVED IMMEDIATELY, NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. PATIENT IS DOING WELL AND HAPPY. Manufacturer Narrative: PRODUCT EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND BENT AND RESIDUE ON LENS SURFACE. IMPLANT/EXPLANT DATE: CORRECTED FROM "(B)(6) 2019" TO "(B)(6) 2019" FOR INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966499&pc=MTA | 2023826-2019-01635 | 2019-01-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -5.5 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8441993&pc=MTA | 2023826-2019-00441 | 2019-01-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, DURING THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8571520&pc=MTA | 2023826-2019-00699 | 2019-01-29 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8911772&pc=MTA | 2023826-2019-01519 | 2019-02-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A 13.2MM, VICMO 13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY, THE LENS WAS REMOVED WITHIN THE SAME SURGERY. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9595577&pc=MTA | 2023826-2020-00080 | 2019-02-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS IMPLANTED AND REMOVED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2020, PATIENT "SWITCHED TO OTHER TYPE OF SURGERY (LASER SURGERY)." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813863&pc=MTA | 2023826-2019-01285 | 2019-03-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. REPORTEDLY, NO PATIENT INJURY OCCURRED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: IN PREVIOUS MDR, " VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS" SHOULD BE CORRECTED TO "VISUAL INSPECTION FOUND THE HAPTIC TORN" CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VSIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8588848&pc=MTA | 2023826-2019-00739 | 2019-03-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO INJURY, THE LENS WAS NOT IMPLANTED. ANOTHER LENS WAS NOT IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: IMPLANTED DATE ADDED. EXPLANTED DATE ADDED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8501346&pc=MTA | 2023826-2019-00586 | 2019-03-12 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 13.2MM, MICL13.2, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) BUT WOULD NOT OPEN UP/ UNABLE TO UNFOLD ONCE IMPLANTED. THE LENS WAS REMOVED AND RETRIED AT LEAST FOUR TIMES, USED OTHER LENS. NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN LENS VIAL. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652145&pc=MTA | 2023826-2019-00947 | 2019-03-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8617670&pc=MTA | 2023826-2019-00818 | 2019-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED DURING INITIAL SURGERY AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. THERE WAS PATIENT CONTACT WITH THE DEVICE BUT NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE AND DEBRIS, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; LENS HAPTIC IS TORN. (B)(4). Manufacturer Narrative: WEIGHT:UNK. ETHNICITY: UNK. RACE:UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8620503&pc=MTA | 2023826-2019-00870 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AFTER IMPLANTING A12.1MM, VICM5_12.1, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 THE LENS REMAINED FOLDED IN THE EYE. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9175922&pc=MTA | 2023826-2019-01832 | 2019-04-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VICMO 12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND BACK-UP LENS IMPLANTED WITHIN THE SAME SURGERY, ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: "THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE." SHOULD BE REMOVED FROM THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660597&pc=MTA | 2023826-2019-00948 | 2019-04-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED WITHIN THE SAME SURGERY FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OPTIC TEAR AND HAPTIC TEAR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8629622&pc=MTA | 2023826-2019-00894 | 2019-04-11 | Death | STAAR SURGICAL COMPANY | 2019-05-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT WITH THE DEVICE BUT NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. LENS WAS RETURNED IN TWO PIECES. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660624&pc=MTA | 2023826-2019-00955 | 2019-04-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEBRIS ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125530&pc=MTA | 2023826-2019-01775 | 2019-04-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VICMO 12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO PATIENT'S LEFT EYE (OS). THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED DURING INITIAL SURGERY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824534&pc=MTA | 2023826-2020-00357 | 2019-05-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1, -05.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8899047&pc=MTA | 2023826-2019-01424 | 2019-06-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8967192&pc=MTA | 2023826-2019-01604 | 2019-06-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8967184&pc=MTA | 2023826-2019-01597 | 2019-06-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876496&pc=MTA | 2023826-2019-01400 | 2019-06-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. A PIECE OF THE LENS OPTIC AND ONE HAPTIC WERE TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120726&pc=MTA | 2023826-2019-01786 | 2019-06-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING DELIVERY INTO THE EYE. THE REPORTER STATED THAT THERE WAS NO PATIENT INJURY AND THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: D9- DEVICE RETURNED TO MANUFACTURER IN PREVIOUS MDR SHOULD BE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN / HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125522&pc=MTA | 2023826-2019-01772 | 2019-06-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON 27/JUN/2019 A VICMO 12.6, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED DURING THE INITIAL SURGERY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8805166&pc=MTA | 2023826-2019-01253 | 2019-07-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. A SAME MODEL/LENGTH/DIOPTER LENS WAS IMPLANTED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND A LENS HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9440943&pc=MTA | 2023826-2019-02334 | 2019-07-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT UNK. ETHNICITY: UNK. RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9337449&pc=MTA | 2023826-2019-02179 | 2019-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO_12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - VICMO_12.6 SHOULD BE CORRECTED TO VICMO_13.2. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9668841&pc=MTA | 2023826-2020-00237 | 2019-07-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD), THE WAS IMPLANTED UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND THE LENS TORE. WITHIN THE SAME SURGERY THE LENS WAS REPLACED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9826524&pc=MTA | 2023826-2020-00503 | 2019-07-11 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912811&pc=MTA | 2023826-2019-01514 | 2019-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2019. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT WAS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND AN OPTIC TEAR AND HAPTIC TEAR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9583801&pc=MTA | 2023826-2020-00098 | 2019-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, DURING THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9711480&pc=MTA | 2023826-2020-00253 | 2019-07-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -17.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THE EXCHANGE LENS WAS IMPLANTED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- "EXCHANGE LENS WAS IMPLANTED ON 17/SEPT/2019" SHOULD BE CORRECTED TO "EXCHANGE LENS WAS IMPLANTED ON (B)(6) 2019" IN THE INITIAL MDR. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120649&pc=MTA | 2023826-2019-01785 | 2019-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING DELIVERY INTO THE EYE. THE REPORTER STATED THAT THERE WAS NO PATIENT INJURY AND THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TEAR AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8937553&pc=MTA | 2023826-2019-01553 | 2019-07-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: "VISUAL INSPECTION FOUND OPTIC TORN" IN PREVIOUS MDR IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9538854&pc=MTA | 2023826-2020-00005 | 2019-07-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: CORRECTED DATA: B5 - LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE WHICH RESOLVED THE PROBLEM. ADDITIONAL INFORMATION: B5 - IT WAS REPORTED ON (B)(6) 2020 THAT "THE PATIENT HAD REFRACTION WITH ASTIGMAT BUT USER DECIDED TO USE A SPHERICAL LENS". THE REPORTER INDICATED THEY BELIEVED THE LENS TEAR WAS DUE TO BAD LOADING. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED IN MICRO-CENTRIFUGE VIA WITH MOISTURE ON LENS, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, RESIDUE AND DEBRIS ON THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9588775&pc=MTA | 2023826-2020-00075 | 2019-07-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A 13.2MM, VICMO13.2, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN Manufacturer Narrative: CORRECTED DATA: B4 - (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019. G4 - 17-DEC-2019 SHOULD BE CORRECTED TO 18-DEC-2019 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131788&pc=MTA | 2023826-2019-01758 | 2019-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTICS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9188605&pc=MTA | 2023826-2019-01908 | 2019-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120438&pc=MTA | 2023826-2019-01741 | 2019-08-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER TORE/BROKE DURING INJECTION. THE LENS WAS REMOVED INTRAOPERATIVELY ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2019. REPORTEDLY, THE ICL WAS STUCK AT THE TOP OF THE INJECTOR BETWEEN THE CARTRIDGE AND THE FOAM TIP PLUNGER. IT IS REPORTED THAT THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR AND DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND TEARS IN THE HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373331&pc=MTA | 2023826-2019-02231 | 2019-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: G4 - DATE RECEIVED BY MANUFACTURER: 03-NOV-2019 SHOULD BE CORRECTED TO 04-NOV-2019 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10496578&pc=MTA | 2023826-2020-01975 | 2019-08-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS REMOVED AND EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K-THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824529&pc=MTA | 2023826-2020-00355 | 2019-09-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, INSIDE A CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B3 - DATE OF EVENT: (B)(6) 2019 SHOULD BE CORRECTED TO "UNKNOWN". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9194689&pc=MTA | 2023826-2019-01926 | 2019-09-06 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, WAS BROKEN WHEN REMOVED FROM THE VIAL. IT APPEARED THE CORNER OF THE HAPTIC WAS CAUGHT BETWEEN THE LID AND VIAL WHEN VIAL WAS OPENED. PART OF THE HAPTIC WAS TORN OFF WHEN REMOVED FROM VIAL. THE SURGEON IMPLANTED THE LENS IN THE PATIENTS LEFT EYE (OS), WITH NO PATIENT INJURY AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THE FAILURE IS LIKELY DUE TO OPERATOR ERROR DURING THE FINAL INSPECTION AND FINAL PACKAGING STEPS OF THE MANUFACTURING PROCESS FOR THIS LENS. THERE IS NO INDICATION THAT SUGGESTS A DEFICIENCY IN THE MANUFACTURING PROCESS OR A TREND OF PROCESS ERRORS. OPERATORS WHO PERFORMED THESE STEPS OF THE ASSOCIATED LOT HAVE BEEN RETRAINED AND MADE AWARE OF THIS OCCURRENCE. FURTHER INVESTIGATION OR ACTION IS NOT REQUIRED AT THIS TIME. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781719&pc=MTA | 2023826-2020-00403 | 2019-09-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819169&pc=MTA | 2023826-2020-00455 | 2019-09-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND STRETCHED OUT WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9602337&pc=MTA | 2023826-2019-02094 | 2019-09-27 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS SURFACE. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306720&pc=MTA | 2023826-2018-02094 | 2019-09-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: INITIAL MDR IS NOT APPLICABLE. REPORT NUMBER WAS SUBMITTED WITH WRONG MFR REPORT YEAR (2018). REFERENCE MFR #2023826-2019-02094 FOR RE-SUBMISSION OF INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957608&pc=MTA | 2023826-2020-00781 | 2019-10-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OPTIC TORN/HAPTIC TORN AND RESIDUE ON LENS. CORRECED DATA B5- STATEMENT "THE LENS WAS IMPLANTED,REMOVED,REPLACED INTRAOPERATIVELY" IN INITIAL MDR IS NOT APPLICABLE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2019 DUE TO LENS TORE DURING INJECTION. ON (B)(6) 2019 A NEW SAME MODEL LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9671826&pc=MTA | 2023826-2020-00245 | 2019-12-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS UPSIDE DOWN, LENS WAS BROKEN WHEN THE SURGEON PULLED IT OUT TO REMOVE IT DURING THE SAME SURGERY. THERE WAS NO PATIENT INJURY AND A REPLACEMENT LENS WAS IMPLANTED WITHIN THE SAME SURGERY. PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988389&pc=MTA | 2023826-2020-00944 | 2019-12-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT 13.2MM, VICMO13.2, -15.00 DIOPTER IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957611&pc=MTA | 2023826-2020-00783 | 2019-12-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842805&pc=MTA | 2023826-2020-00554 | 2019-12-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A 13.7MM, VICMO13.7, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE (OD). THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9696125&pc=MTA | 2023826-2020-00292 | 2019-12-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICH13.2, +9.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING DELIVERY INTO THE EYE. IT WAS REPORTED THAT THE LENS WAS REMOVED WITHIN THE SAME SURGERY, THERE WAS NO PATIENT INJURY. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. THE SURGEON PLANS TO IMPLANT A SMALLER SIZE LENGTH. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES WITH THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: D6B- "(B)(6) 2019" SHOULD BE IN THE INITIAL MDR. H1- "SERIOUS INJURY" SHOULD BE CORRECTED TO "MALFUNCTION" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10334907&pc=MTA | 2023826-2020-01587 | 2019-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME INITIAL SURGERY AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY AND THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSOECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916637&pc=MTA | 2023826-2020-00739 | 2019-12-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA : D9- "23-JUN-2020" SHOULD BE CORRECTED TO "01-MAY-2020" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364454&pc=MTA | 2023826-2020-01642 | 2019-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER TORE DURING INJECTION INTO THE PATIENT'S RIGHT (OD) EYE. THIS OCCURRED ON (B)(6) 2019. IT WAS REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY REPORTED. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE OPTIC AND HAPTIC WERE TORN. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9747641&pc=MTA | 2023826-2020-00322 | 2020-01-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO LENS TORE/BROKE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824594&pc=MTA | 2023826-2020-00493 | 2020-01-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842755&pc=MTA | 2023826-2020-00508 | 2020-02-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER TORE DURING INJECTION INTO THE PATIENT'S LEFT (OS) EYE. THIS OCCURRED ON (B)(6) 2020. IT WAS REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY REPORTED. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2020 AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9793363&pc=MTA | 2023826-2020-00429 | 2020-02-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE UPON INSERTION. THE LENS WAS REMOVED AND REPLACED WITH THE BACKUP LENS. DOCTOR FELT IT WAS THE INJECTOR THAT TORE THE LENS, UNSURE OF MODEL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: A2-DATE OF BIRTH IS CORRECTED TO (B)(6) 1974. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC AND OPTIC. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9857166&pc=MTA | 2023826-2020-00578 | 2020-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), AND THE LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889396&pc=MTA | 2023826-2020-00595 | 2020-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED, REMOVED, AND REPLACED, A 13.2MM VICM5 13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND FOREIGN DEBRIS ON THE LENS AND THE LENS HAS A CURVED SHAPE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9855308&pc=MTA | 2023826-2020-00498 | 2020-02-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED DURING INITIAL SURGERY AND AN EXCHANGE LENS WAS IMPLANTED ON (B)(6) 2020. "PATIENT IS DOING VERY WELL NOW." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS WAS STUCK TO "THE SPONGE UNREASONABLY, LENS TORE WHEN TRYING TO RELEASE." Manufacturer Narrative: CORRECTED DATA: B5 - THE REPLACEMENT LENS DATE OF "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020". CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9855309&pc=MTA | 2023826-2020-00497 | 2020-02-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED DURING INITIAL SURGERY AND AN EXCHANGE LENS WAS IMPLANTED ON (B)(6) 2020. "PATIENT IS DOING VERY WELL NOW." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS WAS STUCK TO "THE SPONGE UNREASONABLY, LENS TORE WHEN TRYING TO RELEASE." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938268&pc=MTA | 2023826-2020-00823 | 2020-02-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, - 14.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND PATIENT IS DOING WELL. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938274&pc=MTA | 2023826-2020-00799 | 2020-02-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. THE PATIENT IS DOING WELL. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941229&pc=MTA | 2023826-2020-00801 | 2020-02-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE LENS GOT STUCK TO FOAM TIP PLUNGER. THE PATIENT'S POST-OP CONDITION IS STABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BENT AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889397&pc=MTA | 2023826-2020-00607 | 2020-02-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED, A 12.6MM VICMO 12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913408&pc=MTA | 2023826-2020-00774 | 2020-03-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE DOCTOR FELT THAT A 13.2MM, MICL 13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS NOT LOADED CORRECTLY. LENS DID NOT FEEL RIGHT INJECTING INTO THE EYE. HE THOUGHT THAT MAYBE THERE WAS DAMAGE TO THE LENS AND DID NOT WANT TO USE IT AND IMPLANTED THE BACKUP LENS. THE LENS MADE CONTACT WITH THE PATIENTS EYE BUT THERE WAS NO PATIENT INJURY. THE DOCTOR FEELS IT WAS A LOADER ERROR. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197358&pc=MTA | 2023826-2020-01331 | 2020-03-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. INTRAOPERATIVELY, THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335018&pc=MTA | 2023826-2020-01589 | 2020-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME INITIAL SURGERY AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY AND THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075078&pc=MTA | 2023826-2020-01093 | 2020-04-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), AND THE LENS WAS TORN DURING DELIVERY. THERE WAS PATIENT CONTACT BUT NO INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6-LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: D6 - (B)(6)2020 SHOULD BE REMOVED, AS IT IS NOT APPLICABLE. D7 - (B)(6) 2020 SHOULD BE REMOVED, AS IT IS NOT APPLICABLE. CLAIM#:(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093677&pc=MTA | 2023826-2020-01073 | 2020-04-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO121, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10147954&pc=MTA | 2023826-2020-01246 | 2020-04-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A 13.2MM, VICM5_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE, NO PATIENT INJURY. A REPLACEMENT LENS WAS LATER IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4- 27-APR-2020 SHOULD BE CORRECTED TO 13-MAY-2020 IN THE INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093685&pc=MTA | 2023826-2020-01072 | 2020-05-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN, DEFORMED AND STRETCHED OUT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158178&pc=MTA | 2023826-2020-01238 | 2020-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS, WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476279&pc=MTA | 2023826-2020-01933 | 2020-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222175&pc=MTA | 2023826-2020-01404 | 2020-05-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE LOADING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE INJECTOR, THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. REPORTEDLY, NO PATIENT CONTACT OCCURRED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTIONS FOUND HAPTIC TORN. CORRECTED DATA: B5- THE DATE IN WHICH THE LENS TORE IS CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10456865&pc=MTA | 2023826-2020-01911 | 2020-05-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A 13.2MM, VICM5_13.2, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT (OD) EYE. THE REPORTER STATED " LENS GOT STUCK IN THE PLUNGER WHILE INJECTING THE LENS AND GOT TORN ON IMPLANTATION.". THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. IT WAS REPORTED THAT ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10213606&pc=MTA | 2023826-2020-01386 | 2020-06-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE AND THE PROBLEM RESOLVED. NO PATIENT INJURY, CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310454&pc=MTA | 2023826-2020-01510 | 2020-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN, AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401661&pc=MTA | 2023826-2020-01695 | 2020-06-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. A LENS WAS IMPLANTED ON (B)(6) 2020, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310579&pc=MTA | 2023826-2020-01462 | 2020-06-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401673&pc=MTA | 2023826-2020-01702 | 2020-06-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VICM5 12.6, -05.00 DIOPTER, IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION, 60% OF LENS REMAINED IN CARTRIDGE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2020, A SAME SIZE LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. REPORTEDLY, "NURSE DID NOT FOLD AND CHARGE AS RECOMMENDED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THE LENS RECEIVED FOR DEVICE EVALUATION WAS NOT THE CORRECT LENS. THE PREVIOUS SUPPLEMENTAL MDR IS NOT APPLICABLE TO THIS CLAIM. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKE.TED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10381109&pc=MTA | 2023826-2020-01660 | 2020-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -13.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE INTO TWO PIECES DURING IMPLANTATION. HALF OF THE LENS HAD PATIENT CONTACT AND THE OTHER HALF GOT STUCK IN THE CARTRIDGE, THERE WAS NO PATIENT INJURY. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED WITH MOISTURE, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR ON THE OPTIC AND HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401653&pc=MTA | 2023826-2020-01691 | 2020-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. A SAME SIZE LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401656&pc=MTA | 2023826-2020-01693 | 2020-07-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VICMO13.2, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2020, A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10484168&pc=MTA | 2023826-2020-02004 | 2020-07-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020 AND REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. USER ERROR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10485131&pc=MTA | 2023826-2020-02013 | 2020-08-05 04:00 | Malfunction | STAAR SUGICAL COMPANY | 2020-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITH BACK UP LENS WITHIN THE SAME SURGERY DUE TO LENS STUCK IN CARTRIDGE. THERE WAS NO PATIENT INJURY. CAUSE UNKNOWN, THE REPORTER STATED " THE LENS STUCK IN THE CARTRIDGE BECAUSE THE SURGEON TRIED TO PUSH THE LENS WITH THE PLUNGER FROM THE BACK OF THE CARTRIDGE INTO IT. HE WAS NOT USING THE ICL LOADING FORCEPS AS WE'VE RECOMMENDED." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS, WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8731294&pc=MTA | 2023826-2019-01102 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-06-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON 03/DEC/2018 A VICMO 12.1, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND BACK UP LENS IMPLANTED DURING INITIAL SURGERY, ON (B)(6) 2018. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ADDED REPORT TYPE PRODUCT PROBLEM. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586879&pc=MTA | 2023826-2020-00017 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL 12.6, -09.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559370&pc=MTA | 2023826-2020-00035 | 2019-11-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -18.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: H1 - SERIOUS INJURY SHOULD BE CORRECTED TO MALFUNCTION FROM INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8754359&pc=MTA | 2023826-2019-01141 | 2018-10-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS LATER REPLACED AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A TEAR ON THE LENS HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 1494: OFF LABEL-USE (PATIENT'S ACD <3.00MM) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361999&pc=MTA | 2023826-2020-01630 | 2018-12-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H1- "SERIOUS INJURY" SHOULD BE CORRECTED TO "MALFUNCTION" IN INITIAL AND SUPPLEMENTAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8967188&pc=MTA | 2023826-2019-01601 | 2019-02-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: FTP LOT # 1411930. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664567&pc=MTA | 2023826-2019-00990 | 2019-04-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME INITIAL SURGERY AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY AND THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 OFF-LABEL USE SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT.(B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869325&pc=MTA | 2023826-2020-00619 | 2019-05-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-23 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -17.5 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9158792&pc=MTA | 2023826-2019-01841 | 2019-05-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, FROM THE PATIENTS LEFT EYE (OS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTIOUN FOUND HAPTIC TORN. DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938227&pc=MTA | 2023826-2020-00806 | 2019-05-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: A3 - MALE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306696&pc=MTA | 2023826-2019-02000 | 2019-06-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VICMO12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. LENS WAS REMOVED INTRAOPERATIVELY. ON (B)(6) 2019, A LONGER LENGTH REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306734&pc=MTA | 2023826-2019-02003 | 2019-06-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.1MM VICMO12.1, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED INTRAOPERATIVELY. ON (B)(6) 2019, A SAME LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336287&pc=MTA | 2023826-2019-02057 | 2019-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- DEVICE CODE 1494: OFF-LABEL USE; AGE < 21 YEARS OLD SHOULD BE ADDED TO THE INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966500&pc=MTA | 2023826-2019-01637 | 2019-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306698&pc=MTA | 2023826-2019-02001 | 2019-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186496&pc=MTA | 2023826-2019-01862 | 2019-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VICMO 12.6, -05.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED DURING THE INITIAL SURGERY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019. PROBLEM HAS BEEN RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC HAD A TEAR. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889399&pc=MTA | 2023826-2020-00593 | 2019-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VICMO 12.6, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2020, "PATIENT DOES NOT INTEND TO REIMPLANT THE LENS. PATIENT'S CURRENT CONDITION IS GOOD." Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824583&pc=MTA | 2023826-2020-00489 | 2019-12-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9977659&pc=MTA | 2023826-2020-00916 | 2019-12-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "LENS WAS IMPLANTED, IT FLIPPED," AND DURING REMOVAL IT WAS "ACCIDENTALLY DAMAGED. PATIENT ALSO SATISFIED WITH THE POSTOPERATIVE VISION." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9597949&pc=MTA | 2023826-2020-00128 | 2019-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.1MM VICMO12.1, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2019, A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140599&pc=MTA | 2023826-2020-01205 | 2020-03-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS REMOVED AND EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094009&pc=MTA | 2023826-2020-01135 | 2020-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. INTRAOPERATIVELY, THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS USER ERROR. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335272&pc=MTA | 2023826-2020-01561 | 2020-06-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445865&pc=MTA | 2023826-2020-01817 | 2020-07-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491983&pc=MTA | 2023826-2020-01968 | 2020-08-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2020. INTRAOPERATIVELY, THE LENS WAS REPLACED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# 717106.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10484152&pc=MTA | 2023826-2020-02006 | 2020-07-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AND REMOVED WITHIN THE SAME SURGERY DUE TO LENS IMPLANTED UPSIDE DOWN. NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), SHOULD HAVE BEEN INCLUDED IN PREVIOUS MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910964&pc=MTA | 2023826-2020-00689 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; Unintended Movement; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD), THE WAS IMPLANTED UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND THE LENS TORE. WITHIN THE SAME SURGERY THE LENS WAS REPLACED WITH AN ALTERNATE, SHORTER LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISIBLE INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: H6 - DEVICE CODE 3189 CORRECTED TO DEVICE CODE 3026. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7467136&pc=MTA | 2023826-2018-00602 | 2018-04-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Activation Failure | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, AND WAS ATTEMPTING TO TUCK THE HAPTICS UNDER THE IRIS. THE SURGEON WAS ABLE TO TUCK ONE HAPTIC BUT THE PUPIL CONSTRICTED AND SURGEON WAS UNABLE TO TUCK THE OTHER HAPTIC. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY. THIS WAS THE 2ND OF 2 LENSES USED ON THIS PATIENT. THE FIRST LENS, SAME MODEL, WAS USED ON (B)(6) 2018 BUT THE SURGEON WAS UNABLE TO TUCK THE HAPTIC DUE TO PUPIL CONSTRICTED - SEE MFR REPORT # 2023826-2018-00603. THE SURGEON DID NOT IMPLANT A 3RD LENS. THE REPORTER INDICATED THE CAUSE OF THIS EVENT WAS PATIENT RELATED. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6464737&pc=MTA | 2023826-2017-00555 | 2017-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, WAS IMPLANTED AND REMOVED DUE TO THE LENS FLIPPING. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WIL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC, SCRATCH ON THE OPTIC, AND A MISSING PIECE OF THE HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTER FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197900&pc=MTA | 2023826-2018-00070 | 2017-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -10.5/+5.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED DURING THE SAME SURGERY DUE TO THE LENS BEING UPSIDE DOWN. THE LENS TORN DURING REMOVAL. THERE WAS NO PATIENT INJURY. ANOTHER LENS WAS IMPLANTED ON (B)(6) 2017 AND IT WAS SUCCESSFUL. IT IS REPORTED THAT POOR LOADING OF THE LENS IN THE CARTRIDGE MIGHT BE THE CONTRIBUTING FACTOR THAT LED TO THE LENS BEING IMPLANTED UPSIDE DOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7216213&pc=MTA | 2023826-2018-00141 | 2017-12-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, FLIPED UPON INSERTION. THE LENS WAS REMOVED AND THE BACK UP LENS WAS IMPLANTED WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. THE REPORTER STATED THAT THE CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A TEAR IN THE OPTIC. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7252889&pc=MTA | 2023826-2018-00164 | 2018-01-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A MICL13.2, -15.5 DIOPTER, IMPLANTABLE COLLAMER LENS, AND THE LENS UNFOLDED UPSIDE DOWN. THE LENS WAS REMOVED DURING INITIAL SURGERY AND THE BACK UP LENS WAS USED. THERE WAS NO PATIENT INJURY. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7253425&pc=MTA | 2023826-2018-00167 | 2018-01-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL13.2, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REMOVED AND EXCHANGED WITH THE BACK UP LENS DURING THE SAME SURGERY. THERE WAS NO PATIENT INJURY AND IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: IMPLANTED UPSIDE DOWN IN THE PATIENT'S LEFT EYE (OS). IT WAS REPORTED THAT THE CAUSE OF EVENT WAS SURGEON ERROR. (B)(4) Manufacturer Narrative: LENS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7364242&pc=MTA | 2023826-2018-00450 | 2018-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+5.0/092 (SPHERE/CYLINDER/AXIS), AND THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS DAMAGED WHEN REMOVED, WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481807&pc=MTA | 2023826-2018-00625 | 2018-03-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -8.5/3.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS IMPLANTED UPSIDE-DOWN AND WAS EXPLANTED IN THE SAME SURGERY. THE LENS WAS DAMAGED DURING EXPLANTATION. THERE WAS NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED A WEEK LATER AND IT IS ALL FINE. IT IS REPORTED THAT THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TORN WITH PRESENCE OF RESIDUE ON THE LENS SURFACE. THE LENS WAS MISSING THE TORN PIECE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6685164&pc=MTA | 2023826-2017-01107 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS,-6.5 DIOPTER, AND THE LENS WENT INTO THE PATIENTS EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7539113&pc=MTA | 2023826-2018-00761 | 2018-02-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. DURING IMPLANTATION, THE FOAM TIP CAUGHT THE OPTIC AND HAPTIC AND; CAUSING A TEAR TO THE LENS. THE LENS WAS INTRAOPERATIVELY EXCHANGED WITH ANOTHER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS OKAY. Manufacturer Narrative: CORRECTED DATA: CORRECT PATIENT IDENTIFIER IS (B)(6). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, RESIDUE ON LENS, AND TORN PIECE MISSING. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6064990&pc=MTA | 2023826-2016-01560 | 2016-09-01 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure; Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS INSERTING A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, AND NOTED THE LENS WAS DAMAGED. THE LENS WAS PARTIALLY INSERTED AND WAS REMOVED WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE LENS MAY NOT HAVE BEEN PULLED DOWN THE CARTRIDGE BARREL ENOUGH AND THE FOAM TIP PLUNGER CAUGHT AND DAMAGED THE LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH PIECES TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5710174&pc=MTA | 2023826-2016-00804 | 2016-05-12 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Scratched Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON LOADED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, ON (B)(6) 2016 AND NOTICED THE LENS WAS SCRATCHED AS HE WAS ATTEMPTING TO INSERT THE LENS IN THE PATIENT'S EYE. THE CAUSE OF THE EVENT WAS UNKNOWN. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION SHOWED A PIECE OF ONE HAPTIC WAS TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951506&pc=MTA | 2023826-2016-01280 | 2016-08-11 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Scratched Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, AND NOTED THERE WAS A SCRATCH ON THE LENS. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6031391&pc=MTA | 2023826-2016-01433 | 2016-09-19 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Scratched Material | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND THE LENS WAS NOTED TO HAVE A SCRATCH ON THE OPTICS OF THE LENS. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: PREVIOUSLY WRONG MODEL NUMBER. CORRECT MODEL NUMBER IS VICM5_13.2. REPORT SOURCE- USER FACILITY WAS PREVIOUSLY WRONG INFORMATION. CORRECT INFORMATION IS DISTRIBUTOR. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, THE HAPTIC BENT, AND FIBERS ON LENS SURFACE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632886&pc=MTA | 2023826-2016-00564 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Sticking | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, AND THE HAPTIC WAS STUCK TOGETHER. THE LENS WAS REMOVED WITH NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704639&pc=MTA | 2023826-2015-00435 | 2015-03-17 | Malfunction | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Sticking; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT THE LENS WAS NOT INJECTING PROPERLY AND WAS STICKING IN THE CARTRIDGE. THE LENS WAS EJECTED ONTO THE STERILE FIELD AND THE LENS WAS FOUND TO BE TORN. THERE WAS NO PATIENT CONTACT OR INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THE SURGEON DID NOT FEEL THERE WAS AN ISSUE WITH THE LENS OR INJECTION SYSTEM. Manufacturer Narrative: (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5079603&pc=MTA | 2023826-2015-01215 | 2015-08-13 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Sticking; Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON ATTEMPTED TO LOAD A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE SFC-45 FP CARTRIDGE, BUT THE LENS APPEARED TO BE STUCK AND TORE. THE SURGEON DIDN'T KNOW IF THE EVENT WAS DUE TO THE LENS OR CARTRIDGE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. SEE MFR. 2023826-2015-01220 FOR THE CARTRIDGE. Manufacturer Narrative: EVALUATION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, CARTRIDGE LOT SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: LENS AND INJECTION SYSTEM NOT RETURNED. METHOD - WORK ORDER SEARCH, CARTRIDGE LOT NUMBER SEARCH RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. A CARTRIDGE LOT NUMBER SEARCH WAS PERFORMED AND ONE OTHER COMPLAINT WAS FOUND. CONCLUSION - BASED ON THE COMPLAINT HISTORY, LENS WORK ORDER SEARCH AND CARTRIDGE LOT NUMBER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS AND INJECTION SYSTEM NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4618628&pc=MTA | 2023826-2015-00335 | 2015-02-10 | Malfunction | STAAR SURGICAL COMPANY | 2015-03-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS BUT THE LENS TORE DURING DELIVERY. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT AGE AND WEIGHT: UNK. IMPLANT AND EXPLANT DATES: NA. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVALUATION: RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119417&pc=MTA | 2023826-2015-01296 | 2015-02-13 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON NOTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+6.0/X089 DIOPTER, WAS TORN BEFORE LOADING. THE LENS WAS NOT IMPLANTED AND THERE WAS NO REPORTED INJURY. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. IMPLANT AND EXPLANT DATES: NA. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. : EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5004868&pc=MTA | 2023826-2015-01082 | 2015-06-23 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE FACILITY RETURNED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS ( -7.0 DIOPTER) TO THE MANUFACTURER TORN. THE FACILITY REPORTED THE LENS WAS NOT IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, MICL WAS NOT IMPLANTED DUE TO UNSPECIFIED REASON. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ON THE PATIENT ID CARD WAS WRITTEN:" OPENED NOT IMPLANTED" THEREFORE, IT IS UNKNOWN IF THE LENS HAD A PATIENT CONTACT OR NOT. DESPITE MULTIPLE ATTEMPTS NO ADDITIONAL INFORMATION HAS BEEN RECEIVED TO DATE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: IMPLANT DATE: UNK. EXPLANT DATE: UNK. METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5155148&pc=MTA | 2023826-2015-01352 | 2015-06-24 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/X098 DIOPTER, AND NOTED THE LENS HAD TORN. THE LENS WAS REMOVED WITH NO REPORTED INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5102460&pc=MTA | 2023826-2015-01258 | 2015-08-26 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER AND THE LENS WAS NOTED TO HAVE TORN. THE LENS TOUCHED THE EYE AND WAS REMOVED WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE: (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS CODE: (EVALUATION RESULT): PER MEDICAL REVIEW - LENS TEARS CAN OCCUR AT ANY STAGE FROM MANUFACTURE TO INTRAOPERATIVE MANIPULATION. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS THAT SUGGESTS A CONTRIBUTORY FACTOR IN THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - 81 (OTHER): LENS NOT RETURNED. METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218802&pc=MTA | 2023826-2015-01457 | 2015-08-31 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS GOT STUCK DURING DELIVERY INTO THE EYE AND TORE. THE LENS WAS REMOVED AN EXCHANGED FOR ANOTHER ICL, SAME MODEL AND DIOPTER SIZE. NO ADVERSE EVENT WAS REPORTED. LAST POST-OP VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -08.00. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5190310&pc=MTA | 2023826-2015-01405 | 2015-09-15 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5240287&pc=MTA | 2023826-2015-01573 | 2015-10-21 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER AND THE LENS TORE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, INTRAOPERATIVE ICL EXCHANGE WAS PERFORMED TO ADDRESS DAMAGE OF THE LENS ("TORE DURING INSERTION"). INCISION WAS NOT ENLARGED AND NO OTHER TISSUE DAMAGE WAS REPORTED. ANOTHER ICL, SAME SIZE AND DIOPTER, WAS SUCCESSFULLY IMPLANTED. ACCORDING TO THE REPORT, BY A TECHNICIAN, DATED ON (B)(6) 2015 THE CAUSE OF THE EVENT WAS UNKNOWN. NO REPORTED POSTOPERATIVE SEQUELAE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS THAT SUGGESTS A CONTRIBUTORY FACTOR IN THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378473&pc=MTA | 2023826-2015-01703 | 2015-12-02 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER AND THE LENS TORE. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULTS: (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS - (EVALUATION RESULT): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE IS LIKELY DUE TO THE END USER HANDLING OF THE LENS. THE MISSING PORTION OF THE LENS IS EASY TO DETECT UNDER IN-PROCESS FINAL INSPECTION AND NO DIRECT PHYSICAL CONTACT IS MADE AFTER THE VISUAL INSPECTION. PER MEDICAL REVIEW: REPORTEDLY, INTRAOPERATIVE MICL REMOVAL/EXCHANGE WAS PERFORMED TO ADDRESS LENS DAMAGE ("TEAR") THAT WAS OBSERVED DURING INSERTION. NO REPORT OF INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ACCORDING TO REPORT BY A TECHNICIAN, DATED ON 12/18/15, THERE WAS NO PATIENT INJURY DURING REMOVAL OF DAMAGED ICL AND A BACKUP LENS WAS SUCCESSFULLY IMPLANTED. THE SAME REPORT STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. NO REPORTED POSTOPERATIVE SEQUELAE. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5365449&pc=MTA | 2023826-2016-00036 | 2015-12-07 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS , -12.5 DIOPTER, TORE DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT. THE REPORTER INDICATED THE EVENT WAS NOT DUE TO USER ERROR. THIS WAS ONE OF TWO LENSES USED FOR THIS PATIENT - SEE MFR. REPORT # 2023826-2016-00037. ANOTHER SAME MODEL LENS, BUT DIFFERENT DIOPTER POWER WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE TO THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, INJECTOR LOT SEARCH, CARTRIDGE LOT SEARCH, FOAM TIP PLUNGER LOT SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH, INJECTOR LOT SEARCH, CARTRIDGE LOT SEARCH, FOAM TIP PLUNGER LOT SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. AN INJECTOR LOT SEARCH FOUND ANOTHER SIMILAR COMPLAINT. A CARTRIDGE LOT SEARCH FOUND ANOTHER SIMILAR COMPLAINT. A FOAM TIP PLUNGER LOT SEARCH FOUND ANOTHER SIMILAR COMPLAINT. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, LENS WORK ORDER SEARCH, INJECTOR, CARTRIDGE, FOAM TIP PLUNGER LOT SEARCHES AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323187&pc=MTA | 2023826-2015-01726 | 2015-12-10 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, WAS TORN AND IT WAS UNKNOWN IF THE EVENT HAD OCCURRED DURING THE LOADING PROCESS OR UPON INSERTION. THE REPORTER STATED THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5370916&pc=MTA | 2023826-2016-00058 | 2015-12-17 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, AND THE LENS TORE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE TO THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DARK/RED RESIDUE ON THE LENS SURFACE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5392103&pc=MTA | 2023826-2016-00104 | 2015-12-22 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, AND THE LENS HAPTIC TORE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY, NO INCISION ENLARGEMENT OR SUTURES. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, HAPTIC WAS TORN OFF DURING INSERTION REQUIRING INTRAOPERATIVE LENS REMOVAL/EXCHANGE. ACCORDING TO THE REPORT, BY A SURGICAL TECHNICIAN, DATED ON (B)(6) 2016 THE INCISION WAS NOT ENLARGED FOR REMOVAL OF A DAMAGED LENS AND ANOTHER BACKUP LENS WAS SUCCESSFULLY IMPLANTED. THE SAME REPORT INDICATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN SINCE THE SURGEON HAS NOT BEEN HAVING ISSUES NEITHER WITH LENS NOR INJECTOR SYSTEM IN THE PAST. NO REPORTED POSTOPERATIVE SEQUELAE CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE THE HAPTIC TEAR. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED THE HAPTIC TEAR. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH PIECES TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND VISUAL INSPECTION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5392102&pc=MTA | 2023826-2016-00103 | 2015-12-28 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, AND THE LENS TORE. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE LENS MAY HAVE BEEN TORN DURING THE LOADING PROCESS AND WAS POSSIBLY DUE TO A LOADING ERROR. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, BY A STAAR REPRESENTATIVE, DATED ON (B)(6) 2016, TEAR OF MICL WAS OBSERVED UPON INSERTION REQUIRING INTRAOPERATIVE LENS EXCHANGE. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ANOTHER LENS WAS IMPLANTED SUCCESSFULLY. THE SAME REPORT STATED THAT THERE WAS A POSSIBILITY THAT LENS WAS TORN OFF DUE TO USER ERROR DURING LOADING PROCESS. NO REPORTED POSTOPERATIVE SEQUELAE. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5392121&pc=MTA | 2023826-2016-00106 | 2015-12-30 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, TORE WHILE BEING TAKEN OUT OF THE VIAL. THE LENS WAS TORN BEFORE LOADING AND WAS NOT DUE TO USER ERROR. THERE WAS NO PATIENT CONTACT. THE REPORTER INDICATED THE EVENT WAS PERHAPS A MANUFACTURING DEFECT. ANOTHER SAME MODEL LENS WAS IMPLANTED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC AND ONE HAPTIC TORN. A PIECE OF ONE HAPTIC WAS TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE OR NOT DETECT THE DAMAGE TO THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5439665&pc=MTA | 2023826-2016-00164 | 2016-01-14 | Malfunction | STAAR SURGICAL COMPANY | 2016-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, AND THE LENS TORE. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE SURGEON INDICATED HE HAD MADE THE INCISION TOO SMALL AND THIS CAUSED THE LENS TO TEAR. Manufacturer Narrative: RESULTS: PER MEDICAL REVIEW - REPORTEDLY, INTRAOPERATIVE LENS EXCHANGE WAS PERFORMED TO ADDRESS LENS DAMAGE NOTED DURING INSERTION. INCISION WAS NOT ENLARGED AND NO OTHER TISSUE DAMAGE WAS REPORTED. ANOTHER LENS WAS SUCCESSFULLY IMPLANTED. ACCORDING TO COMPLAINT DESCRIPTION THE CAUSE OF THE EVENT WAS PROCEDURE FACTOR/ERROR ("TOO SMALL INCISION" CREATED BY A SURGEON PRIOR TO LENS IMPLANTATION). (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN. ONE HAPTIC WAS TORN, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5492647&pc=MTA | 2023826-2016-00249 | 2016-01-20 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE CUSTOMER RETURNED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, TO THE MANUFACTURER TORN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER STATED THE LENS TORE ON DELIVERY AND THERE WAS MOST LIKELY PATIENT CONTACT BUT NO INJURY. ANOTHER LENS WAS IMPLANTED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN AND PIECES OF HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE ROOT CAUSE IS LIKELY OFF -LABEL USE OF AN INJECTOR SYSTEM THAT IS NOT THE RECOMMENDED INJECTOR SYSTEM PER THE DFU. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CONTRIBUTE TO DAMAGED LENSES OR RELEASED NONCONFORMING LENSES. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5464098&pc=MTA | 2023826-2016-00207 | 2016-01-27 | Malfunction | STAAR SURGICAL COMPANY | 2016-02-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, TORE AND THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY REVIEW RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5512158&pc=MTA | 2023826-2016-00265 | 2016-02-16 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS ATTEMPTING TO INSERT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, AND THE LENS TORE WHILE TRYING TO ADVANCE IN THE INJECTOR. THERE WAS PATIENT CONTACT BUT NO INJURY. THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: DEVICE HISTORY REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, PRODUCT EVALUATION AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593264&pc=MTA | 2023826-2016-00506 | 2016-02-24 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, AND THE LENS TORE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY, NO INCISION ENLARGEMENT OR SUTURES. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE OR NOT DETECT A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE CAUSED LENS TEAR DURING INSERTION. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5611386&pc=MTA | 2023826-2016-00550 | 2016-03-30 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, TORE AND IT WAS UNKNOWN IF THERE WAS ANY PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5681606&pc=MTA | 2023826-2016-00761 | 2016-04-27 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS LOADING A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, AND NOTED THE LENS WAS TORN. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED IT WAS UNKNOWN IF THE LENS WAS RECEIVED TORN OR IF THE LENS WAS DAMAGED DURING LOADING. THE REPORTER STATED THE EVENT WAS NOT DUE TO A LOADING ERROR. Manufacturer Narrative: CONCLUSION: (UNABLE TO CONFIRM COMPLAINT). (B)(4). Manufacturer Narrative: PT AGE AND WEIGHT: UNK. IMPLANT AND EXPLANT DATES: NA. RESULTS CODE(S): (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS STUCK IN THE CARTRIDGE. THERE WAS NO VISIBLE DAMAGE TO THE CARTRIDGE. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724435&pc=MTA | 2023826-2016-00813 | 2016-05-10 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A PIECE OF HAPTIC FROM A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, WAS TORN OFF AS SHE WAS ADVANCING THE LENS FORWARD FROM THE CARTRIDGE WITH A FORCEP. THERE WAS NO PATIENT CONTACT. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF HAPTIC WAS TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5788410&pc=MTA | 2023826-2016-00897 | 2016-06-13 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S EYE AND THE HAPTIC BROKE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO THE SURGEON IMPROPERLY LOADED THE LENS. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5758686&pc=MTA | 2023826-2016-00876 | 2016-06-15 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED SHE WAS ADVANCING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, FORWARD FROM THE CARTRIDGE WITH A FORCEP ON (B)(6) 2016 AND THE LENS TORE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. CLAIM #(B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834405&pc=MTA | 2023826-2016-00999 | 2016-07-01 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON OPENED A VIAL AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, AND FOUND THE LENS TORN. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH A TEAR. VISUAL INSPECTION FOUND A SCRATCH AND HAPTIC TORN/BROKEN/BENT/DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834572&pc=MTA | 2023826-2016-01000 | 2016-07-01 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON OPENED A VIAL AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, AND FOUND THE LENS TORN. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. Event Description: THE -13.00 DIOPTER IS INCORRECT, CORRECT DIOPTER IS -11.0. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND LENS TEAR. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909201&pc=MTA | 2023826-2016-01158 | 2016-07-11 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, TORE WHILE BEING LOADED INTO THE INJECTOR. THERE WAS NO PATIENT CONTACT. THE REPORTER WAS UNABLE TO PROVIDE ANY ADDITIONAL INFORMATION. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870490&pc=MTA | 2023826-2016-01045 | 2016-07-13 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, AND THE LENS HAPTIC WAS NOTED TO BE DAMAGED. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5880926&pc=MTA | 2023826-2016-01100 | 2016-07-19 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON LOADED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, BUT THE LENS WOULD NOT LOAD CORRECTLY AND WAS TORN. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL AND DIOPTER LENS WAS IMPLANTED. THE REPORTER WAS UNABLE TO PROVIDE ANY ADDITIONAL INFORMATION. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED INSIDE A CARTRIDGE. THERE WAS NO VISIBLE DAMAGE TO THE CARTRIDGE. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5898665&pc=MTA | 2023826-2016-01142 | 2016-07-27 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS INSERTING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, AND THE LENS WAS TORN IN THE INJECTOR. THERE WAS NO PATIENT CONTACT OR INJURY. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919705&pc=MTA | 2023826-2016-01162 | 2016-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS INSERTING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, AND THE LENS WAS TORN. THE REPORTER STATED IT WAS UNKNOWN IF THERE WAS ANY PATIENT CONTACT BUT THERE WAS NO INJURY. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: PT AGE AND WEIGHT: UNK. IMPLANT AND EXPLANT DATES: UNK. EVENT PROBLEM AND EVALUATION CODES: RESULTS CODE(S): (OPERATIONAL PROBLEM): - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919701&pc=MTA | 2023826-2016-01190 | 2016-07-29 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, WAS TORN GOING THROUGH THE INJECTOR. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5965489&pc=MTA | 2023826-2016-01302 | 2016-08-08 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON ATTEMPTED TO LOAD A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, BUT THE "GRASPER" TORE THE LENS. THERE WAS NO PATIENT CONTACT OR INJURY. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5957287&pc=MTA | 2023826-2016-01288 | 2016-08-18 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS LOADING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, AND NOTED THE LENS WAS TORN WHILE PULLING THE LENS FORWARD THROUGH THE CARTRIDGE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: PT AGE AND WEIGHT: UNK. IMPLANT AND EXPLANT DATES: NA. EVENT PROBLEM AND EVALUATION CODES: RESULTS CODE(S): (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5965488&pc=MTA | 2023826-2016-01298 | 2016-08-24 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGICAL COORDINATOR ATTEMPTED TO PULL A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, DOWN THROUGH THE CARTRIDGE BARREL AND NOTED THE LENS WAS TORN. THE REPORTER DECIDED NOT TO CONTINUE WITH THIS LENS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS LOADED AND IMPLANTED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES TORN OFF AND MISSING FROM BOTH HAPTICS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885868&pc=MTA | 2023826-2015-00796 |  | Malfunction | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS AND THE LENS TORE. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE PATIENT IS VERY HAPPY AND THEIR VISION IS 20/20. Manufacturer Narrative: (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: TEARS OR NICKS CAN OCCUR AT ANY STAGE FROM MANUFACTURE TO SURGICAL MANIPULATION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: AFTER A REVIEW OF THE DHR, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5189884&pc=MTA | 2023826-2015-01395 |  | Malfunction | STAAR SURGICAL COMPANY | 2015-10-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, AND THE SURGEON NOTED THE LENS WAS TORN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE CUSTOMER INDICATED THE SOURCE OF THE BLUE DOT WAS UNKNOWN BUT THE LENS MOST LIKELY WAS NOT RECEIVED WITH A BLUE DOT AS THE SURGEON WOULD HAVE NOTED IT UPON RECEIPT. THE BLUE DOT PROBABLY OCCURRED DURING OR AFTER THE SURGICAL PROCEDURE. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC AND BOTH HAPTICS TORN, WITH PIECES TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. A BLUE DOT WAS NOTED ON THE LENS. BASED ON THE COMPLAINT HISTORY, VISUAL INSPECTION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS - (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS THAT SUGGESTS A CONTRIBUTORY FACTOR IN THE COMPLAINT. PER MEDICAL REVIEW: REPORTEDLY, LENS DAMAGE ("TEAR") WAS OBSERVED UPON INSERTION REQUIRING INTRAOPERATIVE LENS REMOVAL/EXCHANGE. NO REPORTED INCISION ENLARGEMENT OR OTHER TISSUE DAMAGE. ACCORDING TO REPORT BY A NURSE, DATED ON (B)(6) 2015, DAMAGED LENS WAS REMOVED WITHOUT ANY PATIENT INJURY AND ANOTHER LENS WAS SUCCESSFULLY IMPLANTED. THE SAME REPORT STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. NO REPORTED POSTOPERATIVE SEQUELAE. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5290352&pc=MTA | 2023826-2015-01647 |  | Malfunction | STAAR SURGICAL COMPANY | 2015-12-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, AND THE LENS TORE. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INFORMATION PROVIDED BY THE FACILITY, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. PER MEDICAL REVIEW - REPORTEDLY, MICL WAS DAMAGED (TORN) UPON INSERTION REQUIRING INTRAOPERATIVE LENS REMOVAL. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. NO REPORTED POSTOPERATIVE SEQUELAE. DESPITE MULTIPLE ATTEMPTS NO ADDITIONAL INFORMATION HAS BEEN RECEIVED TO DATE. CONCLUSION - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5464095&pc=MTA | 2023826-2016-00206 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-02-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, TORE AND THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW - REPORTEDLY, INTRAOPERATIVE MICL REMOVAL WAS PERFORMED TO ADDRESS UNSPECIFIED LENS TEAR. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE DURING LENS REMOVAL. NO REPORTED POSTOPERATIVE SEQUELAE. NO ADDITIONAL INFORMATION WAS RECEIVED TO DATE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5518778&pc=MTA | 2023826-2016-00304 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER WAS NOTED TO BE TORN PRIOR TO REMOVING IT FROM THE LENS VIAL. THE LENS WAS NOT USED AND THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553203&pc=MTA | 2023826-2016-00413 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, AND THE LENS WAS NOTED TO BE TORN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO A LOADING ERROR. Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. PROCESS EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5758705&pc=MTA | 2023826-2016-00875 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-06-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS ADVANCING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, FORWARD FROM THE CARTRIDGE WITH A FORCEP AND THE LENS TORE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323176&pc=MTA | 2023826-2015-01723 | 2015-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER BUT THE LENS RIPPED. THE REPORTER STATED IT WAS UNKNOWN IF THERE WAS ANY PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE CUSTOMER REPORTED THERE WAS NO PATIENT CONTACT OR INJURY AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE TO THE LENS, OR INCORRECT DIMENSIONS OF THE LENS RESULTING IN A LOSS OF MECHANICAL INTEGRITY. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. PER MEDICAL REVIEW - REPORTEDLY, MICL (13.2 MM) WAS DAMAGED ("RIPPED") DURING SURGERY REQUIRING INTRAOPERATIVE LENS EXCHANGE. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. NO REPORTED POSTOPERATIVE SEQUELAE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND TWO LENSES RETURNED IN THE VIAL. THE LENSES WERE STUCK TOGETHER, IN A DRY CONDITION AND THE LENSES COULD NOT BE EVALUATED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5304879&pc=MTA | 2023826-2015-01697 |  | Malfunction | STAAR SURGICAL COMPANY | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, MISFOLDED AND WAS DAMAGED. THERE WAS NO PATIENT CONTACT OR INJURY AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULTS (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF THE LENS OPTIC AND ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE SURGERY FACILITY BUT NONE HAD BEEN FORTHCOMING. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4655095&pc=MTA | 2023826-2015-01120 | 2015-08-04 | Malfunction | STAAR SURGICAL COMPANY | 2015-04-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/099 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS FLIPPED IN THE EYE AND UPON REMOVAL THE LENS TORE. THERE WAS NO REPORTED PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. THE LENS WAS REHYDRATED AND BOTH THE LENGTH AND WIDTH WERE RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ETC.) TO PREVENT ANY OF THESE COMPLICATIONS FORM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5370942&pc=MTA | 2023826-2016-00059 | 2015-12-17 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, AND THE LENS FLIPPED IN THE EYE. THE LENS WAS REMOVED AND RELOADED. THE SURGEON INSERTED THE LENS A 2ND TIME AND THE LENS TORE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE THE LENS TO UNFOLD INCORRECTLY. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED THE INCORRECT ORIENTATION. PER MEDICAL REVIEW - REPORTEDLY, LENS WAS INSERTED UPSIDE DOWN AND WAS TORN OFF DURING SURGEON`S SECOND ATTEMPT TO INSERT IT AFTER RE-LOADING THE SAME LENS. LENS WAS REMOVED/EXCHANGED INTRAOPERATIVELY. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE TO THE PATIENT. ACCORDING TO THE REPORT, BY A TECHNICIAN, DATED ON (B)(6) 2016 THE CAUSE OF THE EVENT WAS UNKNOWN. THE SAME REPORT STATED THAT THERE WAS NO PATIENT INJURY AND THAT BACKUP LENS WAS SUCCESSFULLY IMPLANTED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A LARGE PORTION OF THE LENS OPTIC AND ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5518780&pc=MTA | 2023826-2016-00298 | 2016-02-24 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS FLIPPED. THE LENS TORE WHILE BEING REMOVED DURING THE SAME SURGERY, WITH NO PATIENT INJURY, NO ENLARGED INCISION OR SUTURES. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047786&pc=MTA | 2023826-2015-01117 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/3.0/94 DIOPTER, AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITH NO APPARENT INJURY. ANOTHER SAME MODEL LENS LENS WAS IMPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5429394&pc=MTA | 2023826-2016-00128 | 2015-06-09 | Malfunction | STAAR SURGICAL COMPANY | 2016-02-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS REMOVED WITH NO REPORTED INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODE/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ACD BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD - PROCESS EVALUATION: WORK ORDER SEARCH EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047792&pc=MTA | 2023826-2015-01140 | 2015-07-30 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED A POSSIBLE CAUSE OF THE EVENT WAS DUE TO A LOADING ERROR, THERE WAS NO ISSUE WITH THE LENS OR THE INJECTION SYSTEM. Manufacturer Narrative: PER MEDICAL REVIEW - FACILITY REPORTED THAT LENS WAS INSERTED UPSIDE DOWN IN THE EYE. REPORT NO MALFUNCTION OF THE INSERTION SYSTEM, AND WORK-ORDER SEARCH SHOWED ONE SIMILAR COMPLAINT. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. THE CAUSE OF THE LENS TO NOT UNFOLD CORRECTLY, AS STATED BY THE REPORTING FACILITY, IS POSSIBLY DUE TO IMPROPER LOADING. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066057&pc=MTA | 2023826-2015-01212 | 2015-08-10 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, AND NOTED THE LENS FLIPPED UPSIDE DOWN INSIDE THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY, NO ENLARGED INCISION OR SUTURES. THE REPORTER STATED THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: PER MEDICAL REVIEW - THE LENS "FLIPPED UP UPSIDE DOWN DURING INSERTION". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF THE FLIP WAS DUE TO IMPROPER LENS LOADING, OR SURGEON HANDLING EITHER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY. BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW, EVALUATION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5102457&pc=MTA | 2023826-2015-01218 | 2015-08-25 | Malfunction | STAAR SURGICAL COMPANY | 2015-09-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER AND THE LENS FLIPPED. THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: PER MEDICAL REVIEW - THE SURGEON REPORTED THAT THE LENS FLIPPED UPSIDE DOWN DURING INSERTION. THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF THE FLIP WAS DUE TO IMPROPER LENS LOADING, OR SURGEON HANDLING. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW, EVALUATION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5138905&pc=MTA | 2023826-2015-01260 | 2015-09-08 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, AND THE LENS FLIPPED. THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, VISUAL INSPECTION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - REPORTEDLY, ICL WAS INSERTED UPSIDE DOWN REQUIRING INTRAOPERATIVE LENS EXCHANGE. INCISION WAS NOT ENLARGED AND NO OTHER TISSUE DAMAGE WAS REPORTED. ANOTHER ICL WAS IMPLANTED SUCCESSFULLY WITH NO INJURY TO THE PATIENT. ACCORDING TO THE REPORT, BY A TECHNICIAN, DATED ON (B)(6) 2015 THE CAUSE OF THE EVENT WAS UNKNOWN SINCE THE LENS "FLIPPED UPON INSERTION" CAUSING TO BE PLACED UPSIDE DOWN IN THE EYE. THE TECHNICIAN STATED THAT STAAR INJECTION SYSTEM WAS USED DURING THE SURGERY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5138850&pc=MTA | 2023826-2015-01261 | 2015-09-09 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: RESULTS: (EVALUATION RESULT): EVALUATION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - REPORTEDLY, ICL WAS INSERTED UPSIDE DOWN REQUIRING INTRAOPERATIVE LENS EXCHANGE. INCISION WAS NOT ENLARGED AND NO OTHER TISSUE DAMAGE WAS REPORTED. ANOTHER ICL WAS IMPLANTED SUCCESSFULLY WITH NO INJURY TO THE PATIENT. ACCORDING TO THE REPORT, BY A TECHNICIAN, DATED ON (B)(6) 2015 THE CAUSE OF THE EVENT WAS UNKNOWN SINCE THE SURGEON NOTED THAT LENS WAS PLACED UPSIDE DOWN IN THE EYE UPON INSERTION. THE TECHNICIAN STATED THAT STAAR INJECTION SYSTEM WAS USED DURING THE SURGERY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323205&pc=MTA | 2023826-2015-01719 | 2015-09-20 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS WAS IMPLANTED INSIDE OUT. THERE WAS NO REPORTED PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/13.3. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5439663&pc=MTA | 2023826-2016-00163 | 2016-01-14 | Malfunction | STAAR SURGICAL COMPANY | 2016-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS FLIPPED. THE LENS WAS REMOVED WITH NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER STATED THE EVENT OCCURRED DURING THE SAME SURGERY AND THE BACKUP LENS WAS IMPLANTED. THERE WAS NO PATIENT INJURY, NO ENLARGED INCISION OR SUTURES. Manufacturer Narrative: METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE OR NOT DETECT A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE CAUSED INVERSION, FRONT TO BACK, OF THE LENS DURING INSERTION. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553065&pc=MTA | 2023826-2016-00414 | 2016-03-09 | Injury | STAAR SURGICAL COMPANY | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, AND THE LENS WAS NOTED TO BE UPSIDE DOWN IN THE PATIENT'S EYE. THE SURGEON HAD SOME DIFFICULTY REMOVING THE LENS DURING THE SAME SURGERY BUT THERE WAS NO PATIENT INJURY. THE SURGEON DECIDED NOT TO IMPLANT ANOTHER MICL LENS AT THAT TIME AND THE SURGERY WAS ABORTED. ANOTHER MICL LENS WAS IMPLANTED AT A LATER DATE WITH NO PROBLEM AND THE PATIENT IS DOING WELL. Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5611364&pc=MTA | 2023826-2016-00543 | 2016-03-24 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER. THE LENS FLIPPED AND WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870489&pc=MTA | 2023826-2016-01046 | 2016-07-07 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE PATIENT MOVED WHILE THE SURGEON WAS INSERTING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, AND THE LENS FLIPPED IN THE PATIENT'S EYE. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS PATIENT RELATED AND THERE WAS NO PRODUCT MALFUNCTION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073837&pc=MTA | 2023826-2018-01603 | 2018-05-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VTICMO 13.2, -13.00/3.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FLIPPED. THE LENS WAS REMOVED AND EXCHANGED DURING INITIAL SURGERY WITH NO PATIENT INJURY. REPORTER STATES UPSIDEDOWN WHEN IMPLANTING THE LENS. THE LENS TORE WHEN REMOVING. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND RESIDUE ON THE LENS. THE REPORTER INDICATED THAT ON (B)(6) 2018 A VTICMO13.2, -13.00/3.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FLIPPED UPSIDE DOWN DURING IMPLANTATION. THE LENS WAS REMOVED IN THE SAME SURGERY AND TORE DURING REMOVAL WITH NO PATIENT INJURY. A DIFFERENT LENS WAS LATER IMPLANTED. THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7707568&pc=MTA | 2023826-2018-01069 | 2018-05-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE SURGEON STATED THAT A 12.6MM MICL12.6, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS INJECTED INTO THE PATIENT'S LEFT EYE (OS) AND IT FLIPPED OVER IN THE AC (ANTERIOR CHAMBER). THE LENS WAS REMOVED FROM THE ANTERIOR CHAMBER AND ANOTHER LENS WAS IMPLANTED DURING THE SAME SURGERY. THE LENS TORE DURING REMOVAL BUT THERE WAS NO PATIENT INJURY. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED DRY, STUCK TO LENS VIAL. BSS WAS ADDED TO REMOVE LENS FROM VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC AND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7871125&pc=MTA | 2023826-2018-01362 | 2018-06-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, AND THE LENS EJECTED UPSIDE DOWN. THE LENS WAS REMOVED DURING THE SAME SURGERY, TO BE RE-INSERTED AND THE LENS TORE. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870375&pc=MTA | 2023826-2018-01376 | 2018-08-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. AFTER INSERTION, THE LENS "FLIPPED IN THE EYE." INTRAOPERATIVELY, THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS. REPORTEDLY, NO PATIENT INJURY OCCURRED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7954103&pc=MTA | 2023826-2018-01482 | 2018-09-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, AND THE LENS FLIPPED UPSIDE DOWN IN THE PATIENTS EYE. THE SURGEON DID NOT LIKE HOW THE LENS LOOKED, SO DECIDED TO REMOVE THE LENS. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM AND THE PATIENT IS DOING WELL. THE REPORTER WAS UNABLE TO PROVIDE ANY ADDITIONAL INFORMATION. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A POUCH. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8025966&pc=MTA | 2023826-2018-01587 | 2018-10-03 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS INJECTED INTO THE PATIENTS RIGHT EYE (OD) UPSIDE DOWN ON (B)(6) 2018 AND DURING REMOVAL THE LENS TORE. THE LENS WAS REMOVED AND REPLACED WITH THE BACK UP LENS DURING INITIAL SURGERY WITH NO PATIENT INJURY.CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8203029&pc=MTA | 2023826-2018-01826 | 2018-11-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A 13.2MM VTICM5_13.2, -12.00/1.5/098 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS RIGHT EYE (OD). THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO BEING IMPLANTED UPSIDE DOWN. IT WAS REPORTED THE LENS WAS DAMAGED WHEN THE SURGEON REMOVED THE LENS. AND THAT SURGEON WILL IMPLANT A NEW LENS AT A LATER DATE TO RESOLVE THE PROBLEM. "NO PATIENT INJURY IS REPORTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255277&pc=MTA | 2023826-2019-00075 | 2018-12-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED, REMOVED AND EXCHANGED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, WITHIN THE SAME SURGERY, DUE TO THE LENS FLIPPED WHEN IT WAS FULLY INSERTED. UPON REMOVAL OF THE LENS, THE LENS TORE. THERE WAS NO PATIENT INJURY. THE PROBLEM WAS RESOLVED WHEN THE LENS WAS EXCHANGED FOR THE BACKUP LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656464&pc=MTA | 2023826-2019-00945 | 2019-01-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, DURING THE SAME SURGERY DUE TO THE LENS FLIPPED OVER DURING IMPLANTATION. THE LENS WAS BROKEN WHILE BEING REMOVED FROM THE EYE DURING THE SAME SURGERY. THERE WAS NO PATIENT INJURY. ANOTHER ICL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8438526&pc=MTA | 2023826-2019-00430 | 2019-02-07 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, AND NOTED THE LENS WENT INTO THE PATIENTS EYE UPSIDE DOWN. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8485050&pc=MTA | 2023826-2019-00514 | 2019-02-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, - 11.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS DAMAGED WHEN REMOVED. THERE WAS NO PATIENT INJURY. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC TORN OFF. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824592&pc=MTA | 2023826-2020-00492 | 2019-03-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. DURING THE PROCESS OF INSERTING THE LENS INTO THE EYE,THE LENSL TURNED OVER AND WAS ACCIDENTALLY TORN DURING THE ADJUSTMENT. THE LENS WAS REMOVED AND EXCHANGED WITH THE SAME LENGTH LENS ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8686757&pc=MTA | 2023826-2019-00979 | 2019-05-14 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 THE SURGEON INSERTED A 13.2MM VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE LENS "UNFOLDED IN THE WRONG ORIENTATION." THE LENS WAS DAMAGED DURING REMOVAL AND BACK UP LENS WAS IMPLANTED DURING THE INITIAL SURGERY. THE PATIENT'S VISION IS 6/6 UA. CAUSE OF THE EVENT IS REPORTED AS "WRONG ORIENTATION OF THE LENS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912825&pc=MTA | 2023826-2019-01507 | 2019-05-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REPORTEDLY, THE LENS FLIPPED DURING INJECTION AND WAS UPSIDE DOWN. INTRA-OPERATIVELY THE LENS WAS REMOVED AND THE LENS TORE DURING EXPLANT. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. "NO PATIENT INJURY" IS REPORTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8798414&pc=MTA | 2023826-2019-01131 | 2019-06-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM MICL 12.6, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PER THE REPORTER, "THE LENS OPENED UPSIDE DOWN", THE LENS WAS REMOVED AND BACK UP LENS IMPLANTED DURING INITIAL SURGERY WITH NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: "VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN" SHOULD BE CORRECTED TO "VISUAL INSPECTION FOUND THE HAPTIC TORN". (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8798411&pc=MTA | 2023826-2019-01127 | 2019-06-06 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. "LENS FLIPPED IN EYE". LENS WAS REMOVED AND BACK UP LENS IMPLANTED DURING INITIAL SURGERY. SURGEON DOES NOT KNOW WHY THE LENS FLIPPED; NO LOADING ISSUES WERE REPORTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813784&pc=MTA | 2023826-2019-01206 | 2019-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -4.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2019. THE SURGEON INDICATED THAT AFTER INSERTION, THE LENS "UNFOLDED UPSIDE DOWN" IN THE EYE. INTRAOPERATIVELY, THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS. REPORTEDLY, NO PATIENT INJURY. "PATIENT DOING WELL, WITH NO EVIDENCE OF PROBLEMS IN THE EYE." Manufacturer Narrative: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565426&pc=MTA | 2023826-2020-00008 | 2019-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM MICL 13.7, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS FLIPPED UPSIDE DOWN ON INSERTION INTO PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS REMOVED AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. REPORTEDLY, "LENS FLIPPED UPSIDE DOWN ON INSERTION, WITH LENS 3/4 OF THE WAY INJECTED. LENS DAMAGED ON REMOVAL." Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9639511&pc=MTA | 2023826-2020-00190 | 2019-12-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -2.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REPORTEDLY, THE PATIENT MOVED SUDDENLY AT THE TIME OF IMPLANT, CAUSING THE LENS TO BE IMPLANTED UPSIDE DOWN. INTRA-OPERATIVELY THE LENS WAS REMOVED AND THE LENS TORE DURING EXPLANT. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. "NO PATIENT INJURY" IS REPORTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271242&pc=MTA | 2023826-2020-01475 | 2020-04-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5 12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN AND REMOVED FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PATIENT CONTACT (LENS TOUCHED THE EYE). NO PATIENT INJURY. THE LENS WAS EXCHANGED WITH A SIMILAR LENS INTRAOPERATIVELY AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10125666&pc=MTA | 2023826-2020-01216 | 2020-05-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN (UNFOLDED UPSIDE DOWN) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITH THE BACKUP LENS WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. PROBLEM RESOLVED, NO ISSUES, VISION DOING WELL. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189318&pc=MTA | 2023826-2020-01276 | 2020-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, +05.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS FLIPPED UPSIDE DOWN IN THE EYE UPON INSERTION. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE REPLACEMENT LENS RESOLVED THE PROBLEM AND THE PATIENT IS SEEING VERY WELL. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CORRECTED DATA: D10: RETURNED TO MANUFACTURER: 08-JUN-2020. (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: UPDATED TO ADVERSE EVENTS. B2: UPDATED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. H1: UPDATED TO SERIOUS INJURY. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323184&pc=MTA | 2023826-2015-01724 |  | Malfunction | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER BUT THE LENS WAS INSERTED INTO THE EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: PER MEDICAL REVIEW - REPORTEDLY, MICL WAS INSERTED UPSIDE DOWN REQUIRING INTRAOPERATIVE LENS EXCHANGE. ANOTHER LENS WAS SUCCESSFULLY IMPLANTED. ACCORDING TO THE REPORT, DATED ON (B)(6) 2015, INCISION WAS NOT ENLARGED AND THERE WAS NO OTHER INJURY TO THE PATIENT. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE THE LENS TO UNFOLD INCORRECTLY. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED THE INCORRECT ORIENTATION. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325938&pc=MTA | 2023826-2015-01745 |  | Malfunction | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER AND THE LENS WENT INTO THE PATIENT'S EYE BACKWARDS. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE FOR THIS COMPLAINT COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH WOULD HAVE CAUSED THE LENS TO BE IMPLANTED UPSIDE DOWN. PHYSICIAN REPORT DOES NOT INDICATE ANY ADVERSE EVENT OR CONDITION OR ANY IMPROPER HANDLING THAT WOULD HAVE CAUSED THE LENS TO BE IMPLANTED UPSIDE DOWN. PER MEDICAL REVIEW: REPORTEDLY, MICL WAS INSERTED BACKWARDS REQUIRING INTRAOPERATIVE LENS REMOVAL/EXCHANGE. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ACCORDING TO REPORT, DATED ON (B)(6) 2015, THERE WAS NO PATIENT INJURY DURING LENS REMOVAL AND BACKUP LENS WAS SUCCESSFULLY IMPLANTED. THE SAME REPORT STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. NO REPORTED POSTOPERATIVE SEQUELAE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5392152&pc=MTA | 2023826-2016-00102 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE TO THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. PER MEDICAL REVIEW: REPORTEDLY, MICL WAS INSERTED UPSIDE DOWN REQUIRING INTRAOPERATIVE LENS REMOVAL/EXCHANGE. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ACCORDING TO THE REPORT BY A TECHNICIAN, DATED ON (B)(6) 2016, THE CAUSE OF THE EVENT WAS UNKNOWN. THE SAME REPORT STATED THAT THERE WAS NO PATIENT INJURY AND THAT BACKUP LENS WAS IMPLANTED WITHOUT PROBLEMS. NO REPORTED POSTOPERATIVE SEQUELAE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN A SMALL AMOUNT OF LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783188&pc=MTA | 2023826-2018-01214 | 2018-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENTS LEFT EYE (OS) UPSIDE DOWN . THE LENS WAS REMOVED AND THE BACK UP LENS WAS IMPLANTED WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY, CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: THE AWARE DATE IS 07/16/18 WHICH IS CORRECTED FROM 07/17/18 IN THE INITIAL MDR. THE DATE SHOULD HAVE BEEN 07/16/18 IN THE INITIAL MDR. (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7830755&pc=MTA | 2023826-2018-01200 | 2018-07-30 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, AND THE LENS FLIPPED UPON DEPLOYMENT AND TORE. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY. ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED IN THE SAME SURGERY AND THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A ZIP LOC BAG. VISUAL INSPECTION FOUND ONE HAPTIC TORN, WITH RESIDUE ON THE LENS. CORRECTED DATA: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, AND THE LENS FLIPPED UPON DEPLOYMENT. THE LENS WAS DAMAGED UPON REMOVAL FROM THE PATIENTS EYE, DURING THE SAME SURGERY WITH NO PATIENT INJURY. ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED IN THE SAME SURGERY AND THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918795&pc=MTA | 2023826-2019-01491 | 2019-03-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; No Apparent Adverse Event | No Consequences Or Impact To Patient; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED WITHIN THE SAME SURGERY, FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: HALOES WERE ALSO REPORTED. UDI#: N/A. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OPTIC TORN/HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163041&pc=MTA | 2023826-2020-01258 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL 13.7, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH AN INCISION ENLARGEMENT INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. REPORTEDLY, "DIFFICULT LOADING, LENS TRAILING HAPTIC STOCK UNDER THE FOAM TIP PLUNGER. NO DAMAGE TO THE EYE." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS VIAL. VISUAL OBSERVATION FOUND LENS OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). H1- IN INITIAL MDR AND SUPPLEMENTAL MDR SHOULD BE CORRECTED TO SERIOUS INJURY. H6-HEALTH IMPACT IMPACT: (B)(4) : INCISION ENLARGEMENT. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7834998&pc=MTA | 2023826-2018-01316 | 2018-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2, -14.50/+3.5/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE SO THE SURGEON REMOVED THE LENS WITHIN THE SAME SURGERY, THERE WAS NO PATIENT INJURY. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DIOPTER CORRECTED TO: "-14.00/+3.5/173. LENS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS TORN INTO TWO PIECES. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 -14.00/+3.5/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE SO THE SURGEON REMOVED THE LENS WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2018 A NEW LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IIS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7943005&pc=MTA | 2023826-2018-01455 | 2018-08-06 | Injury | STAAR SURGICAL COMPANY | 2018-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Consequences Or Impact To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE FRONT PART OF A VICMO12.6, -04.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS TORN WHEN IT WAS IMPLANTED INTO PATIENTS LEFT EYE ON (B)(6) 2018. IT WAS REPORTED THAT THERE WAS PATIENT CONTACT WITH NO INJURY. THE SURGEON DECIDED TO EXCHANGE THE LENS ON (B)(6) 2018 WHEN NEW LENS ARRIVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT; PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066032&pc=MTA | 2023826-2018-01655 | 2018-05-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. ON THE SAME DATE, THE LENS WAS REMOVED AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. IT IS UNKNOWN AS TO WHETHER THIS OCCURRED INTRA-OPERATIVELY OR NOT. "NO PATIENT INJURY" IS INDICATED.Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT (OD) EYE, THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. INTRAOPERATIVELY, THE LENS WAS REMOVED AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. "NO PATIENT INJURY" IS INDICATED. THE SURGEON REPORTS THAT NOTHING UNUSUAL WAS NOTED DURING THE SURGERY TO CAUSE THIS EVENT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INPSECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT.(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8051080&pc=MTA | 2023826-2018-01594 | 2018-10-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Consequences Or Impact To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VICMO 12.1, -02.25 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, ON (B)(6) 2018. A REPLACEMENT LENS WAS IMPLANTED. PATIENT IS DOING WELL AND IS "EXTREMLY" HAPPY. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON LENS AND PORTION OF LENS WAS RETURNED. CORRECTED DATA: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VICMO 12.1, -02.25 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED WITHIN THE SAME SURGERY ON (B)(6) 2018. A REPLACEMENT LENS WAS LATER IMPLANTED ON (B)(6) 2018. PATIENT IS DOING WELL AND IS EXTREMELY HAPPY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). "WORD" ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164365&pc=MTA | 2023826-2019-01870 | 2019-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Consequences Or Impact To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER TORE/BROKE DURING INJECTION. THIS OCCURRED ON (B)(6) 2019. THE LENS TOUCHED THE EYE BUT WAS NOT FULLY IMPLANTED. REPORTEDLY, NO PATIENT INJURY OCCURRED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. THE REPORTER STATES THAT CURRENTLY, THE PATIENT IS WAITING FOR A NEW ICL TO BE IMPLANTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE IN INITIAL MDR. 'THE LENS TOUCHED THE EYE BUT WAS NOT FULLY IMPLANTED' CORRECTED TO 'REPORTEDLY, THERE WAS PATIENT CONTACT AND THE LENS WAS EXPLANTED' IN INITIAL MDR. TYPE OF REPORTABLE EVENT: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. PATIENT CODE 2199 CORRECT TO PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION, EXPLANT) IN INITIAL MDR. DEVICE CODE 3189 CORRECTED TO DEVICE CODE 2993 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5352444&pc=MTA | 2023826-2016-00001 | 2015-06-30 | Injury | STAAR SURGICAL COMPANY | 2016-01-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 -18.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS TORE DURING INSERTION INTO THE EYE AND THE LENS WAS REMOVED. ANOTHER LENS, SAME MODEL AND SAME DIOPTER SIZE WAS IMPLANTED ON (B)(6) 2015. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS CASE IS AN ADVERSE EVENT, NOT PRODUCT MALFUNCTION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00. DEVICE EVALUATED BY MANUFACTURER: NO, DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5309192&pc=MTA | 2023826-2015-01632 | 2015-10-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -10.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED. A BACKUP LENS, SAME MODEL AND DIOPTER WAS IMPLANTED. NO ADVERSE EVENT REPORTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: PMA 510(K0: - CORRECTED AWARE DATE - (B)(6) 2016 (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.50. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540382&pc=MTA | 2023826-2016-00378 | 2016-01-22 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO REPORTED INJURY. ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE OPTIC AND THE HAPTIC WERE TORN. CONCLUSION CODE - AS PER DFU OF THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL HAVING AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.8MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5495686&pc=MTA | 2023826-2016-00260 | 2015-12-24 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material; Unintended Movement | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS DAMAGED WHILE REMOVING FROM THE EYE WITH NO REPORTED INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A SAME MODEL LENS. THIS RESOLVED THE PROBLEM. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6678860&pc=MTA | 2023826-2017-01066 | 2017-05-09 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENTS RIGHT EYE AND NOTICED THE LENS WAS TORN. PORTION OF THE LENS HAPTIC WAS FOUND IN THE DELIVERY DEVICE. THE INCISION WAS ENLARGED TO 3.2MM TO REMOVE THE LENS, NO SUTURES WERE REQUIRED. THE REPORTED STATED THAT THE DELIVERY SYSTEM EJECTED THE LENS SLOWER THAN USUAL BUT THAT THE CAUSE OF THE EVENT WAS UNKNOWN. THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES, THE PATIENT LEFT THE OPERATING ROOM IN GOOD CONDITION. Manufacturer Narrative: CORRECTED DATA: THE DEVICE WAS RETURNED TO MANUFACTURER ON 03-JUL-2017. (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801531&pc=MTA | 2023826-2017-01346 | 2017-07-14 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.00/1.00/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT IS EXPERIENCING REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ACCORDING TO THE REPORTER THE LENS IS UPSIDE DOWN AND AN EXCHANGE IS PLANNED FOR (B)(6) 2017. Manufacturer Narrative: EXPLANTED DATE: (B)(6) 2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS FOUND TO BE 20/20. (B)(4). CORRECTED DATA: THE STATEMENT "THE PATIENT IS EXPERIENCING REFRACTIVE SURPRISE" IS CORRECTED TO "THE SURGEON NOTED REFRACTIVE SURPRISE". THE CLAIM NUMBER IN THE INITIAL SHOULD BE "(B)(4)". Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE US. RESULT: WORK ORDER SEARCH WAS PERFORMED- NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197806&pc=MTA | 2023826-2018-00101 | 2017-12-19 | Injury | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A TEAR ON A MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, WAS NOTICED DURING INSERTION. THE REPORTER STATED THAT PATIENT CONTACT IS UNKNOWN BUT THAT THERE WAS NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES. THE REPORTER STATED THAT THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER WAS UNABLE TO PROVIDE FURTHER INFORMATION. Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND PORTION OF A HAPTIC MISSING AND A HAPTIC TEAR. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261924&pc=MTA | 2023826-2018-00166 | 2017-12-21 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL12.6, -7.5 DIOPTER, WAS IMPLANTED ON (B)(6) 2017 AND THEN REMOVED AND EXCHANGED ON (B)(6) 2017 AS A RESULT OF SHALLOW ANGLE AND NORMAL INTRAOCULAR PRESSURE. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY DURING PROCEDURE. IN THE SURGEONS OPINION THE LENS MALFUNCTION. THE LENS WAS EXCHANGED WITH WITH A SAME MODEL AND DIOPTER LENS. Manufacturer Narrative: (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7294652&pc=MTA | 2023826-2018-00272 | 2018-01-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS TORN/BROKEN DURING THE SAME SURGERY. THERE WAS NO INJURY AND THE PATIENT WAS IN GOOD CONDITION. CORRECTED DATA: (B)(4). Manufacturer Narrative: CORRECTED INFORMATION: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293176&pc=MTA | 2023826-2015-01650 | 2015-09-17 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -12.0 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. LOW VAULT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01649 FOR RIGHT EYE. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND IN THE WORK ORDER SEARCH. (B)(4). Manufacturer Narrative: DIOPTER: -12.00. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261525&pc=MTA | 2023826-2018-00163 | 2017-12-18 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Scratched Material | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS IMPLANTING A MICL13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE AND A SCRATCH WAS NOTICED. THE LENS WAS REMOVED AND THE BACK UP LENS WAS USED. THE REPORTER STATED THAT THERE WAS NO PATIENT INJURY BECAUSE IT WAS STILL IN THE UNFOLDING PROCESS SO IT WAS ONLY HALF WAY INTO THE EYE. IT WAS REPORTER THAT THE EVENT CAUSE WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON WAS IMPLANTING A MICL13.2 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017, WHEN HE NOTICED A SCRATCH. THE LENS WAS REMOVED AND THE BACK-UP LENS WAS USED. IT IS UNCERTAIN IF THE LENS REMOVAL OCCURRED INTRAOPERATIVELY OR EXTRAOPERATIVELY. THE REPORTER STATED THAT THERE WAS NO PATIENT INJURY BECAUSE IT WAS STILL IN THE UNFOLDING PROCESS SO IT WAS ONLY HALF WAY INTO THE EYE. IT WAS REPORTED THAT THE EVENT CAUSE WAS UNKNOWN. EXPLANTED DATE: IT IS CORRECTED TO UNK. (B)(4). Manufacturer Narrative: EXPLANT DATE: IT IS CORRECTED FROM "UNK" TO "18-DEC-2017". (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMLIAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5662052&pc=MTA | 2023826-2016-00698 | 2016-02-17 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/+1.00/X094 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS TORE DURING DELIVERY INTO THE EYE. DURING THE SAME PROCEDURE THE LENS WAS REMOVED FROM THE EYE AND EXCHANGED FOR ANOTHER SAME MODEL WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/20. NO ADVERSE EVENT REPORTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE. (B)(4). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14/+1.5/94 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO A LENS TEAR/BREAK DURING INJECTION. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR ANOTHER SAME MODEL AND SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5662051&pc=MTA | 2023826-2016-00697 | 2016-03-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/+2.00/X064 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS GOT STUCK IN THE INJECTOR TIP AND PLUNGER AND TORE. DURING THE SAME PROCEDURE THE LENS WAS REMOVED FROM THE EYE AND EXCHANGED FOR ANOTHER SAME MODEL WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2016; UCVA WAS 20/20. NO ADVERSE EVENT WAS REPORTED. Manufacturer Narrative: CORRECTED DATA: THIS CASE IS AN ADVERSE EVENT. THE LENS WAS EXCHANGED FOR A SIMILAR LENS WITH A -15/+2,5/96 DIOPTER ON (B)(6) 2016. (B)(4). Manufacturer Narrative: (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED OBSERVING THE OPTIC WAS TORN/SPLIT, HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON THE LENS SURFACE (UNIDENTIFIED). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE CORRECT LENS POWER IS -15.00/+2.50/X064 DIOPTER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6058822&pc=MTA | 2023826-2016-01478 | 2016-06-22 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/90 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND THE LENS WAS NOTED TO BE TORN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LENS THE SAME DIOPTER AND LENGTH AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN. WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THIS CASE IS AN ADVERSE EVENT, NOT A PRODUCT PROBLEM. THE CORRECT IMPLANT DATE PER MEDICAL REVIEW IS (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5654689&pc=MTA | 2023826-2016-00639 | 2016-04-07 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult To Position; Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, AND NOTED THE LENS WAS TOO LONG FOR THE PATIENT'S EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. ANOTHER LENS WAS NOT IMPLANTED. THE SURGEON FELT THE LENS MAY HAVE BEEN MISLABELED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS INSERTED AND REMOVED FROM THE LEFT EYE (OS). THE INCISION WAS ENLARGED TO REMOVE THE LENS AND SUTURES WERE REQUIRED TO CLOSE THE INCISION. THE REPORTER INDICATED THE ICL WOULD NOT POSITION CORRECTLY AND KEPT ROTATING OUT OF POSITION. DURING REMOVAL, THE ICL CAUGHT THE IRIS WHICH CAUSED DEPIGMENTATION AND ATROPHY TEMPORALLY. THE PATIENT EXPERIENCED PHOTOSENSITIVITY. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5981803&pc=MTA | 2023826-2016-01322 | 2015-10-15 | Injury | STAAR SURGICAL COMPANY | 2016-09-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO 12.6 IMPLANTABLE COLLAMER LENS - -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059195&pc=MTA | 2023826-2016-01487 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, PIGMENT DISPERSION, AND VERY SHALLOW AC. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED BROKEN IN THE BAG. LENS RETURNED IN PIECES AND UNABLE TO EVALUATE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6926706&pc=MTA | 2023826-2017-01555 | 2017-08-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.0/083 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED AND REMOVED DURING THE SAME SURGERY AND THERE WAS NO PATIENT INJURY. CORRECT DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SHORTER LENS. CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7448874&pc=MTA | 2023826-2018-00585 | 2018-02-20 | Injury | STAAR SURGICAL COMPANY | 2018-04-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. THE PATIENT'S POST-OP UCVA WAS REPORTED TO BE 20/15. (B)(4) Manufacturer Narrative: IMPLANTED DATE: IT IS CORRECTED FROM "(B)(6) 2018" TO "(B)(6) 2018". EXPLANTED DATE: IT IS CORRECTED FROM "UNK" TO "(B)(6) 2018". (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A HAPTIC TEAR. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7890067&pc=MTA | 2023826-2018-01348 | 2018-08-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM MICL 13.7 , -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) AND REMOVED IT ON (B)(6) 2018. DURING INITIAL SURGERY DUE TO OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. ADDITIONAL INFORMATION REPORTED; PATIENT HAS RESIDUAL ASTIGMATISM WITH PLANS FOR PRK (PHOTOREFRACTIVE KERATECTOMY) IN 1 MONTH. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8383748&pc=MTA | 2023826-2019-00344 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE DOCTOR FEELS THAT A 12.6MM, MICL12.6, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS THAT WAS IMPLANTED ON (B)(6) 2019 IS TOO LARGE AND WILL BE REMOVING IT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8383903&pc=MTA | 2023826-2019-00345 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -13.0 DIOPTER IMPLANTABLE COLLAMER LENS THAT WAS IMPLANTED ON (B)(6) 2018 IS SCHEDULED TO BE REMOVED DUE TO SIZING ISSUE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: EVENT: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2018 AND EXPLANTED ON (B)(6) 2019 DUE TO SIZING ISSUE. DEVICE EVALUATED BY MFR: LENS WAS RETURNED DRY, IN CARTRIDGE CASE. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC WERE TORN. THERE WAS RESIDUE ON LENS, THE LENS WAS RETURNED IN FOUR PIECES. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789591&pc=MTA | 2023826-2015-00550 | 2015-04-29 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO BEING TOO LARGE FOR THE PATIENT'S EYE. THERE WAS NO REPORTED INJURY AND THE LENS WAS EXCHANGED FOR ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS EXCHANGED FOR A SHORTER LENS. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE MICL LENS WAS INSERTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED WITHIN THE SAME SURGERY. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5217461&pc=MTA | 2023826-2015-01439 | 2015-06-15 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED. NO OTHER LENS WAS IMPLANTED. NO ADVERSE EVENT WAS REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.00. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5217468&pc=MTA | 2023826-2015-01417 | 2015-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND EXCHANGED FOR ANOTHER ICL, SAME MODEL AND DIOPTER SIZE. NO ADVERSE EVENT WAS REPORTED. LAST POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: (B)(4). THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND OPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.00. DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5216032&pc=MTA | 2023826-2015-01450 | 2015-10-12 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+2.50/X009 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED. NO ADVERSE EVENT WAS REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN AND PEN MARKS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -15.50/+2.50/X009. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323202&pc=MTA | 2023826-2015-01720 | 2015-09-20 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS WAS IMPLANTED INSIDE OUT. THERE WAS NO REPORTED PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/13.3. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7932092&pc=MTA | 2023826-2018-01477 | 2018-08-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Consequences Or Impact To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS FLIPPED DURING INSERTION AND UPON REMOVAL THE LENS TORE. INTRAOCULAR INJECTION WAS REPORTED. A SAME MODEL AND DIOPTER LENS WAS ORDERED AND IMPLANTED ONE WEEK LATER. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5243913&pc=MTA | 2023826-2015-01530 | 2015-10-10 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS IMPLANTED UPSIDE DOWN. THE LENS WAS REMOVED AND DURING REMOVAL FROM THE EYE, THE LENS TORE. THE LENS WAS EXCHANGED FOR ANOTHER LENS, SAME MODEL AND SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. REPORTER INDICATED THE PATIENT WAS UN-COOPERATIVE AND SUDDEN MOVEMENT LED TO UPSIDE DOWN AND LENS TEAR DURING REMOVAL OF LENS. NO ADVERSE EVENTS REPORTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND LIQUID ON LENS SURFACE. THE PREVIOUS STATEMENT OF IT BEING A PRODUCT PROBLEM IS INCORRECT. THE CORRECT STATEMENT IS THAT THERE WAS AN ADVERSE EVENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -10.00. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7855722&pc=MTA | 2023826-2018-01354 | 2018-04-06 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Information | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -7.5/1.5/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -7.50/1.5/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. LATER, A NEW SHORTER LENS WAS IMPLANTED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). "LENS WAS RETURNED DRY" NOT INCLUDED IN THE SUPPLEMENTAL MDR#1. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540077&pc=MTA | 2023826-2016-00360 | 2015-11-11 | Malfunction | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 -09.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. THE TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND WAS REPLACED WITH ANOTHER SAME MODEL AND DIOPTER SIZE LENS. NO ADVERSE EVENT REPORTED. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN AND LENS HAVING FOREIGN MATERIAL ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -09.00. METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549343&pc=MTA | 2023826-2016-00368 | 2016-03-05 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 --06.00/+3.50/X122 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS TORN WHILE IMPLANTING INTO THE EYE. THE ICL WAS REMOVED DURING THE SAME PROCEDURE. NO INJURY TO THE PATIENT. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -06.00/+3.50/X122 (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5215711&pc=MTA | 2023826-2015-01481 | 2015-06-24 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE. THERE WAS NO REPORTED PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218850&pc=MTA | 2023826-2015-01419 | 2015-06-24 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AN EXCHANGED FOR ANOTHER ICL, SAME MODEL AND DIOPTER SIZE. NO ADVERSE EVENT WAS REPORTED. LAST POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: CORRECTED DATA: (B)(4) - PREVIOUS CODE NOT APPLICABLE, THERE IS NO STATEMENT DECLARING NO PATIENT INJURY. (B)(4). PREVIOUS CODE NOT APPLICABLE, THERE IS NO BREAK OR TORN MATERIAL ASSOCIATED WITH A MATERIAL INTEGRITY ISSUE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. DEVICE EVALUATED BY MANUFACTURER? NO- DEVICE NOT RETURNED. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WAS ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5217475&pc=MTA | 2023826-2015-01420 | 2015-08-15 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AN EXCHANGED FOR ANOTHER ICL, SAME MODEL AND DIOPTER SIZE. NO ADVERSE EVENT WAS REPORTED. UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES WHICH WAS UNABLE TO BE EVALUATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -04.50. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6196969&pc=MTA | 2023826-2016-01803 | 2016-09-22 | Injury | STAAR SURGICAL COMPANY | 2016-12-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -3.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND IT WAS NOTED THAT THE LENS WAS TORN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LENS THE SAME SIZE, MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248742&pc=MTA | 2023826-2015-01547 | 2015-10-15 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material; Unintended Movement | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+4.5/061 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REMOVED ON (B)(6) 2015 DUE TO THE LENS UNFOLDED UPSIDE DOWN. THE LENS TORE WHEN REMOVED . THERE WAS NO REPORTED PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE STATUS OF THE EYE IS EXCELLENT. Manufacturer Narrative: CORRECTED DATA: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632998&pc=MTA | 2023826-2016-00541 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Failure to Unfold or Unwrap; Activation Failure | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -3.00/+1.50/X173 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE. THE LENS DID NOT FULLY UNFOLD IN THE EYE. LENS WAS REMOVED AND EXCHANGED WITH ANOTHER SAME LENS MODEL BUT SMALLER DIOPTER SIZE. NO INJURY TO THE PATIENT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. UNIDENTIFIED FOREIGN MATERIAL WAS PRESENT ON THE LENS SURFACE. AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.96MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613093&pc=MTA | 2023826-2016-00542 | 2016-03-10 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Failure to Unfold or Unwrap; Activation Failure | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -04.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (B)(6) 2016. THE LENS WOULD NOT UNFOLD IN THE EYE. THE LENS WAS REMOVED AND EXCHANGED WITH A SIMILAR LENS. THIS RESOLVED THE PROBLEM. NO ADVERSE EVENT WAS REPORTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE. NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7744579&pc=MTA | 2023826-2018-01159 | 2018-06-14 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -15.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED INTRAOPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS SURFACE. CORRECTED DATA: (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7747988&pc=MTA | 2023826-2018-01154 | 2018-05-30 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND THE PATIENT WAS IN GOOD CONDITION AFTER THE SURGERY. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A ,MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). DATE RETURNED SHOULD BE CORRECTED TO 08/30/2018 IN PREVIOUS MDR. DATE RECEIVED SHOULD BE CORRECTED TO 09/14/2018 IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8203046&pc=MTA | 2023826-2018-01883 | 2018-11-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VICM5_12.1, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING LOADING. THERE WAS NO PATIENT CONTACT (LENS DID NOT TOUCH THE EYE). ON (B)(6) 2018 AN ALTERNATE LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: PATIENT CODE CORRECTED TO: "2199" ONLY, IN THE INITIAL MDR. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876513&pc=MTA | 2023826-2019-01390 | 2019-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE BACK HAPTIC OF A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, WAS NOTED TO BE TORN OFF PRIOR TO INJECTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS INSERTED AND REMOVED DURING THE SAME SURGERY DUE TO THE LENS WAS NOTED TO BE TORN. THERE WAS NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANTED: UNK. EXPLANTED: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7749004&pc=MTA | 2023826-2018-01157 | 2018-04-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; No Apparent Adverse Event | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2018. THE PATIENT'S SUSPENSORY LIGAMENT WAS FLABBY, WHICH MAY CAUSE A DIPLACEMENT OF THE LENS SO PATIENT IS NOT SUITABLE FOR IMPLANTING LENS. ON (B)(6) 2018 IT WAS REPORTED THAT THE ICL WAS DAMAGED, PATIENT CURRENT STATUS: IN GOOD CONDITION. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE LENS WAS NOT DAMAGED AS INITIALLY REPORTED. (B)(4) Manufacturer Narrative: INFORMATION ADDED: REPORTEDLY, "SUSPENSORY LIGAMENT [WAS] WEAK IN SUPPORT AND THE ARCH OF THE [LENS] WAS LOW. THE SURGEON WAS AFRAID OF DAMAGING THE LENS SO HE REMOVED IT." DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136376&pc=MTA | 2023826-2017-01935 | 2017-11-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Consequences Or Impact To Patient; No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS ATTEMPTING TO LOAD A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/080 (SPHERE/CYLINDER/AXIS), AND THE LENS BROKE. THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: IN THE INITIAL MDR, (B)(6) 2017 CORRECTED TO (B)(6) 2017. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS), LENS MISSING A PIECE OF HAPTIC, AND LENS HAVING A CURVED SHAPE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7971671&pc=MTA | 2023826-2018-01492 | 2018-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON FOUND A VICM5_13.2, -06.00 DIOPTER, INTRAOCULAR LENS DAMAGE (LENS TEAR/BREAK) BEFORE IMPLANTING. THERE WAS NO PATIENT CONTACT AND NO PATIENT INJURY. THE REPLACEMENT LENS (SAME MODEL AND DIOPTER) WAS IMPANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: DEVICE RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. RESULT CODE UPDATED (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7974255&pc=MTA | 2023826-2018-01493 | 2018-05-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON FOUND A VICM5_13.2, -08.50 DIOPTER, INTRAOCULAR LENS DAMAGE (LENS TEAR/BREAK) BEFORE IMPLANTING. THERE WAS NO PATIENT CONTACT AND NO PATIENT INJURY. THE REPLACEMENT LENS ( SAME MODEL AND DIOPTER) WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: CORRECTED TO "(B)(6)". ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT WITH PRESENCE OF RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794656&pc=MTA | 2023826-2018-01198 | 2018-07-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO LOAD A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00 DIOPTER AND THE LENS TORE WHILE THE SURGEON WAS PULLING THE LENS VIA LOADING FORCEPS INSIDE THE CARTRIDGE. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL LENS WAS LOADED AND IMPLANTED. THE PATIENT IS FINE AND THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8438229&pc=MTA | 2023826-2019-00464 | 2019-02-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT WITH THE DEVICE BUT NO PATIENT INJURY. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10147987&pc=MTA | 2023826-2020-01247 | 2019-12-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A 12.6MM, VICM5_12.6, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE, NO PATIENT INJURY. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CORRECTED DATA: B4 - (B)(6) 2020. B5 - THE REPORTER INDICATED THAT ON (B)(6) 2019 A 12.6MM VICM5_12.6; -13.00 DIOPTER IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT CONTACT/INJURY. ON (B)(6) 2019 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913112&pc=MTA | 2023826-2020-00773 | 2020-03-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS GOT STUCK IN THE CARTRIDGE DURING LOADING PROCESS. THERE WAS NO PATIENT CONTACT AND THE CAUSE OF THE EVENT IS UNKNOWN. THE PATIENT HAS RECEIVED A REPLACEMENT LENS AND IS SATISFIED. Manufacturer Narrative: THE REPORTER STATED THAT THE LENS WAS IMPLANTED , REMOVED AND REPLACED WITHIN THE SAME SURGERY. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H6 - PATIENT CODE 2645 CORRECTED TO PATIENT CODE 2199. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073812&pc=MTA | 2023826-2018-01675 | 2018-08-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient; No Patient Involvement; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO13.7, -8.00/.05/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING DELIVERY. IT WAS REPORTED THAT THERE WAS NO PATIENT CONTACT AND NO PATIENT INJURY. BACK UP LENS WAS USED AND DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577306&pc=MTA | 2023826-2017-00861 |  | Injury | STAAR SURGICAL COMPANY AG | 2017-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | No Information | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. THE LENS WAS RETURNED IN BIOHAZARD BAG. NO FURTHER INFORMATION AVAILABLE. MULTIPLE ATTEMPTS TO CONTACT REPORTER HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5055150&pc=MTA | 2023826-2015-01099 | 2015-08-02 | Injury | STAAR SURGICAL COMPANY | 2015-09-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/05/122 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. SURGERY IS PENDING TO EXPLANT AND EXCHANGE THE LENS DUE EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/05/122. NO CODE AVAILABLE (VAULTING). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WAS NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066131&pc=MTA | 2023826-2015-01176 | 2015-08-05 | Injury | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.50/+1.00/X081 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS NOTED ONE WEEK POST-OP. A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. SEE MFR # 2023826-2015-01177 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -5.50/+1.00/X081. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066174&pc=MTA | 2023826-2015-01177 | 2015-08-05 | Injury | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS NOTED ONE WEEK POST-OP. A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. SEE MFR # 2023826-2015-01176 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -5.00. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5991808&pc=MTA | 2023826-2016-01359 | 2016-07-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+2.5/150 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND SPOT ON SURFACE; MEASUREMENT WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5991827&pc=MTA | 2023826-2016-01361 | 2016-09-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+3.0/122 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO REMOVE AND EXCHANGE FOR A SMALLER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5991829&pc=MTA | 2023826-2016-01362 | 2016-09-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO REMOVE AND EXCHANGE FOR A SMALLER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT CODE (B)(4) - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CONCLUSION CODE AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) UNDER 3.0MM. CORRECTED DATA: THE CORRECT REPORT SOURCE COUNTRY IS (B)(6). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7855444&pc=MTA | 2023826-2018-01353 | 2018-03-22 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.5/1.5/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7857553&pc=MTA | 2023826-2018-01359 | 2018-06-16 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO OBSERVATION OF LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS WAS RETURNED TORN IN TWO PIECES WHICH WERE STUCK TOGETHER ALONG WITH RESIDUE ON THE LENS SURFACE. CORRECTED DATA: ORIGINAL MDR STATED "ON (B)(6) /2018 THE LENS WAS EXCHANGED FOR A LONGER LENS" WHICH SHOULD BE CORRECTED TO "ON (B)(6) 2018 THE LENS WAS EXPLANTED. ON (B)(6) 2018 A REPLACEMENT LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED." (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE (B)(4): SECONDARY SURGICAL INTERVENTION, EXPLANT (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9334502&pc=MTA | 2023826-2019-02168 | 2019-05-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENT'S RIGHT EYE (OS) AND "SMALL VAULT (70UM) WAS OBSERVED. REPORTER INDICATED CAUSE OF EVENT WAS "SIZING ISSUE." A 13.2MM VICM5_13.2 WAS IMPLANTED INTO THE PATIENT'S RIGHT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- "RIGHT EYE" SHOULD BE CORRECTED TO "LEFT EYE" IN THE INITIAL MDR. B5- "A 13.2MM VICM5_13.2 WAS IMPLANTED INTO THE PATIENT'S RIGHT." SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. MODEL #: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4636966&pc=MTA | 2023826-2015-00386 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT. THE LENS WILL BE EXPLANTED AND EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER: NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVALUATION CODES: METHOD: MEDICAL REVIEW RESULTS: PER MEDICAL REVIEW - ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE IDENTIFIED. AS REPORTED ON THE MEDICAL REVIEW, EXCESSIVE VAULTING IS A CONSEQUENCE OF WRONG LENS USE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF WHITE TO WHITE MEASUREMENTS BASED ON TECHNIQUES, IMPROPER SULCUS TO SULCUS MEASUREMENT, AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS, IRREGULAR SULCUS OR CILIARY SULCUS CYST). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5984001&pc=MTA | 2023826-2016-01329 |  | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS WAS TOO SHORT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. DATE OF EVENT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890690&pc=MTA | 2023826-2018-01300 |  | Injury | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information | P030016 | 1 | Event Description: RECEIVED A REPORT WHERE THE SURGEON HAS BEEN HAVING DIFFICULTY WITH THE SIZING OF IMPLANTABLE COLLAMER LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: EXPIRATION DATE UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7835570&pc=MTA | 2023826-2018-01301 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information | P030016 | 1 | Event Description: RECEIVED A REPORT WHERE THE SURGEON HAS BEEN HAVING AN ISSUE WITH SELECTING THE LENS SIZES FOR IMPLANTABLE COLLAMER LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7835580&pc=MTA | 2023826-2018-01304 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information | P030016 | 1 | Event Description: RECEIVED A REPORT WHERE THE SURGEON HAD QUESTIONS IN REGARDS TO VAULTING OF IMPLANTABLE COLLAMER LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7835599&pc=MTA | 2023826-2018-01309 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information | P030016 | 1 | Event Description: RECEIVED A REPORT WHERE THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS. THE SURGEON HAD A QUESTION ON ICL EXCHANGE FOR A SMALLER ICL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7835609&pc=MTA | 2023826-2018-01308 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information | P030016 | 1 | Event Description: RECEIVED A REPORT WHERE THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS IN A -13.00 PATIENT WITH VAULT ISSUES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4757687&pc=MTA | 2023826-2015-00529 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-05-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2015. THE LENS HAD AN EXCESSIVE VAULTING AND THE SURGEON WAS PLANNING TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: THE REPORTER PROVIDED ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, BLURRY VISION, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS INCORRECT VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PER CUSTOMER COMPLAINT QUESTIONNAIRE COMPLETED 05/22/2015, THE SURGEON INDICATED PATIENT'S BSCVA WAS 20/20 WITH NO PROBLEMS DURING THE LAST VISIT. PLEASE DISREGARD THE FOLLOWING STATEMENT: "THE PATIENT UNCORRECTED VISUAL ACUITY WAS 20/20." DATE RECEIVED BY MFR - 05/22/2015. PER MEDICAL REVIEW, THE SURGEON NOTES THAT THE CAUSE OF THE EVENT WAS USER ERROR DUE TO PRE-OP MISCALCULATION (" POOR PREDICTION OF SULCUS DIAMETER FROM PRE-OP MEASUREMENTS"). (B)(4). Manufacturer Narrative: (B)(4). (OTHER): LENS NOT RETURNED. EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: PER MEDICAL REVIEW - WRONG LENS SIZE WAS THE CAUSE OF THE EVENT. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: (B)(4). RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4870131&pc=MTA | 2023826-2015-00715 | 2015-05-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAS AN EXCESSIVE VAULT AND WILL BE EXPLANTED FOR A SHORTER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. PER MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4922994&pc=MTA | 2023826-2015-00822 | 2015-05-18 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO EXPLANT THE LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247899&pc=MTA | 2023826-2015-01551 | 2015-07-24 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A LOW VAULT AND THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5400086&pc=MTA | 2023826-2016-00077 | 2015-08-17 | Injury | STAAR SURGICAL COMPANY | 2016-01-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5400111&pc=MTA | 2023826-2016-00076 | 2015-08-17 | Injury | STAAR SURGICAL COMPANY | 2016-01-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5399894&pc=MTA | 2023826-2016-00078 | 2015-10-13 | Injury | STAAR SURGICAL COMPANY | 2016-01-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT. THE INTRAOCULAR PRESSURE IS NORMAL. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: ADVERSE EVENT TO PRODUCT PROBLEM; PREVIOUSLY, B1 WAS INCORRECTLY REPORTED AS AN ADVERSE EVENT. THE INFORMATION PRESENTED IN THE CLAIM SHOWS THAT IT IS A PRODUCT PROBLEM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225076&pc=MTA | 2023826-2015-01511 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A LOW VAULT AND THE PLAN IS TO EXPLANT THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS REMAINS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5375633&pc=MTA | 2023826-2016-00067 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+3.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON THE LENS SURFACE. DATE OF EVENT WAS INCORRECT. THE DATE OF EVENT IS (B)(6) 2015. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5597060&pc=MTA | 2023826-2016-00487 | 2016-01-15 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629107&pc=MTA | 2023826-2015-00383 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAS AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT THE LENS AND EXCHANGE FOR A SHORTER LENS. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CAE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4) - NO INFORMATION; SIZE INCORRECT FOR PATIENT; (VAULTING); (IMPLANTED). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4889340&pc=MTA | 2023826-2015-00803 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. THE SURGEON IS PLANNING TO EXPLANT THE LENS DUE TO EXCESSIVE VAULTING AND THE LENS DID NOT WORK FOR THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). LENS IMPLANTED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5880928&pc=MTA | 2023826-2016-01099 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S EYE IN (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND WILL BE EXPLANTED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806367&pc=MTA | 2023826-2015-00582 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND NOTED EXCESSIVE VAULT IN THE POST-OPERATIVE PERIOD. THIS WAS ASSOCIATED WITH A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED AND THE SURGEON IS PLANNING TO REMOVE AND EXCHANGE THIS LENS FOR A SHORTER LENS. Manufacturer Narrative: (B)(4). (OTHER): LENS REMAINS IMPLANTED. EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH, MEDICAL REVIEW RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4516764&pc=MTA | 2023826-2015-00152 | 2014-11-05 | Injury | STAAR SURGICAL COMPANY | 2015-02-13 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND LOW VAULT WAS NOTED. THE LENS WAS ROTATED AND REPOSITIONED. THE SURGEON IS PLANNING ON EXPLANTING THE LENS FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT WEIGHT: UNK. EXPLANT DATE: NA. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN (B)(4) OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENT'S ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION ECT.). ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5240537&pc=MTA | 2023826-2015-01520 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2015-11-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -5.0/+4.5/100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A LOW VAULT AND LENS ROTATION. THE SURGEON IS PLANNING TO EXPLANT THE LENS AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686541&pc=MTA | 2023826-2016-00773 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S EYE. THE LENS WILL BE EXPLANTED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR. REPORT # 2023826-2016-00774 FOR THE OTHER EYE. Manufacturer Narrative: (B)(4). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686542&pc=MTA | 2023826-2016-00774 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S EYE. THE LENS WILL BE EXPLANTED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR. REPORT # 2023826-2016-00773 FOR THE OTHER EYE. Manufacturer Narrative: (B)(4). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CLAIM # (B)(4). THE LENS WAS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972065&pc=MTA | 2023826-2016-01319 |  | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS IN A PATIENT AND SOON AFTER REPORTED COMPLICATIONS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4613564&pc=MTA | 2023826-2015-00327 | 2014-12-29 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS ROTATED AND REMAINS IMPLANTED. THE REPORTER INDICATED THE ROTATION WAS NOT ASSOCIATED WITH A LOW VAULT. Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE SURGEON REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S UCVA IS OPTIMAL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) -DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4625075&pc=MTA | 2023826-2015-00334 | 2015-02-19 | Injury | STAAR SURGICAL COMPANY | 2015-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS WAS DRIFTING DOWN INTO THE PATIENT'S LINE OF VISION. THE LENS WILL BE EXPLANTED AND EXCHANGED FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO DISLOCATION/SUBLUXATION (DECENTRATION) OF THE LENS. THERE WERE NO ADDITIONAL COMPLICATIONS RELATED TO THE EVENT. THE REPORTER INDICATED THE CAUSE WAS DUE TO THE ICL WAS TOO SHORT (INSUFFICIENT LENGTH). THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW RESULTS: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE IDENTIFIED. AS REPORTED ON THE MEDICAL REVIEW, IT IS LIKELY, THE EVENT WAS A CONSEQUENCE OF WRONG LENS USE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, OR A MISMATCH OF WHITE TO WHITE AND OR SULCUS TO SULCUS MEASUREMENT). CLAIM #(B)(4). Manufacturer Narrative: PT WEIGHT.: UNK EVENT DATE: UNK. EXPLANT DATE: UNK. INITIAL REPORTER: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GIVEN THE REPORTED INFORMATION, IT IS VERY LIKELY THAT THE SMALLER LENS SIZE CAUSED DECENTRATION OF THE LENS. ACCORDING TO THE FMEA (FAILURE MODES AND EFFECT ANALYSIS), THE WRONG SIZING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED ON DIFFERENT METHOD UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN INDIVIDUAL BASIS AND PATIENT CONDITION). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: (B)(4). RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. THE INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER WERE NOT RETURNED. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4625076&pc=MTA | 2023826-2015-00333 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS WAS DRIFTING DOWN INTO THE PATIENT'S LINE OF VISION. THE LENS WILL BE EXPLANTED AND EXCHANGED FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO DISLOCATION/SUBLUXATION (DECENTRATION) OF THE LENS AND A CATARACT HAD DEVELOPED. THE PATIENT EXPERIENCED DIPLOPIA AND HALOS. THE REPORTER INDICATED THE CAUSE WAS DUE TO THE ICL WAS TOO SHORT (INSUFFICIENT LENGTH). THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED ON (B)(6) 2015. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE IDENTIFIED. AS REPORTED ON THE MEDICAL REVIEW, IT IS LIKELY, THE THE EVENT WAS A CONSEQUENCE OF WRONG LENS USE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, OR A MISMATCH OF WHITE TO WHITE AND OR SULCUS TO SULCUS MEASUREMENT). (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GIVEN THE REPORTED INFORMATION, IT IS VERY LIKELY THE SMALL SIZE OF ICL LENS CONTRIBUTED TO THE DISLOCATION/SUBLUXATION OF LENS WHICH CAUSED THE PATIENT'S SYMPTOMS OF HALO AND DIPLOPIA. ACCORDING TO THE FMEA (FAILURE MODES AND EFFECT ANALYSIS), THE EVENT WAS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED ON DIFFERENT METHOD UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT ( IF UBM USED) ;POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN INDIVIDUAL BASIS AND PATIENT CONDITION). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: PT WEIGHT: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. (B)(4). LENS NOT RETURNED. EVALUATION, METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. THE INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER WERE NOT RETURNED. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5478181&pc=MTA | 2023826-2016-00091 | 2015-11-17 | Injury | STAAR SURGICAL COMPANY | 2016-03-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, - 18.00/+2.50/X079 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. AFTER IMPLANTATION, REFRACTIVE SURPRISE WAS NOTED. THE ICL REMAINS IMPLANTED. LAST OFFICE VISIT ON (B)(6) 2015, BCVA: 20/20, UCVA: 20/25. SEE MFR. REPORT #2023826-2016-00092 FOR RIGHT EYE. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00/+2.50/X079. DEVICE EVALUATED BY MANUFACTURER?: NO. ICL REMAINS IMPLANTED. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). THE ICL REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5478184&pc=MTA | 2023826-2016-00092 | 2015-11-17 | Injury | STAAR SURGICAL COMPANY | 2016-03-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/+3.00/X094 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. AFTER IMPLANTATION, REFRACTIVE SURPRISE WAS NOTED. THE ICL REMAINS IMPLANTED. LAST OFFICE VISIT ON (B)(6) 2015, BCVA: 20/20, UCVA: 20/30. SEE MFR. REPORT #2023826-2016-00091 FOR LEFT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -17.00/+3.00/X094. DEVICE EVALUATED BY MANUFACTURER?: NO - ICL REMAINS IMPLANTED.(B)(4). CONCLUSIONS CODE(S): CONCLUSION- CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6001931&pc=MTA | 2023826-2016-01399 | 2016-05-17 | Injury | STAAR SURGICAL COMPANY | 2016-10-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NO EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: SHALLOW ANGLES WERE REPORTED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND FOREIGN MATERIAL NOT POSSIBLE TO SPECIFY ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5991812&pc=MTA | 2023826-2016-01358 | 2016-07-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -5.50/+4.5/150 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/ SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (SPOT) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6616593&pc=MTA | 2023826-2017-00944 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-06-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED DUE TO LOW VAULTING OF THE LENS, THE PLAN IS TO EXPLANT AND EXCHANGE. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: PREVIOUS: STATED LENS REMAINS IMPLANTED. UPDATED: LENS WAS EXCHANGED ON (B)(6) 2017. ADDITIONAL INFORMATION REGARDING LENS EXCHANGE IS UNKNOWN (I.E. LENS SIZE, PROBLEM RESOLUTION). DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: "STUDY" IN FOLLOW UP #1 WAS CHECKED BY MISTAKE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7857365&pc=MTA | 2023826-2018-01358 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO_12.6 IMPLANTABLE COLLAMER LENS (-8.5 DIOPTER) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO LOW VAULT ON (B)(6) 2018. A LONGER REPLACEMENT LENS WAS IMPLANTED IN A LATER SURGERY ON (B)(6) 2018; PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7855987&pc=MTA | 2023826-2018-01355 | 2018-06-26 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -15/1.5/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED, WITH FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. LENS HAPTIC WAS CAUGHT IN VIAL CAP. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO_12.6, -15/1.5/79 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A SHORTER REPLACEMENT LENS WAS IMPLANTED POST-OPERATIVELY ON (B)(6) 2018 TO RESOLVE THE ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7857161&pc=MTA | 2023826-2018-01357 | 2018-07-04 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A DEFORMED HAPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: VISUAL INSPECTION- "LENS WAS CAUGHT ON LENS VIAL CAP" NOT INCLUDED IN SUPPLEMENT MDR#1. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574376&pc=MTA | 2023826-2015-00264 | 2014-10-30 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND REMOVED IT DUE TO LOW VAULT. THE LENS WAS NOT REPLACED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVALUATION: EVALUATION METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5355940&pc=MTA | 2023826-2016-00032 | 2015-10-05 | Injury | STAAR SURGICAL COMPANY | 2016-01-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2, 2.50/+6.00/X072 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WAS NOTED AND THE LENS WAS REMOVED FROM THE EYE. NO INFORMATION IF LENS WAS REPLACED. SUPPLEMENTAL WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS RECEIVED. Manufacturer Narrative: (B)(4). THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND DRIED AND DISCOLORED FOREIGN MATERIAL ON LENS SURFACE BUT NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: 2.50/+6.00/X072. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593300&pc=MTA | 2023826-2016-00486 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/078 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND REMAINS IMPLANTED. Manufacturer Narrative: THE PATIENT HAD A LENS EXCHANGE WITH A MULTIFOCAL IOL. NO IMPACT ON PATIENT'S VISION REPORTED. PREVIOUS REPORT SOURCE, USER FACILITY, IS INCORRECT. THE CORRECT REPORT SOURCE IS DISTRIBUTOR. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593301&pc=MTA | 2023826-2016-00485 | 2016-01-27 | Injury | STAAR SURGICAL COMPANY | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+0.5/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND REMAINS IMPLANTED. Manufacturer Narrative: PATIENT HAD A LENS EXCHANGE WITH MULTIFOCAL IOL. THERE IS NO IMPACT ON PATIENT'S VISION REPORTED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). THE PATIENT IS OVER THE AGE OF (B)(6) YEARS OLD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596433&pc=MTA | 2023826-2016-00491 | 2016-02-01 | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND WAS REPOSITIONED. THE PROBLEM WAS RESOLVED AND THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CONCLUSION: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT'S AGE BELOW 21 OR OVER 45 YEARS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5659927&pc=MTA | 2023826-2016-00664 | 2016-03-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/058 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. IT IS PLANNED TO EXCHANGE FOR A LONGER LENS. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL INFORMATION: DESCRIBE EVENT OR PROBLEM - THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS. THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 AS OF (B)(6) 2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4). EVENT PROBLEM AND EVALUATION CODES: (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5937655&pc=MTA | 2023826-2016-01238 | 2016-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/1.50/89 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. IT WAS INDICATED THE PATIENT EXPERIENCED EXCESSIVE VAULTING, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.91MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5710824&pc=MTA | 2023826-2016-00672 | 2016-05-06 | Injury | STAAR SURGICAL COMPANY | 2016-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+3.0/X160 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED ON (B)(6) 2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED IN THE PRESENCE OF KERATOCONUS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816236&pc=MTA | 2023826-2016-00970 | 2016-06-02 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/115 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. THE LENS IS PLANNED TO EXCHANGE FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+0.5/115 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED TO HAVE A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS, HOWEVER THE ICL VAULT DID NOT IMPROVE. PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.95 MM AT THE TIME OF IMPLANTATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816119&pc=MTA | 2023826-2016-00972 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/179 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE LENS IS PLANNED TO EXCHANGE FOR A SHORTER LENS. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/30. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938278&pc=MTA | 2023826-2016-01253 | 2016-07-03 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND WILL BE EXPLANTED. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE REPORTER CORRECTED THE INITIAL IMPLANT DATE TO (B)(6) 2016. THE SURGEON NOTICED THE LOW VAULT ON (B)(6) 2016. THE REPORTER STATED THAT THE INITIAL LENS WAS EXCHANGED WITH A LONGER LENS (SN: (B)(4) ON (B)(6) 2016. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869094&pc=MTA | 2023826-2016-01040 | 2016-08-02 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM513.7 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT. THE LENS WAS PLANNED TO EXCHANGE FOR A SHORTER LENS. Manufacturer Narrative: LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THE PROBLEM WAS RESOLVED. PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 DURING THEIR LAST VISIT ON (B)(6) 2016. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE. ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869102&pc=MTA | 2023826-2016-01039 | 2016-08-02 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM513.7 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT. THE LENS WAS PLANNED TO EXCHANGE FOR A SHORTER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A SHORTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2016, THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/20. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND THE PRESENCE OF FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5958117&pc=MTA | 2023826-2016-01275 | 2016-08-18 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS REMAINS IMPLANTED AND IS PLANNED FOR EXCHANGED IN OCTOBER DUE TO LOW VAULT. THE PATIENTS POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.94MM AT THE TIME OF IMPLANTATION. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. LOW VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED AS THE SURGEON DECIDED AGAINST A LENS EXCHANGE SURGERY AND INSTEAD DECIDED TO CLOSELY MONITOR THE PATIENT. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 AS OF (B)(6) 2016. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5478719&pc=MTA | 2023826-2016-00230 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WILL BE EXPLANTED DUE TO VAULT IS TOO STEEP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE CUSTOMER REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT IS DOING VERY WELL.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. POST-OP THE LENS HAD A GOOD VAULT AND PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE SURGEON FELT THE CAUSE OF THE EVENT WAS DUE TO A CALCULATION PROBLEM. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW, AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4) NEW INCISION. SECONDARY SURGERY. EXPLANTED. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955103&pc=MTA | 2023826-2016-01262 |  | Injury | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT, WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES.THE LENS REMAINS IMPLANTED BUT WILL BE REMOVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806369&pc=MTA | 2023826-2015-00583 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND NOTED EXCESSIVE VAULT IN THE POST-OPERATIVE PERIOD. THIS WAS ASSOCIATED WITH A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED AND THE SURGEON IS PLANNING TO REMOVE AND EXCHANGE THIS LENS FOR A SHORTER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT AGE AND WEIGHT: UNK. EVENT DATE: UNK. IMPLANT AND EXPLANT DATES: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED.. METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: CONCLUSION: PER DFU, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH A KERATOCONUS EYE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, AND THE HAPTIC WAS OBSERVED TO BE TORN/BROKEN/BENT/DEFORMED. UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: ANGLE CLOUSURE WAS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577175&pc=MTA | 2023826-2017-00854 | 2017-03-28 | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Information; No Code Available | N/A | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2017 THAT A PATIENT FROM (B)(6) HAD A "SIZING ISSUE." ATTEMPTS TO OBTAIN MORE INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7856947&pc=MTA | 2023826-2018-01356 | 2018-03-02 | Injury | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -7.5/4/75 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO OBSERVATION OF LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. CORRECTED DATA: ORIGINAL MDR STATED "ON 06/07/18 THE LENS WAS EXCHANGED" WHICH SHOULD BE CORRECTED TO "ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF LOW VAULT WITH LENS ROTATION. ON (B)(6) 2018 A LONGER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED." (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5132184&pc=MTA | 2023826-2015-01232 | 2016-05-04 | Injury | STAAR SURGICAL COMPANY | 2015-10-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+4.50/X077 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE PATIENT EXPERIENCED EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, IRIS ATROPHY, PUPIL OVALIZATION, AND SYNECHIA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596430&pc=MTA | 2023826-2016-00492 | 2016-02-10 | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/X139 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND LENS ROTATION. THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR A LONGER LENS. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS RECEIVED: THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS WHICH RESOLVED THE PROBLEM. ON PATIENT'S LAST VISIT (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. CONCLUSION: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS PATIENT'S AGE BELOW 21 OR OVER 45 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909069&pc=MTA | 2023826-2016-01187 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+0.5/165 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT, LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016 FOR A SHORTER LENS OF THE SAME MODEL AND SIMILAR DIOPTER WHICH RESOLVED THE PROBLEM. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 ON (B)(6) 2016. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5696626&pc=MTA | 2023826-2016-00671 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING A LOW VAULT, WITH LENS ROTATION. THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED SEVERAL WEEKS AFTER IMPLANTATION. CORRECTED DATA: REPORT SOURCE IS FROM (B)(6), NOT (B)(6). (B)(4). PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED TO PATIENTS WITH PREVIOUS OR PRE-EXISTING OCULAR DISEASE THAT WOULD PRECLUDE POST-OPERATIVE VISUAL ACUITY OF 20/60 OR BETTER. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6632235&pc=MTA | 2023826-2017-00955 |  | Injury | STAAR SURGICAL COMPANY | 2017-06-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Information; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THE SURGEON IS PLANNING TO EXPLANT THE LENS AT A LATER DATE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6438974&pc=MTA | 2023826-2017-00497 | 2017-02-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Activation Failure | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. IT WAS DOCUMENTED THAT AT THE FIRST INSTANCE, THE LENS UNFOLDED, AND WAS IMPLANTED UPSIDE DOWN. DURING EXPLANT, THE LENS TORE. ON THE SAME SURGERY, THE LENS WAS EXCHANGED, AND THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE AND DEBRIS AND RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356132&pc=MTA | 2023826-2018-00429 | 2018-01-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Activation Failure | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD), THE LENS DID NOT UNFOLD. REPORTEDLY, THE FOLDED LENS WAS INSERTED INTO THE ANTERIOR CHAMBER AND IT DID NOT OPEN BUT REMAINED FOLDED. THIS OCCURRED ON (B)(6) 2018. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A REPLACEMENT LENS. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5755928&pc=MTA | 2023826-2016-00867 | 2016-06-08 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Activation, Positioning or Separation Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS INSERTING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S EYE AND THE LENS WOULD NOT COME OUT FROM THE CARTRIDGE. THERE WAS PATIENT CONTACT WITH THE CARTRIDGE BUT NOT THE LENS. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6063728&pc=MTA | 2023826-2016-01503 | 2016-10-04 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Activation, Positioning or Separation Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, AND THE LENS ROLLED OVER IN THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY, NO ENLARGED INCISION OR SUTURES. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6415466&pc=MTA | 2023826-2017-00340 | 2016-11-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED AND REMOVED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, ON (B)(6) 2016. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6769262&pc=MTA | 2023826-2017-01249 | 2017-01-24 | Injury | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: RECEIVED A LENS IMPLANT CARD FROM THE CUSTOMER FOR A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER. THE LENS WAS REPORTED AS "EXCHANGE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899599&pc=MTA | 2023826-2017-01482 | 2017-08-04 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED AND REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER. THERE WAS NO PATIENT INJURY AND THE BACK UP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: DEVICE AVAILABLE FOR EVALUATION: THE DATE THE DEVICE WAS RETURNED TO THE MANUFACTURER IS CORRECTED FROM "11-SEP-2017" TO "18-SEP-2017". CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CARTRIDGE CASE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC BENT. THERE WAS EVIDENCE OF CLEAR AND PURPLE RESIDUE ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6867191&pc=MTA | 2023826-2017-01434 | 2017-08-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS REMOVED DURING THE SAME SURGERY DUE TO THE PATIENT'S SUSPENSORY LIGAMENT WAS FLABBY DURING THE OPERATION, WHICH MAY CAUSE A DISPLACEMENT OF THE LENS. THERE WAS NO REPORT OF ANY PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS AND THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/33. THE REPORTER INDICATED THE EVENT WAS PATIENT RELATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457974&pc=MTA | 2023826-2020-01900 |  | Malfunction | STAAR SUGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_-8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020 AS A REPLACEMENT LENS. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. HOWEVER, " LOW VAULT " WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). LENS ROTATION WAS ALSO REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- LATER REPORTER INDICATED THAT "VAULTING IS FINE NOW, PATIENT IS GOOD" CAUSE OF THE EVENT WAS A "SIZING ISSUE" AND THE REPLACEMENT LENS RESOLVED THE PROBLEM. CORRECTED DATA: B5- "VICM5_" SHOULD BE CORRECTED TO "VICM5_13.7." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6298880&pc=MTA | 2023826-2017-00211 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THEIR LEFT EYE (OS). THE PATIENT REPORTED NEEDS READING GLASSES. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6298928&pc=MTA | 2023826-2017-00210 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THEIR RIGHT EYE (OD). THE PATIENT REPORTED NEEDS READING GLASSES. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6415468&pc=MTA | 2023826-2017-00341 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED AND REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574062&pc=MTA | 2023826-2015-00145 | 2015-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED SHE REMOVED THE LENS FROM THE VIAL WITH FORCEPS AND THE LENS WAS GIVEN TO THE SURGEON TO LOAD. IN THE PROCESS OF LOADING THE LENS, THE SURGEON NOTED, UNDER THE MICROSCOPE, THERE WAS A SMUDGE ON THE LENS OPTIC. HE DECIDED NOT TO USE THE LENS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THE LENS WAS RECEIVED DEFECTIVE. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, PACKAGING OR SHIPPING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, THE MOST LIKELY ROOT CAUSE OF THE EVENT WAS DUE TO HANDLING AND MANIPULATION OF THE LENS BY THE USER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4599020&pc=MTA | 2023826-2015-00316 | 2015-02-10 | Malfunction | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS BUT THE LENS BACK AND TRAILING HAPTICS LOOKED INCORRECT. THE LENS WAS REMOVED WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). HAPTICS LOOKED INCORRECT. DEVICE EVALUATED BY MANUFACTURER? NO: THE LENS BOX WAS RETURNED BUT THERE WAS NO LENS OR VIAL IN THE BOX. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, PACKAGING OR STERILIZATION PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4695675&pc=MTA | 2023826-2015-00332 | 2015-03-09 | Malfunction | STAAR SURGICAL COMPANY | 2015-04-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) BUT DID NOT LIKE HOW THE LENS WAS UNFOLDING. THE LENS WAS PARTIALLY IN THE EYE AND WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE SURGEON DECIDED NOT TO USE THIS LENS AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO DOCTOR LOADING ERROR, THERE WAS NO PROBLEM WITH THE LENS OR INJECTOR. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DRIED SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THE REPORT FROM A MEMBER OF THE FACILITY SUGGESTS THAT MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING. A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS IDENTIFIED AS THE CAUSE OF THE COMPLAINT. THE CAUSE OF THE LENS TO NOT UNFOLD CORRECTLY, AS ALSO STATED BY THE FACILITY, IS POSSIBLY DUE TO IMPROPER LOADING OR HANDLING PRACTICES IN THE OPERATING ROOM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743981&pc=MTA | 2023826-2015-00473 | 2015-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS FROM THE VIAL AND WAS INSPECTING THE LENS, TO PREPARE FOR LOADING. THE SURGEON NOTED A SMALL NOTCH MISSING FROM THE LENS. HE DECIDED NOT TO USE THE LENS, THE LENS WAS NOT LOADED AND THERE WAS NO PATIENT CONTACT OR INJURY. THE REPORTER INDICATED THE SURGEON FELT THE LENS WAS DEFECTIVE. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, PACKAGING OR STERILIZATION PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVALUATION: RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789460&pc=MTA | 2023826-2015-00597 | 2015-04-11 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-05-21 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IMPLANTED UPSIDE DOWN IN THE PATIENT'S RIGHT (OD) EYE . THE LENS WAS EXCHANGE FOR ANOTHER SAME MODEL LENS AND NO PATIENT INJURY WAS REPORTED. PATIENT'S LAST BCVA WAS 20/20 ON (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CLAIM # (B)(4). Manufacturer Narrative: EVALUATION: METHOD: LENS WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED LENS SHOWED IT WAS RETURNED DRY AND A HAPTIC WAS TORN. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW: WHILE THE LENS IS BEING INJECTED INTO THE ANTERIOR CHAMBER THE SURGEON MUST CHECK FOR LENS LANDMARKS TO ENSURE PROPER LENS ORIENTATION. IF THIS IS NOT DONE, THE LENS MAY END UP BEING IMPLANTED UPSIDE-DOWN. ONCE THE SURGEON REALIZES OF THE IMPROPER ORIENTATION, THE LENS MUST BE MOVED TO AVOID COMPLICATION AND A NEW LENS SHOULD BE IMPLANTED. THIS IS ROUTINELY DONE USING THE SAME INCISION. THE MOST PROBABLE CAUSE OF THIS EVENT WAS USER'S ERROR WHILE LOADING/INJECTING THE LENS IN THE ANTERIOR CHAMBER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789537&pc=MTA | 2023826-2015-00598 | 2015-04-11 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-05-21 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IMPLANTED UPSIDE DOWN IN THE PATIENT'S RIGHT (OD) EYE. THE LENS WAS EXCHANGE FOR ANOTHER SAME MODEL LENS AND NO PATIENT INJURY WAS REPORTED. PATIENT'S LAST BCVA WAS 20/20 . Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4803687&pc=MTA | 2023826-2015-00622 | 2015-04-28 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT THE SURGEON NOTED, UNDER THE MICROSCOPE, ONE OF THE HAPTICS DID NOT LOOK RIGHT. THE SURGEON WAS NOT SURE IF THE LENS HAD TORN, SO DECIDED NOT TO USE THE LENS. THERE WAS PATIENT CONTACT WITH THE CARTRIDGE TIP BUT NO CONTACT WITH THE LENS. THERE WAS NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: (B)(4). LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVALUATION: RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, HAPTIC OF MICL DIDN`T APPEAR RIGHT TO SURGEON DURING INSERTION SO SHE DECIDED TO INTRAOPERATIVE EXCHANGE THE LENS. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ACCORDING TO REPORT, DATED ON (B)(6) 2015, ONLY CARTRIDGE HAD A PATIENT CONTACT NOT THE LENS AND THE CAUSE OF THE EVENT WAS UNKNOWN. THE SAME REPORT STATED THAT THERE WAS NO PATIENT INJURY AND THAT ANOTHER LENS WAS IMPLANTED AS A BACKUP. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917238&pc=MTA | 2023826-2015-00800 | 2015-06-15 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS FROM THE VIAL AND NOTED THE LENS DID NOT LOOK RIGHT, THE SURGEON STATED THE LENS LOOKED "WARPED." THE SURGEON DECIDED NOT TO USE THE LENS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: (B)(4) - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT; (WARPED). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. VISUAL INSPECTION OF THE RETURNED LENS FOUND NO VISIBLE DAMAGE TO THE LENS, AND HENCE THE COMPLAINT "LENS RECEIVED DAMAGED" WAS NOT CONFIRMED. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5399892&pc=MTA | 2023826-2016-00079 | 2015-10-13 | Injury | STAAR SURGICAL COMPANY | 2016-01-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT. THE INTRAOCULAR PRESSURE IS NORMAL. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6085639&pc=MTA | 2023826-2016-00891 | 2016-04-08 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, HAD HYPERPIGMENTATION ON THE CRYSTAL OPTICAL SURFACE. THE LENS WAS NOT IMPLANTED AND A BACKUP LENS WAS USED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS SCRATCHED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680327&pc=MTA | 2023826-2016-00768 | 2016-04-21 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, BUT THE LENS WOULD NOT OPEN. THE LENS WAS PARTIALLY INSERTED INTO THE EYE AND WAS REMOVED. THERE WAS NO PATIENT INJURY REPORTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO THE LENS WAS NOT LOADED CORRECTLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: (B)(4). RESULTS CODE(S): (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4509435&pc=MTA | 2023826-2015-00116 |  | Malfunction | STAAR SURGICAL COMPANY | 2015-02-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE FACILITY RETURNED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS TO THE MANUFACTURER STUCK IN A SFC-45 FP CARTRIDGE. THE FACILITY REPORTED THE LENS WAS FOLDED BUT NOT INSERTED DUE TO CRACKED PT. LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER STATED THE SURGEON WAS PERFORMING A YAG ON THE PATIENT AND THE LASER CRACKED THE PATIENT'S NATURAL LENS. THE SURGEON DECIDED TO WAIT AND NOT TO IMPLANT THE ICL. THE ICL WAS LOADED AND WAS NOT USED. THERE WAS NO PATIENT CONTACT. THE LENS DID NOT CAUSE ANY PATIENT INJURY. Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. IMPLANT DATE: NA. EXPLANT DATE: NA. INITIAL REPORTER: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS STUCK IN A CARTRIDGE. THERE WAS NO VISIBLE DAMAGE TO THE CARTRIDGE. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4728604&pc=MTA | 2023826-2015-00464 |  | Malfunction | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WAS FOUND TO BE DEFECTIVE PRIOR TO LOADING AND WAS NOT IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT; DEFECTIVE. EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: AFTER A REVIEW OF THE DHR, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4),
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5730049&pc=MTA | 2023826-2016-00820 | 2016-05-23 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Bent | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S EYE AND NOTED THE HAPTIC WAS BENT AND DID NOT FIT RIGHT. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789345&pc=MTA | 2023826-2015-00596 | 2014-11-11 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE HAPTIC OF THE 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS BROKE DURING INJECTION INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS REMOVED AND REPLACED WITH ANOTHER ONE. PATIENT LAST UCVA WAS 20/50. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. CLAIM # (B)(4). Manufacturer Narrative: EVALUATION CODES: METHOD: - LENS WORK ORDER SEARCH . A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION (USER ERROR): BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, THE MOST PROBABLE CAUSE OF THE LENS TEARING DURING INJECTION IS USER OR TECHNICAL ERROR. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN/ BROKEN. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575803&pc=MTA | 2023826-2016-00442 | 2015-12-24 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -13.00/+2.50/X090 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND REPLACED WITH BACKUP LENS, SAME MODEL AND SIZE DURING THE SAM PROCEDURE. THIS RESOLVED THE PROBLEM. NO ADVERSE EVENT WAS REPORTED. LAST VISIT ON (B)(6) 2016, UCVA WAS 20/16. Manufacturer Narrative: (B)(4). DEVICE EVALUATION-LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED, AND LENS TEAR. CLAIM # (B)(4). Manufacturer Narrative: DIOPTER: -13.00/+2.50/X090. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575872&pc=MTA | 2023826-2016-00443 | 2016-01-14 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 -08.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED ON (B)(6) 2016 DURING THE SAME PROCEDURE AND REPLACED WITH ANOTHER LENS, SAME MODEL AND SIZE THE NEXT DAY (B)(6) 2016. THIS RESOLVED THE PROBLEM. NO ADVERSE EVENT WAS REPORTED. LAST VISIT ON (B)(6) 2016, UCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL (PLASTIC) ON THE LENS SURFACE. (B)(4). THERE WAS NO BREAK IN THE MATERIAL DUE TO MATERIAL INTEGRITY. (B)(4). Manufacturer Narrative: DIOPTER: -08.00. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5174775&pc=MTA | 2023826-2015-01384 | 2015-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: (B)(4). IT WAS REPORTED THAT THE LENS TORE DURING INJECTION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5288025&pc=MTA | 2023826-2015-01637 | 2015-08-18 | Malfunction | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, AND NOTED THE LENS HAD TORN/BROKEN. THE LENS WAS REMOVED WITH NO REPORTED INJURY AND WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5621847&pc=MTA | 2023826-2016-00519 | 2016-01-23 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, AND THE LENS TORE/BROKE. THERE WAS NO REPORT OF ANY PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5648789&pc=MTA | 2023826-2016-00668 | 2016-03-30 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY. THE LENS WAS REMOVED WITH NO APPARENT PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS DURING THE SAME PROCEDURE AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN, BENT, AND HAIR ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831912&pc=MTA | 2023826-2016-00923 | 2016-06-22 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.92MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5835230&pc=MTA | 2023826-2016-00950 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT PATIENT INJURY. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THIS IS AN ADVERSE EVENT, NOT A PRODUCT PROBLEM. DEVICE CODE (B)(4) PREVIOUS CODES NOT APPLICABLE, MATERIAL INTEGRITY ISSUE NOT REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5766315&pc=MTA | 2023826-2016-00872 | 2016-06-26 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00/-2.50/170 DIOPTER, AND THE LENS TORE/BROKE WHILE INJECTING. THE REPORTER INDICATED THE HAPTIC WAS CATCHING ON THE FOAM TIP PLUNGER. THERE WAS NO PATIENT CONTACT. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. "THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12/+2.00/080 DIOPTER, AND THE LENS TORE/BROKE WHILE INJECTING. THE REPORTER INDICATED THE HAPTIC WAS CATCHING ON THE FOAM TIP PLUNGER. THERE WAS NO PATIENT CONTACT. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20." NO MATERIAL INTEGRITY ISSUE REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. CLAIM #(B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655185&pc=MTA | 2023826-2016-00702 | 2016-04-05 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Improper or Incorrect Procedure or Method; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN AND BENT. CORRECTED DATA: (B)(4). THE REPORTER STATES THAT AN ERROR IN LENS LOADING CAUSED THE TEAR IN THE LENS. CONCLUSION CODE: THE REPORTER STATES THAT A USE ERROR WAS MADE DURING LENS LOADING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578348&pc=MTA | 2023826-2015-00224 | 2013-08-27 | Malfunction | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED INJURY AND ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED. THE EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, THE MOST PROBABLE CAUSE OF THE EVENT IS USER OR TECHNICAL ERROR. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4552435&pc=MTA | 2023826-2015-00132 | 2014-07-30 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-02-26 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS BUT THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED PATIENT INJURY AND EXCHANGED FOR ANOTHER ICL. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4) - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT; TORN MATERIAL; BREAK. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED LENS SHOWED A PIECE OF A HAPTIC PLATE WAS TORN OFF AND MISSING. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND PRODUCT EVALUATION, THE MOST PROBABLE CAUSE OF LENS TEARING DURING INJECTION IS A USER OR TECHNICAL ERROR. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5015607&pc=MTA | 2023826-2015-01040 | 2014-12-22 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-08-19 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS WAS UNSTABLE AFTER IT WAS INSERTED INTO THE PATIENT'S LEFT (OS) EYE. THE LENS WAS TORN DURING REPOSITIONING AND WAS REMOVED AND REPLACED WITH ANOTHER SAME MODEL LENS. THIS RESOLVED THE PROBLEM. PATIENT'S LAST UCVA WAS 20/24. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5046692&pc=MTA | 2023826-2015-01141 | 2014-12-29 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND THE LENS TORE/BROKE DURING INSERTION. THE LENS WAS STABLE IN THE EYE DESPITE THE BROKEN HAPTIC, HOWEVER, IT WAS GOING DOWN SO THE LENS WAS EXCHANGED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4695703&pc=MTA | 2023826-2015-00423 | 2015-02-18 | Malfunction | STAAR SURGICAL COMPANY | 2015-04-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM (B)(4) IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE LENS WAS REMOVED WITH NO REPORTED INJURY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVALUATION CODES: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4695704&pc=MTA | 2023826-2015-00427 | 2015-02-18 | Malfunction | STAAR SURGICAL COMPANY | 2015-04-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND THE LENS (TRAILING HAPTIC) TORE/BROKE. THERE WAS PATIENT CONTACT AND NO REPORTED INJURY. Manufacturer Narrative: (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVALUATION: RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, THE MOST PROBABLE ROOT CAUSE OF THE EVENT IS USER OR TECHNICAL ERROR. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5040396&pc=MTA | 2023826-2015-01110 | 2015-07-08 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE DURING INJECTION. THERE WAS NO REPORTED INJURY OR COMPLICATIONS. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY WITH THE RESULTS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613015&pc=MTA | 2023826-2016-00527 | 2016-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+3.0/085 DIOPTER, AND THE LENS TORE/BROKE. THE LENS WAS REMOVED AND THERE WAS NO REPORT OF ANY PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE OPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7704702&pc=MTA | 2023826-2018-01050 | 2018-01-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Contamination | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, WITHIN THE SAME SURGERY, DUE TO FOREIGN BODY ADHESION (BLACK DOT). THE FOREIGN BODY WAS NOTED AFTER LENS INSERTION. THERE WAS NO REPORT OF ANY PATIENT INJURY. ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172147&pc=MTA | 2023826-2018-01845 | 2018-09-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Contamination /Decontamination Problem | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS FOREIGN BODY ADHESION WHICH WAS FOUND DURING IMPLANT. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. THE SURGEON STATES HE BELIEVES THE "FOREIGN BODY WAS ORIGINALLY ADHESIVE ON THE LENS." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552496&pc=MTA | 2023826-2019-00663 | 2019-03-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Contamination /Decontamination Problem | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER FROM THE VIAL, USING THE MICROSCOPE, HE DISCOVERED A BLACK FOREIGN BODY ON THE LENS. THIS OCCURRED ON (B)(6) 2019. THE SURGEON DID NOT USE THE LENS AND CLOSED THE VIAL CAP. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND A BLACK PARTICLE ON THE LENS. METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. METHOD CODE 4112-THE PARTICULATE WAS ISOLATED AND TESTED (B)(6) 2019. FOURIER TRANSFORM INFRARED SPECTROSCOPY (FTIR) RESULTS FOUND THE SPECTRUM TO BE MASKED BY THE PRESENCE OF ONE OR MORE INORGANIC COMPONENTS WHICH PREVENTED A DEFINITIVE DEFINITION. THE BEST SPECTRAL MATCHES INCLUDED A VARIETY OF EPOXY RESINS AND A WEAKER MATCH WAS MADE WITH CHLORINATED POLYETHYLENE. THE PARTICULATE WAS FOUND TO BE APPROXIMATELY 288UM LONG AND 222UM WIDE. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372306&pc=MTA | 2023826-2019-00313 | 2019-01-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Contamination; Difficult to Fold, Unfold or Collapse | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) AND NOTED FOREIGN MATERIAL ON THE LENS. THE LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) AND HAD AN ISSUE GETTING THE LENS (ICL) TO UNFOLD. AFTER TRYING A FEW TECHNIQUES HE FELT IT WAS BEST TO REMOVE THE LENS WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED SUCCESSFULLY. DATE OF REPORT: 25-JAN-2019. DATE RECEIVED BY MANUFACTURER: 25-JAN-2019. DEVICE CODE - 1254. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6109410&pc=MTA | 2023826-2016-01633 | 2016-07-08 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Crack | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS NOT IMPLANTED DUE TO A LENS CRACK BEFORE INJECTION INTO THE EYE. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL AND SIZE; THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT/DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DRIED DISCOLORED MATERIAL). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6280828&pc=MTA | 2023826-2017-00163 | 2016-10-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-26 | MTA | INTRAOCULAR LENS | Crack; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, HAD CRACKED JUST WHEN THE LENS WAS COMING OUT FROM THE INJECTOR. THIS OCCURRED ON (B)(6) 2016, BEFORE INJECTING INTO THE LEFT EYE (OS). THE LENS WAS NOT IMPLANTED AND THERE IS NO REPORTED PATIENT INJURY. THIS LENS HAS BEEN RETURNED AND EVALUATED. Manufacturer Narrative: (B)(4). NO REPORTED MATERIAL INTEGRITY ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATION - PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN WITH DRIED, DISCOLORED MATERIAL ON THE LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191976&pc=MTA | 2023826-2016-01783 | 2016-11-21 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Delivered as Unsterile Product | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON NOTED A SUBSTANCE/RESIDUE ON A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, AND THE LENS WAS NOT USED. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER?: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS RE-EXAMINED UNDER 20X MAGNIFICATION AND NO DEFECTS WERE DETECTED ON THE LENS. IF THE SUBSTANCE IS NOW IN SOLUTION, WE CANNOT ISOLATE IT BECAUSE IT IS NOT CLEAR WHAT THE PARTICULATE IS OR HOW IT LOOKS LIKE. METHOD: WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6312918&pc=MTA | 2023826-2017-00147 | 2016-12-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contamination with Chemical or Other Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON DID NOT IMPLANT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS. THE LENS WAS NOT IMPLANTED DUE TO IT HAVING " BLACK WASTE." THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (BLACK DOTS). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: METHOD CODE: (B)(4). PARTICULATE EVALUATION: FOREIGN MATERIAL EVALUATED AND FOUND NO PROBABLE SOURCE IN MANUFACTURING PROCESS OR ENVIRONMENT. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7188510&pc=MTA | 2023826-2018-00067 | 2017-12-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON NOTICED MICROSCOPIC BLACK FLAKES ON A MICL13.2, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS WHEN THEY OPENED THE LENS VIAL AND WERE PREPARING TO LOAD THE LENS IN THE CARTRIDGE FOR THE INJECTION ON (B)(6) 2017. THE SURGEON DESCRIBED OPENING THE LENS VIAL WITH HEMOSTATS, REMOVING THE METAL RING, AND REMOVING THE GREY STOPPER WITH A STERILE GLOVED HAND. THE LENS WAS POURED OUT OF THE VIAL ONTO THE INJECTOR CARTRIDGE TRAY WITH THE ASSISTANCE OF THE INJECTOR PLUNGER TO GUIDE THE LENS OUT OF THE VIAL. WHEN THE LENS WAS PICKED UP WITH ICL LOADING FORCEPS, THE BLACK FLAKES WERE NOTICED. THERE WAS NO PATIENT CONTACT. THEY USED A NEW CARTRIDGE/INJECTOR AND A NEW ICL WAS LOADED AND USED TO PERFORM THE SURGERY. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SPECIMEN CUP AND AN EMPTY CARTRIDGE IN THE CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7188506&pc=MTA | 2023826-2018-00065 | 2017-12-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON NOTICED MICROSCOPIC BLACK FLAKES ON A MICL13.2, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS WHEN THEY OPENED THE LENS VIAL AND WERE PREPARING TO LOAD THE LENS IN THE CARTRIDGE FOR THE INJECTION ON (B)(6) 2017. THE SURGEON DESCRIBED OPENING THE LENS VIAL WITH HEMOSTATS, REMOVING THE METAL RING, AND REMOVING THE GREY STOPPER WITH A STERILE GLOVED HAND. THE LENS WAS POURED OUT OF THE VIAL ONTO THE INJECTOR CARTRIDGE TRAY WITH THE ASSISTANCE OF A WEK-CELL TO GUIDE THE LENS OUT OF THE VIAL. WHEN THE LENS WAS PICKED UP WITH ICL LOADING FORCEPS, THE BLACK FLAKES WERE NOTICED. THERE WAS NO PATIENT CONTACT. THEY USED A NEW CARTRIDGE/INJECTOR AND A NEW ICL WAS LOADED AND USED TO PERFORM THE SURGERY. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS/PARTICLES IN LENS VIAL. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267545&pc=MTA | 2023826-2018-00216 | 2017-12-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. AFTER IMPLANT THE SURGEON STATES HE FOUND A "TINY BLACK-COLORED FOREIGN SUBSTANCE ON THE LENS SURFACE." REPORTEDLY, THE SURGEON ATTEMPTED TO REMOVE THE SUBSTANCE BUT WAS UNSUCCESSFUL. THE SURGEON REMOVED THE LENS INTRA-OPERATIVELY. ON THE SAME DATE, IN A SECOND SURGERY, AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR RESIDUE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7429500&pc=MTA | 2023826-2018-00549 | 2018-03-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. AFTER THE LENS WAS IMPLANTED THE SURGEON FOUND THAT A WHITE PARTICLE WAS FLOATING IN THE SOLUTION OF THE VIAL. SURGEON IS CONCERN THAT THE FOREIGN BODY MIGHT CAUSE ANY COMPLICATION BECAUSE THE LENS IS IN THE EYE. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: VIAL WAS RETURNED WITH LIQUID. VISUAL INSPECTION FOUND A PARTICULATE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456989&pc=MTA | 2023826-2018-00597 | 2018-03-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. UPON UNFOLDING OF THE LENS, DEPOSITS AND GLITTERING/OILY RESIDUE WAS SEEN ON THE OPTIC. IT WAS FLUSHED WITH BSS, BUT ONLY SOME OF THE RESIDUE CAME OFF AND THE LENS REMAINED PARTIALLY COVERED. THE LENS WAS INTRAOPERATIVELY EXCHANGED WITH ANOTHER LENS. UPON POST-OP EXAMINATION, THE RESIDUE IS STILL VISBLE ON THE PERIPHERAL OPTIC, BUT NOT VISUALLY SIGNIFICANT FOR THE PATIENT. Manufacturer Narrative: THE PATIENT STILL HAS THE TRACES ON THE LENS INSIDE THE EYE BUT NO VISUAL IMPACTS AND NO INFLAMMATION. THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO OILY/GLITTERING RESIDUE NOTED ON THE LENS. THE DELIVERY SYSTEM WAS RETURNED LOOSE IN LENS BOX. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON/IN DEVICE (RESIDUE ON DEVICE). TYPE OF REPORTABLE EVENT IN INITIAL MDR SHOULD BE MALFUNCTION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: NO OILY / GLITTERING RESIDUE ON THE PRODUCT. CORRECTED DATA: DEVICE EVALUATION: THE DELIVERY SYSTEM WAS RETURNED LOOSE IN A LENS BOX. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE AND FOREIGN MATERIAL ON / IN DEVICE (RESIDUE ON DEVICE). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7208090&pc=MTA | 2023826-2018-00133 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON NOTICED A MICROSCOPIC BLACK SPECK ON A MICL13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS. THE LENS WAS POURED OUT OF THE VIAL INTO THE PACKAGING OF THE CARTRIDGE AND THAT'S WHEN THE SPECK WAS NOTICED. IT WAS REPORTED THAT THE ICL HAD TOUCHED NOTHING BUT THE PLUNGER WHICH WAS DIRECTLY OPENED AND USED TO SLIDE THE ICL OUT. THE SPECK WAS LOCATED ON THE SURFACE OF THE LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS INSPECTED UNDER 20X ON THE MICROSCOPE AND NO BLACK SPEC WAS NOTED. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075957&pc=MTA | 2023826-2020-01108 | 2020-04-19 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE PATIENT REPORTED " HAD LENS IMPLANTS FOR BOTH EYES OVER 10 YEARS AGO. ON SUNDAY, THE RIGHT LENS HAS BECOME DETACHED FROM MY IRIS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8208780&pc=MTA | 2023826-2019-00003 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | UNK | 1 | Event Description: THE REPORTER INDICATED THAT THE PATIENT PRESENTED TO THE SURGEON STATING THAT A TORIC ICL HAD BEEN IMPLANTED AT ANOTHER FACILITY AND HAD ROTATED. NO ADDITIONAL INFORMATION IS ABLE TO BE OBTAINED. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8048832&pc=MTA | 2023826-2018-01600 | 2018-10-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -09.00/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THAT THE INSIDE PAITENT LABEL HAD AN INCORRECT SERIAL NUMBER, THE LENS WAS IMPLANTED, THE AXIS IS ALIGNED ACCORDING TO THE LABEL ON THE BOX NOT THE INSIDE LABEL. THE PATIENT OUT COME IS GOOD. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122178&pc=MTA | 2023826-2018-01667 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO 12.6 -9.50/+1.50/102, SERIAL (B)(4), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). IT WAS REPORTED THAT THE PATIENT CHART LABEL PROVIDED INSIDE THE PACKAGING BOX SHOW DISCREPANT LENS INFORMATION AS TO VTICMO 12.6 -9.50/+1.50/100, SERIAL # (B)(4). THE REPORTER STATES "THE DOCTOR FOUND THE LENS POWER TO BE SIMILAR TO WHAT HE ORDERED, SO HE IMPLANTED ACCORDINGLY". THE LENS REMAINS IMPLANTED AND IT WAS REPORTED THAT THERE IS NO NEGATIVE FEEDBACK OR COMPLAINTS ABOUT THE PATIENT'S CONDITION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122361&pc=MTA | 2023826-2018-01668 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO 12.6 -12.00/+1.50/083, SERIAL (B)(4), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). IT WAS REPORTED THAT THE PATIENT CHART LABEL PROVIDED INSIDE THE PACKAGING BOX SHOW DISCREPANT LENS INFORMATION AS TO VTICMO 12.6 -12.00/+1.5/082, SERIAL # (B)(4). THE REPORTER STATES "THE DOCTOR FOUND THE LENS POWER TO BE SIMILAR TO WHAT HE ORDERED, SO HE IMPLANTED ACCORDINGLY". THE LENS REMAINS IMPLANTED AND IT WAS REPORTED THAT THERE IS NO NEGATIVE FEEDBACK OR COMPLAINTS ABOUT THE PATIENT'S CONDITION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122438&pc=MTA | 2023826-2018-01669 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO 12.6 -9.50/+1.50/100, SERIAL (B)(4), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). THE REPORTER STATES THAT THE SAME MODEL, POWER AND SERIAL NUMBER OF THIS LENS (VTICMO 12.6 -9.50/+1.50/100, SERIAL # (B)(4)) WAS SHOWN ON THE PATIENT CHART LABEL INSIDE THE BOX FOR ANOTHER LENS WITH VTICMO 12.6 -9.50/+1.50/102, SERIAL # (B)(4) INDICATED ON THE EXTERNAL BOX LABEL. THE LENS REMAINS IMPLANTED AND IT WAS REPORTED THAT THERE IS NO NEGATIVE FEEDBACK OR COMPLAINTS ABOUT THE PATIENT'S CONDITION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8566359&pc=MTA | 2023826-2019-00715 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-04-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. IT WAS REPORTED THAT THE INSIDE PATIENT LABEL HAD AN INCORRECT SERIAL NUMBER, THE LENS WAS IMPLANTED, THE AXIS IS ALIGNED ACCORDING TO THE LABEL ON THE BOX NOT THE INSIDE LABEL. "WE WERE NOT INFORMED ANY NEGATIVE INFORMATION FROM DOCTOR AND PATIENT." LENS REMAINS IMPLANTED, DOCTOR RECOMMENDED NOT TO EXPLANT/EXCHANGE THE LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6812227&pc=MTA | 2023826-2017-01342 | 2014-09-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2014, AS THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -18.0/+2.0/127 INTO THE PATIENT'S EYE, HE DISCOVERED THE LENS BROKE. THE SURGEON STATED THAT THE HAPTICS OF THE LENS WERE BROKEN ON ONE SIDE WHEN THE LENS CAME OUT OF THE "SHOOTER." THE LENS WAS NOT IMPLANTED AND A NEW LENS WAS ORDERED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6509194&pc=MTA | 2023826-2017-00605 | 2015-04-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON PREPARED TO IMPLANT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+3.5/083 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS), IT BECAME "UNSTABLE." A DIFFERENT LENS OF THE SAME TYPE AND SIZE AND SIMILAR DIOPTER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. IT IS UNKNOWN IF THERE WAS PATIENT CONTACT; TO DATE, ATTEMPTS AT CUSTOMER CONTACT HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN THE LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODES ADDED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [ACD < 3.0MM (2.9MM)]. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6419639&pc=MTA | 2023826-2017-00447 | 2015-06-09 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON NOTED A DEFECT ON A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, AND DECIDED NOT TO USE THE LENS. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. THE REPORTER INDICATED THE LENS WAS TORN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6112598&pc=MTA | 2023826-2016-01664 | 2015-08-17 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6500531&pc=MTA | 2023826-2017-00657 | 2015-10-31 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON TRIED TO IMPLANT A 13.2MM VTICMO13.2 -12.5/1.5/082 DIOPTER IMPLANTABLE COLLAMER LENS. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. LENS WAS NOT IMPLANTED. NO PATIENT INJURY REPORTED. Manufacturer Narrative: THIS IS A PRODUCT PROBLEM NOT AN ADVERSE EVENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)46)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6597937&pc=MTA | 2023826-2017-00813 | 2016-01-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT SURGEON ATTEMPTED AN IMPLANT OF A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). LENS TEAR/BREAK BEFORE INJECTION INTO THE EYE WAS NOTED. LENS WAS NOT IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND FOREIGN MATERIAL ON LENS SURFACE. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6522664&pc=MTA | 2023826-2017-00716 | 2016-03-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE SURGEON ATTEMPTED TO INSERT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS WITH A -10.50/+4.5/075 DIOPTER IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE BEFORE INJECTION IN THE EYE. LENS WAS NOT IMPLANTED. NO PATIENT CONTACT. Manufacturer Narrative: MANUFACTURED IN THE U.S., BUT NOT MARKETED. PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6327785&pc=MTA | 2023826-2017-00237 | 2016-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12/+1/84 DIOPTER, WAS TORN BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. A BACK-UP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. DURING THE PATIENT'S LAST VISIT ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6330729&pc=MTA | 2023826-2017-00239 | 2016-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6/2/084 DIOPTER, TORE/BROKE BEFORE INJECTION/DELIVERY IN THE PATIENT'S LEFT EYE (OS). THE SURGEON STATED "EVERYTHING IS FINE, PATIENT IS HAPPY." ON (B)(6) 2016, PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6592865&pc=MTA | 2023826-2017-00824 | 2016-05-01 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON TRIED TO IMPLANT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (NOT SPECIFIED). METHOD: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6566822&pc=MTA | 2023826-2017-00820 | 2016-05-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.5/2.0/84 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6605330&pc=MTA | 2023826-2017-00918 | 2016-06-09 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/3.5/89 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS TORE/BROKE DURING LOADING INTO THE CARTRIDGE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, THE HAPTIC TORN/BROKEN/BENT/DEFORMED, SCRATCHES ON THE LENS SURFACE, AND A BLUE SPOT. (B)(4), VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN PATIENTS UNDER THE AGE OF 21 YEARS OLD. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6125978&pc=MTA | 2023826-2016-01674 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/84 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND THE LENS WAS NOTED TO BE TORN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS THE SAME MODEL BUT A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: EVENT WAS INCORRECTLY CLASSIFIED AS A PRODUCT PROBLEM. EVENT WAS INCORRECTLY CLASSIFIED AS A MALFUNCTION. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6131781&pc=MTA | 2023826-2016-01688 | 2016-07-25 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS TEAR DURING DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED FOR AN IDENTICAL MODEL AND DIOPTER. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATIONS - LENS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6214639&pc=MTA | 2023826-2016-01822 | 2016-08-02 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, HAD TORN DURING THE ICL LOADING PROCESS. THIS OCCURRED ON (B)(6) 2016, AS THE SURGEON WAS PREPARING TO IMPLANT THE LENS INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS NOT IMPLANTED AND ANOTHER LENS WAS USED. Manufacturer Narrative: THE PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/22. AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH(ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.95MM AT THE TIME OF IMPLANTATION. THE CORRECT LENS DIOPTER IS -12.0/+3.0/88. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE HAPTIC TORN. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6131764&pc=MTA | 2023826-2016-01689 | 2016-08-29 | Injury | STAAR SURGICAL COMPANY | 2016-11-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS TEAR DURING DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED FOR AN IDENTICAL MODEL AND DIOPTER. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN AND OPTIC TORN/SPLIT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365375&pc=MTA | 2023826-2017-00284 | 2016-08-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50/2.5/90 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS BROKE DURING INJECTION AND WAS EXCHANGED FOR THE SAME MODEL AND DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS TORN INTO TWO PIECES, TORN THROUGH THE HAPTICS AND OPTIC . A PIECE OF ONE HAPTIC WAS TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6590441&pc=MTA | 2023826-2017-00821 | 2016-08-30 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.5/2.0/103 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE (BEFORE CORNEAL INCISION). THE SURGERY WAS DEFERRED, NO DAMAGE ON PATIENT SIDE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN/BENT/DEFORMED, AND FOREIGN MATERIAL ON LENS SURFACE (NOT SPECIFIED). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6260240&pc=MTA | 2023826-2017-00102 | 2016-09-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.50/+3.0/082 DIOPTER, TORE/BROKE BEFORE INJECTION INTO THE EYE. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6111694&pc=MTA | 2023826-2016-01657 | 2016-09-20 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VICM5 13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6219312&pc=MTA | 2023826-2017-00008 | 2016-09-29 | Injury | STAAR SURGICAL COMPANY | 2017-01-03 | MTA | INTRAOCULAR LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS HAD TORN BEFORE INJECTION INTO THE EYE. THIS OCCURRED ON (B)(6) 2016, AS THE SURGEON WAS PREPARING TO IMPLANT THE LENS INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS NOT IMPLANTED. AS OF THE DATE OF MDR REPORTING, THE LENS HAS NOT BEEN RETURNED AND THEREFORE NO PRODUCT EVALUATION HAS OCCURRED. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. A WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. THIS CASE IS AN ADVERSE EVENT, NOT A PRODUCT MALFUNCTION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6111674&pc=MTA | 2023826-2016-01661 | 2016-09-30 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, TORE BEFORE INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT AND NO APPARENT INJURY. ANOTHER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE EVENT WAS CAUSED BY USE ERROR. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FIBERS ON THE LENS SURFACE. REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6111691&pc=MTA | 2023826-2016-01656 | 2016-10-04 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE PLUNGER WAS DAMAGED DURING LOADING OF A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, AND THE LENS WAS TORN. THE LENS WAS NOT IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION: (B)(6) 2016. ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021366&pc=MTA | 2023826-2017-01675 | 2016-10-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VICMO13.2, -17.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE BEFORE INJECTION INTO THE EYE. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2017, ANOTHER LENS WAS IMPLANTED AND IT RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND THE HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6176017&pc=MTA | 2023826-2016-01768 | 2016-10-18 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE MANUFACTURER RECEIVED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, FROM THE CUSTOMER DAMAGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DESCRIBE EVENT OR PROBLEM: ADDITIONAL INFORMATION RECEIVED - THE REPORTER STATED THE LENS WAS DAMAGED DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO A LOADING/TECH ERROR AND THERE WERE NO PROBLEM WITH THE LENS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6197129&pc=MTA | 2023826-2016-01810 | 2016-10-21 | Injury | STAAR SURGICAL COMPANY | 2016-12-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND IT WAS NOTED THAT THE LENS WAS TORN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LENS THE SAME SIZE, MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CORRECTED DATA: PREVIOUS REPORT AS "PRODUCT PROBLEM" IS INCORRECT. THE CORRECT REPORT IS "ADVERSE EVENT". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140275&pc=MTA | 2023826-2016-01709 | 2016-10-24 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, TORE DURING LOADING. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS USED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS INSERTED AND WAS NOTED TO BE TORN AFTER INSERTION. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND REPLACED WITH ANOTHER SAME MODEL/DIOPTER LENS. THE POST-OP BSCVA WAS 20/30-1, WITH PHOTOPHOBIA. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF LENS OPTIC AND HAPTICS TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF REDDISH COLOR RESIDUE ON THE LENS SURFACE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6298440&pc=MTA | 2023826-2017-00214 | 2016-10-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS ATTEMPTING TO INSERT A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER ,AND NOTICED THE LENS WAS TORN. THERE WAS NO PATIENT CONTACT OR INJURY AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6111700&pc=MTA | 2023826-2016-01652 | 2016-10-27 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, AND THE SURGEON NOTED THE LENS HAD TORN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY, NO ENLARGED INCISION OR SUTURES. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULT CODE(S): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6179654&pc=MTA | 2023826-2016-01772 | 2016-11-03 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, TORE WHILE LOADING AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS USED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULT CODE(S): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN AND SCRATCHED. A PIECE OF ONE HAPTIC WAS TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172504&pc=MTA | 2023826-2016-01762 | 2016-11-12 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, WAS BEING PULLED OUT OF THE INJECTOR AND RIPPED. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER?: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED IN LIQUID. METHOD: LENS WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6179664&pc=MTA | 2023826-2016-01778 | 2016-11-14 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, AND IT WAS NOTED THE LENS WAS TORN. THE LENS WAS REMOVED WITH NO PATIENT INJURY. THE BACKUP LENS WAS USED. THE REPORTER INDICATED THE LENS WAS TORN DURING LOADING. Manufacturer Narrative: RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF THE HAPTIC TORN OFF AND UNKNOWN RESIDUE ON THE LENS SURFACE. THE LENS WAS RETURNED DRY. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227521&pc=MTA | 2023826-2017-00025 | 2016-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THERE WAS AN ISSUE WITH THE LENS IN THE INJECTOR AND THE LENS WAS REMOVED. THERE WAS NO PATIENT INJURY AND THE LENS TORE UPON REMOVAL. THE BACKUP LENS WAS IMPLANTED. THE PATIENT IS DOING WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION: THE REPORTER INDICATED A 13.2MM VICMO13.2 IMPLANTABLE LENS, -12.50 DIOPTER, TORE BEFORE INJECTION INTO THE EYE AND WAS NOT IMPLANTED. THE BACKUP LENS WAS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC MISSING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6214842&pc=MTA | 2023826-2016-01846 | 2016-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS ATTEMPTING TO IMPLANT A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12/1.5/103 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) AND IT WAS NOTED THAT THE LENS WAS TORN. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6246925&pc=MTA | 2023826-2017-00062 | 2016-11-21 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-12 | MTA | INTRAOCULAR LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/080 DIOPTER, HAD TORN/BROKEN BEFORE INJECTION INTO THE LEFT (OS) EYE. THIS OCCURRED ON (B)(6) 2016. THE LENS WAS NOT IMPLANTED. AS OF THE DATE OF MDR SUBMISSION, THIS LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/080 DIOPTER, HAD TORN/BROKEN BEFORE INJECTION INTO THE LEFT (OS) EYE. THIS OCCURRED ON (B)(6) 2016. THERE WAS NO PATIENT CONTACT AND THE LENS WAS NOT IMPLANTED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: DATE RECEIVED BY MANUFACTURER ON INITIAL REPORT SHOULD BE 12/02/2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6432865&pc=MTA | 2023826-2017-00491 | 2016-11-22 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/078 DIOPTER, RUPTURED DURING THE LOADING PROCESS AND BECAME STUCK IN THE CARTRIDGE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PARTIAL PIECE OF LENS RETURNED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6234309&pc=MTA | 2023826-2017-00038 | 2016-11-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/080 DIOPTER, CRACKED DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: THE EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BENT. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6211551&pc=MTA | 2023826-2016-01845 | 2016-11-25 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND IT WAS NOTED THAT THE LENS WAS TORN DURING IMPLANTATION. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND THERE WAS NO REPORT OF ANY PATIENT INJURY. Manufacturer Narrative: CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND OPTIC TORN, AND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227519&pc=MTA | 2023826-2017-00026 | 2016-12-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON HAD TROUBLE LOADING A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, AND THE LENS WAS DAMAGED. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS IN THE PROCESS OF BEING PULLED FORWARD INTO THE CARTRIDGE TIP WHEN A PIECE OF THE ICL WAS TORN OFF BETWEEN THE LEADING FOOTPLATES BY THE FORCEPS. THE ICL, CARTRIDGE AND FORCEPS WERE REPLACED AND THE SURGERY WAS COMPLETED UNEVENTFULLY. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447914&pc=MTA | 2023826-2017-00540 | 2016-12-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6/+4.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2016 THE LENS WAS TORN BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BE RETURNED YET. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS AND A CARTRIDGE WERE RETURNED IN CARTRIDGE CASE INSIDE A BAG. THE LENS WAS RETURNED DRY AND VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS AND RESIDUE ON LENS SURFACE. THE CARTRIDGE WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE ON PRODUCT WITH NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6281310&pc=MTA | 2023826-2017-00178 | 2016-12-14 | Injury | STAAR SURGICAL COMPANY | 2017-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6)2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL AND SIZE, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/20. THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE LENS HAPTIC BENT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #(B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6265501&pc=MTA | 2023826-2017-00094 | 2016-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE HAPTIC OF A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, WAS TORN WHILE BEING LOADED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED STUCK IN CARTRIDGE. VISUAL INSPECTION OF THE CARTRIDGE FOUND NO VISIBLE DAMAGE TO THE DEVICE, DRY RESIDUE ON CARTRIDGE, AND THE LENS STUCK INSIDE THE DEVICE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6370469&pc=MTA | 2023826-2017-00302 | 2016-12-22 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. AS OF 03/01/2017, THE NEW LENS HAS YET TO BE IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A CENTRIFUGE VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DRY CLEAR RESIDUE). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS FOUND CRACKED BEFORE THE SURGERY BEGAN AND IT WAS REPORTED THAT THERE WAS NO PATIENT CONTACT. THE PATIENT WAS FRUSTRATED FOR WAITING OF BACK UP LENS BEING PREPARED AND LEFT THE FACILITY WITHOUT HAVING THE SURGERY. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365317&pc=MTA | 2023826-2017-00277 | 2016-12-26 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8 DIOPTER, TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A NEW LENS WITH THE SAME MODEL AND SIMILAR DIOPTER ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: LENS RETURNED DRY IN A SMALL CENTRIFUGE VIAL WITH CLEAR, STICKY SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED, CLEAR AND DARK RESIDUE ON LENS, AND MISSING PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6408307&pc=MTA | 2023826-2017-00415 | 2016-12-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON LOADED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE CARTRIDGE BUT DID NOT LIKE HOW THE LENS WAS LOADED, HE FELT THERE WAS SOMETHING WRONG WITH THE LENS. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER WAS UNABLE TO PROVIDE ANY ADDITIONAL INFORMATION, SO THE CAUSE IS UNKNOWN. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6378771&pc=MTA | 2023826-2017-00323 | 2017-01-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON LOADED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, BUT DID NOT LIKE HOW THE LENS WAS SITTING IN THE CARTRIDGE, SO DECIDED NOT TO USE IT. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS LOADED AND IMPLANTED WITH NO PROBLEM, USING THE SAME CARTRIDGE. THE REPORTER STATED THE SURGEON FELT THERE WAS SOMETHING WRONG WITH THE LENS, BUT WAS UNABLE TO IDENTIFY THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6281925&pc=MTA | 2023826-2017-00180 | 2017-01-04 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, AND THE LENS TORE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY, NO ENLARGED INCISION OR SUTURES. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THAT ONLY HALF OF THE LENS WAS RETURNED, THE OTHER HALF WAS MISSING. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6296434&pc=MTA | 2023826-2017-00207 | 2017-01-04 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE HAPTIC OF A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, WAS TORN WHILE BEING LOADED AND THERE WAS NO PATIENT CONTACT. THE BACK UP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PIECES OF THE HAPTIC MISSING. WORK ORDER SEARCH : NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6426153&pc=MTA | 2023826-2017-00459 | 2017-01-10 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE CUSTOMER RETURNED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, DAMAGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE REPORTER STATED THE SURGEON ATTEMPTED TO INSERT THE LENS BUT THE LENS TORE. THERE WAS PATIENT CONTACT BUT NO INJURY. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531864&pc=MTA | 2023826-2017-00485 | 2017-01-10 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON THE SAME SURGERY THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THE REPORTER STATED THAT HE THOUGHT "BAD LOADING" WAS THE CAUSE OF THE TEAR. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: CONCLUSION: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6618999&pc=MTA | 2023826-2017-00926 | 2017-01-21 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.1MM VICMO_12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. THE LENS WAS TORN/BROKEN BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED. THE REPORTER SAID THE LENS TORE DURING SETTING PROCESS, AND NO PATIENT CONTACT. ON (B)(6)2017 THE ALTERNATIVE LENS WITH SAME SIZE AND DIOPTER WAS IMPLANTED, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN SMALL VIAL. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC, DEBRIS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6549353&pc=MTA | 2023826-2017-00774 | 2017-01-23 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD), IT TORE/BROKE. THE LENS WAS IMPLANTED. THIS OCCURRED ON (B)(6) 2017. THE SURGEON STATES THERE WAS NOT ENOUGH VISCOELASTIC USED DURING SETUP AND SO THE LENS TORE. AN ALTERNATE LENS WAS IMPLANTED DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: PREVIOUS STATED: NOT ENOUGH VISCOELASTIC WAS USED DURING SET-UP AND SO THE LENS TORE. CORRECTED DATA: NOT ENOUGH BSS WAS USED DURING SET-UP AND SO THE LENS TORE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT A PIECE OF HAPTIC WAS TORN AND FOREIGN MATERIAL (CLEAR RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6341611&pc=MTA | 2023826-2017-00242 | 2017-01-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, WAS TORN DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE EVENT OCCURRED WHEN THE LENS WAS BEING PULLED INTO THE CARTRIDGE TIP BY THE POSITIONING FORCEPS. THE REPORTER WAS UNABLE TO PROVIDE ANY ADDITIONAL INFORMATION, SO THE CAUSE IS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER; VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF DRY CLEAR RESIDUE ON THE OPTIC. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6457271&pc=MTA | 2023826-2017-00574 | 2017-01-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON ATTEMPTED TO INSERT A MICL13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS BUT THE LENS WAS NOT IMPLANTED DUE TO THE PATIENT'S PUPIL CONSTRICTED TO 2MM BEFORE NASAL ICL FOOTPLATES WERE PLACED. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS DISCARDED. Manufacturer Narrative: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CONCLUSION: PATIENT'S PUPIL CONSTRICTED TO 2MM BEFORE NASAL ICL FOOTPLATES WERE PLACED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6619016&pc=MTA | 2023826-2017-00942 | 2017-01-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -2.0/+6/86 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS GOT TORN BEFORE INJECTION INTO THE EYE, LENS WAS NOT IMPLANTED. ON THE SAME DAY, THE BACK-UP LENS WITH LONGER SIZE 13.2MM WAS USED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS RETURNED IN LENS VIAL WITH LIQUID. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC. NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6460501&pc=MTA | 2023826-2017-00559 | 2017-01-26 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME SURGERY, THE LENS WAS EXCHANGED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A LARGER SIZE AND SAME DIOPTER LENS DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THE PROBLEM WAS RESOLVED. LENS WAS RETURNED DRY, IN SMALL VIAL. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC, CLEAR AND BROWN SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6474269&pc=MTA | 2023826-2017-00611 | 2017-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS TORN/BROKEN BEFORE INJECTION INTO THE EYE, AND WAS NOT IMPLANTED. THE BACK-UP LENS WAS USED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH WAS PERFORMED, NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6430483&pc=MTA | 2023826-2017-00480 | 2017-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: A 13.2MM MICL IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, WAS RETURNED WITH A PIECE OF HAPTIC MISSING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING, IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A HAPTIC PIECE MISSING. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6444902&pc=MTA | 2023826-2017-00519 | 2017-02-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS HAVING PROBLEMS LOADING A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER. THE SURGEON ATTEMPTED SEVERAL TIMES BUT THE LENS WOULD NOT SIT WELL IN THE CARTRIDGE, AND THE LENS WAS DAMAGED. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS LOADED AND IMPLANTED WITH NO PROBLEM. THE REPORTER INDICATED THE CAUSE WAS UNKNOWN BUT THERE WAS NO PROBLEM WITH THE LENS OR INJECTOR SYSTEM. Manufacturer Narrative: DEVICE EVALUATION - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF HAPTIC TORN OFF AND MISSING. THE LENS OPTIC WAS SCRATCHED. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6423184&pc=MTA | 2023826-2017-00458 | 2017-02-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, WAS NOT USED DUE TO A LOADING ERROR. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED IN 2 PIECES. THE LENS WAS RETURNED IN LIQUID, WITH LIGHT RESIDUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6483121&pc=MTA | 2023826-2017-00603 | 2017-02-13 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2017 THE LENS WAS TORN/BROKEN BEFORE INJECTION INTO THE EYE, AND WAS NOT IMPLANTED. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY, IN LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN, AND MISSING PIECE. LENS, CARTRIDGE, AND FOAM TIP PLUNGER WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6483420&pc=MTA | 2023826-2017-00622 | 2017-02-13 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-11 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. DURING INJECTION/DELIVERY, THE LENS TORE/BROKE AND IT WAS REMOVED FROM THE EYE ON AND EXCHANGED FOR A LENS OF THE SAME TYPE, SIZE, AND DIOPTER. SURGEON STATED THAT LENS INJECTION WAS A LITTLE MORE DIFFICULT THAN BEFORE. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND MISSING A PIECE OF HAPTIC. THE LENS WAS TORN IN HALF. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. OFF-LABEL USE (ACD BELOW 3.0MM AT THE TIME OF IMPLANT). (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6503628&pc=MTA | 2023826-2017-00652 | 2017-02-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SAME SIZE AND DIOPTER LENS DUE LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THERE WAS A RESISTANCE DURING INJECTION THE LENS INTO THE EYE, AND THE LENS GOT TORN BECAUSE IT WAS NOT ALIGNED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN TWO PIECES, AND IN SMALL VIAL. VISUAL INSPECTION FOUND HAPTIC TORN OFF, DEBRIS AND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE.(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS, INJECTION SYSTEM WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6465439&pc=MTA | 2023826-2017-00589 | 2017-03-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, WAS RETURNED DAMAGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY. VISUAL INSPECTION FOUND BROWNISH AND CLEAR DRY RESIDUE ON LENS AND A PIECE OF HAPTIC MISSING. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6565208&pc=MTA | 2023826-2017-00793 | 2017-03-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO REPORTED INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO BAD LOADING OF THE LENS. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6473610&pc=MTA | 2023826-2017-00610 | 2017-03-09 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, AND THE LENS TORE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER INDICATED THE CAUSE WAS UNKNOWN. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6486098&pc=MTA | 2023826-2017-00614 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATES THERE WAS A LITTEL DRAG DURING INJECTION INTO THE EYE. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DURING THE SAME SURGERY, THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THERE WAS A LITTLE DRAG DURING INJECTION INTO THE EYE. PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES ON (B)(6) 2017 WERE 20/20. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS - CLEAR RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6573265&pc=MTA | 2023826-2017-00819 | 2017-03-10 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR THE SAME DIOPTER BUT LARGER SIZE DUE THE LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THE CUSTOMER STATED ACCORDING TO THE SURGEON, THE LENS WAS IMPLANTED UPSIDE DOWN, AND THEN WAS EXPLANTED. DURING RE-LOADING THE LENS, THE SURGEON SAW THE HAPTIC WAS TORN, SO THE BACK-UP LENS USED. THE IS NO ISSUE ON SIZING FOR THE PATIENT, THE CUSTOMER ALSO MENTIONED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN SMALL VIAL. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. LENS, AND INJECTOR WORK ORDER SEARCHES: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6618993&pc=MTA | 2023826-2017-00927 | 2017-03-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICMO_13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS TORN/BROKEN BEFORE INJECTION INTO THE EYE, AND WAS NOT IMPLANTED. THE REPROTER SAID THE LENS TORE DURING SETTING PROCESS, AND NO PATIENT CONTACT. ON (B)(6) 2017 THE ALTERNATIVE LENS WAS IMPLANTED, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN SMALL VIAL. VISUAL INSPECTION FOUND HAPTIC TORN, CLEAR RESIDUE ON LENS SURFACE. OFF LABEL USE (OVER 45 YEARS OF AGE AT THE TIME OF IMPLANT) . CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6619902&pc=MTA | 2023826-2017-00922 | 2017-03-30 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/030 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD), IT TORE WHILE LOADING. THIS OCCURRED ON (B)(6) 2017. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, WITH DEBRIS AND CLEAR AND WHITE RESIDUE COVERING THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. VIAL WAS SEALED WITH TAPE INSTEAD OF A STOPPER. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6557223&pc=MTA | 2023826-2017-00779 | 2017-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/095 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS TORN/BROKEN BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED. THE SURGEON SAID THE LEADING PART, HELD BY THE FORCEPS, OF THE LENS WAS TORN IN THE BEGINING OF SUPER FUNNEL CARTRIDGE DURING LOADING PROCESS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE CLAIM IS A PRODUCT PROBLEM, NOT AN ADVERSE EVENT. THE CLAIM CONCERNS A MALFUNCTION, NOT A SERIOUS INJURY. INITIAL MDR STATES (B)(6) 2017. (B)(6) 2017 IS THE CORRECT DATE.(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6620418&pc=MTA | 2023826-2017-00934 | 2017-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 12.6 MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD), IT TORE PRIOR TO INJECTION. THIS OCCURRED ON (B)(6) 2017. THE LENS WAS NOT IMPLANTED. SURGERY WAS POSTPONED UNTIL THE FOLLOWING DAY, IN WHICH A REPLACEMENT LENS WAS IMPLANTED SUCCESSFULLY. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS HAS BEEN RETURNED AND EVALUATED. LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR FOREIGN RESIDUE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637170&pc=MTA | 2023826-2017-00963 | 2017-04-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/84 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) "201". IN THE SAME SURGERY, THE LENS WAS EXCHANGED DUE TO LENS TEAR BREAK DURING INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. THE REPORTER STATED THERE WAS A RESISTANCE WHILE PUSHING THE LENS WITH FOAM TIP PLUNGER BECAUSE OF "OVD'S SHORTAGE" (NOT ENOUGH VISCOELASTIC). AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC AND CLEAR RESIDUE ON LENS SURFACE. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13/+3/84 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON "(B)(6) 201" IS INCORRECT DATE; CORRECTED DATE IS (B)(6) 2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION: BELOW 21 YEARS OF AGE AT THE TIME OF IMPLANT. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721298&pc=MTA | 2023826-2017-01179 | 2017-04-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/3.0/91 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. A LENS TEAR/BREAK OCCURED BEFORE INJECTION INTO THE EYE. A LENS CRACK OCCURED. AN ALTERNATE LENS WAS IMPLANTED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/16 AND THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/16. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE LENS WAS COMPLETELY ROLLED UP. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/3.0/91 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. A LENS TEAR/BREAK OCCURED BEFORE INJECTION INTO THE EYE. AN ALTERNATE LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/16 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/16. (B)(4). Manufacturer Narrative: PMA 510(K): - THE PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6699349&pc=MTA | 2023826-2017-01126 | 2017-04-15 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, INTENDING TO INJECT INTO THE PATIENT'S LEFT EYE (OS), IT TORE/BROKE DURING LOADING. THIS OCCURRED ON (B)(6) 2017. THE REPORTER STATED THAT THE LENS BROKE WHILE LOADING THE LENS ONTO THE INJECTOR. THE CAUSE FOR THE TEAR/BREAK IS UNKNOWN AT THIS TIME. ON (B)(6) 2017 A REPLACEMENT LENS OF THE SAME SIZE AND SIMILAR DIOPTER WAS IMPLANTED SUCCESSFULLY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, WITH CLEAR RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC WAS BROKEN AND BENT. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673610&pc=MTA | 2023826-2017-01068 | 2017-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 IN THE SAME SURGERY DUE TO LENS TEAR DURING INJECTION. THE LENS WAS EXCHANGED WITH ANOTHER LENS OF SIMILAR MODEL AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON THE LENS. VISUAL INSPECTION FOUND OPTIC TORN AND A MISSING PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6734572&pc=MTA | 2023826-2017-01131 | 2017-05-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VICMO13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE LENS TORN DURING LOADING, BEFORE INJECTION. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6754126&pc=MTA | 2023826-2017-01243 | 2017-05-12 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-31 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DURING THE SURGERY, THE DOCTOR FOUND THAT THE LENS WAS FLIPPED INSIDE THE EYE AND PROCEEDED TO IMMEDIATELY REMOVE IT. THE LENS BROKE IN THE PROCESS. THE LENS WAS EXCHANGED INTRAOPERATIVELY FOR ANOTHER LENS OF THE SAME MODEL, SIZE, AND DIOPTER. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, HAPTIC TORN AND A PIECE OF HAPTIC MISSING ALONG WITH CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6685111&pc=MTA | 2023826-2017-01102 | 2017-05-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON HAD TWO 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENSES, -7.0 DIOPTER, THAT TORE DURING FOLDING. THE PRIMARY LENS HAD PATIENT CONTACT AND THE BACKUP LENS DID NOT. THE REPORTER STATED THAT THE CAUSE OF THE EVENT COULD HAVE BEEN DUE TO A LOADING ISSUE. THERE WAS NO PATIENT INJURY AND THE SURGEON DECIDED TO IMPLANT THE BACKUP LENS OF THE OPPOSITE EYE TO COMPLETE THE SURGERY. THE REPORTER WAS NOT SURE WHICH LENS WAS THE PRIMARY LENS THAT HAD PATIENT CONTACT BECAUSE THE LABEL WAS REMOVED WITH THE WRAP. Manufacturer Narrative: LENS WAS RETURNED IN LENS VIAL, THERE WAS NO LIQUID IN THE VIAL. VISUAL INSPECTION FOUND A TEAR ON THE OPTIC AND A MISSING PIECE OF HAPTIC. ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670145&pc=MTA | 2023826-2017-01028 | 2017-05-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3/94 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXPLANTED DUE TO LENS TEAR/BEAK DURING INJECTION INTO THE EYE. THE CUSTOMER STATED THAT HE THOUGHT THE INADEQUATELY HYDRATED FOAM TIP IN BSS IS CONTRIBUTING FACTOR TO THE LENS TEAR. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIALS ON LENS SURFACE (CLEAR RESIDUE) AND LENS WAS RETURNED IN A CURVED SHAPE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6769305&pc=MTA | 2023826-2017-01268 | 2017-06-01 | Malfunction | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017 AND NOTICED THE LENS WAS TORN WHEN IT UNFOLDED IN THE ANTERIOR CHAMBER. ACCORDING TO THE SURGEON THERE WAS SOME DIFFICULTY DURING INJECTION. IT WAS REPORTED THAT THE CAUSE OF THE EVENT MIGHT BE RELATED TO THE FTP (FOAM TIP PLUNGER) NOT BEING HYDRATED ENOUGH OR INAPPROPRIATE FOLDING OF THE LENS. THE LENS WAS EXCHANGED ON THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, HAPTIC BENT AND A PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694764&pc=MTA | 2023826-2017-01111 | 2017-06-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6691804&pc=MTA | 2023826-2017-01125 | 2017-06-09 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -1.25 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD), HE NOTICED AN 'INDENT' ON THE LENS OPTIC DURING LOADING, PRIOR TO INJECTION. THIS OCCURRED ON (B)(6) 2017. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: ADDITIONAL NARRATIVE: ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6863565&pc=MTA | 2023826-2017-01429 | 2017-06-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, AND THE SURGEON NOTED THE LENS WAS DAMAGED. THE LENS WAS PARTIALLY INSERTED AND WAS REMOVED WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THE LENS WAS MOST LIKELY DAMAGED DURING THE LOADING PROCESS (THE FOAM TIP PLUNGER PINCHED THE LENS), AS THE LENS WAS BEING PULLED THROUGH THE CARTRIDGE. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890306&pc=MTA | 2023826-2017-01475 | 2017-06-19 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTICED THAT THE LENS WAS BROKEN DURING IMPLANTATION. THE LENS WAS EXCHANGED FOR A SIMILAR LENS DURING THE SAME SURGERY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6726494&pc=MTA | 2023826-2017-01208 | 2017-06-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, TORE UPON INSERTION, WHEN ADVANCING FROM THE CARTRIDGE. THERE WAS PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6820643&pc=MTA | 2023826-2017-01372 | 2017-06-23 | Malfunction | STAAR SURGICAL COMPANY | 2017-08-25 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, TORE/BROKE UPON INSERTION IN THE PATIENT'S RIGHT EYE (OD). A BACKUP LENS WAS USED AND THE REPORTER STATED THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS -13.5 DIOPTER TORE WHEN LOADING. THERE WAS NO PATIENT CONTACT AND A BACK-UP LENS WAS USED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6737867&pc=MTA | 2023826-2017-01218 | 2017-06-26 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-24 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED AND EXCHANGED INTRAOPERATIVELY ON (B)(6) 2017 DUE TO A LENS TEAR DURING INJECTION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED AND EXCHANGED INTRAOPERATIVELY ON (B)(6) 2017 DUE TO A LENS TEAR DURING INJECTION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THIS EVENT WAS STATED AS "NOT ENOUGH HYDRATION OF THE FOAM TIP". CONCLUSION: THE CAUSE OF THIS EVENT WAS STATED AS "NOT ENOUGH HYDRATION OF THE FOAM TIP". (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND PIECES OF THE HAPTIC MISSING AND CLEAR RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6737269&pc=MTA | 2023826-2017-01235 | 2017-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-24 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. THE SURGEON "HAD TO USE THE BACKUP LENS FOR THE PHAKIC PATIENT IN THEATRE YESTERDAY. THE FIRST LENS TORN ON APPLICATION." Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -15.5/2.5/069 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON "HAD TO USE THE BACKUP LENS FOR THE PHAKIC PATIENT IN THEATRE YESTERDAY." THE FIRST LENS TORE ON APPLICATION ON (B)(6) 2017 AND HAD TO BE EXCHANGED WITH THE BACK UP LENS OF THE SAME MODEL AND DIOPTER. THE PROBLEM WAS RESOLVED. THE POST-OP STATED THAT THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) AND UNCORRECTED VISUAL ACUITY (UCVA) WERE BOTH FOUND TO BE 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. (B)(4). Manufacturer Narrative: THE FIRST LENS TORE ON APPLICATION ON (B)(6) 2017 AND HAD TO BE EXCHANGED WITH THE BACK-UP LENS OF THE SAME MODEL AND DIOPTER INTRAOPERATIVELY. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6845961&pc=MTA | 2023826-2017-01398 | 2017-06-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS ABOUT TO IMPLANT A MICL12.6, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. WHILE THE LENS WAS BEING LOADED, THE LENS TORE. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850309&pc=MTA | 2023826-2017-01395 | 2017-07-10 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, AND THE LENS WAS TORN DURING THE LOADING PROCESS. THE HAPTIC WAS TORN WHILE PUSHING THE LENS WITH THE FOAM TIP PLUNGER. THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890686&pc=MTA | 2023826-2017-01468 | 2017-07-17 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON TRIED TO IMPLANT A 13.2 MM VICM5_13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). DURING IMPLANTATION, THE LENS BROKE AS IT WAS BEING INJECTED. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872821&pc=MTA | 2023826-2017-01441 | 2017-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTICED THAT THE LENS TORE DURING INJECTION. THE SURGEON BELIEVES THE DEHYDRATION OF THE FOAM TIP CAUSE THE TEAR. THE LENS WAS EXCHANGED FOR A SIMILAR LENS DURING THE SAME SURGERY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC AND OPTIC MISSING AND WHITE RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899576&pc=MTA | 2023826-2017-01512 | 2017-07-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT BEFORE THE SURGEON WAS ABLE TO IMPLANT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THIS OCCURRED ON (B)(6) 2017. REPORTEDLY, THE LENS DID NOT TOUCH THE PATIENT'S EYE. ON (B)(6) 2017, THE SURGEON THEN IMPLANTED A BACK-UP LENS OF THE SAME SIZE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND CLEAR RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6818713&pc=MTA | 2023826-2017-01325 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS LOADING A MICL13.2, -15.5 DIOPTER, IMPLANTABLE COLLAMER LENS ON (B)(6) 2017. IT WAS NOTICED THAT THE LENS TORE WHILE LOADING. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE AND DEBRIS AND CLEAR RESIDUE ON LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6813301&pc=MTA | 2023826-2017-01366 | 2017-07-27 | Injury | STAAR SURGICAL COMPANY | 2017-08-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, TORE WHILE BEING LOADED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335041&pc=MTA | 2023826-2018-00342 | 2017-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -9.50/+4.0/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. PATIENT WAS DISSATISFIED WITH VISUAL ACUITY AS IT IS 20/25. THE SURGEON IS CURRENTLY CHECKING ON THE PATIENT'S CONDITION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089422&pc=MTA | 2023826-2017-01828 | 2017-08-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANTED A 12.6 MM VTICMO12.6, -15.0/+1.5/168 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. WHILE LOADING THE LENS IN THE CARTRIDGE, THE LENS GOT TORN. THE SURGEON ABORTED THE SURGERY. ON (B)(6) 2017, ANOTHER LENS WAS IMPLANTED AND IT WAS SUCCESSFUL WITH THE PATIENT DOING WELL. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE VIAL. VISUAL INSPECTION FOUND THE LENS MISSING PIECES OF ITS HAPTIC AND PRESENCE OF CLEAR RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201646&pc=MTA | 2023826-2018-00088 | 2017-08-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2017. ON (B)(6) 2017 A SAME SIZE, TYPE, AND DIOPTER LENS WAS SUCCESSFULLY IMPLANTED. REPORTEDLY, THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS EXPLANTED INTRAOPERATIVELY ON (B)(6) 2017. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6863590&pc=MTA | 2023826-2017-01433 | 2017-08-22 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, WAS TORN DURING LOADING, AS THE LENS WAS BEING FORWARDED. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899564&pc=MTA | 2023826-2017-01504 | 2017-08-30 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. REPORTEDLY, THE LENS TORE WITH THE FORCEPS, IT WAS NOT IMPLANTED, AND IT DID NOT TOUCH THE PATIENT'S EYE. THE SURGEON THEN IMPLANTED A BACK-UP LENS OF THE SAME SIZE. THE SURGERY WAS COMPLETED ROUTINELY. THIS OCCURRED ON (B)(6) 2017. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE LENS MISSING A PIECE OF HAPTIC. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899554&pc=MTA | 2023826-2017-01514 | 2017-08-31 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM MICL13.2 DIOPTER -9.5 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE, THE LENS TORE. REPORTEDLY, THE DOCTOR NOTICED THE TEAR AS IT OCCURRED WHILE HE WAS PULLING IT THROUGH THE INJECTION CARTRIDGE. IT WAS NOT IMPLANTED, AND IT DID NOT TOUCH THE PATIENT'S EYE. THE SURGEON THEN IMPLANTED A LENS OF THE SAME SIZE, MODEL, AND DIOPTER. THIS OCCURRED ON (B)(6) 2017. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: DEVICE EVALUATION: SHOULD HAVE BEEN RECORDED ON SUPPLEMENTAL #1. THE DELIVERY SYSTEM WAS RETURNED IN A DEVICE TRAY. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE, LENS HAVING FOREIGN MATERIAL ON DEVICE (WHITE RESIDUE ON CARTRIDGE), AND LENS STUCK IN CARTRIDGE. CLAIM #: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN CARTRIDGE. VISUAL INSPECTION FOUND THE LENS RETURNED IN CARTRIDGE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021355&pc=MTA | 2023826-2017-01669 | 2017-09-12 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR ANOTHER LENS DURING THE SAME SURGERY DUE TO THE LENS BEING TORN. THE PROBLEM WAS RESOLVED. ACCORDING TO THE SURGEON, THE TEAR MAY BE DUE TO POOR LOADING OF THE LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6997297&pc=MTA | 2023826-2017-01648 | 2017-09-19 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. THE SURGEON WAS HAVING DIFFICULTY LOADING THE LENS AS THE LENS KEPT FOLDING. AFTER THE IMPLANTATION, THE SURGEON NOTICED THAT THE LENS TORE. THE LENS WAS EXCHANGED INTRAOPERATIVELY WITH A BACK-UP LENS AND THE PROBLEM WAS RESOLVED Manufacturer Narrative: ADDITIONAL INFORMATION: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPES WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND A PIECE MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069117&pc=MTA | 2023826-2017-01787 | 2017-09-23 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VICMO13.2, -2.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. BEFORE IMPLANTATION THE LENS WAS ACCIDENTALLY DAMAGED. THE LENS WAS NOT IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE UNSUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6997998&pc=MTA | 2023826-2017-01646 | 2017-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS THE NEXT DAY, AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE ICL FRONT LOADING FORCEP. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND RED RESIDUE ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7007950&pc=MTA | 2023826-2017-01700 | 2017-10-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A MICL13.2, -8.0 DIOPTER IN THE PATIENT'S EYE ON (B)(6) 2017. WHILE LOADING THE LENS, THE SURGEON FELT THAT THE FORCEP SLIPPED DURING LOADING. THE SURGEON WAS UNSURE WHETHER THAT THE LENS TORE, SO HE USED A BACK-UP LENS INSTEAD AND IT WAS SUCCESSFUL. THERE WAS NO PATIENT CONTACT WITH THE FIRST LENS. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050896&pc=MTA | 2023826-2017-01713 | 2017-10-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS BROKEN BEFORE INJECTION INTO THE EYE. ATTEMPTS TO OBTAIN ADDITION INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTER CONFIRMED THAT THE LENS DID NOT HAVE ANY PATIENT CONTACT AND THAT THE EVENT WAS CAUSED BY USER/TECHNIQUE ERROR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7048928&pc=MTA | 2023826-2017-01778 | 2017-10-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE LENS WAS TORN DURING DELIVERY ON (B)(6) 2017.THE SURGEON STATED THAT HE BELIEVES THE FOAM TIP PLUNGER (FTP) MAY NOT HAVE EXPANDED FULLY AS THE ICL WAS ADVANCED DOWN THE INJECTOR. AS A RESULT, THE ICL BECAME PARTIALLY STUCK BETWEEN THE FTP AND THE INJECTOR AND DID NOT ADVANCE PROPERLY. THE LENS WAS PLACED INTO THE EYE BUT WAS TORN AT THE TRAILING HAPTIC/ FOOTPLATE. THE ICL WAS THEN REMOVED AND A BACK-UP LENS WAS SUCCESSFULLY IMPLANTED INTRAOPERATIVELY. SURGEON STATES THAT THE WHOLE PROCESS WAS COMPLETED IN LESS THAN 30 MINUTES. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE IMPLANTABLE COLLAMER LENS, MICL 12.6 OF -13.0 DIOPTER WAS TORN DURING DELIVERY ON (B)(6) 2017. THE SURGEON STATED THAT HE BELIEVES THE FOAM TIP PLUNGER (FTP) MAY NOT HAVE EXPANDED FULLY AS THE ICL WAS ADVANCED DOWN THE INJECTOR. AS A RESULT THE ICL BECAME PARTIALLY STUCK BETWEEN THE FTP AND THE INJECTOR AND DID NOT ADVANCE PROPERLY. THE LENS WAS PLACED INTO THE EYE BUT WAS TORN AT THE TRAILING HAPTIC/ FOOTPLATE. THE ICL WAS THEN REMOVED AND A BACK-UP LENS WAS SUCCESSFULLY IMPLANTED INTRAOPERATIVELY. SURGEON STATES THAT THE WHOLE PROCESS WAS COMPLETED IN LESS THAN 30 MINUTES. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: USAGE OF DEVICE: "INITIAL USE OF DEVICE" SHOULD HAVE BEEN CHECKED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND TEARS IN THE HAPTIC AND A PIECE OF HAPTIC MISSING. METHOD CODE : NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7068475&pc=MTA | 2023826-2017-01797 | 2017-10-31 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017, AS THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S LEFT (OS) EYE, HE DISCOVERED THE LENS TORE. REPORTEDLY, THE LENS TORE IN THE CARTRIDGE. THE LENS BROKE BEFORE INJECTION AND WAS NOT IMPLANTED. AN ALTERNATE LENS WAS IMPLANTED SUCCESSFULLY "AT THE SAME TIME" AND ON THE SAME DATE AS THE ORIGINAL LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114226&pc=MTA | 2023826-2017-01871 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE LEFT EYE (OS) ON (B)(6) 2017. DURING IMPLANTATION, THE PATIENT'S EYE ROLLED. THE SURGEON EXPLANTED AND EXCHANGED THE LENS INTRAOPERATIVELY WITH A BACK-UP LENS AND IT WAS SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: DATE OF THIS REPORT: (B)(6) 2017 CORRECTED TO (B)(6) 2017. SERIAL #: (B)(4) CORRECTED TO (B)(4). CLAIM# (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114229&pc=MTA | 2023826-2017-01872 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE RIGHT EYE (OD) ON (B)(6) 2017. DURING IMPLANTATION, THE PATIENT'S EYE ROLLED. THE SURGEON EXPLANTED AND EXCHANGED THE LENS INTRAOPERATIVELY WITH A BACK-UP LENS AND IT WAS SUCCESSFUL. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320719&pc=MTA | 2023826-2018-00317 | 2017-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S LEFT EYE (OS), THE LENS BROKE. INTRA-OPERATIVELY THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. THIS OCCURRED ON (B)(6) 2017. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438478&pc=MTA | 2023826-2018-00539 | 2017-11-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -11.5/2.5/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS INTRAOPERATIVELY EXPLANTED DUE TO THE LENS BEING TORE DURING INJECTION. THE PATIENT IS FINE AND DOES NOT WANT TO HAVE ANOTHER OPERATION. THE REPORTER BELIEVES THAT LACK OF HYDRATION MIGHT CONTRIBUTE TO THE TEAR. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE TORN PIECE MISSING. CORRECTED INFORMATION: LENS WAS IMPLANTED ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: LENS WAS IMPLANTED ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183433&pc=MTA | 2023826-2018-00051 | 2017-11-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICM5_13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). DURING THE ICL LOADING, THE LENS TORE. THE SURGEON SAID THA TTHE LENS WAS TORN WHILE PUSHING THE LENS WITH THE FOAM TIP PLUNGER DUE TO ITS RESISTANCE. AFTERWARDS, THE SURGEON THEN PERFORMED ANOTHER SURGERY WITH A SIMILAR LENS INTO THE PATIENT. ON (B)(6) 2017, THE PATIENT'S POST-OP BCVA WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND CLEAR AND WHITE RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7186670&pc=MTA | 2023826-2018-00079 | 2017-11-24 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), AND THE LENS TORE/BROKE. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC AND RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180675&pc=MTA | 2023826-2018-00030 | 2017-11-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.7 MM VICM5_13.7, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. BEFORE IMPLANTATION THE LENS TORE/BROKE. THE LENS WAS NOT IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE UNSUCCESSFUL. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID; IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201663&pc=MTA | 2023826-2018-00098 | 2017-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -18.00 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2017. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND PIECES OF THE HAPTIC MISSING AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268472&pc=MTA | 2023826-2018-00190 | 2017-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE 12.6 MM VICM5_12.6, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS GOT ADHERED TO THE FORCEPS AND TORE WHEN THE SURGEON LOADED IT INTO THE CARTRIDGE BEFORE IMPLANTING THE LENS INTO THE EYE ON (B)(6) 2017. ANOTHER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE/DEBRIS ON LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: ANOTHER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD). DEVICE EVALUATED BY MANUFACTURER: IN THE INITIAL MDR, CORRECTED FROM YES TO NO. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7285998&pc=MTA | 2023826-2018-00188 | 2017-12-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. IT WAS NOTICED THAT THE LENS TORE DURING INJECTION. THE LENS WAS INTRAOPERATIVELY EXCHANGED WITH A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: PRODUCT EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL OBSERVATION FOUND THE LENS MISSING A PIECE OF HAPTIC AND RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268462&pc=MTA | 2023826-2018-00176 | 2018-01-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE THE SURGEON ATTEMPTED TO LOAD A 13.2 MM MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS ON (B)(6) 2018, BUT IT WAS LOADED INCORRECTLY. THERE WAS NOT ANYTHING UNUSUAL THAT LED TO THE PROBLEM. THERE WAS NO PATIENT CONTACT WITH THE DEVICE. Manufacturer Narrative: CORRECTED DATA: AGE OF THE PATIENT AT THE TIME OF THE SURGERY IS CORRECTED FROM " (B)(6)" TO "(B)(6)". CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300910&pc=MTA | 2023826-2018-00282 | 2018-01-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS IMPLANTING A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), IT TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. THERE WAS PATIENT CONTACT WITH THIS LENS HOWEVER THE LENS WAS NOT FULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL OBSERVATION FOUND THE LENS WITH NO VISIBLE DAMAGE. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS; -13.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) WHEN IT TORE/BROKE BEFORE IMPLANTATION INTO THE EYE ON (B)(6) 2018. THERE WAS PATIENT CONTACT BUT THE LENS WAS NOT IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: A BACK UP LENS WAS IMPLANTED AT A LATER TIME AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571616&pc=MTA | 2023826-2018-00854 | 2018-01-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/093 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED INTRAOPERATIVELY, THE SURGEON NOTING LENS TEAR/BREAK DURING INJECTION INTO THE EYE AND "BROKEN LENS." THIS OCCURRED ON (B)(6) 2018. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE/DEBRIS ON LENS). CORRECTED DATA: REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378406&pc=MTA | 2023826-2018-00471 | 2018-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. INTRAOPERATIVELY THE LENS WAS EXPLANTED. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CORRECTED DATA: TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7346784&pc=MTA | 2023826-2018-00422 | 2018-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, BUT THE LENS TORE/BROKE BEFORE INJECTION INTO THE PATIENT'S EYE. THERE WAS NO PATIENT CONTACT AND ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7440971&pc=MTA | 2023826-2018-00568 | 2018-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM VICM5_13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS, THE LENS TORE. REPORTEDLY, THE LENS "GOT SQUEEZED INTO THE INJECTOR DURING LOADING" AND THIS LED TO THE LENS TEAR. THE INTENT WAS TO IMPLANT THE LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647591&pc=MTA | 2023826-2018-00985 | 2018-02-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE REPORTER INDICATED THE LENS GOT STUCK ON THE PLUNGER. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO REPORTED INJURY. THE CORNEA WAS CLEAR. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT; AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7410050&pc=MTA | 2023826-2018-00516 | 2018-03-08 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO LOAD A 12.6 MM VTICMO12.6, -11.5/+2/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE CARTRIDGE ON (B)(6) 2018 WHEN THE LENS TORE. THE LENS WAS BEING SLID SOFTLY WITH THE FORCEPS THRU THE NARROWEST PART OF THE CARTRIDGE WHEN IT SUDDENLY TORE FROM ONE OF THE CENTRAL HOLES TO THE HOLE IN THE DISTAL HAPTIC AND TO THE DISTAL LEFT HAPTIC. A REPLACEMENT LENS IS PLANNED FOR THE PATIENT. THE REPORTED BELIEVED THAT THE LENS HAD A LITTLE SCRATCH FROM MANUFACTURER. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL OBSERVATION FOUND THE LENS WITH THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7508583&pc=MTA | 2023826-2018-00720 | 2018-03-14 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, AND THE LENS TORE DURING DELIVERY INTO THE PATIENT'S EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE TORN HAPTIC PIECE MISSING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7487578&pc=MTA | 2023826-2018-00583 | 2018-03-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON FOUND A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/091 (SPHERE/CYLINDER/AXIS), WAS CRACKED AFTER OPENING THE LENS VIAL. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. THERE WAS NO MEDICAL OR SURGICAL INTERVENTION NECESSARY AND THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484264&pc=MTA | 2023826-2018-00642 | 2018-03-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON PLANNED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/139 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2018. HOWEVER, PRIOR TO IMPLANT, THE HAPTIC TORE. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2018. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE TORN HAPTIC PIECE IS MISSING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535059&pc=MTA | 2023826-2018-00775 | 2018-04-08 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE LENS MISSING THE PIECE OF TORN PIECE. THERE WAS ALSO PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535110&pc=MTA | 2023826-2018-00777 | 2018-04-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7617159&pc=MTA | 2023826-2018-00924 | 2018-04-19 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-19 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+3.0/095 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2018. REPORTEDLY, THE LENS GOT STUCK IN THE PLUNGER AND THE HAPTIC TORE. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639917&pc=MTA | 2023826-2018-00955 | 2018-04-24 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE. THE LENS WAS IMPLANTED, THEN EXPLANTED INTRAOPERATIVELY ON (B)(6) 2018. REPORTEDLY, THE SURGEON DID NOT USE THE FORCEPS TO GRASP THE LENS BUT USED THE FOAM TIP PLUNGER TO PUSH THE LENS TO THE FRONT OF THE INJECTOR. IT IS POSSIBLE THAT THIS IS THE REASON THE LENS TORE. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: FAILURE TO FOLLOW INSTRUCTIONS (FORCEPS NOT USED TO GRASP LENS) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639946&pc=MTA | 2023826-2018-00954 | 2018-05-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), HE DISCOVERED THAT THE LENS WAS UPSIDE DOWN. REPORTEDLY, THE SURGEON "OVERFILLED" VISCOELASTIC INTO THE ANTERIOR CHAMBER. WHEN THE SURGEON EXTRACTED THE LENS, IT TORE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS A SMALL PIECE OF THE LENS RETURNED. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6176021&pc=MTA | 2023826-2016-01767 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE MANUFACTURER RECEIVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, FROM THE CUSTOMER DAMAGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS DAMAGED DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT OR INJURY. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO A LOADING ERROR, THERE WAS NO ISSUE WITH THE PRODUCT. EVENT PROBLEM AND EVALUATION CODES: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6226914&pc=MTA | 2023826-2017-00024 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/087 DIOPTER, TORE/BROKE BEFORE INJECTION INTO THE EYE. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN, HAPTIC WAS BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6298355&pc=MTA | 2023826-2017-00195 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS "WAS DAMAGED AND THE BACK UP LENS HAD TO BE USED." NO OTHER INFORMATION IS AVAILABLE AT THIS TIME, ATTEMPTS ARE ON-GOING. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, WAS DAMAGED DURING SURGERY WHILST LOADING AND BACK UP LENS HAD TO BE USED. THREE ATTEMPTS WERE MADE FOR MORE INFORMATION, NONE WAS FORTHCOMING. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6322385&pc=MTA | 2023826-2017-00192 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-02-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE, THE LENS TORE. LENS WAS NOT IMPLANTED. Manufacturer Narrative: THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6323511&pc=MTA | 2023826-2017-00230 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, WAS TORN WHILE BEING LOADED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TEAR ON THE HAPTIC. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6421840&pc=MTA | 2023826-2017-00440 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, AS THE SURGEON PREPARED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.0/105 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD), THAT THE "LENS WAS RECEIVED DEFECTIVE." A BACK-UP LENS WAS USED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN AND THE OTHER HAPTIC HAD A PIECE TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6426360&pc=MTA | 2023826-2017-00473 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, WAS TORN WHILE BEING INSERTED AND IT WAS UNKNOWN IF THERE WAS ANY PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID, WITH CLEAR RESIDUE ON THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447950&pc=MTA | 2023826-2017-00542 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S RIGHT EYE (OD). THE CUSTOMER STATES THAT THE LENS DID NOT HAVE ENOUGH PRIMER, THEREFORE IT GOT STUCK AND DAMAGE IN THE CARTRIDGE, AND IT WAS NOT IMPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS WAS NOT RETAINED, UNABLE TO RETURN. Manufacturer Narrative: CORRECTED DATA: (B)(6) 2017. PREVIOUS - STATED THE LENS WAS IMPLANTED. CORRECTED - THE LENS WAS DAMAGE WITHIN THE CARTRIDGE. NO PATIENT CONTACT. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. RESULT CODE: WORK ORDER SEARCH PERFORMED, NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487514&pc=MTA | 2023826-2017-00617 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574203&pc=MTA | 2023826-2017-00835 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, AND THE LENS WAS NOTED TO BE TORN. THERE WAS PATIENT CONTACT BUT NO INJURY. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS STUCK IN AN SFC-45 FP CARTRIDGE. THERE WAS NO VISIBLE DAMAGE TO THE CARTRIDGE. THERE WAS EVIDENCE OF CLEAR RESIDUE ON CARTRIDGE. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6602175&pc=MTA | 2023826-2017-00897 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION OF THE LENS FOUND A PIECE OF A HAPTIC MISSING. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6679387&pc=MTA | 2023826-2017-01098 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, WAS TORN DURING LOADING AND WAS NOT IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6734723&pc=MTA | 2023826-2017-01233 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED SURGEON LOADED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, AND NOTED THE LENS WAS WRINKLED. THE SURGEON DECIDED NOT TO USE THE LENS AND REQUESTED THE BACKUP LENS, WHICH WAS IMPLANTED WITH NO PROBLEM. THERE WAS NO PATIENT CONTACT AND THE CAUSE OF THE LENS DAMAGE WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6853773&pc=MTA | 2023826-2017-01416 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED SURGEON LOADED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, AND NOTED THE LENS WAS WRINKLED. THE SURGEON DECIDED NOT TO USE THE LENS AND REQUESTED THE BACKUP LENS, WHICH WAS IMPLANTED WITH NO PROBLEM. THERE WAS NO PATIENT CONTACT AND THE CAUSE OF THE LENS DAMAGE WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: A WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS WERE FOUND. (B)(4) Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6997274&pc=MTA | 2023826-2017-01649 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | COLLAMER ULTRAVIOLET-ABSORBING POSTERIOR CHAMBER THREE PIECE FOLDABLE IOL | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P990013 | 1 | Event Description: THE REPORTER INDICATED THAT THE CQ2015A, +23.0 DIOPTER, INTRAOCULAR LENS GOT STUCK IN THE CARTRIDGE. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE PRODUCT WAS RETURNED LOOSE IN THE LENS BOX. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE AND THE LENS SUCK IN THE CARTRIDGE. CLAIM #: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7165895&pc=MTA | 2023826-2018-00025 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, WAS CLIPPED WITH THE INJECTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7286007&pc=MTA | 2023826-2018-00250 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: A 12.6 MM MICL12.6, -15.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS RETURNED DAMAGED TO THE MANUFACTURER. ATTEMPTS AT OBTAINING ADDITION INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND CLEAR RESIDUE ON LENS SURFACE AND THE HAPTIC MISSING A PIECE. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378739&pc=MTA | 2023826-2018-00465 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICMO12.6, -12.00/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S EYE. IT WAS NOTICED THAT THE LENS WAS TORN IN THE MIDDLE. THE SURGEON PLANS TO EXCHANGE THE LENS. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO [OR #] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7448149&pc=MTA | 2023826-2018-00574 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-04-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER WAS UNABLE TO PROVIDE ANY INFORMATION FOR A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, THAT WAS RETURNED TO THE MANUFACTURER DAMAGED. THE REPORTER STATED IT WAS UNKNOWN IF THERE WAS ANY PATIENT CONTACT AND NO MORE INFORMATION WAS AVAILABLE. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6109400&pc=MTA | 2023826-2016-01645 | 2016-08-08 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Activation Failure | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTICED THE LENS WAS FOLDED AND DID NOT IMPLANT THE LENS. THE BACKUP LENS WAS IMPLANTED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (SMALL PARTICULATES). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7286002&pc=MTA | 2023826-2018-00187 | 2017-12-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -11.00/1.5/X088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. AFTER IMPLANTATION, LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS NOTICED AS WELL AS THE LENS BEING TORN DURING INJECTION. A LENS EXCHANGE SURGERY IS PLANNED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7434208&pc=MTA | 2023826-2018-00556 | 2018-03-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Mechanical Jam | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/107 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE SURGEON REPORTED THE LENS GOT JAMMED IN THE FOAM TIP PLUNGER WHEN IMPLANTING AND WAS TORN. THE LENS WAS REMOVED AND EXCHANGED DURING THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS REMOVED DURING INITIAL SURGERY. ANOTHER LENS WAS IMPLANTED AT A LATER DATE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5667306&pc=MTA | 2023826-2016-00576 | 2015-02-10 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.50/+2.0/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED AND OPTIC TORN/SPLIT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5733036&pc=MTA | 2023826-2016-00829 | 2016-04-05 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THERE WAS NO APPARENT PATIENT INJURY REPORTED AND ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED. THE EXCHANGED LENS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655327&pc=MTA | 2023826-2016-00707 | 2016-04-11 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. CORRECTED DATA: DEVICE CODE: (B)(4) IS INCORRECT. THE LENS TEAR WAS NOT DUE TO A MATERIAL INTEGRITY ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817530&pc=MTA | 2023826-2016-00921 | 2016-06-15 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/133 DIOPTER, FROM THE VIAL, AND THE LENS WAS NOTED TO BE IN A TORN CONDITION. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS RETURNED DRY IN LENS/CASE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938280&pc=MTA | 2023826-2016-01248 | 2016-06-21 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, AND THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED AND ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT WAS (B)(6) AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5805681&pc=MTA | 2023826-2016-00901 | 2016-07-04 | Injury | STAAR SURGICAL COMPANY | 2016-07-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY REPORTED. THE LENS WAS EXCHANGED WITH ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN AND FOREIGN MATERIAL ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6124479&pc=MTA | 2023826-2016-01641 | 2016-07-13 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND THE LENS WAS NOTED TO BE TORN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A SHORTER LENS THE SAME MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND HAPTIC BROKEN AND FIBERS ON LENS SURFACE. METHOD CODE; WORK ORDER SEARCH PERFORMED AND NO SIMILAR COMPLAINTS FOUND. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5869117&pc=MTA | 2023826-2016-01036 | 2016-07-29 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4) IS INCORRECT. THE LENS TEAR WAS NOT DUE TO A MATERIAL INTEGRITY ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955100&pc=MTA | 2023826-2016-01261 | 2016-08-03 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, AND THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY AND WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS. THE EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5996366&pc=MTA | 2023826-2016-01373 | 2016-08-25 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS - -17.50/3.50/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS TEARING DURING INJECTION INTO THE EYE. THE SURGEON IMPLANTED THE BACK-UP LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: CORRECTED DATA: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE BROKEN. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK . (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059282&pc=MTA | 2023826-2016-01525 | 2016-09-27 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS - -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO THE LENS TEARING DURING THE DELIVERY INTO THE EYE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: CONCLUSION: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. CORRECTED DATA: (B)(4) PREVIOUS CODE NOT APPLICABLE, NO MATERIAL INTEGRITY PROBLEM REPORTED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS BENT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248526&pc=MTA | 2023826-2015-01474 | 2015-08-11 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Slipped; Positioning Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER AND THE LENS "SLID OUT" OF THE PATIENT'S EYE. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULTS (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE LENS HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. THE CAUSE OF THE LENS TO SLIP OUT OF THE EYE, IS POSSIBLY DUE TO IMPROPER LOADING OR HANDLING PRACTICES IN THE OPERATING ROOM. PER MEDICAL REVIEW - REPORTEDLY, LENS "SLID OUT " OF THE EYE REQUIRING INTRAOPERATIVE LENS EXCHANGE. NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. THE BACKUP LENS WAS SUCCESSFULLY IMPLANTED. GIVEN THE INFORMATION ABOUT THE LENS THAT SLIPPED OUT OF AN EYE IT IS VERY LIKELY THAT USER ERROR (PREMATURE LENS INJECTION) HAVE CAUSED/CONTRIBUTED TO THE EVENT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, IT IS VERY LIKELY THAT USER ERROR (PREMATURE LENS INJECTION) HAVE CAUSED/CONTRIBUTED TO THE EVENT. (B)(4). Manufacturer Narrative: METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6265535&pc=MTA | 2023826-2017-00097 | 2016-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult or Delayed Positioning; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, DID NOT FOLD WELL INTO THE CARTRIDGE DURING THE LOADING PROCESS. THE SURGEON DID NOT LIKE HOW THE LENS WAS INJECTING, SO THE BACKUP LENS WAS LOADED AND IMPLANTED WITH NO PROBLEM. THERE WAS PATIENT CONTACT WITH THE FIRST LENS BUT NO INJURY. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS VIAL. VISUAL INSPECTION OF THE LENS FOUND LENS HAD DRY RESIDUE AND PARTICULATES. EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S)-WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6772663&pc=MTA | 2023826-2017-01252 | 2017-07-04 | Malfunction | STAAR SURGICAL COMPANY | 2017-08-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Difficult or Delayed Positioning; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO REPORT OF ANY PATIENT INJURY. ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED WHEN THE LENS WAS INJECTING INTO THE EYE, THE LENS BECAME TRAPPED BETWEEN THE FOAM TIP PLUNGER AND THE CARTRIDGE TIP AND WOULD NOT RELEASE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BENT AND A PIECE OF THE OTHER HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824665&pc=MTA | 2023826-2020-00449 | 2020-02-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Fold, Unfold or Collapse | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, MICL 13.7, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS TO BE IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). "SURGEON DID NOT LIKE THE WAY THE LENS WAS LOADED OR UNFOLDING AS IT REACHED THE END OF THE CARTRIDGE." THE BACK-UP LENS WAS INSTEAD IMPLANTED. REPORTEDLY THE INITIAL "LENS NEVER CAME IN CONTACT WITH THE PATIENT." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- LENS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4). Manufacturer Narrative: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393946&pc=MTA | 2023826-2017-00363 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-03-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Fold, Unfold or Collapse | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS HAVING DIFFICULTY UNFOLDING THE TRAILING HAPTIC OF AN MICL IMPLANTABLE COLLAMER LENS, UPON INJECTION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6444891&pc=MTA | 2023826-2017-00520 | 2017-03-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Fold, Unfold or Collapse; Device Operates Differently Than Expected; Material Deformation | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, BUT THE LENS WOULD NOT UNFOLD. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. UPON REMOVAL, THE LENS HAPTIC WAS NOTED TO BE CURVED AND BENT UPWARDS. THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT MAY HAVE BEEN DUE TO AGGRESSIVE MANIPULATION IN ORDER TO GET THE LENS TO UNFOLD. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842043&pc=MTA | 2023826-2015-00667 | 2015-05-19 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-06-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Insert | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS BUT THE LENS APPEARED TO BE STUCK AND ONLY CAME HALF WAY OUT OF THE CARTRIDGE. THE LENS TOUCHED THE PATIENT'S EYE BUT WAS NOT FULLY INSERTED. THE SURGEON IMPLANTED THE BACKUP LENS WITH NO PROBLEM, USING A NEW INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER. THE REPORTER INDICATED THE SURGEON FELT THE EVENT WAS DUE TO AN OVERLY USED INJECTOR AND THE FOAM TIP PLUNGER, WHICH WAS NOT CONNECTED PROPERLY BY THE TECHNICIAN. THE SURGEON DID NOT FEEL THERE WAS A PROBLEM WITH THE LENS OR INJECTION SYSTEM. Manufacturer Narrative: (B)(4)- NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT; DIFFICULT TO INSERT. EVALUATION METHOD: WORK ORDER SEARCH, CARTRIDGE LOT NUMBER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. A CARTRIDGE LOT NUMBER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND DRIED SURGICAL RESIDUE ON THE LENS SURFACE AND WAS UNABLE TO DETECT ANY DAMAGE TO THE LENS. AN SFC-45 FP CARTRIDGE WAS RETURNED WITH NO VISIBLE DAMAGE, THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE INJECTOR AND FOAM TIP PLUNGER WERE NOT RETURNED. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, CARTRIDGE LOT NUMBER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THERE IS NO INFORMATION TO DETERMINE THE CAUSE OF INCOMPLETE EGRESS OF THE LENS FROM THE CARTRIDGE, OR IF IT WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING. A CARTRIDGE LOT NUMBER SEARCH SHOWED NO SIMILAR COMPLAINTS. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, CARTRIDGE LOT NUMBER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8081325&pc=MTA | 2023826-2018-01715 | 2018-10-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Open or Remove Packaging Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGICAL STAFF HAD DIFFICULTY OPENING THE ICL LENS VIAL DURING SURGERY. THERE WAS ONLY ONE VIAL INVOLVED AND THERE WAS NO INJURY TO THE SURGICAL STAFF. THE LENS WAS IMPLANTED. THE PRODUCT WILL NOT BE RETURNED. THE PROBLEM WAS RESOLVED WITH ADDITIONAL TRAINING PERFORMED BY THE SALES REP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD CODE 4110: LENS WORK ORDER SEARCH COMPLETED. NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724427&pc=MTA | 2023826-2016-00814 | 2016-05-18 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult To Position; Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S EYE AND THE LENS WOULD NOT CENTERED CORRECTLY. THE LENS TORE UPON REMOVAL. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED WITHIN THE SAME SURGERY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817657&pc=MTA | 2023826-2016-00978 | 2016-06-28 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Entrapment of Device | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A PIECE OF HAPTIC FROM A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, WAS TORN OFF AND TRAPPED INSIDE A CARTRIDGE AS THE SURGEON WAS PULLING THE LENS FORWARD. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS TRAPPED INSIDE THE CARTRIDGE. THERE WAS NO VISIBLE DAMAGE TO THE CARTRIDGE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140268&pc=MTA | 2023826-2016-01710 | 2016-10-31 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Entrapment of Device | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, BECAME STUCK IN THE LOADER. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULTS CODES: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS STUCK INSIDE AN SFC-45 FP CARTRIDGE. THE CARTRIDGE TIP WAS SPLIT AND A FOAM TIP PLUNGER WAS STUCK INSIDE. THERE WAS EVIDENCE OF DRY SURGICAL RESIDUE. CONCLUSIONS CODE: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069171&pc=MTA | 2023826-2017-01796 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Entrapment of Device; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A MICL13.7,-10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON COULD NOT ADVANCE THE LENS IN THE INJECTOR. THE SURGEON BELIEVED THAT THE LENS COULD HAVE BEEN DAMAGED AND CAUSE IT TO BE STUCK. THERE WAS NO PATIENT CONTACT WITH THE DEVICE. DURING THE SAME SURGERY; A BACK-UP LENS WAS USED AND IT WAS SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: DATE OF BIRTH CORRECTED FROM (B)(6) TO (B)(6). TYPE OF REPORTABLE EVENT CORRECTED FROM SERIOUS INJURY TO MALFUNCTION. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS MISSING PIECES OF ITS HAPTIC. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT.(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7542831&pc=MTA | 2023826-2018-00793 | 2017-12-29 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Entrapment of Device; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/089 (SPHERE/CYLINDER/AXIS), AND THE LENS BECAME STUCK IN THE PLUNGER. THE LENS WAS REMOVED FROM THE EYE, WITHIN THE SAME SURGERY, WITH NO APPARENT INJURY. UPON REMOVING FROM THE CARTRIDGE, THE LENS TORE. THE SURGERY WAS ABORTED. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5745388&pc=MTA | 2023826-2016-00852 | 2016-05-25 | Malfunction | STAAR SURGICAL COMPANY | 2016-06-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Failure to Unfold or Unwrap | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016 AND THE LENS WOULD NOT UNFOLD. THE LENS WAS REMOVED WITHIN THE SAME SURGERY. THE SURGEON HAD DIFFICULTY PULLING THE LENS FORWARD FROM THE CARTRIDGE WITH A FORCEP PRIOR TO INSERTION. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. THE SURGEON INDICATED THAT THE MANUFACTURER'S INJECTION SYSTEM WAS NOT THE CAUSE OF THE EVENT. Manufacturer Narrative: RESULTS: LENS STUCK INSIDE THE VIAL AND COULD NOT BE EVALUATED. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5664131&pc=MTA | 2023826-2016-00663 | 2016-03-12 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Failure to Unfold or Unwrap; Unintended Movement | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WOULD NOT UNFOLD AFTER INJECTION AND HAD ROTATED (REVERSED). THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN/BENT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917242&pc=MTA | 2023826-2015-00801 | 2015-06-18 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Folded; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS FROM THE VIAL, TO PREPARE FOR LOADING, AND NOTED ONE OF THE HAPTICS LOOKED "FLOPPY". THE SURGEON INDICATED ONE OF THE HAPTICS WAS FOLDED OVER AND HE ATTEMPTED TO UNFOLD THE HAPTIC BUT THE LENS DID NOT LOOK RIGHT. THE SURGEON DECIDED NOT TO USE THE LENS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER INDICATED THEY FELT THE LENS WAS DEFECTIVE. Manufacturer Narrative: (B)(4). METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC BENT. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356140&pc=MTA | 2023826-2018-00430 | 2018-02-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON DAMAGED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/094 (SPHERE/CYLINDER/AXIS) DURING THE LOADING PROCESS. THE WAS DISCOVERED WHILE IMPLANTING THE LENS INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2018. THE LENS WAS REMOVED INTRAOPERATIVELY. REPORTEDLY, THE DAMAGE WAS DUE TO USER ERROR. THE SURGEON PLANS TO IMPLANT A REPLACEMENT LENS IN THE FUTURE. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PATIENT'S UCVA IS 20/20 1 DAY POST OP. THE PATIENT CONDITION IS GOOD. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6817560&pc=MTA | 2023826-2017-01369 | 2017-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, AND THE LENS WAS OBSERVED TO HAVE TORN AT RIGHT LEADING FOOTPLATE, DUE TO POSSIBLE LOADING ERROR. THE LENS WAS REMOVED FROM EYE BEFORE FULLY DELIVERED INTO EYE. THE BACKUP LENS WAS IMPLANTED SUCCESSFULLY. THERE WAS NO REPORT OF PATIENT INJURY. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7236034&pc=MTA | 2023826-2018-00145 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON USED A VISCOELASTIC THAT WAS TEARING ICL IMPLANTABLE COLLAMER LENS DURING THE LOADING PROCESS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON USED A VISCOELASTIC THAT TORE THE ICL; IMPLANTABLE COLLAMER LENS. THE REPORTER ALSO STATED THAT "THE ICL IS PUSHED BACK OUT OF THE IMPLANTER TOWARDS THE PLUNGER ENTRANCE END DUE TO THE HIGH VISCOSITY ON THE HYALINE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAVE BEEN FORTHCOMING. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7236067&pc=MTA | 2023826-2018-00146 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON USED A VISCOELASTIC THAT WAS TEARING ICL IMPLANTABLE COLLAMER LENS DURING THE LOADING PROCESS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED INFORMATION: THE REPORTER INDICATED THE SURGEON USED A VISCOELASTIC THAT WAS TEARING THE ICL IMPLANTABLE COLLAMER LENS. "THE ICL IS PUSHED BACK OUT OF THE IMPLANTER TOWARDS THE PLUNGER ENTRANCE DUE TO THE HIGH VISCOSITY ON THE Z HYALIN". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7236074&pc=MTA | 2023826-2018-00144 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON USED A VISCOELASTIC THAT WAS TEARING ICL IMPLANTABLE COLLAMER LENS DURING THE LOADING PROCESS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC MISSING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7236113&pc=MTA | 2023826-2018-00147 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON USED A VISCOELASTIC THAT WAS TEARING ICL IMPLANTABLE COLLAMER LENS DURING THE LOADING PROCESS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON USED A VISCOELASTIC THAT TORE THE ICL; IMPLANTABLE COLLAMER LENS. THE REPORTER ALSO STATED THAT "THE ICL IS PUSHED BACK OUT OF THE IMPLANTER TOWARDS THE PLUNGER ENTRANCE END DUE TO THE HIGH VISCOSITY ON THE Z- HYALINE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAVE BEEN FORTHCOMING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7236117&pc=MTA | 2023826-2018-00148 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON USED A VISCOELASTIC THAT WAS TEARING ICL IMPLANTABLE COLLAMER LENS DURING THE LOADING PROCESS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED INFORMATION: THE REPORTER INDICATED THE SURGEON USED A VISCOELASTIC THAT WAS TEARING THE ICL IMPLANTABLE COLLAMER LENS. "THE ICL IS PUSHED BACK OUT OF THE IMPLANTER TOWARDS THE PLUNGER ENTRANCE DUE TO THE HIGH VISCOSITY ON THE Z HYALIN". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6091609&pc=MTA | 2023826-2016-01629 | 2016-09-18 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Improper or Incorrect Procedure or Method; Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, 5.0/+2.5/093 DIOPTER, AND THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THERE WAS NO PATIENT CONTACT AND THE PROBLEM WAS RESOLVED BY IMPLANTING ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO TECHNIQUE ERROR/SURGEON MISHANDLING. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4419700&pc=MTA | 2023826-2015-00062 | 2015-01-05 | Malfunction | STAAR SURGICAL COMPANY | 2015-01-14 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS AND NOTED THE LENS WAS TOO LONG FOR THE PATIENT'S EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE PATIENT RETURNED THE NEXT DAY AND A SHORTER LENS (12.1MM) WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. Manufacturer Narrative: (B)(4) - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT; SIZE INCORRECT FOR PATIENT. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS" BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: PER MEDICAL REVIEW - ACCURATE DETERMINATION OF ANTERIOR SEGMENT DIMENSIONS IS KEY TO THE SAFETY OF IMPLANTED ICLS. UBM WAS USED AS THE SIZING METHOD, NOT WTW AND ACD AS RECOMMENDED PER THE FDA APPROVED DFU. SAFETY AND EFFECTIVENESS OF UBM METHODOLOGY AS WELL AS APPROPRIATE NOMOGRAM FOR SIZING ARE YET TO BE ESTABLISHED AT THIS TIME. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4962172&pc=MTA | 2023826-2015-01007 | 2015-06-08 | Injury | STAAR SURGICAL COMPANY | 2015-08-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND LOW VAULT WAS OBSERVED. . THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. (B)(4). LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6012358&pc=MTA | 2023826-2016-01409 | 2015-08-18 | Injury | STAAR SURGICAL COMPANY | 2016-10-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS - -9.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT HAS COMPLAINED OF LOW VAULT. THE LENS REMAINS IMPLANTED. THE SURGEON WILL EXCHANGE THE CURRENT LENS FOR A LONGER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024723&pc=MTA | 2023826-2016-01431 | 2016-09-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. Manufacturer Narrative: ON (B)(6) 2016, PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. REPORT SOURCE IS A DISTRIBUTOR, NOT A USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6199701&pc=MTA | 2023826-2016-01806 | 2016-09-30 | Injury | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: REPORT TYPE WAS INCORRECTLY CLASSIFIED AS A MALFUNCTION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO KNOWN PATIENT IMPACT. EXCESSIVE VAULTING OBSERVED DURING INITIAL PROCEDURE. LENS EXCHANGED DURING THE SAME SURGERY. (B)(4). CLAIM EVALUATED BY REPORTER RESPONSE TO COMPLAINT QUESTIONNAIRE. PER DFU, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. PATIENT WAS (B)(6) AT THE TIME OF IMPLANT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089900&pc=MTA | 2023826-2016-01608 | 2016-10-12 | Injury | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS NOTED TO BE TOO BIG AND WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY OR COMPLICATIONS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER?: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY WITH RESIDUE. METHOD: WORK ORDER SEARCH. RESULT: A WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DATE OF BIRTH - UNK. (B)(4). CLAIM # 429083
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089911&pc=MTA | 2023826-2016-01609 | 2016-10-12 | Injury | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS NOTED TO BE TOO BIG AND WAS REMOVED DURING THE SAME SURGERY WITH NO PATIENT INJURY OR COMPLICATIONS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER?: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN AND A PIECE OF HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY WITH RESIDUE. METHOD: WORK ORDER SEARCH. RESULT: A WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6625975&pc=MTA | 2023826-2017-00948 | 2016-11-17 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, DIOPTER -8.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON THE SAME DATE, DURING THE SAME SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT OF THE LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE OPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721306&pc=MTA | 2023826-2017-01178 | 2016-11-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/2.0/98 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED IMMEDIATELY ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. SEE CLAIM# (B)(4) FOR PROBLEM RESOLUTION. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED AND DEVICE EVALUATION IS EXPECTED, BUT NOT YET PERFORMED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: ACD < 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6397791&pc=MTA | 2023826-2017-00378 | 2016-11-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/141 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED, SURGEON PLANS TO EXCHANGE. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6397794&pc=MTA | 2023826-2017-00379 | 2016-11-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/049 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED, SURGEON PLANS TO EXCHANGE. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6397786&pc=MTA | 2023826-2017-00371 | 2016-12-26 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS REMAIN IMPLANTED. PATIENT'S LAST VISIT ON (B)(6) 2016 UNCORRECTED VISUAL ACUITY IS 20/99. Manufacturer Narrative: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS (-5.50 DIOPTER) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. LOW VAULTING WAS OBSERVED ONE WEEK AFTER IMPLANTATION. THE LENS REMAINS IMPLANTED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY ON (B)(6) 2016 WAS 20/99. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6397863&pc=MTA | 2023826-2017-00377 | 2016-12-26 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS REMAIN IMPLANTED. PATIENT'S LAST VISIT ON (B)(6) 2016 UNCORRECTED VISUAL ACUITY IS 20/25. Manufacturer Narrative: PATIENT EXPERIENCE LOW VAULTING AND THE LENS REMAINS IMPLANTED. PER MEDICAL REVIEW, LOW VAULTING WAS OBSERVED ONE WEEK AFTER IMPLANTATION. PER CUSTOMER QUESTIONNAIRE FORM, NO OTHER COMPLICATIONS ASSOCIATED WITH LOW VAULTING WAS OBSERVED. THE SURGEON CONSIDERS AN EXCHANGE FOR A LONGER LENS LENGTH. DESPITE MULTIPLE ATTEMPTS, NO ADDITIONAL INFORMATION OR CONFIRMATION OF LENS REMOVAL OR EXCHANGE WAS RECEIVED. CORRECTED DATA: IS NOT UNKNOWN, CORRECT DATA IS THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6335875&pc=MTA | 2023826-2017-00185 | 2017-01-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.50/4.5/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. EXCESSIVE VAULTING OCCURED WITH A VAULT VALUE OF 1400 MICROMETERS. THE LENS IS PLANNED TO BE EXCHANGED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6335881&pc=MTA | 2023826-2017-00186 | 2017-01-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-16 | MTA | ICL (INTRAOCULAR COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/4.5/80 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. EXCESSIVE VAULTING OCCURRED WITH A VAULT VALUE OF (B)(4) MICROMETERS. THE LENS IS PLANNED TO BE EXCHANGED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE'S WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6308559&pc=MTA | 2023826-2017-00225 | 2017-01-12 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON WAS IMPLANTING A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, WHEN HE NOTICED THERE WAS NOT ENOUGH ROOM IN THE CHAMBER. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND THE SURGERY WAS CANCELLED. Manufacturer Narrative: NO LENS WAS RETURNED FOR EVALUATION, ONLY IOL PACKAGING BOX WAS RETURNED. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447610&pc=MTA | 2023826-2017-00523 | 2017-01-18 | Injury | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/98 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR SHORTER LENS DUE TO EXCSSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISISON, THE LENS WAS LOST, UNABLE TO RETURNED. Manufacturer Narrative: ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SMALL VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6498936&pc=MTA | 2023826-2017-00651 | 2017-02-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+4.0/84 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6585683&pc=MTA | 2023826-2017-00869 | 2017-02-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+3.5/089 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DATE AND DURING THE SAME SURGERY, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND LENS IS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711990&pc=MTA | 2023826-2017-01171 | 2017-05-13 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -14.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DAY, DURING THE SAME SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE DEVICE HAS BEEN RETURNED FOR EVALUATION BUT NOT YET EVALUATED. Manufacturer Narrative: PREVIOUS: LENS HAS BEEN RETURNED AND NOT YET EVALUATED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6708264&pc=MTA | 2023826-2017-01170 | 2017-05-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/084 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DAY AND DURING THE SAME SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, AND CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670148&pc=MTA | 2023826-2017-01043 | 2017-05-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND PIECE OF HAPTIC MISSING, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6690849&pc=MTA | 2023826-2017-01114 | 2017-05-26 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, DURING THE SAME SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER AND SAME DIOPTER LENS, DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: PREVIOUS: LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED.-DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC SLIGHTLY BENT, NO OTHER VISIBLE DAMAGE.(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694738&pc=MTA | 2023826-2017-01122 | 2017-06-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/92 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. IN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE DEVICE WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC SLIGHTLY BENT. THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135215&pc=MTA | 2023826-2017-01918 | 2017-06-15 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021464&pc=MTA | 2023826-2017-01673 | 2017-06-22 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -6.5/+1.0/94 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING DURING THE SAME SURGERY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6769237&pc=MTA | 2023826-2017-01253 | 2017-07-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS REMOVED DURING THE SAME SURGERY DUE TO LOW VAULTING. THERE WAS NO REPORT OF ANY PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS AND THE PATIENT IS DOING VERY WELL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN ITS SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899575&pc=MTA | 2023826-2017-01509 | 2017-07-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 DIOPTER -12.0/+2.0/081 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DATE, INTRAOPERATIVELY, THE LENS WAS EXCHANGED WITH A SMALLER, DIFFERENT DIOPTER LENS DUE TO EXCESSIVE VAULT. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850684&pc=MTA | 2023826-2017-01405 | 2017-07-27 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, SO THE SURGEON EXPLANTED THE LENS INTRAOPERATIVELY, RESOLVING THE PROBLEM. THERE IS NO PLAN TO EXCHANGE THE LENS. Manufacturer Narrative: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: CORRECTED DATA: TYPE OF REPORTABLE EVENT: IT SHOULD BE CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION" IN THE INITIAL AND SUPPLEMENTAL MDR #1. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899572&pc=MTA | 2023826-2017-01506 | 2017-08-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 DIOPTER -10.5/+2.0/091 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DATE, INTRAOPERATIVELY, THE LENS WAS EXCHANGED WITH A SMALLER, DIFFERENT DIOPTER LENS DUE TO EXCESSIVE VAULT. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS EXCHANGED FOR A SMALLER, SIMILAR DIOPTER LENS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BENT AND A PIECE OF THE HAPTIC MISSING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT (20 YRS), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.99MM)]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7033907&pc=MTA | 2023826-2017-01737 | 2017-08-31 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING, WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT WAS SATISFIED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918212&pc=MTA | 2023826-2017-01536 | 2017-09-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/X090 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. CLAIM REPORT STATES AN EXCESSIVE VAULT VALUE OF 1080 UM WAS NOTED, INCIDENT DATE OF (B)(6) 2017. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: IMPLANTED DATE: THE DATE OF IMPLANTATION WAS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2017" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: UPDATED INFORMATION: REPORTEDLY, THE PATIENT IS DOING WELL-NO HIGH IOP, NO ANGLE CLOSURE, PUPIL REACTION IS NORMAL. THE SURGEON HAS DECIDED TO NOT EXCHANGE THE LENS AND WILL MONITOR THE PATIENT'S STATUS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6950066&pc=MTA | 2023826-2017-01573 | 2017-09-10 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), AND NOTED THE LENS HAD A LOW VAULT. THE LENS WAS REMOVED DURING THE SAME OPERATION, WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: TYPE OF REPORTABLE EVENT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION" IN THE PREVIOUS SUPPLEMENTAL. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7848736&pc=MTA | 2023826-2018-01310 | 2017-10-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2, -13.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017 AS A REPLACEMENT OF THE FIRST IMPLANTED LENS TO CORRECT LOW VAUT. THIS EXCHNAGE DID NOT RESOLVE THE PROBLEM. IT WAS REPORTED THAT THE PATIENT'S POSTOPERATIVE ARCH HEIGH WAS 100 MICRONS, SO PROBLEM WAS NOT SOLVED. THE PATIENT IS UNWILLING TO OPERATE AGAIN, PATIENT IS IN GOOD CONDITION. THERE WERE NO OTHER ACCIDENTS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7165904&pc=MTA | 2023826-2018-00026 | 2017-11-03 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND THE PATIENT IS UNDER OBSERVATION IF THE VAULT IS FURTHER COMPROMISED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED ON (B)(6) 2017. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7165962&pc=MTA | 2023826-2018-00021 | 2017-12-05 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+0.5/049 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: CORRECTED DATA: "ADVERSE EVENT" IS CORRECTED TO " PRODUCT PROBLEM". TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7166019&pc=MTA | 2023826-2018-00022 | 2017-12-05 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+0.5/152 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438484&pc=MTA | 2023826-2018-00545 | 2017-12-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9/2/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS INTRAOPERATIVELY REMOVED DUE TO EXCESSIVE VAULTING. ON (B)(6) 2017, A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300917&pc=MTA | 2023826-2018-00278 | 2018-01-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.50/+2.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. INTRA-OPERATIVELY THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF DEBRIS ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647434&pc=MTA | 2023826-2018-00972 | 2018-01-03 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, AND THE LENS HAD EXCESSIVE VAULTING. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENS DURING THE SAME SURGERY WITH NO REPORT OF ANY INJURY. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON THE LENS. CORRECTED DATA: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267522&pc=MTA | 2023826-2018-00211 | 2018-01-07 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/103 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULT; INCIDENT DATE OF (B)(6) 2018. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: SHOULD HAVE BEEN RECORDED IN INITIAL MDR. SURGEON STATES THAT "PATIENT DEVELOPED A VERY HIGH VAULT POSTOPERATIVELY. WE NEED TO EXPLANT THE LENS AND EXCHANGE WITH A 0.5MM SMALLER OVERALL DIAMETER". CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378399&pc=MTA | 2023826-2018-00468 | 2018-01-19 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -11.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL CONTINUE TO BE MONITORED. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378401&pc=MTA | 2023826-2018-00470 | 2018-01-19 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -11.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL CONTINUE TO BE MONITORED. NO ICL EXPLANT/EXCHANGE IS PLANNED AT THIS TIME. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516812&pc=MTA | 2023826-2018-00695 | 2018-01-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENGTH LENS DURING THE SAME SURGERY (INTRAOPERATIVELY) DUE TO OBSERVATION OF EXCESSIVE VAULT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE ON THE LENS SURFACE, LENS HAS A CURVED SHAPE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: "THE PROBLEM WAS RESOLVED." IS ADDED TO THE INITIAL MDR. DATE RETURNED TO MANUFACTURER: 08-MAY-2018. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696914&pc=MTA | 2023826-2018-01047 | 2018-01-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES "STILL PARTLY HIGH VAULT" AND "HIGH VAULT BUT BETTER THAN PREOP". REFERENCE MFR#: 2023826-2018-00654 FOR PREVIOUS LENS EXCHANGE. ATTEMPTS TO OBTAIN PATIENT'S CURRENT STATUS HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.,. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438482&pc=MTA | 2023826-2018-00538 | 2018-02-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -17.50/+1.0/66 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS INTRAOPERATIVELY EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: REPORTER STATED THAT, IN HIS OPINION, THE CAUSE OF THE EVENT IS THAT STAAR RECOMMENDED AN INCORRECT LENS SIZE FOR THIS PATIENT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7514427&pc=MTA | 2023826-2018-00728 | 2018-02-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -03.00 DIOPTER, DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THERE WAS NO REPORT OF PATIENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEFORMED ALONG WITH PRESENCE OF RESIDUE ON THE LENS SURFACE. THE HAPTIC WAS CAUGHT IN THE LENS VIAL CAP. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672646&pc=MTA | 2023826-2019-01047 | 2018-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO LOW VAULTING. A LONGER LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7885150&pc=MTA | 2023826-2018-01399 | 2018-05-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +05.50 DIOPTER, DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN, BENT AND DEFORMED. THERE WAS EVIDENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740039&pc=MTA | 2023826-2018-01149 | 2018-06-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL CONTINUE TO BE MONITORED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: CORRECTED DATA: DIOPTER CORRECTED TO "-8.50" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8076044&pc=MTA | 2023826-2018-01652 | 2018-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/100 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE DOCTOR JUDGED THAT THE SIZE WAS UNSUITABLE. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED ANOTHER SAME MODEL LENS BUT DIFFERENT AXIS WAS IMPLANTED AS EXCHANGE LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7745144&pc=MTA | 2023826-2018-01162 | 2018-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.00/+2.5/139 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT WAS OBSERVED. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED, IT WAS REPORTED THAT THE SURGEON PLANS TO EXCHANGE THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716093&pc=MTA | 2023826-2018-01106 | 2018-06-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/139 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS AN EXCHANGE LENS BUT THE PROBLEM WAS NOT RESOLVED, THE LENS STILL HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT OR EXCHANGE THE LENS. SEE MFR. REPORT 2023826-2018-01098 FOR THE INITIAL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7897643&pc=MTA | 2023826-2018-01429 | 2018-06-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TICM513.2, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: ORIGINAL MDR REPORTED "13.2MM TICM513.2" WHICH SHOULD BE CORRECTED TO "13.2MM VICM5_13.2". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740024&pc=MTA | 2023826-2018-01151 | 2018-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -8.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL CONTINUE TO BE MONITORED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884858&pc=MTA | 2023826-2018-01267 | 2018-07-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -09.0.DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TEAR, HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824061&pc=MTA | 2023826-2018-01276 | 2018-07-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8404826&pc=MTA | 2023826-2019-00354 | 2018-07-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+4.0/100 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7943011&pc=MTA | 2023826-2018-01456 | 2018-08-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENSDURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7857657&pc=MTA | 2023826-2018-01366 | 2018-08-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD A HIGH VAULT . THE LENS REMAINS IMPLANTED AND THE PATIENT IS BEING MONITORED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8072015&pc=MTA | 2023826-2018-01641 | 2018-08-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, WITHIN THE SAME SURGERY, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237904&pc=MTA | 2023826-2019-00079 | 2018-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/098 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002559&pc=MTA | 2023826-2018-01536 | 2018-09-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/171 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THERE WAS NO APPARENT INJURY. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND ONE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073867&pc=MTA | 2023826-2018-01687 | 2018-09-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS), REMOVED THE LENS AND EXCHANGED LENS FOR A LONGER LENGTH LENS WITHIN THE SAME SURGERY ON (B)(6) /2018, DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564655&pc=MTA | 2023826-2019-00691 | 2018-10-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8208777&pc=MTA | 2023826-2019-00006 | 2018-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. INTRAOPERATIVELY, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564643&pc=MTA | 2023826-2019-00687 | 2018-10-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -05.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TEAR. (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510(K) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8065897&pc=MTA | 2023826-2018-01619 | 2018-10-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5¬¬_13.2, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018 AND REMOVED WITHIN THE SAME SURGERY DUE TO OBSERVATION OF EXCESSIVE VAULT. A REPLACEMENT LENS, SAME MODEL SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354939&pc=MTA | 2023826-2019-00295 | 2018-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 12.1MM, VICM5_12.1,-9.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS WITHIN THE SAME SURGERY DUE TO LOW VAULT. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237900&pc=MTA | 2023826-2019-00078 | 2018-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/099 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITHIN THE SAME SURGERY FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227440&pc=MTA | 2023826-2019-00049 | 2018-11-26 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -2.0/+4.0/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE SURGEON STATES THAT THE "SURGERY WAS COMPLEX AND [THE] DOCTOR FINALLY DECIDED [TO] NOT IMPLANT THE ICL. IN THEIR OPINION THE ICL PUSHED THE IRIS IN THE SURGERY BECAUSE THE ICL IS TOO BIG FOR THE EYE". SEE MFR REPORT 2023826-2019-00046 FOR INITIALLY IMPLANTED LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMENSIONAL INSPECTION WAS WITHIN LENS SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497512&pc=MTA | 2023826-2019-00576 | 2018-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO "LENS SIZE ISSUE." AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON THE SAME DATE, IN A SECONDARY SURGERY. THE SURGEON REPORTS "THE FIRST IMPLANTED LENS LOOKED SHORTER THAN [THE] STANDARD SIZE OF 12.6 LENS." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. METHOD CODE ADDED, RESULT CODE UPDATED, CONCLUSION CODE UPDATED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272691&pc=MTA | 2023826-2019-00153 | 2018-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.00/+2.5/143 (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING (THE LENS WAS PUSHING THE IRIS) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON PLANS TO IMPLANT A SHORTER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO DEVICE (CALIPER MALFUNCTION). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION WAS WITHIN LENS SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8401840&pc=MTA | 2023826-2019-00372 | 2018-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM,VICMO13.2, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO OBSERVATION OF EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTICS BENT AND DEFORMED. THERE WAS PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664911&pc=MTA | 2023826-2019-00999 | 2018-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO OBSERVATION OF LOW VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772304&pc=MTA | 2023826-2019-01199 | 2019-01-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -05.50 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8848615&pc=MTA | 2023826-2019-01313 | 2019-01-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8527703&pc=MTA | 2023826-2019-00628 | 2019-02-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON THE SAME DAY DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. PLEASE DISREGARD PREVIOUS DEVICE EVALUATION STATEMENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8523238&pc=MTA | 2023826-2019-00626 | 2019-02-21 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 AS A REPLACEMENT OF THE FIRST IMPLANTED LENS TO CORRECT LOW VAULT. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. IT WAS REPORTED THAT THE SURGEON WILL MONITOR THE PATIENT CLOSELY, NO FURTHER INTERVENTION HAS BEEN DECIDED AS OF RIGHT NOW. THE PATIENTS CONDITION IS FINE, CURRENT BCVA AND UCVA IS 20/20. STATUS OF THE EYE WAS REPORTED AS "OK". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8527912&pc=MTA | 2023826-2019-00629 | 2019-03-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM,VICM5_12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON THE SAME DAY DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8848618&pc=MTA | 2023826-2019-01315 | 2019-03-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652210&pc=MTA | 2023826-2019-00946 | 2019-04-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED WITHIN THE SAME SURGERY FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9158790&pc=MTA | 2023826-2019-01848 | 2019-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, FROM THE PATIENTS RIGHT EYE (OD), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H1- "SERIOUS INJURY" SHOULD BE CORRECTED TO "MALFUNCTION" IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843651&pc=MTA | 2023826-2019-01333 | 2019-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912809&pc=MTA | 2023826-2019-01504 | 2019-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8826245&pc=MTA | 2023826-2019-01296 | 2019-07-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED AND THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: CORRECTED DATA: SERIAL NUMBER: "(B)(4)" SHOULD BE CORRECTED TO "(B)(4)" IN THE INITIAL MDR. A NEW WORK ORDER SEARCH WAS PERFORMED FOR THE CORRECT SERIAL NUMBER. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOTS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8826247&pc=MTA | 2023826-2019-01297 | 2019-07-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: SERIAL NUMBER: "(B)(4)" SHOULD BE CORRECTED TO "(B)(4)" IN THE INITIAL MDR. A NEW WORK ORDER SEARCH WAS PERFORMED FOR THE CORRECT SERIAL NUMBER. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOTS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9148148&pc=MTA | 2023826-2019-01852 | 2019-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8921744&pc=MTA | 2023826-2019-01527 | 2019-07-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED, AND THE LENS REMAINS IMPLANTED. THE REPORTER STATED " THE VAULT WAS 277 MICROMETERS ONE WEEK AFTER SURGERY (B)(6) 2019), BUT AFTER THAT THE VAULT DECREASED." SURGEON CONTINUES TO MONITOR THE PATIENT AND IS CONSIDERING EXPLANTATION OR EXCHANGE DEPENDING ON THE SITUATION." THERE IS NO PROBLEM WITH THE PATIENTS CONDITION SO FAR. CAUSE OF THIS EVENT IS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATE OF BIRTH IN INITIAL MDR. UDI NUMBER CORRECTED TO N/A FROM UNK. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310527&pc=MTA | 2023826-2020-01506 | 2019-07-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. THIS IS THE EXCHANGED LENS, SEE MFR. REPORT # 2023826-2020-01505 FOR INITIAL LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8921550&pc=MTA | 2023826-2019-01469 | 2019-07-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient |  | 1 | Event Description: THE REPORTER INDICATED THAT ONE OF THE SURGEONS PATIENTS IMPLANTED WITH AN ICL ON (B)(6) 2019, "APPEARED TO HAVE A LOWER THAN EXPECTED VAULT IN EACH EYE." THE SURGEON WILL "MONITOR AND WILL BE IN TOUCH AFTER THE PATIENT'S ONE WEEK VISIT." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT: DATE OF EVENT IS UNK. IMPLANT DATE IS (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8921552&pc=MTA | 2023826-2019-01470 | 2019-07-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient |  | 1 | Event Description: THE REPORTER INDICATED THAT ONE OF THE SURGEONS PATIENTS IMPLANTED WITH AN ICL ON (B)(6) 2019, "APPEARED TO HAVE A LOWER THAN EXPECTED VAULT IN EACH EYE." THE SURGEON WILL "MONITOR AND WILL BE IN TOUCH AFTER THE PATIENT'S ONE WEEK VISIT." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT IS UNK. IMPLANT DATE IS (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9447905&pc=MTA | 2023826-2019-02379 | 2019-08-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED ANOTHER SAME MODEL, LENGTH AND DIOPTER LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9148062&pc=MTA | 2023826-2019-01851 | 2019-08-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9145568&pc=MTA | 2023826-2019-01828 | 2019-08-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEFORMED WITH RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842718&pc=MTA | 2023826-2020-00514 | 2019-09-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 - H3: "DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS" SHOULD BE CORRECTED TO "DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS". CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824501&pc=MTA | 2023826-2020-00349 | 2019-09-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. DUE TO LOW VAULT. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS 13.2 WAS STILL SMALL AND WAS REMOVED. PER THE REPORTER ON (B)(6) 2020, PATIENT "NOT READY FOR RE IMPLANTATION AT PRESENT. AND THE PATIENT'S CURRENT CONDITION IS GOOD." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252737&pc=MTA | 2023826-2019-02036 | 2019-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICM5_ 12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO OBSERVATION OF EXCESSIVE VAULT. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957589&pc=MTA | 2023826-2020-00779 | 2019-10-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9618770&pc=MTA | 2023826-2020-00135 | 2019-10-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9826514&pc=MTA | 2023826-2020-00477 | 2019-11-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - "06-APR-2020" SHOULD BE CORRECTED TO "01-06-2020" IN THE PREVIOUS MDR. H1- "SERIOUS INJURY" SHOULD BE CORRECTED TO "MALFUNCTION" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197353&pc=MTA | 2023826-2020-01333 | 2019-12-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -16.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9855339&pc=MTA | 2023826-2020-00499 | 2019-12-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS INSUFFICIENT SIZE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON LENS AND LENS CURVED IN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9589387&pc=MTA | 2023826-2020-00102 | 2019-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE SURGEON WILL MONITOR THE PATIENT FOR ANY CHANGES. Manufacturer Narrative: CORRECTED DATA: B3 - DATE OF EVENT: (B)(6) 2019 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308778&pc=MTA | 2023826-2020-01516 | 2019-12-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2019 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416093&pc=MTA | 2023826-2020-01718 | 2020-01-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416107&pc=MTA | 2023826-2020-01719 | 2020-01-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789262&pc=MTA | 2023826-2020-00415 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9747685&pc=MTA | 2023826-2020-00319 | 2020-01-27 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE PLAN IS TO IMPLANT A LONGER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9747686&pc=MTA | 2023826-2020-00320 | 2020-01-27 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE PLAN IS TO IMPLANT A LONGER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9850111&pc=MTA | 2023826-2020-00570 | 2020-02-19 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308783&pc=MTA | 2023826-2020-01512 | 2020-02-25 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, THE SURGEON IS "PLANNING ON EXCHANGING LENS FOR THE NEXT SIZE UP." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308803&pc=MTA | 2023826-2020-01513 | 2020-02-25 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, THE SURGEON IS "PLANNING ON EXCHANGING LENS FOR THE NEXT SIZE UP." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214888&pc=MTA | 2023826-2020-01354 | 2020-03-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, FROM THE PATIENTS RIGHT EYE (OD) WITHIN THE SAME SURGERY, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310571&pc=MTA | 2023826-2020-01459 | 2020-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140595&pc=MTA | 2023826-2020-01209 | 2020-04-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271252&pc=MTA | 2023826-2020-01476 | 2020-06-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL 12.6, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS SIZING ISSUE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364460&pc=MTA | 2023826-2020-01643 | 2020-06-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364441&pc=MTA | 2023826-2020-01640 | 2020-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CLAIM #: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND DEBRIS ON THE LENS SURFACE AND THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10456859&pc=MTA | 2023826-2020-01898 |  | Malfunction | STAAR SUGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_-9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AS A REPLACEMENT LENS. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. HOWEVER, " CALM STILL LOW VAULT " WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LATER REPORTER INDICATED THAT "VAULTING IS FINE NOW, PATIENT IS GOOD" CAUSE OF THE EVENT WAS A "SIZING ISSUE" AND THE REPLACEMENT LENS RESOLVED THE PROBLEM. CORRECTED DATA: B5 - "VICM5_"SHOULD BE CORRECTED TO "VICM5_13.7" CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136389&pc=MTA | 2023826-2017-01939 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.0/+4.5/002 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THE LENS HAD A HIGH VAULT BUT THE OUTCOME WAS GOOD. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201641&pc=MTA | 2023826-2018-00087 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON REPORTS AN EXCESSIVE VAULT; PLANS TO EXCHANGE THE LENS. NO ADDITIONAL INFORMATION AVAILABLE. Manufacturer Narrative: UPDATED INFORMATION RECEIVED: ON 27-JAN-2018, STAAR SURGICAL RECEIVED NOTIFICATION FROM CLAIM REPORTER THAT THIS IS A DUPLICATE REPORT FOR THE PATIENT INVOLVED IN MFR REPORTS 2023826-2018-00045 AND 2023826-2018-00046 (CLAIM FILE# (B)(4)). CLAIM HAD BEEN REPORTED AS ONE CASE, HOWEVER IS ACTUALLY ONE PATIENT, BOTH EYES AFFECTED AND ALREADY REPORTED. MDR WAS SUBMITTED PRIOR TO DISCOVERY OF DUPLICATE STATUS. CLAIM# (B)(4). Manufacturer Narrative: PMA 510(K): - THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7346748&pc=MTA | 2023826-2018-00421 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7487513&pc=MTA | 2023826-2018-00669 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571623&pc=MTA | 2023826-2018-00855 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN BMC OPTHALMOLOGY TITLED, 'ANALYSIS OF INTRAOCULAR POSITIONS OF POSTERIOR IMPLANTABLE COLLAMER LENS BY FULL-SCALE ULTRASOUND BIOMICROSCOPY.' THE ARTICLE REFERENCES 18 EYES WITH EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571630&pc=MTA | 2023826-2018-00853 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN BMC OPTHALMOLOGY TITLED, 'ANALYSIS OF INTRAOCULAR POSITIONS OF POSTERIOR IMPLANTABLE COLLAMER LENS BY FULL-SCALE ULTRASOUND BIOMICROSCOPY.' THE ARTICLE REFERENCES 29 EYES WITH INSUFFICIENT VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8071358&pc=MTA | 2023826-2018-01678 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-11-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: IT WAS REPORTED THAT A 13.7MM, VTICMO13.7, -17/+2.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT RIGHT EYE (OD) AS A REPLACEMENT LENS DUE TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM (VAULT WAS A BIT LOW, BUT IT WAS THE BIGGEST LENS). ON (B)(6) 2018 IT WAS REPORTED THAT THE PROBLEM RESOLVED, THE PATIENT IS OK AND NO ADDITIONAL TREATMENTS WILL BE PERFORMED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8204316&pc=MTA | 2023826-2018-01946 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: RIGHT EYE (OD): THE REPORTER STATED THAT WHEN THE VAULT OF THE LEFT EYE TURNED OUT TO BE TOO HIGH, THEY CHOSE THE LENGTH OF 12.6MM FOR THE RIGHT EYE (OD), THEN THE VAULT OF THE RIGHT EYE (OD) WAS STILL LOW ENOUGH (1000UM). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BENN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: RIGHT EYE (OD): THE REPORTER STATED THAT, "WHEN THE VAULT OF THE LEFT EYE TURNED OUT TO BE TOO HIGH, THEY CHOSE THE LENGTH OF 12.6MM FOR THE RIGHT EYE (OD), THEN THE VAULT OF THE RIGHT EYE (OD) WAS STILL LOW ENOUGH (1000UM)". THE DIOPTER FOR THE RIGHT EYE (OD) IS -11.00/+0.5/008 (SPHERE/CYLINDER/AXIS). NO OTHER INFORMATION HAS BEEN PROVIDED. CLAIM#: (B)(4). Manufacturer Narrative: NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8591098&pc=MTA | 2023826-2019-00774 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-08 | MTA | UNK | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT EXCESSIVELY HIGH VAULT WAS OBSERVED. PATIENT IS HAPPY WITH VISION BUT DOCTOR IS RELUCTANT TO LEAVE IT THAT WAY. BECAUSE IT IS SO DRAMATICALLY DIFFERENT THAN THE OTHER EYE, DOCTOR SUSPECTS THAT THERE IS AN ISSUE WITH THE HAPTICS CAUGHT ON IRIS OR SOMETHING. PATIENT MISSED HER LAST CHECK UP SO NO MEDICAL THERAPY HAS BEEN TRIED. PATIENT WAS DOING WELL 20/20 AND PRESSURE WAS FINE. SURGEON WAS HOPING TO CONFIRM THE VAULT AT THE 1-MONTH VISIT BUT PATIENT RESCHEDULED FOR JUNE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306670&pc=MTA | 2023826-2019-02053 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2019, "VA IS GOOD, VAULT IS LOW. AFTER DOCTOR EXPLAINED EVERYTHING TO THE PATIENT," SURGEON IS "WILLING TO DO CLOSE FOLLOW UP INSTEAD OF EXPLANTING LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9664317&pc=MTA | 2023826-2020-00231 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). PER THE REPORTER ON (B)(6) 2020, "LENS NEEDS TO BE EXCHANGED FOR A LARGER SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - "THE REPORTER STATED THAT THERE IS NO COMPLAINT AGAINST THE LENS IN THIS EYE", SHOULD HAVE BEEN INCLUDED IN PREVIOUS MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9668830&pc=MTA | 2023826-2020-00241 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.25 DIOPTER INTO THE PATIENT'S LEFT EYE (OS). INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. ON AN UNKNOWN DATE A SHORTER LENS WAS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: PREVIOUS CLAIM # IN H10 IS CORRECTED TO CLAIM# (B)(4). ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBEL DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.25 DIOPTER, IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105256&pc=MTA | 2023826-2020-01082 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | UNK | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S EYE. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. PER THE SURGEON ON (B)(6) 2020, SEEKING CONSULT "ABOUT THE ICL VAULT IN PATIENT WITH A RESIDUAL REFRACTIVE ERROR." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138554&pc=MTA | 2023826-2020-01182 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. "OVERVAULTING" WAS OBSERVED. PER THE SURGEON ON (B)(6) 2020, "THE LENS IS SIMPLY OVER-VAULTED. VISION AND IOP ARE GOOD, BUT THE LONG TERM RISK IS TOO GREAT TO LEAVE THIS LENS IN THE PATIENTS EYE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT A 13.7 MM MICL13.7 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). "OVERVAULTING" WAS OBSERVED. PER THE SURGEON ON (B)(6) 2020, "THE LENS IS SIMPLY OVER-VAULTED. VISION AND IOP ARE GOOD, BUT THE LONG TERM RISK IS TOO GREAT TO LEAVE THIS LENS IN THE PATIENT'S EYE." THE LENS WAS LATER EXPLANTED ON AN UNKNOWN DATE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING." IN THE INITIAL MDR. D9 - DATE RETURNED TO MANUFACTURE CORRECTED TO 15 JUNE 2020. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA B5 - "LENS WAS EXPLANTED ON AN UNKNOWN DATE." SHOULD BE ADDED TO SUPPLEMENTAL MDR#1. ADDITIONAL MANUFACTURE NARRATIVE H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138564&pc=MTA | 2023826-2020-01181 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. "OVERVAULTING" WAS OBSERVED. PER THE SURGEON ON (B)(6) 2020, "THE LENS IS SIMPLY OVER-VAULTED. VISION AND IOP ARE GOOD, BUT THE LONG TERM RISK IS TOO GREAT TO LEAVE THIS LENS IN THE PATIENTS EYE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT A 13.7 MM MICL13.7 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). "OVERVAULTING" WAS OBSERVED. PER THE SURGEON ON (B)(6) 2020, "THE LENS IS SIMPLY OVER-VAULTED. VISION AND IOP ARE GOOD, BUT THE LONG TERM RISK IS TOO GREAT TO LEAVE THIS LENS IN THE PATIENT'S EYE." THE LENS WAS LATER EXPLANTED ON AN UNKNOWN DATE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING." IN THE INITIAL MDR. D9 - DATE RETURNED TO MANUFACTURE CORRECTED TO 15 JUNE 2020. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "LENS WAS EXPLANTED ON AN UNKNOWN DATE." SHOULD BE ADDED TO SUPPLEMENTAL MDR#1. H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184448&pc=MTA | 2023826-2020-01346 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AS A REPLACEMENT LENS TO CORRECT LOW VAULT. LENS DID NOT RESOLVE THE PROBLEM, TOO SHORT. FOR STATUS OF THE EYE THE REPORTER STATED " QUIET EYE BUT STILL LOW VAULT ". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184653&pc=MTA | 2023826-2020-01348 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020 AS A REPLACEMENT LENS TO CORRECT LOW VAULT. LENS DID NOT RESOLVE THE PROBLEM, TOO SHORT FU. FOR STATUS OF THE EYE THE REPORTER STATED " QUIET EYE LOW-MEDIUM VAULT ". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10225639&pc=MTA | 2023826-2020-01384 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-07-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER STATED " AFTER AN UNEVENTFUL IMPLANTATION OF THE ICL IN THE FIRST EYE OF MY PATIENT, TODAY 24 HOURS LATER DESPITE THE FACT THAT THE PATIENT IS NORMOTENSIVE, CORNEA IS CLEAR, THE CENTRAL VAULT IS WAY TOO LARGE." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10236861&pc=MTA | 2023826-2020-01443 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-07-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -13.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AS A REPLACEMENT LENS TO CORRECT LOW VAULT. LENS DID NOT RESOLVE THE PROBLEM ( VA 2020/ IOP 8MMHG, ANGLE OPEN AND VAULT 200 MICRON). THE REPORTER STATED "SURGEON WOULD LIKE TO WAIT AND FOLLOW UP THIS EYE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10368636&pc=MTA | 2023826-2020-01676 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: ON 06 JUL 2020 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN CLINICAL AND EXPERIMENTAL OPTOMETRY, ' HORIZONTAL AND ANTERIOR CHAMBER DIAMETER FOR PHAKIC INTRAOCULAR LENS SIZING.' THE PAPER REPORTED SUBOPTIMAL VAULT IN FOUR EYES AND EXCESSIVE VAULT IN ONE EYE, HOWEVER NO COMPLICATIONS WERE REPORTED IN ANY OF THESE EYES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429073&pc=MTA | 2023826-2020-01836 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457339&pc=MTA | 2023826-2020-01883 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON STATES THERE IS A HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR UNK ENTRY. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4645446&pc=MTA | 2023826-2015-00401 | 2014-12-12 | Injury | STAAR SURGICAL COMPANY | 2015-03-31 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, THE LENS WAS EXCHANGED WITHIN THE SAME PROCEDURE FOR A SMALLER LENS AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST UCVA WAS 20/20 AS OF (B)(6) 2014. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY THE MANUFACTURER: NO: LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5072370&pc=MTA | 2023826-2015-01127 | 2015-02-09 | Injury | STAAR SURGICAL COMPANY | 2015-09-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -13.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND EXPLANTED DUE TO INADEQUATE VAULTING ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP VISIT ON (B)(6) 2015, UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4839165&pc=MTA | 2023826-2015-00673 | 2015-04-14 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT (OD) EYE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTED ASSOCIATED WITH EXCESSIVE VAULT . THE WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4). EVALUATION CODES METHOD (OTHER) - LENS WORK ORDER SEARCH RESULTS (OTHER) A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION (OTHER) BASED ON THE COMPLAINT INFORMATION AND WORK ORDER SEARCH, WE WERE ABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4962360&pc=MTA | 2023826-2015-01009 | 2015-07-10 | Injury | STAAR SURGICAL COMPANY | 2015-08-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE CUSTOMER REPORTED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS UPSIDE DOWN IN THE PATIENT'S RIGHT EYE (OD) AND LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE EVENT WAS DUE TO A USER'S ERROR. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. METHOD 86 - LENS WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954762&pc=MTA | 2023826-2015-01003 | 2015-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED AND REMOVED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS DURING THE SAME SURGERY DUE TO THE LENS WAS TOO BIG (EXCESSIVE VAULT) FOR THE PATIENT'S EYE. THERE WAS NO PATIENT INJURY, NO INCISION ENLARGEMENT OR SUTURES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. THE SURGEON FELT THE EVENT WAS DUE TO A SIZING ISSUE. Manufacturer Narrative: (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - ACCURATE DETERMINATION OF ANTERIOR SEGMENT DIMENSIONS IS KEY TO THE SAFETY OF IMPLANTED ICLS. THERE IS NO INFORMATION AS TO HOW LENS SIZE WAS DETERMINED FOR THIS PATIENT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5667075&pc=MTA | 2023826-2016-00570 | 2015-11-23 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND NOTED THE LENS HAD AN EXCESSIVE VAULT. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. IT WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4772714&pc=MTA | 2023826-2015-00542 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-05-14 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN BOTH OF THE PATIENT'S EYES AND NOTICED ROTATION OF BOTH LENSES AFTER IMPLANTATION. ICL VAULT WAS NOTICED TO BE LOW IN BOTH EYES AND THE ROTATION WAS ASSOCIATED WITH UNEXPECTED REFRACTIVE OUTCOME. THE SURGEON JUDGED THE LENS TO BE OFF-AXIS BY 27 DEGRESS CW IIN THE RIGHT EYE AND 20 DEGRESS CW IN THE LEFT EYE. BOTH LENSES REMAIN IMPLANTED. THE SURGEON IS PLANNING TO EXCHANGE BOTH LENSES. IT IS UNKNOWN IF THE LENSES WILL BE LONGER WITH A DIFFERENT POWER OR ONLY LONGER DUE TO INCREASE VAULTING. SEE MFR# 2023826-2015-00543 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. PT INFORMATION: UNK. EVENT DATE AND IMPLANT DATE: UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: LENS WORK ORDER SEARCH & DEVICE HISTORY REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE DEVICE HISTORY REVIEW DETERMINED THAT NOTHING IN THE MANUFACTURING OF THE LENS COULD HAVE CAUSED THE REPORTED EVENT. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND DEVICE HISTORY REVIEW, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). LENS REMAIN IMPLANTED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: ROTATION WAS ASSOCIATED WITH UNEXPECTED REFRACTIVE OUTCOME. POST-OP REFRACTION WAS SLIGHTLY MYOPIC (SEQ OF -0.25 D IN THE RIGHT EYE AND -0.75 IN THE LEFT EYE) THE SURGEON JUDGED THE LENS TO BE OFF AXIS BY 27 DEGREES CW IN THE RIGHT EYE AND BY 20 DEGREES CW IN LEFT EYE. THE SURGEON IS PLANNING TO EXCHANGE THESE LENSES WITH NEW ONES (IT IS UNKNOWN IF THE NEW LENSES WILL BE LONGER AND WITH DIFFERENT POWER OR ONLY LONGER TO INCREASE VAULTING). LOW VAULT AND ROTATION MAY BE DUE TO SHORT LENSES. THE CAUSE OF THE RESIDUAL REFRACTION MAY BE MULTIFACTORIAL, INCLUDING OFF-AXIS POSITION SECONDARY TO ROTATION, INACCURATE PREOPERATIVE MEASUREMENTS, SURGICALLY INDUCED CHANGES, TILTING, ETC. CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION,ECT). ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6501514&pc=MTA | 2023826-2017-00654 | 2017-03-17 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -14.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT BEGAN TO EXPERIENCE LOW VAULT WITH LENS ROTATION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: SURGEON AND PATIENT HAVE DECIDED TO NOT REMOVE THE LENS. PATIENT WILL VISIT SURGE ON TO MONITOR THE VAULTING. (B)(4). NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6938692&pc=MTA | 2023826-2017-01547 | 2017-09-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -7.5/+4.0/60 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LOW VAULTING AND LENS ROTATION WERE NOTICED. THE SURGEON DECIDED TO LEAVE THE LENS IN THE EYE AND WILL HAVE STRICT MONITORING OF THE VAULT. THE PATIENT'S VISION IS STABLE AND GOOD (20/20). Manufacturer Narrative: METHOD CODE: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTED DATA: DATE OF BIRTH IS CORRECTED FROM (B)(6) TO (B)(6). (B)(4). Manufacturer Narrative: CORRECTED DATA: THE TYPE OF EVENT IS CORRECTED FROM "ADVERSE EVENT" TO "PRODUCT PROBLEM". TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516588&pc=MTA | 2023826-2019-02473 | 2019-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). "SIGNIFICANT INCREASE OF VAULT EARLY-POSTOP." WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B3- SHOULD BE UNKNOWN IN THE INITIAL MDR. B5- PER UPDATED INFO "NO EVENTS" REPORTED FOR LEFT EYE (OS). NO PRODUCT ISSUE. EXCESSIVE VAULT IS FOR RIGHT EYE (OD). H6- DEVICE CODE 1583 SHOUD BE CORRECTED TO "CODE 3189" IN THE INITIAL MDR. H6- CONCLUSION CODE 4315 SHOULD BE CORRECTED TO "CODE 67- NO PROBLEM DETECTED)" IN THE INITIAL MDR. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8447842&pc=MTA | 2023826-2019-00479 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, MICL13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO PATIENT'S EYE ON (B)(6) 2019. REPORTEDLY THE SURGEON "WILL DETERMINE IF THE LENS NEEDS TO BE EXCHANGED DUE TO VAULT". LENS REMAINS IMPLANTED. "THERE ARE NO COMPLICATIONS WITH PATIENT AS OF NOW." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THERE IS CONCERN FOR ASSESSMENT OF VAULTING AND THE SURGEON WILL MONITOR TO DETERMINE IF AN EXCHANGE IS NECESSARY. THE LENS REMAINS IMPLANTED AND THE PATIENT IS DOING WELL. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116178&pc=MTA | 2023826-2020-01155 | 2019-09-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -12.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. WITHIN THE SAME SURGERY THE LENS WAS EXPLANTED DUE TO LOW VAULT. REPORTEDLY, PREOPERATIVE EXAMINATION SHOWED THE PATIENT HAD A LARGE WTW AND ACD MEASUREMENT. THE SURGEON TOLD THE PATIENT THE LENS MAY NOT BE APPROPRIATE. AFTER IMPLANT THE SURGEON FOUND THE VAULT WAS LOW AND THE LENS WAS REMOVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED WITH THE HAPTIC BENT, DEFORMED AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10432997&pc=MTA | 2023826-2020-01801 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VICMO13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT RIGHT EYE (OD) AS A REPLACEMENT LENS DUE TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM (VAULT IS STILL NOT IDEAL). ON (B)(6) 2020 THE REPORTER STATED " THE PATIENT'S CONDITION IS GOOD NOW EXCEPT THE LOW VAULT (ABOUT 200). TO EXPLAIN WHY: EXCHANGED LENS DIDN'T RESOLVE THE PROBLEM, THE DOCTOR SAID MAYBE THE EYES (BOTH OD&OS) OF THE PATIENT'S CILIARY SULCUS MORPHOLOGY WERE NOT GOOD WHICH CAUSED THE LOW VAULT. IN THE END, THE SURGEON DECIDED TO FOLLOW UP THIS CASE." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6414012&pc=MTA | 2023826-2017-00410 | 2017-02-17 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Incomplete or Missing Packaging | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. THE SURGEON FOUND THAT THE SERIAL NUMBER LABEL ON THE LENS CASE ((B)(4)) DIDN'T MATCH THE SERIAL NUMBER LABEL ON THE OUTSIDE OF THE BOX ((B)(4)). LENS REMAINS IMPLANTED, WILL NOT BE RETURNED. REFERENCE MDR 2023826-2017-00411. Manufacturer Narrative: DATE OF EVENT ON INITIAL REPORT IS INCORRECT; CORRECT DATE IS UNKNOWN THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. THE SURGEON FOUND THAT THE SERIAL NUMBER LABEL ON THE LENS CASE ((B)(4)) DID NOT MATCH THE SERIAL NUMBER LABEL ON THE OUTSIDE OF THE BOX ((B)(4)). THE LENS REMAINS IMPLANTED AND WILL NOT BE RETURNED. REFERENCE MDR 2023826-2017-00450. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. RESULT: WORK ORDER SEARCH PERFORMED, ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6425335&pc=MTA | 2023826-2017-00450 | 2017-02-17 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Incomplete or Missing Packaging | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. THE SURGEON FOUND THAT THE SERIAL NUMBER LABEL ON THE LENS CASE ((B)(4)) DIDN'T MATCH THE SERIAL NUMBER LABEL ON THE OUTSIDE OF THE BOX ((B)(4)). LENS REMAINS IMPLANTED, WILL NOT BE RETURNED. REFERENCE MDR 2023826-2017-00410. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH PERFORMED, ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4857482&pc=MTA | 2023826-2015-00704 | 2015-03-31 | Malfunction | STAAR SURGICAL COMPANY | 2015-06-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT THE SURGEON DID NOT LIKE HOW THE LENS LOOKED. THE LENS WAS REMOVED WITH NO INJURY TO THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - IN SPITE OF THREE ATTEMPTS TO GATHER INFORMATION, THERE IS NO INFORMATION TO DETERMINE WHAT HAPPENED DURING SURGERY. IT IS NOT KNOWN IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM OR THE LENS, BUT A WORK-ORDER INJECTION SYSTEMS SEARCH SHOWED NO SIMILAR COMPLAINTS. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6378775&pc=MTA | 2023826-2017-00324 | 2016-11-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON REMOVED AND INSPECTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, PRIOR TO LOADING, AND NOTED THERE WAS SOMETHING WRONG WITH THE LENS. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE SURGEON FELT THE LENS WAS DEFECTIVE BUT WAS UNABLE TO PROVIDE WHAT THE DEFECT WAS. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696908&pc=MTA | 2023826-2018-01045 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/100 (SPHERE/CYLINDER/AXIS) WAS RETURNED DAMAGED WITH NO ADDITIONAL INFORMATION PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE PTIC AND HAPTIC TORN AND THE HAPTIC BENT. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475821&pc=MTA | 2023826-2019-00533 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-04-02 | MTA | UNK | Insufficient Information | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE INT J OPHTHALMOL IN FEBRUARY 2019 TITLED, 'VISUAL PERFORMANCE AFTER IMPLANTATION OF TWO TYPES OF PHAKIC FOLDABLE INTRAOCULAR LENSES FOR CORRECTION OF HIGH MYOPIA.' THE ARTICLE STATES THAT ONE EYE OF THE ICL V4C GROUP HAD FINE ANTERIOR SUBCAPSULAR LENS OPACITY WHICH WAS CAUSED BY INADVERTENT IATROGENIC TOUCH OF CRYSTALLINE LENS DURING IMPLANTATION. IT WAS PRESENTED DURING EARLY FOLLOW UP AND REMAINED STABLE THROUGHOUT THE WHOLE FOLLOW UP PERIOD WITHOUT EFFECT ON CDVA" ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475842&pc=MTA | 2023826-2019-00532 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-04-02 | MTA | UNK | Insufficient Information | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE INT J OPHTHALMOL IN FEBRUARY 2019 TITLED, 'VISUAL PERFORMANCE AFTER IMPLANTATION OF TWO TYPES OF PHAKIC FOLDABLE INTRAOCULAR LENSES FOR CORRECTION OF HIGH MYOPIA.' THE ARTICLE STATES THAT PIGMENT DISPERSION OCCURRED IN FOUR EYES OF THE ICL V4C GROUP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965297&pc=MTA | 2023826-2019-01561 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE IN 2019. WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY THERE IS A POSSIBILITY THE LENS MAY BE EXPLANTED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1426. CLAIM#: (B)(4). Manufacturer Narrative: EVENT DATE: 2019. IMPLANT DATE: 2019. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965305&pc=MTA | 2023826-2019-01563 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE IN 2019. WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY THERE IS A POSSIBILITY THE LENS MAY BE EXPLANTED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1426. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: 2019. IMPLANTED: 2019. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965308&pc=MTA | 2023826-2019-01564 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE IN 2019. WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY THERE IS A POSSIBILITY THE LENS MAY BE EXPLANTED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 3189 IS NOT APPLICABLE IN INITIAL MDR. DEVICE CODE 1426 SHOULD HAVE BEEN REPORTED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DIOPTER CORRECTED TO -4.0 IN INITIAL MDR. RESULT CODE 213 CORRECTED TO RESULT CODE 114 IN INITIAL MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY UNK. RACE: UNK. DATE OF EVENT: 2019. IMPLANTED: 2019. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965311&pc=MTA | 2023826-2019-01565 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -04.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE IN 2019. WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY THERE IS A POSSIBILITY THE LENS MAY BE EXPLANTED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: RESULT CODE 213 IN ORIGINAL MDR IS CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 3189 IS CORRECTED TO 1426 FOR OPACITY ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4820212&pc=MTA | 2023826-2015-00641 | 2015-04-27 | Malfunction | STAAR SURGICAL COMPANY | 2015-06-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Material Twisted/Bent; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS BUT THE LENS APPEARED TWISTED AND WAS NOT POSITIONED PROPERLY IN THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: (B)(4). RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER WERE NOT RETURNED FOR EVALUATION. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS COMPLAINT FILE STATES THAT THE LENS APPEARED TWISTED, AND WAS NOT PROPERLY POSITIONED IN THE EYE. THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, PRODUCT EVALUATION AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6864973&pc=MTA | 2023826-2017-01458 | 2017-08-17 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Mechanical Jam | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 AS THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0, INTO THE PATIENT'S LEFT (OS) EYE, IT GOT STUCK TO THE PLUNGER AND THE LENS BROKE. THERE WAS NO PATIENT CONTACT WITH THE LENS; BACK-UP LENS INSERTED DURING THE SAME SURGERY. Manufacturer Narrative: ATTEMPTED IMPLANT DATE SHOULD BE (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SPECIMIN CUP. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6159876&pc=MTA | 2023826-2016-01746 | 2016-09-21 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Mechanical Jam; Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: RECEIVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, FROM THE CUSTOMER STUCK INSIDE AN SFC-45 FP CARTRIDGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE REPORTER STATED THE LENS WAS LOADED AND PREPARED FOR SURGERY BUT WAS NOT USED DUE TO PATIENT RELATED MEDICAL ISSUE. THERE WAS NO PATIENT CONTACT AND THE SURGERY HAD NOT BEGUN/ABORTED UNTIL A LATER DATE, WHEN ANOTHER ICL LENS WILL BE IMPLANTED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER?: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED IN AN SFC-45 FP CARTRIDGE. THERE WAS NO VISIBLE DAMAGE TO THE CARTRIDGE. THE WAS EVIDENCE OF CLEAR DRY RESIDUE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6323331&pc=MTA | 2023826-2016-01643 | 2016-07-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-10 05:00 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) . THE SURGEON NOTICED A CRACK IN THE LENS; THE DEVICE WAS NOT USED. THE LENS WAS EXCHANGED FOR AN IDENTICAL MODEL AND DIOPTER. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6513916&pc=MTA | 2023826-2016-01667 | 2016-10-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-22 04:00 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS TEAR DURING DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED FOR AN IDENTICAL MODEL AND DIOPTER. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716250&pc=MTA | 2023826-2018-01102 | 2018-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO APPARENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7799788&pc=MTA | 2023826-2018-01250 | 2018-02-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON PREPARED TO IMPLANT A 12.6MM VICM5_12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. BEFORE/DURING LOADING, THE LENS TORE/BROKE. THE LENS WAS NOT IMPLANTED. INTRAOPERATIVELY AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED STUCK IN SFC-45 CARTRIDGE. CARTRIDGE RETURNED IN PLASTIC BAG WITH CLEAR SURGICAL RESIDUE ON PRODUCT AND NO VISIBLE DAMAGE TO DEVICE. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721260&pc=MTA | 2023826-2018-01089 | 2018-03-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -15.5 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. INTRAOPERATIVELY, THE LENS WAS REMOVED. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783475&pc=MTA | 2023826-2018-01209 | 2018-03-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2018. REPORTEDLY, ON (B)(6) 2018 AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE REPORTER STATES "THE DEHYDRATION OF THE FOAM TIP" IS A POSSIBLE CAUSE OF THE LENS TEAR. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. THERE WAS RESIDUE ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716197&pc=MTA | 2023826-2018-01103 | 2018-03-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/087 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO APPARENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2018. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906909&pc=MTA | 2023826-2018-01442 | 2018-03-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S EYE ON (B)(6) 2018. INTRAOPERATIVELY, THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS. THE REASON FOR THE EXCHANGE IS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720848&pc=MTA | 2023826-2018-01125 | 2018-03-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/097 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO APPARENT INJURY. A DIFFERENT MODEL LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720688&pc=MTA | 2023826-2018-01113 | 2018-03-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. IT WAS UNKNOWN IF THERE WAS ANY PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720866&pc=MTA | 2023826-2018-01126 | 2018-04-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/096 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. IT IS UNKNOWN IF THERE WAS ANY PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696909&pc=MTA | 2023826-2018-01033 | 2018-04-24 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -17.5 INTO THE PATIENT'S LEFT EYE (OS), THE LENS BROKE. THIS OCCURRED ON (B)(6) 2018. REPORTEDLY, THE LENS WAS NOT IMPLANTED. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7716219&pc=MTA | 2023826-2018-01104 | 2018-05-14 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO APPARENT INJURY. THE SURGEON DID NOT IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721236&pc=MTA | 2023826-2018-01088 | 2018-05-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. INTRAOPERATIVELY, AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND THE LENS HAS DEBRIS AND RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720926&pc=MTA | 2023826-2018-01115 | 2018-06-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT, AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783504&pc=MTA | 2023826-2018-01205 | 2018-06-25 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS REMOVED INTRA-OPERATIVELY. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE REPORTER STATES THAT IT'S POSSIBLE THE "DEHYDRATION OF THE FOAM TIP" CAUSED THE TEAR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824537&pc=MTA | 2023826-2018-01258 | 2018-07-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED DURING THE SAME SURGERY WITH NO APPARENT INJURY. ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THE LENS WAS ALSO RETURNED IN TWO PIECES AND STUCK TOGETHER. THERE WAS RESIDUE ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853019&pc=MTA | 2023826-2018-01339 | 2018-08-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, AFTER IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S LEFT EYE (OS), THE SURGEON DISCOVERED THAT THE LENS HAD BEEN BROKEN BEFORE/DURING LOADING. THE LENS WAS IMPLANTED ON (B)(6) 2018 AND WAS REMOVED INTRAOPERATIVELY. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. THE SURGEON STATES THE LENS WAS BROKEN IN THE CENTER OF THE OPTICAL ZONE. REPORTEDLY, SHE REVIEWED THE SURGICAL VIDEO THAT SHOWS THE LENS WAS BROKEN PRIOR TO REMOVAL FROM THE VIAL. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8081331&pc=MTA | 2023826-2018-01704 | 2018-09-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+3.5/149 (SPHERE/CYLINDER/AXIS), WAS TORN/BROKEN BEFORE INJECTION INTO THE EYE. THERE WAS NO REPORT OF PATIENT CONTACT OR INJURY. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL OBSERVATION FOUND THE OPTIC TORN AND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254833&pc=MTA | 2023826-2019-00061 | 2018-10-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -16.5 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED ON (B)(6) 2018 AND WAS REMOVED AND REPLACED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS USER ERROR. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8158154&pc=MTA | 2023826-2018-01795 | 2018-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS WAS DAMAGED DURING THE LOADING PROCESS. ANOTHER LENS (13.2MM) WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8332643&pc=MTA | 2023826-2019-00258 | 2018-12-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16/1.5/14 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8591894&pc=MTA | 2023826-2019-00424 | 2019-01-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICM5 13.2, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. DAMAGE WAS NOTED WHILE REMOVING FROM VIAL. NO PATIENT CONTACT. REPLACEMENT LENS IMPLANTED ON (B)(6) 2019, "WITHOUT PROBLEMS." "PATIENT IS OK." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. LENS RETURNED IN TWO PIECES. CLAIM#: (B)(4). Manufacturer Narrative: RESULT CODE 114: CODE SHOULD BE ADDED TO PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699734&pc=MTA | 2023826-2019-01053 | 2019-01-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO 12.1, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. DAMAGE WAS NOTED WHILE REMOVING LENS FROM VIAL. NO PATIENT CONTACT. REPLACEMENT LENS IMPLANTED INTO PATIENTS LEFT EYE (OS) DURING INITIAL SURGERY ON (B)(6) 2019, PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECITION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8544924&pc=MTA | 2023826-2019-00596 | 2019-03-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO13.7, -17.00 DIOPTER, IMPLANTABLE COLLARMER LENS TORE DURING LOADING. DAMAGE WAS NOTED AFTER OPENING VIAL, WHEN PULLING LENS INTO CARTRIDGE. CAUSE OF THE EVENT IS REPORTED AS "LENS FELT SOFT/TORE WHILE PULLING INTO CARTRIDGE.' Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED DRY IN A CARTRIDGE. VISUAL INSPECTION FOUND LENS RETURNED IN CARTRIDGE. THE LENS WAS REMOVED FROM THE CARTRIDGE TO VERIFY THE COMPLAINT DESCRIPTION. LENS IS CURLED UP AND UNABLE TO EVALUATE. CARTRIDGE WAS RETURNED WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO DEVICE AND RESDUE ON DEVICE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843638&pc=MTA | 2023826-2019-01331 | 2019-04-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00 DIOPTER, WAS TORN/BROKEN AFTER OPENING THE VIAL. ANOTHER SAME MODEL/LENGTH/DIOPTER LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS TORN DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL,LENGTH,DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANTED: NA. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692878&pc=MTA | 2023826-2019-01069 | 2019-04-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. INTRAOPERATIVELY, THE LENS WAS REMOVED AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. REPORTEDLY, NO PATIENT INJURY OCCURRED. Manufacturer Narrative: B2-ITEM PREVIOUSLY CHECKED IN ERROR. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE OPTIC WAS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8755969&pc=MTA | 2023826-2019-01164 | 2019-04-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, INTO THE CARTRIDGE AND FOUND A FOLD IN THE LENS. THE SURGEON DISCONTINUED USING THAT LENS AND THERE WAS NO PATIENT CONTACT. A NEW LENS WAS USED/IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8755977&pc=MTA | 2023826-2019-01165 | 2019-04-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, INTO THE CARTRIDGE AND FOUND A FOLD IN THE LENS. THE SURGEON DISCONTINUED USING THAT LENS AND THERE WAS NO PATIENT CONTACT. A NEW LENS WAS USED/IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671663&pc=MTA | 2023826-2019-01014 | 2019-05-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER TORE/BROKE BEFORE/DURING LOADING. THE LENS WAS NOT IMPLANTED BUT IT DID CONTACT THE PATIENT'S LEFT EYE (OS). THIS OCCURRED ON (B)(6) 2019. Manufacturer Narrative: B5 - ADDITIONAL INFORMATION: A REPLACEMENT LENS OF THE SAME PARAMETERS WAS IMPLANTED ON AN UNSPECIFIED DATE WHICH RESOLVED THE PROBLEM. H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE OPTIC AND THE HAPTIC TORN WITH RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125504&pc=MTA | 2023826-2019-01767 | 2019-05-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON 19/MAY/2019 A VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED DURING INITIAL SURGERY. A REPLACEMENT LENS WAS IMPLANTED ON 07/JUN/2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPLANTED DATE: (B)(6) 2019. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429127&pc=MTA | 2023826-2020-01829 | 2019-07-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE ATTEMPTING TO LOAD A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0, THE LENS TORE/BROKE. THE DAMAGE WAS NOTED AFTER OPENING THE VIAL. THIS OCCURRED ON (B)(6) 2019. AT A LATER DATE AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8910148&pc=MTA | 2023826-2019-01449 | 2019-07-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO 13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. A BACK-UP LENS WAS IMPLANTED DURING INITIAL LENS AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9088385&pc=MTA | 2023826-2019-01708 | 2019-08-08 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VICM5 12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, ON (B)(6) 2019. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019. PER REPORTER, "SECOND LENS IMPLANTATION WAS SMOOTH AND PATIENT IS DOING WELL AND HAPPY." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC TORN AND THE LENS TORN IN TWO PIECES STUCK TOGETHER. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9221391&pc=MTA | 2023826-2019-01972 | 2019-09-06 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-22 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, BEFORE/DURING LOADING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5, THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. THE LENS TEAR/BREAK WAS NOTED AFTER OPENING THE VIAL. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782005&pc=MTA | 2023826-2020-00342 | 2019-10-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICM5 13.2, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. DAMAGE WAS NOTED DURING INJECTION INTO PATIENT'S RIGHT EYE (OS). THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: CORRECTED DATA: B3-DATE OF EVENT IS CORRECTED TO (B)(6) 2019. B4-DATE OF THIS REPORT IS CORRECTED TO(B)(6) 2020. B5:THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. DAMAGE WAS NOTED DURING INJECTION INTO PATIENT'S LEFT EYE (OS). THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2019. D6-IMPLANT DATE: (B)(6) 2019. D7-EXPLANT DATE: (B)(6) 2019. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559367&pc=MTA | 2023826-2020-00037 | 2019-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER GOT STUCK IN THE INJECTOR. WHEN REMOVING THE LENS, THE LENS WAS TORN INTO PIECES. ANOTHER LENS WAS IMPLANTED. THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9806038&pc=MTA | 2023826-2020-00445 | 2020-01-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: A 13.7MM, MICL13.7, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS RETURNED BACK TO US DAMAGE. " NOT USED" WAS HANDWRITTEN ON THE LENS PACKAGING BOX. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: D10- LENS RETURNED TO MANUFACTURER DATE 24-FEB-2020 IN INITIAL MDR CORRECTED TO 30-JAN-2020. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9741495&pc=MTA | 2023826-2020-00338 | 2020-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND A REPLACEMENT LENS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE.THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10032661&pc=MTA | 2023826-2020-01024 | 2020-02-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO 13.2, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING ON (B)(6) 2020. DAMAGE WAS NOTED AFTER OPENING VIAL. NO PATIENT CONTACT. CAUSE OF EVENT IS REPORTED AS DEVICE. PER REPORTER, "PROBABLY THE LENS WAS BROKEN IN THE CARTRIDGE (EXPERIENCED ICL SURGEON). QUIET EYE, SECOND SURGEON IN THE FUTURE PLANNED." Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CARTRIDGE MODEL-SFC-45; LOT# 1442149 AND FOAM TIP PLUNGER MODEL-FTP; LOT#- 1434034 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - CARTRIDGE WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9875292&pc=MTA | 2023826-2020-00638 | 2020-02-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT UPON PULLING A 12.6MM, MICL12.6, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS THROUGH THE CARTRIDGE IT APPEARED THAT THE SHAPE OF THE "LEADING EYELET WAS NOT SMOOTH." IT WAS REPORTED THAT IT WAS NOTED AS THE LENS WAS BEING INJECTED INTO THE EYE, SO THE LENS AND CARTRIDGE WERE REMOVED FROM THE EYE AND THE BACKUP LENS WAS USED. DOCTOR WAS CONCERN THAT THE ROUGH EDGE COULD RUB AGAINST POSTERIOR IRIS TISSUE. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: D10-RETURNED TO MANUFACTURER DATE 12-MAR-2020 IN INITIAL MDR IS CORRECTED TO 07-MAR-2020. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271247&pc=MTA | 2023826-2020-01491 | 2020-04-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THERE WAS NO PATIENT CONTACT. DURING THE INITIAL SURGERY ON (B)(6) 2020, A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184782&pc=MTA | 2023826-2020-01322 | 2020-04-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICM5 12.1, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING ON (B)(6) 2020. DAMAGE WAS NOTED AFTER OPENING THE VIAL. A LONGER LENGTH REPLACEMENT LENS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10463784&pc=MTA | 2023826-2020-01916 | 2020-05-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED AND EXCHANGED WITHIN THE SAME SURGERY. THE REPORTER STATED "COULD NOT HAVE EXPECTED EYESIGHT. IT IS POSSIBLE THAT LENS HAD STAIN." IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM, " EYESIGHT WAS IMPROVED AS EXPECTED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LATER IT WAS REPORTED THAT THE DOCTOR TOOK EARLY ACTION BECAUSE HE FOUND THE STAIN OR SCRATCH ON THE SURFACE OF THE LENS; AND MAY HAVE NOTICED WHEN IMPLANTING INTO THE PATIENT'S EYE. THE STAIN IS NOT VISIBLE WITHOUT MAGNIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS, ALSO THE LENS WAS SCRATCHED WITH A TEAR ON OPTIC. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457347&pc=MTA | 2023826-2020-01882 | 2020-05-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE. INTRAOPERATIVELY THE LENS WAS EXCHANGED DUE TO A "SLIGHT SCRATCH" BEING FOUND ON THE LENS NEAR THE OPTICAL PART. THE PROBLEM WAS RESOLVED. THIS OCCURRED ON (B)(6) 2020. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND FOREIGN RESIDUE AND THE LENS HAS A CURVED SHAPE. ON 2X A SMALL SCRATCH WAS NOTED NEAR THE OPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10404418&pc=MTA | 2023826-2020-01706 | 2020-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE A 12.6MM MICL12.6, -11.5 DIOPTER, IMPLANTABLE COLANDER LENS, MEANT FOR THE PATIENT'S EYE ON (B)(6) 2020 HAD "A BROWN SPOT ON THE LENS. LENS WAS NEVER INSERTED, BACK-UP LENS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, PATIENT IS "RECOVERING AT A NORMAL RATE." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. FOREIGN MATERIAL ON LENS SURFACE. LENS WAS OBSERVED @ 2X NO BROWN SPOT WAS NOTED. CLAIM#(B)(4). Manufacturer Narrative: CORRECTED DATA: B4: IN INITIAL MDR SHOULD BE CORRECTED TO 24-JUL-2020. B5: THE STATEMENT " BACKUP LENS IMPLANTED" IS NOT APPLICABLE. G4: IN INITIAL MDR SHOULD BE CORRECTED TO 24-JUL-2020. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10484165&pc=MTA | 2023826-2020-02003 | 2020-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE BEFORE/DURING LOADING. IT WAS REPORTED THAT THE DAMAGE WAS NOTED "AFTER INJECTED". PATIENT CONTACT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS OF -18.00 DIOPTER WAS BROKEN WHILE LOADING IT ON 22-07-2020. REPORTER BELIEVES "IT WAS WHILE GRASPING THE LENS WITH FRONT LOADING FORCEPS." THERE WAS NO PATIENT CONTACT. CORRECTED DATA: H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783534&pc=MTA | 2023826-2018-01212 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: TORN LENS RECEIVED RETURNED. NO INFORMATION AVAILABLE REGARDING LENS DETAILS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: TORN LENS RECEIVED RETURNED. UNKNOWN IF THERE WAS PATIENT CONTACT OR PATIENT INJURY. NO INFORMATION AVAILABLE REGARDING LENS DETAIL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172168&pc=MTA | 2023826-2018-01870 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-17 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: A VICMO13.2, 13.2MM IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 WAS RETURNED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: MFR REPORT INITIATED IN ERROR. CURRENT LENS IS ASSOCIATED WITH MFR REPORT# 2023826-2018-01853. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: PLEASE DISREGARD PREVIOUS STATEMENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8184770&pc=MTA | 2023826-2018-01882 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: IT WAS REPORTED THAT A TORIC IMPLANTABLE COLLAMER LENS CAME OUT DAMAGE FROM THE CARTRIDGE IN THE LOADING PROCESS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURER DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8563075&pc=MTA | 2023826-2019-00710 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-04-29 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: A DAMAGED SPHERICAL IMPLANTABLE COLLAMER LENS WAS RETURNED, INCORRECTLY LABELED AS A TORIC LENS. ATTEMPTS TO OBTAIN INFORMATION REGARDING THE REASON FOR THE RETURN OR CORRECT IDENTIFICATION OF THE LENS HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136936&pc=MTA | 2023826-2019-01815 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-30 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. REPORTEDLY, THE LENS MAY HAVE BEEN IMPLANTED UPSIDE DOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, AND RACE: UNK. DATE OF EVENT UNK. PROCODE: UNK. MODEL #, IMPLANT DATE: UNK. DEVICE MANUFACTURER: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9843168&pc=MTA | 2023826-2020-00572 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS STUCK IN INJECTOR DURING SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9873680&pc=MTA | 2023826-2020-00633 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-03-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THERE WAS DIFFICULTY WITH THE LENS. THE REPORTER STATED " I UNDERSTAND IT WAS DUE TO DIFFICULTY INSERTING IT INTO THE EYE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: MODEL, SERIAL AND EXPIRATION DATE SHOULD BE REMOVED AS IT IS N/A. DEVICE MANUFACTURER DATE SHOULD BE REMOVED AS IT IS N/A. METHOD CODE: 4110 SHOULD BE REMOVED. RESULT CODE : 213 SHOULD BE REMOVED. WORK ORDER SEARCH INFORMATION SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10339263&pc=MTA | 2023826-2020-01603 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-07-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: RECEIVED A 12.6MM, MICL12.6,-11.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN CARTRIDGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS WAS RETURNED STUCK IN CARTRIDGE. NO VISIBLE DAMAGE TO DEVICE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - CARTRIDGE EVALUATION: NO VISIBLE DAMAGE FOUND TO THE DEVICE. SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10368836&pc=MTA | 2023826-2020-01677 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON HAS TO CHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427095&pc=MTA | 2023826-2019-02298 |  | Malfunction | STAAR SURGICAL COMPANY AG | 2019-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM, ICM125V4, -07.0 DIOPTER LENS INTO PATIENT'S EYE IN 2010. PER REPORTER, SURGEON WANTS TO REMOVE THE SPHERICAL LENS AND EXCHANGE WITH A TORIC LENS. PER THE REPORTER, A SAIC (INTRACORNEAL RING SEGMENT) HAD ALREADY BEEN PLACED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8484801&pc=MTA | 2023826-2019-00537 |  | Malfunction | Unknown Manufacturer | 2019-04-04 | MTA | UNK | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THEY WOULD BE RETURNING A DAMAGED LENS. PATIENT CONTACT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9747629&pc=MTA | 2023826-2020-00318 | 2020-01-22 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Ejection Problem | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED WHILE THE SURGEON WAS INSERTING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, - 8.0 DIOPTER, THE FOAM TIP PLUNGER OVERRODE THE LENS AND WHILE TRYING TO DISENGAGE, THE LENS WAS RIPPED. THERE WAS PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN LENS VIAL, WLITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122172&pc=MTA | 2023826-2018-01747 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Insufficient Information | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: PLANNING TO DO LEFT (OS) EYE WITH ONE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE CORRECTED TO: "3189" IN THE INITIAL MDR, PREVIOUS CODE NOT APPLICABLE. RESULT CODE CORRECTED TO: "3221" IN THE INITIAL MDR , PREVIOUS CODE NOT APPLICABLE. CONCLUSION CODE CORRECTED TO: "67" IN THE INITIAL MDR, PREVIOUS CODE NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308809&pc=MTA | 2023826-2020-01528 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Insufficient Information | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: A IMPLANTABLE COLLAMER LENS WAS RETURNED WITH ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED WITH MOISTURE, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A DAMAGE LENS ( TORN OPTIC AND HAPTIC). CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - A IMPLANTABLE COLLAMER LENS WAS RETURNED WITH DAMAGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NO INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. G4 - DATE IN INITIAL MEDWATCH REPORT SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813858&pc=MTA | 2023826-2019-01288 | 2019-01-15 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT. REFERENCE MFR REPORT# 2023826-2019-01287 FOR INITIAL IMPLANT. REPORTEDLY, THE LENS REMAINS IMPLANTED, THE PATIENT IS STABLE, AND THE SURGEON IS OBSERVING THE SITUATION. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652108&pc=MTA | 2023826-2019-00920 | 2019-03-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. INTRA-OPERATIVELY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8930921&pc=MTA | 2023826-2019-01462 | 2019-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS; RESIDUE ON LENS. DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8937678&pc=MTA | 2023826-2019-01551 | 2019-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS IN THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912820&pc=MTA | 2023826-2019-01503 | 2019-07-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189451&pc=MTA | 2023826-2020-01286 | 2019-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL. CONDITION OF LENS WAS DRY AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9850090&pc=MTA | 2023826-2020-00567 | 2019-11-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALULATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559333&pc=MTA | 2023826-2020-00045 | 2019-12-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: THE REPORTER STATES, "THE VAULT IN THE OD SEEMS TO BE LESS THAN BEFORE." AS OF 21-JAN-2020 THE VAULT IS AT MORE THAN 1000 MICRONS. THE IRIDO-CORNEAL ANGLE IS LESS OPEN COMPARED TO OS BUT IS NOT CLOSED. SUBJECTIVELY THE PATIENT IS "FEELING VERY HAPPY." UCVA AND BCVA IS NOTED AS 20/25. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842751&pc=MTA | 2023826-2020-00513 | 2020-01-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT. LENS WAS REPLACED INTRAOPERATIVELY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094017&pc=MTA | 2023826-2020-01126 | 2020-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -8.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093996&pc=MTA | 2023826-2020-01127 | 2020-04-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457335&pc=MTA | 2023826-2020-01891 | 2020-04-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843599&pc=MTA | 2023826-2019-01337 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-07-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. SEE MFR REPORT# 2023826-2019-01016 FOR PREVIOUSLY IMPLANTED LENS. Manufacturer Narrative: IN ORIGINAL MDR THE REPORTER INDICATED THAT THE EXCHANGE LENS STILL HAD LOW VAULT; HOWEVER, VIA EMAIL AN UPDATED COMPLAINT QUESTIONNAIRE INDICATED THAT THE EXCHANGE RESOLVED THE PROBLEM. THIS IS NOT A COMPLAINT AGAINST THE DEVICE. DEVICE CODE 1583: VAULTING IN ORIGINAL MDR IS NOT APPLICABLE. EXCHANGE SOLVED THE PROBLEM. DEVICE CODE 1494: OFF LABEL (ACD <3.00MM) IN ORIGINAL MDR IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587404&pc=MTA | 2023826-2020-00115 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -12.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE SURGEON INTENDS TO CONTINUE TO FOLLOW UP WITH THE PATIENT. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308820&pc=MTA | 2023826-2020-01524 | 2019-06-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. DURING THE INITIAL SURGERY ON (B)(6) 2019, A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164363&pc=MTA | 2023826-2019-01868 | 2019-07-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER TORE/BROKE DURING LOADING THE LENS INTO THE CARTRIDGE. THIS OCCURRED ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON THE SAME DATE AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: CORRECTED DATA: CONCOMITANT MEDICAL PRODUCT: INJECTOR MODEL-MSI-PF; LOT#-1439800 SHOULD BE CORRECTED TO INJECTOR MODEL-MSI-PF; LOT#-UNK AS IT IS NOT APPLICABLE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941302&pc=MTA | 2023826-2019-01583 | 2019-07-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER TORE/BROKE DURING LOADING THE LENS INTO THE NOZZLE. THIS OCCURRED ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON THE SAME DATE AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND HAPTIC BENT AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910951&pc=MTA | 2023826-2020-00679 | 2020-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS IMPLANTED UPSIDE DOWN. INTRA-OPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS AND THE PROBLEM WAS RESOLVED. THE SURGEON SUSPECTS THERE WAS A PROBLEM DURING LOADING. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM#(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870925&pc=MTA | 2023826-2019-01404 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Packaging Problem; Inaccurate Information | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE FOR A PACKAGED 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5, THE PACKAGE WAS LABELED WITH A DIFFERENT SERIAL NUMBER THAN THE PATIENT LABELS INCLUDED INSIDE THE BOX. NOTE-BOTH SERIAL NUMBERS WERE WITHIN THE SAME WORK ORDER, SAME MODEL, SIZE, AND POWER. REPORTEDLY, THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: DEVICE CODE 3007 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. RESULT CODE 111 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED INTACT. THE BOX WAS NOT OPENED. VISUAL INSPECTION FOUND THAT THE LENS PACKAGE WAS RETURNED UNOPENED AND THAT THE CHART LABEL SERIAL NUMBER DOES NOT MATCH THE BOX SERIAL NUMBER. (B)(4). Manufacturer Narrative: NO PATIENT INVOLVEMENT. DATE OF EVENT: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870929&pc=MTA | 2023826-2019-01403 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Packaging Problem; Inaccurate Information | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S EYE ON (B)(6) 2019. AFTER IMPLANT IT WAS DISCOVERED THAT THE PACKAGE WAS LABELED WITH A DIFFERENT SERIAL NUMBER THAN THE PATIENT LABELS INCLUDED INSIDE THE BOX. NOTE-BOTH SERIAL NUMBERS WERE WITHIN THE SAME WORK ORDER, SAME MODEL, SIZE, AND POWER. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE CODE 3007 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. RESULT CODE 111 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9747623&pc=MTA | 2023826-2020-00317 | 2020-01-15 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER , IN THE PATIENTS LEFT EYE (OS). THE SURGEON WAS UNABLE TO POSITION THE LENS PROPERLY AND WAS UNABLE TO GET THE LENS INTO THE SULCUS. THIS WAS DUE TO PUPILLARY DILATION HAD BECOME INADEQUATE AND WAS NOT RESPONDING TO INTRAOCULAR MEDS. THE LENS WAS REMOVED AND THE PATIENT WAS SENT BACK TO THE PRE-OP AREA FOR MORE DROPS TO DILATE THE PUPIL. AFTER 45 MINUTES, THIS WAS SUCCESSFUL AND THE PATIENT WAS BROUGHT BACK TO THE OR. THE BACKUP LENS WAS IMPLANTED SUCCESSFULLY. THE PATIENT IS HAPPY. THE SURGEON INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT ANATOMY AND IRIS CONFIGURATION. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6262593&pc=MTA | 2023826-2017-00093 | 2016-07-13 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, AND THE LENS FLIPPED IN THE EYE. THE LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN MINIMAL LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND A PIECE OF HAPTIC MISSING. METHOD: WORK ORDER SEARCH. RESULT: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6764487&pc=MTA | 2023826-2017-01168 | 2016-10-27 | Injury | STAAR SURGICAL COMPANY | 2017-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS NOTICED TO BEING FLIPPED AND WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6407107&pc=MTA | 2023826-2017-00407 | 2016-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, AND THE LENS FLIPPED IN THE EYE. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES. THE REPORTER STATED THAT THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6290872&pc=MTA | 2023826-2017-00198 | 2016-11-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, AND NOTICED THE LENS WAS INJECTED IN THE EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACK UP LENS WAS IMPLANTED WITH NO ISSUES. THE REPORTED STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191969&pc=MTA | 2023826-2016-01815 | 2016-11-28 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON HAD DIFFICULTY LOADING A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, AND UPON INSERTION, THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. THE LENS WAS DAMAGED UPON REMOVAL. THE REPORTER STATED THE SURGEON FELT THE LENS WAS DEFECTIVE. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER?: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE MISSING. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6332331&pc=MTA | 2023826-2017-00104 | 2017-01-17 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.00 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. AFTER IMPLANTATION, THE LENS FLIPPED OVER AND WAS IMMEDIATELY (SAME SURGERY) EXCHANGED WITH ANOTHER LENS OF THE SAME SIZE AND DIOPTER. PROBLEM WAS RESOLVED. AS THE LENS WAS REMOVED FROM THE EYE, IT TORE. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. RESULTS: CONTRA-INDICATED USE-PATIENT (B)(6). (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6324390&pc=MTA | 2023826-2017-00232 | 2017-01-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6323345&pc=MTA | 2023826-2017-00231 | 2017-01-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, AND THE LENS FLIPPED IN THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY AND THE SURGERY WAS POSTPONED. ANOTHER LENS WAS IMPLANTED ON (B)(6) 2017 WITH NO ISSUES. PATIENTS POST -OP VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND MISSING PIECES FROM HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6441308&pc=MTA | 2023826-2017-00518 | 2017-03-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.50/082 DIOPTER, AND THE LENS WAS NOTED TO BE FRONT TO BACK UPON INSERTION. THE HAPTIC BROKE WHILE BEING REMOVED DURING THE SAME SURGERY. THERE WAS NO APPARENT PATIENT INJURY. THE SURGERY WAS CANCELLED AND ANOTHER LENS WILL BE IMPLANTED AT A LATER DATE. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/400. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE LENS HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE PATIENTS EYE ONE DAY POST-OP WAS SORE, WITH REDNESS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6488047&pc=MTA | 2023826-2017-00616 | 2017-03-27 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, AND THE LENS WENT INTO THE PATIENTS EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6505532&pc=MTA | 2023826-2017-00679 | 2017-04-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, AND UPON INJECTION, THE LENS FLIPPED IN THE EYE. THE LENS TORE WHILE BEING REMOVED DURING THE SAME SURGERY. THERE WAS NO PATIENT INJURY. THE SURGERY WAS CANCELLED AND ANOTHER LENS WILL BE IMPLANTED AT A LATER DATE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595991&pc=MTA | 2023826-2017-00840 | 2017-05-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, AND THE LENS WAS NOTED TO HAVE FLIPPED. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY, NO ENLARGED INCISION OR SUTURES. THE IRIS HAD COME DOWN SO THE SURGEON DECIDED TO BRING THE PATIENT BACK ON (B)(6) 2017 TO IMPLANT THE BACKUP LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE REPORTER INDICATED THEY USED ANOTHER MANUFACTURERS INJECTOR SYSTEM WITH THIS MODEL LENS AND THIS IS OFF-LABEL USE. Manufacturer Narrative: PREVIOUS: THE REPORTER INDICATED THEY USED ANOTHER MANUFACTURERS INJECTOR SYSTEM WITH THIS MODEL LENS AND THIS IS OFF LABEL USE. CORRECTED: BY THE REPORTER CLARIFICATION THAT THE STAAR INJECTION SYSTEM WAS USED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762442&pc=MTA | 2023826-2017-01263 | 2017-06-13 | Malfunction | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 AS THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5, INTO THE PATIENT'S EYE, IT BECAME TWISTED DURING INJECTION. THE SURGEON TRIED TO ADJUST THE LENS BUT THEN REMOVED IT AND IMPLANTED A BACK UP LENS. REPORTER STATES NO PATIENT INJURY. ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832456&pc=MTA | 2023826-2017-01380 | 2017-07-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, AND THE LENS FLIPPED IN THE EYE. THE LENS WAS CUT IN HALF TO REMOVE AND THERE WAS NO PATIENT INJURY. THE BACK UP LENS WAS IMPLANTED WITH NO ISSUES. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, SERIAL # (B)(4), AND THE LENS FLIPPED IN THE EYE. THE LENS WAS CUT IN HALF TO REMOVE AND THERE WAS NO PATIENT INJURY. THE BACKUP LENS, SERIAL # (B)(4), WAS IMPLANTED WITH NO ISSUES. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. CORRECTED DATA: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A PIECE OF GAUZE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE AND DEBRIS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918294&pc=MTA | 2023826-2017-01533 | 2017-09-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS | Positioning Failure | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.50 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS "FLIPPED"/IMPLANTED "UPSIDE DOWN." AS THE SURGEON REMOVED THE LENS IN ORDER TO EXCHANGE IT, THE LENS WAS "DAMAGED." THIS OCCURRED ON (B)(6) 2017. THE LENS WAS EXCHANGED INTRAOPERATIVELY WITH THE SAME SIZE, DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD NOT BE CHECKED. THE SURGEON REMOVED THE LENS IN ORDER TO EXCHANGE IT, THE LENS WAS DAMAGED WHICH OCCURRED ON (B)(6) 2017. CLAIM #: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND THICK, CLEAR RESIDUE ON THE LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6987325&pc=MTA | 2023826-2017-01637 | 2017-09-05 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. DURING IMPLANTATION, THE SURGEON NOTICED THAT THE LENS UNFOLDED UPSIDE DOWN. THE LENS WAS EXCHANGED DURING THE SAME SURGERY WITH A BACK-UP LENS. THE PATIENT IS FINE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY LOOSE IN TRAY. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE, THE LENS HAS A CURVED SHAPE, AND CLEAR RESIDUE ON LENS . CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT OVER THE AGE OF (B)(6). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6917523&pc=MTA | 2023826-2017-01519 | 2017-09-13 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, AND NOTED THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACK UP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT TYPE EVENT WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. RETURNED PRODUCT EVALUATION FOUND PIECES OF HAPTIC TORN OFF AND MISSING. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7294660&pc=MTA | 2023826-2018-00276 | 2017-11-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, AND THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS TORE AS IT WAS REMOVED DURING THE INITIAL SURGERY, WITH NO APPARENT INJURY. THE SURGEON WAS UNABLE TO IMPLANT THE BACKUP LENS DUE TO THE PUPIL CONSTRICTED. THE SURGERY WAS ABORTED AND ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED ON (B)(6) 2017. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL OBSERVATION FOUND THE LENS MISSING A PIECE OF HAPTIC AND RESIDUE ON LENS SURFACE. THE CARTRIDGE WAS RETURNED IN A DEVICE TRAY. VISUAL INSPECTION OF THE CARTRIDGE FOUND NO VISIBLE DAMAGE TO THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180664&pc=MTA | 2023826-2018-00034 | 2017-12-05 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS | Positioning Failure | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5 INTO THE PATIENT'S LEFT EYE (OS), THE LENS IMPLANTED UPSIDE DOWN. THE LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2017. UPON REMOVAL, THE LENS TORE/BROKE. ON (B)(6) 2017 A SAME SIZE, TYPE, AND DIOPTER LENS WAS SUCCESSFULLY IMPLANTED. REPORTEDLY, THE "EYE IS QUIET" AND THE PATIENT IS "DOING GOOD." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7211135&pc=MTA | 2023826-2018-00116 | 2017-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6 MM VICMO12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE ON (B)(6) 2017. DURING INSERTION, THE LENS GOT STUCK IN THE PLUNGER. THE REPORTER STATED THAT IT WAS DUE TO TECHNIQUE ERROR. THE LENS WAS NOT IMPLANTED. WHILE REMOVING THE LENS FROM THE CARTRIDGE, THE LENS TORE. ON (B)(6) 2017, A SIMILAR LENS WAS IMPLANTED AND THE SURGERY WENT OFF WELL. THE PATIENT IS HAPPY. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, THE HAPTIC BENT AND PRESENCE OF DEBRIS ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268467&pc=MTA | 2023826-2018-00189 | 2017-12-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE 12.1 MM VICM5_12.1, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS GOT JAMMED IN THE FOAM TIP PLUNGER AND TORE WHEN THE SURGEON LOADED IT INTO THE CARTRIDGE BEFORE IMPLANTING THE LENS INTO THE EYE ON (B)(6) 2017. ANOTHER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AND THE PATIENT IS SATISFIED. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A HAPTIC PIECE AND RESIDUE ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: DEVICE EVALUATED BY MANUFACTURER: IN THE INITIAL MDR, CORRECTED FROM YES TO NO. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7364253&pc=MTA | 2023826-2018-00454 | 2018-02-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, AND THE LENS WENT INTO THE EYE INSIDE OUT. THE LENS WAS DAMAGED WHILE BEING REMOVED WITHIN THE SAME SURGERY AND THERE WAS NO REPORT OF ANY INJURY. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE LENS MISSING A PIECE OF HAPTIC AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6407179&pc=MTA | 2023826-2017-00408 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, AND THE LENS FLIPPED IN THE EYE. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES. THE REPORTER STATED THAT THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7161181&pc=MTA | 2023826-2018-00001 | 2017-11-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure; Entrapment of Device; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6 MM VTICMO12.6, -12.50/2.5/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS "HALF IMPLANTED" WHEN THE PLUNGER OVERRODE THE LENS AND GOT STUCK IN THE CARTRIDGE. WHEN THE LENS WAS REMOVED FOR RELOADING, A TEAR IN THE PERIPHERICAL HAPTIC WAS NOTED. THE LENS WAS NOT REIMPLANTED. ON (B)(6) 2017, A SIMILAR LENS WAS IMPLANTED AND IT WAS SUCCESSFUL. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6899522&pc=MTA | 2023826-2017-01515 | 2017-08-30 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-27 | MTA | IMPLANTABLE COLLAMER LENS | Tear, Rip or Hole in Device Packaging; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: REPORTEDLY, AS THE SURGEON WENT TO OPEN THE PACKAGE OF A STERILE 12.6MM MICL12.6 LENS, HE NOTICED THE TYVEK COVERING WAS ALREADY SLIGHTLY OPENED. SURGEON CONTINUED TO IMPLANT THE LENS. DURING IMPLANT INTO THE RIGHT (OD) EYE, THE LENS TORE. HE THEN REMOVED THE LENS AND USED THE BACK-UP LENS. THIS OCCURRED ON (B)(6) 2017. Manufacturer Narrative: ADDITIONAL INFORMATION: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773454&pc=MTA | 2023826-2015-00549 | 2009-02-04 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-05-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE FACILITY RETURNED A LENS IMPLANT CARD FOR A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS AND INDICATED "HAS A TEAR ON THE LENS, DIDN'T USE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). LENS NOT RETURNED. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4488988&pc=MTA | 2023826-2015-00129 | 2014-12-18 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-02-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE FACILITY RETURNED A TORN 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS TO THE MANUFACTURER. INFORMATION WAS PROVIDED WITH THE RETURNED PRODUCT - "OPENED, NOT USED, NOT IN EYE, NICK IN LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER STATED THE SURGEON LOADED THE LENS AND NOTED UNDER THE MICROSCOPE A NICK/TEAR IN THE LENS. THE LENS WAS EJECTED OUT OF THE CARTRIDGE AND THE BACKUP LENS WAS IMPLANTED. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4) - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT; TORN MATERIAL; EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4471653&pc=MTA | 2023826-2015-00102 | 2014-12-22 | Malfunction | STAAR SURGICAL COMPANY | 2015-01-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS AND A SMALL PIECE OF LENS HAPTIC TORE OFF. THERE WAS PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY. THE REPORTER INDICATED THE SURGEON DID NOT FEEL THERE WAS A PROBLEM WITH THE LENS OR THE INJECTION SYSTEM. THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO: LENS WAS DISCARDED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS DISCARDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4522028&pc=MTA | 2023826-2015-00163 | 2015-01-14 | Malfunction | STAAR SURGICAL COMPANY | 2015-02-17 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED AND REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS. THE LENS WAS CUT INTO PIECES TO REMOVE WITH NO FURTHER INTERVENTION NEEDED. THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED SHE WAS NOT SURE IF THE LENS TORE WHILE BEING INSERTED. Manufacturer Narrative: PT AGE AND WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC AND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED DRY. CONCLUSIONS (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - IN SPITE OF THREE ATTEMPTS TO GATHER INFORMATION, THERE IS NO INFORMATION TO DETERMINE WHAT HAPPENED DURING SURGERY. IT IS NOT KNOWN IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE WHETHER THERE WAS IMPROPER LENS LOADING OR SURGEON HANDLING. THE REPORT STATES THAT THE LENS WAS CUT UP DURING EXPLANTATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES THAT IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. A CARTRIDGE LOT NUMBER SEARCH FOUND NO SIMILAR COMPLAINTS. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW, CARTRIDGE LOT NUMBER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4738522&pc=MTA | 2023826-2015-00492 | 2015-02-11 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-04-29 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS AND THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED DURING THE SAME PROCEDURE AND THERE WAS NO PATIENT INJURY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727970&pc=MTA | 2023826-2015-00463 | 2015-03-23 | Malfunction | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND THE LENS TORE UPON INSERTION. THE LENS WAS REMOVED AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. THE SURGEON PLANS ON REPLACING THE LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: CONCLUSION: BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, THE MOST PROBABLE CAUSE OF THE LENS TEARING DURING INJECTION IS USER OR TECHNICAL ERROR. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CASE/ VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND PEN MARKS ON LENS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773358&pc=MTA | 2023826-2015-00548 | 2015-04-14 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS AND NOTED THE LENS WAS TORN. THE LENS TOUCHED THE PATIENT'S EYE AND THERE WAS NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER INDICATED THE LENS DAMAGE MAY HAVE OCCURRED WHILE THE LENS WAS BEING LOADED. Manufacturer Narrative: (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - TEARS OR NICKS CAN OCCUR AT ANY STAGE FROM MANUFACTURE TO SURGICAL MANIPULATION, AND IT IS NOT KNOWN WHEN THE TEAR OCCURRED IN THIS CASE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: AFTER A REVIEW OF THE DHR, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. THE REPORTER STATED THE LENS DAMAGE MOST LIKELY OCCURRED WHILE BEING LOADED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4961785&pc=MTA | 2023826-2015-01005 | 2015-04-23 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-08-03 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS AND THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED DURING THE SAME PROCEDURE AND THERE WAS NO PATIENT INJURY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4929836&pc=MTA | 2023826-2015-00913 | 2015-04-28 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-07-21 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS TORE AS THE SURGEON WAS DELIVERING IT IN THE EYE. THE LENS WAS REMOVED AND NO PATIENT INJURY WAS REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4) - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT; TORN MATERIAL; BREAK. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4841978&pc=MTA | 2023826-2015-00665 | 2015-05-13 | Malfunction | STAAR SURGICAL COMPANY | 2015-06-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS AND NOTED A TEAR/DEFECT ON THE LENS. THE LENS ALSO TRIED TO FLIP WHILE BEING INSERTED. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED BUT THE SURGEON ALSO NOTED A SMALL TEAR/DEFECT ON THAT LENS. THE SURGEON FELT THE TEAR WOULD NOT AFFECT THE PATIENT'S VISION, SO THE LENS REMAINS IMPLANTED - SEE MFR. # 2023826-2015-00666. THE PATIENT IS VERY HAPPY WITH HER VISION. Manufacturer Narrative: (B)(4) - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT; TORN MATERIAL; UNINTENDED MOVEMENT. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. THE INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER WERE NOT RETURNED. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: TEARS OR NICKS CAN OCCUR AT ANY STAGE FROM MANUFACTURE TO SURGICAL MANIPULATION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842072&pc=MTA | 2023826-2015-00666 | 2015-05-13 | Malfunction | STAAR SURGICAL COMPANY | 2015-06-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS AND NOTED A TEAR/DEFECT ON THE LENS. THE SURGEON FELT THE TEAR WOULD NOT AFFECT THE PATIENT'S VISION, SO THE LENS REMAINS IMPLANTED. THERE WAS NO PATIENT INJURY AND THE PATIENT IS VERY HAPPY WITH HER VISION. THIS WAS THE 2ND/BACKUP LENS FOR THIS PATIENT - SEE MFR. # 2023826-2015-00665 FOR THE FIRST LENS. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4871320&pc=MTA | 2023826-2015-00740 | 2015-05-16 | Malfunction | STAAR SURGICAL COMPANY | 2015-06-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE LENS TORE UPON INSERTION. THE LENS WAS REMOVED AND NO PATIENT INJURY WAS REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4975077&pc=MTA | 2023826-2015-01014 | 2015-06-24 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS FROM THE VIAL TO PREPARE FOR LOADING. THE SURGEON CHECKED THE LENS UNDER THE MICROSCOPE AND NOTED THE LEAD HAPTIC HAD A TEAR. THE LENS WAS NOT USED OR LOADED AND THERE WAS NO PATIENT CONTACT. THE REPORTER STATED THE SURGEON FELT THE LENS WAS RECEIVED TORN AND WAS DEFECTIVE. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM CODES PATIENT CODE: (B)(4)- NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT, LABELED. DEVICE CODE: (B)(4)- TORN MATERIAL, UNLABELED. DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS NOT RETURNED. EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171704&pc=MTA | 2023826-2015-01289 | 2015-07-03 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER -13.00 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND ADVERSE EVENT WAS REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.00. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5015483&pc=MTA | 2023826-2015-01060 | 2015-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/+1.5/045 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS TORE DURING INJECTION INTO THE EYE. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -09.50/ +1.5/045. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE NOT RETURNED.(B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5022556&pc=MTA | 2023826-2015-01066 | 2015-07-23 | Malfunction | STAAR SURGICAL COMPANY | 2015-08-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6, -05.50 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED. Manufacturer Narrative: LENS NOT RETURNED. (B)(4). BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919727&pc=MTA | 2023826-2016-01196 | 2016-02-15 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, AND THE LENS WAS TORN DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY AND ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTION DATA: PATIENT'S AGE AT TIME OF EVENT (B)(6) IS INCORRECT, CORRECT AGE IS (B)(6). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISIBLE INPSECTION FOUND LENS HAPTIC TORN, AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS DRIED, DISCOLORED MATERIAL. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5671675&pc=MTA | 2023826-2016-00644 | 2016-03-29 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. Manufacturer Narrative: (B)(4). REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS (B)(6) OF AGE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, THE HAPTIC TORN/BROKEN/BENT/DEFORMED, AND ALSO FOREIGN MATERIALS ON LENS SURFACE SPECIFIED AS SPOTS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676350&pc=MTA | 2023826-2016-00722 | 2016-04-15 | Malfunction | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/080 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) BUT THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS ON (B)(6) 2016 WITH NO APPARENT INJURY. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5781003&pc=MTA | 2023826-2016-00894 | 2016-05-12 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/107 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED WITH ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT AGE BELOW 21 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5840748&pc=MTA | 2023826-2016-01009 | 2016-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, WAS NOT USED DUE TO A BROKEN HAPTIC. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULTS: (OPERATIONAL PROBLEM) - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC WAS TORN OFF AND MISSING. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5840778&pc=MTA | 2023826-2016-01010 | 2016-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, WAS NOT USED DUE TO BROKEN HAPTIC. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC WAS TORN OFF. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6004360&pc=MTA | 2023826-2016-01379 | 2016-07-18 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -10/1/92 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 WHEN THE LENS WAS NOTED TO BE TORN DURING INJECTION AS A PROXIMAL HAPTIC WAS BROKEN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND THE SURGEON CANCELED THE SURGERY DUE TO NO BACK UP LENS. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN AND SPOTS ON LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902553&pc=MTA | 2023826-2016-01161 | 2016-07-27 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON WAS INSERTING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, AND THE INJECTOR OVERRODE AND THE LENS WAS DAMAGED. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923621&pc=MTA | 2023826-2016-01222 | 2016-08-01 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE CUSTOMER NOTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/092 DIOPTER, WAS TORN PRIOR TO LOADING. THE REPORTER INDICATED SINCE THERE WAS NO BACKUP LENS AVAILABLE, THE SURGEON IMPLANTED THE LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THERE WAS NO REPORT OF ANY APPARENT INJURY AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE PATIENT WILL BE MONITORED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6063721&pc=MTA | 2023826-2016-01501 | 2016-09-13 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, TORE DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT OR INJURY. THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6054100&pc=MTA | 2023826-2016-01521 | 2016-10-05 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM MICL12.6 LENS, -8.0 DIOPTER, TORE DURING LOADING. THERE WAS NO PATIENT CONTACT, AND A BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH PIECES MISSING. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074378&pc=MTA | 2023826-2016-01570 | 2016-10-10 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 LENS, -12.00 DIOPTER, TORE DURING LOADING ON (B)(6) 2016. THERE WAS NO PATIENT CONTACT AND A BACKUP LENS WAS USED. THE CAUSE OF THE TEAR IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH A PIECE MISSING. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4636964&pc=MTA | 2023826-2015-00395 |  | Malfunction | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS WAS TORN AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT AGE AND WEIGHT: UNK. EVENT DATE: UNK. IMPLANT AND EXPLANT DATES: UNK. INITIAL REPORTER: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. PER MEDICAL REVIEW: REPORTEDLY, INTRAOPERATIVE LENS EXCHANGE WAS PERFORMED TO ADDRESS LENS DAMAGE ("TEAR"). NO REPORTED INCISION ENLARGEMENT OR ANY OTHER TISSUE DAMAGE. ANOTHER LENS WAS IMPLANTED AS A BACKUP. NO REPORTED POSTOPERATIVE SEQUELAE. DESPITE MULTIPLE ATTEMPTS NO ADDITIONAL INFORMATION HAS BEEN RECEIVED TO DATE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972062&pc=MTA | 2023826-2016-01318 |  | Malfunction | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE CUSTOMER RETURNED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, TO THE MANUFACTURER DAMAGED. THE CUSTOMER INDICATED THE HAPTIC RIPPED WHILE FOLDING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER STATED THE LENS WAS DAMAGED WHILE FOLDING DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS USED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. (B)(4). Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE TORN OFF AND MISSING FROM ONE HAPTIC. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4735118&pc=MTA | 2023826-2015-00485 | 2015-01-04 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-04-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT WHILE BEING INSERTED, THE LENS FLIPPED AND WAS GOING INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS PARTIALLY INSERTED AND WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE SURGEON NOTED A SMALL TEAR ON ONE OF THE LENS HAPTICS WHEN REMOVED. THE REPORTER STATED THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM AND THE PATIENT IS DOING FINE. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. EVALUATION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS "LENS FLIPPED DURING INSERTION, AND WAS GOING INTO THE EYE UPSIDE DOWN ". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE WHETHER MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING EITHER. THE NICK IN THE HAPTIC FLANGE OCCURRED VERY LIKELY DURING EXPLANTATION. THERE IS NO INFORMATION IN THE REPORT ABOUT THE DAMAGE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4555199&pc=MTA | 2023826-2015-00214 | 2015-01-29 | Malfunction | STAAR SURGICAL COMPANY | 2015-02-27 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE LENS TORE WHILE BEING REMOVED AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. EVALUATION OF THE RETURNED PRODUCT FOUND A PIECE OF LENS OPTIC AND ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DRIED SURGICAL RESIDUE ON THE LENS SURFACE. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4598988&pc=MTA | 2023826-2015-00315 | 2015-02-10 | Malfunction | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT THE LENS DID NOT UNFOLD PROPERLY AND WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS TORE WHILE BEING REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DRIED SURGICAL RESIDUE ON THE LENS SURFACE. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS "LENS DID NOT UNFOLD PROPERLY DURING INSERTION, AND WENT INTO EYE UPSIDE DOWN ". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE WHETHER MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING. THE TEAR IN THE HAPTIC FLANGE WAS VERY LIKELY DURING EXPLANTATION. THERE IS NO INFORMATION IN THE REPORT ABOUT THE DAMAGE. A CARTRIDGE LOT NUMBER SEARCH FOUND NO SIMILAR COMPLAINTS. A FOAM TIP PLUNGER LOT NUMBER SEARCH FOUND NO SIMILAR COMPLAINTS. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, CARTRIDGE LOT NUMBER SEARCH, FOAM TIP PLUNGER LOT NUMBER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773382&pc=MTA | 2023826-2015-00533 | 2015-03-17 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS AND THE LENS FLIPPED OVER IN THE PATIENT'S LEFT EYE (OS). THE LENS TORE WHILE BEING REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS DISCARDED BY THE FACILITY. EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH. RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS "FLIPPED UP UPSIDE DOWN DURING INSERTION". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF THE FLIP WAS DUE TO IMPROPER LENS LOADING, OR SURGEON HANDLING. A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES THAT COULD BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. THE CAUSE OF THE LENS TO NOT UNFOLD CORRECTLY, IS POSSIBLY DUE TO IMPROPER LOADING OR HANDLING PRACTICES IN THE OPERATING ROOM. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A POSSIBLE CAUSE OF THE EVENT MAY BE DUE TO IMPROPER LOADING OR HANDLING PRACTICES IN THE OPERATING ROOM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4776266&pc=MTA | 2023826-2015-00532 | 2015-04-07 | Malfunction | STAAR SURGICAL COMPANY | 2015-05-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS BUT THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS TORE WHILE BEING REMOVED, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO : LENS NOT RETURNED. METHOD - WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS : BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) : LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS WAS INSERTED UPSIDE DOWN. THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING EITHER. THE HAPTIC FLANGE WAS REPORTEDLY SEVERED DURING EXPLANTATION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4868631&pc=MTA | 2023826-2015-00736 | 2015-05-12 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-06-25 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS AND THE LENS WENT INTO THE EYE UPSIDE DOWN AND WAS TORN DURING DELIVERY. THE SURGEON ALSO NOTED LENS ROTATION. ENLARGEMENT OF PI OR ADDITION OF NEW ONE. THE LENS WAS EXCHANGED FOR SAME SIZE AND POWER. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613017&pc=MTA | 2023826-2016-00528 | 2016-02-05 | Malfunction | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, AND THE LENS INJECTED INTO THE EYE UPSIDE DOWN. UPON REMOVAL OF THE LENS, THE LENS WAS DAMAGED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4509428&pc=MTA | 2023826-2015-00140 | 2014-09-25 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-02-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS BUT THE LENS OPENED UP UPSIDE DOWN IN THE PATIENT'S EYE. THE LENS WAS REMOVED WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. EVALUATION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS "OPENED UP UPSIDE DOWN IN THE EYE". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING. THE HAPTIC FLANGE WAS VERY LIKELY SEVERED DURING EXPLANTATION. THERE IS NOTHING IN THE REPORT ABOUT THE DAMAGE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW, AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW, AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574341&pc=MTA | 2023826-2015-00248 | 2014-11-14 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-03-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS AND THE LENS WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS REMOVED WITH NO REPORTED PATIENT INJURY AND WAS NOT REPLACED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED A PIECE OF A HAPTIC PLATE WAS TORN OFF AND MISSING AND THERE WAS A CLEAR SURGICAL RESIDUE ON THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - WHEN THE LENS IS BEING INJECTED INTO THE ANTERIOR CHAMBER THE SURGEON MUST CHECK FOR LENS LANDMARKS TO ENSURE PROPER LENS ORIENTATION. IF THIS IS NOT DONE, THE LENS MAY END-UP BEING IMPLANTED UPSIDE DOWN. ONCE THE SURGEON REALIZES THE IMPROPER ORIENTATION, THE LENS MUST BE REMOVED TO AVOID COMPLICATIONS AND A NEW LENS SHOULD BE IMPLANTED. THIS IS NORMALLY DONE USING THE SAME INCISION. THE MOST PROBABLE CAUSE OF THE EVENT WAS USER'S ERROR WHILE LOADING/INJECTING THE LENS IN THE ANTERIOR CHAMBER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600861&pc=MTA | 2023826-2015-00310 | 2014-12-18 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS BUT THE LENS INVERTED AS IT WENT INTO THE PATIENT'S EYE. THE LENS WAS REMOVED WITH NO REPORTED PATIENT INJURY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - WHILE THE LENS IS BEING EJECTED INTO THE ANTERIOR CHAMBER THE SURGEON MUST CHECK FOR LENS LANDMARKS TO ENSURE PROPER LENS ORIENTATION. IF THIS IS NOT DONE, THE LENS MAY END-UP BEING IMPLANTED UPSIDE-DOWN. ONCE THE SURGEON REALIZES OF THE IMPROPER ORIENTATION, THE LENS MUST BE REMOVED TO AVOID COMPLICATIONS AND A NEW LENS SHOULD BE IMPLANTED. THIS IS NORMALLY DONE USING THE SAME INCISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, THE MOST PROBABLE CAUSE OF THIS EVENT WAS USER'S ERROR WHILE LOADING/INJECTING THE LENS IN THE ANTERIOR CHAMBER. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4509607&pc=MTA | 2023826-2015-00144 | 2015-01-14 | Malfunction | STAAR SURGICAL COMPANY | 2015-02-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT THE LENS TURNED UPSIDE DOWN IN THE PATIENT'S EYE. THE LENS WAS REMOVED WITH NO PATIENT INJURY. THE SURGEON DECIDED AT THAT TIME, NOT TO USE THE BACKUP LENS (13.2MM), AS HE FELT THIS SIZE WAS TOO BIG FOR THE PATIENT'S EYE. THEY HAD A 12.6MM ICL LENS IN STOCK, SO THIS LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE INITIAL LENS INSERTION, REMOVAL AND EXCHANGE WERE PERFORMED WITHIN THE SAME SURGERY. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. EVALUATION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN AND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS "WENT INTO THE EYE UPSIDE DOWN /TURNED UPSIDE DOWN WHILE BEING INSERTED IN THE EYE". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTER, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE WHETHER MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING EITHER. THE LENS WAS CHANGED TO A SMALLER SIZE ICL INTRAOPERATIVELY. THE HAPTIC FLANGE WAS VERY LIKELY SEVERED, AND THE OPTIC TORN, DURING EXPLANTATION. THERE IS NO INFORMATION IN THE REPORT ABOUT THE DAMAGE. A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES THAT COULD BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. THE CAUSE OF THE LENS TO NOT UNFOLD CORRECTLY, IS POSSIBLY DUE TO IMPROPER LOADING OR HANDLING PRACTICES IN THE OPERATING ROOM. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A POSSIBLE CAUSE OF THE EVENT WAS DUE TO IMPROPER LOADING OR HANDLING PRACTICES IN THE OPERATING ROOM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704638&pc=MTA | 2023826-2015-00434 | 2015-03-19 | Malfunction | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS BUT THE LENS FLIPPED AND WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER STATED THEY FELT THERE WAS NO PROBLEM WITH THE LENS OR INJECTION SYSTEM. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. THE LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS "FLIPPED AND ENTERED THE EYE UPSIDE DOWN". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS A MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING EITHER. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE IDENTIFIED. AS STATED BY THE REPORTING FACILITY, THERE WAS NO ISSUE WITH THE LENS OR THE INJECTION SYSTEM. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4822524&pc=MTA | 2023826-2015-00646 | 2015-04-21 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-06-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT THE LENS FLIPPED AND WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED DURING THE SAME SURGERY, WITH NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER INDICATED THE EVENT WAS DUE TO A LOADING ERROR. Manufacturer Narrative: (B)(4). RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. EVALUATION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER WERE NOT RETURNED. (B)(4). Manufacturer Narrative: THIS COMPLAINT FILE STATES THAT THE LENS "FLIPPED UP UPSIDE DOWN DURING INSERTION". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF THE FLIP WAS DUE TO IMPROPER LENS LOADING, OR SURGEON HANDLING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4889401&pc=MTA | 2023826-2015-00797 | 2015-06-04 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT THE LENS FLIPPED AND WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED AND THE PATIENT IS DOING WELL. Manufacturer Narrative: (B)(4) - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT; UNINTENDED MOVEMENT. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS COMPLAINT FILE STATES THAT THE LENS "FLIPPED UP UPSIDE DOWN DURING INSERTION". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF THE FLIP WAS DUE TO IMPROPER LENS LOADING, OR SURGEON HANDLING. THE HAPTIC FOOT WAS VERY LIKELY SEVERED DURING EXPLANTATION. THERE IS NOTHING IN THE REPORT ABOUT THE DAMAGE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4889337&pc=MTA | 2023826-2015-00798 | 2015-06-15 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS BUT THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED AND THE PATIENT IS DOING WELL. THE REPORTER STATED THE LENS WAS TORN WHILE BEING REMOVED. Manufacturer Narrative: (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES TORN OFF AND MISSING FROM ONE HAPTIC. THE OTHER HAPTIC WAS TORN. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS COMPLAINT FILE STATES THAT THE LENS WAS INSERTED UP UPSIDE DOWN IN THE EYE. THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4954780&pc=MTA | 2023826-2015-01002 | 2015-07-09 | Malfunction | STAAR SURGICAL COMPANY | 2015-07-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE-DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS "WENT INTO EYE UPSIDE DOWN". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING. EVALUATION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7997720&pc=MTA | 2023826-2018-01557 | 2018-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, DURING THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENTS EYE. ANOTHER SAME MODEL/DIOPTER LENS WAS INSERTED AS A REPLACEMENT LENS AND THE PROBLEM WAS RESOLVED. THERE WAS NO APPARENT INJURY AND THE PATIENT IS DOING GREAT. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC TORN OFF. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372530&pc=MTA | 2023826-2019-00286 | 2019-01-03 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, WAS REQUESTED TO BE RETURNED WITH THE INFORMATION "LENS FLIPPED, NOT IMPLANTED." DATE OF SCHEDULED SURGERY WAS (B)(6) 2019. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8798403&pc=MTA | 2023826-2019-01108 | 2019-06-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. DURING INSERTION, THE LENS FLIPPED IN THE EYE, "AS INSERTED THE LENS APPEARED TO FLIP." INTRAOPERATIVELY, THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS. REPORTEDLY, NO PATIENT INJURY OCCURRED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: CORRECTION: "INTRAOPERATIVELY, THE LENS WAS SUCCESSFULLY EXCHANGED..." IN INITIAL MDR SHOULD BE CHANGED TO "THE LENS WAS SUCCESSFULLY EXCHANGED ON THE SAME DAY. HOWEVER, IT IS UNKNOWN IF THE EXCHANGE OCCURRED DURING THE INITAL SURGERY." (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842798&pc=MTA | 2023826-2020-00504 | 2020-02-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE RIGHT (OD) EYE, IT FLIPPED IN THE EYE. THIS OCCURRED ON (B)(6) 2020. THE LENS WAS IMPLANTED, EXPLANTED, AND REPLACED WITHIN THE SAME SURGERY. THE CAUSE IS REPORTED AS POSSIBLY DUE TO THE LENS WAS "THIN" AS IT WAS -4.0 POWER. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332985&pc=MTA | 2023826-2020-01576 | 2020-07-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, 11.5 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. AFTER INSERTION, THE LENS "FLIPPED IN THE EYE." INTRAOPERATIVELY, THE LENS WAS SUCCESSFULLY REMOVED AND EXCHANGED WITH AN ALTERNATE LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID A VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4601763&pc=MTA | 2023826-2015-00317 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS AND NOTED THE LENS WAS INVERTED (UPSIDE DOWN) IN THE PATIENT'S LEFT EYE (OS). THE LENS WILL BE EXPLANTED AND EXCHANGED FOR ANOTHER ICL LENS.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE VAULT WAS TOO SHALLOW. THE PATIENT IS DOING VERY WELL. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. PER MEDICAL REVIEW - THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM THAT LED TO UPSIDE-DOWN INSERTION, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF MISPLACEMENT OF THE LENS WAS DUE TO IMPROPER LENS LOADING OR SURGEON HANDLING. BASED ON THE COMPLAINT HISTORY, DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVALUATION CODES: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194490&pc=MTA | 2023826-2018-01905 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: J CATARACT REFRACTIVE SURGERY ARTICLE WAS RECEIVED ON 30-NOV-2018, ORIGINALLY PUBLISHED IN JANUARY 2002, WHICH CITED A STUDY TITLED, 'CORRECTION OF PSEUDOPHAKIC ANISOMETROPIA WITH THE STAAR COLLAMER IMPLANTABLE CONTACT LENS.' THE ARTICLE STATES THAT THERE WAS ONE PERIOPERATIVE COMPLICATION IN WHICH AN ICL WAS INSERTED UPSIDE DOWN. REPORTEDLY, THIS LIKELY OCCURRED DUE TO THE ICL BEING INCORRECTLY LOADED INTO THE CARTRIDGE HOWEVER COULD ALSO HAVE BEEN CAUSED BY THE ICL ROTATING DURING INJECTION. THE LENS WAS IMMEDIATELY REMOVED AND REINJECTED WITH THE CORRECT SIDE UP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608206&pc=MTA | 2023826-2020-00165 | 2019-12-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL 13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS HAD A TORN HAPTIC. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S EYE. DURING INSERTION, THE LENS "FLIPPED AS IT WAS BEING INJECTED INTO THE EYE." INTRAOPERATIVELY, THE HAPTIC TORE AS THE LENS WAS BEING REMOVED. PER THE REPORTER ON 25/JAN/2020, "THERE IS NO OFFICIAL COMPLAINT REGARDING THE ICL." CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138816&pc=MTA | 2023826-2020-01187 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE SURGEON REPORTED THAT DURING A IMPLANTABLE COLLAMER LENS SURGERY A LENS WAS "INSERTED INVERTED. THE LENS WAS REMOVED AND REINSERTED CORRECTLY THE SECOND TIME." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B4 - CORRECTED TO 14-MAY-2020 IN INITIAL MDR G3 - CORRECTED TO 14-MAY-2020 IN INITIAL MDR H6 - DEVICE CODE 3189 CORRECTED TO DEVICE CODE 3026 H6 - INVESTIGATION FINDINGS CODE 213 CORRECTED TO INVESTIGATION FINDINGS CODE 114 CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364446&pc=MTA | 2023826-2020-01641 | 2019-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2020 THE LENS WAS EXPLANTED. REPORTEDLY, DURING THE PLANNED EXCHANGE THERE WAS IRIS DYALISIS. THERE IS AN IRIS SUTURE AND IMPLANT OF A LONGER LENS PLANNED. H6-HEALTH EFFECT CLINICAL CODE: 4581 (IRIS DYALISIS). CLAIM #: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS MOISTURE AND RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN; THE LENS WAS RETURNED SPLIT INTO TWO PIECES. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9855361&pc=MTA | 2023826-2020-00534 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS "SIZING ERROR". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- "ADVERSE EVENT" SHOULD BE IN MDR SUPP #1. B2- "REQUIRED INTERVENTION" SHOULD BE IN MDR SUPP #1. B3- "03-DEC-2019" SHOULD BE IN MDR SUPP #1. H6- PATIENT CODE 2692 SHOULD BE CORRECTED TO 4629- SECONDARY SURGICAL INTERVENTION: LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308250&pc=MTA | 2023826-2020-01495 | 2020-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT IS CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- EXPLANT DATE IS (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B5-THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. 6B. (B)(6) 2020 H6- CLINICAL CODE 4582: VAULTING. IMPACT CODE 4629; LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION- LENS RETURNED IN A MICRO- CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8122335&pc=MTA | 2023826-2018-01670 |  | Injury | STAAR SURGICAL COMPANY | 2018-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Markings/Labelling Problem | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO 12.6 -12.00/+1.50/082, SERIAL (B)(4), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). THE REPORTER STATES THAT THE SAME MODEL, POWER AND SERIAL NUMBER OF THIS LENS (VTICMO 12.6 -12.00/+1.50/082, SERIAL # (B)(4) WAS SHOWN ON THE PATIENT CHART LABEL INSIDE THE BOX FOR ANOTHER LENS WITH VTICMO 12.6 -12.00/+1.50/083, SERIAL # (B)(4) INDICATED ON THE EXTERNAL BOX LABEL. THE LENS REMAINS IMPLANTED AND IT WAS REPORTED THAT THERE IS NO NEGATIVE FEEDBACK OR COMPLAINTS ABOUT THE PATIENT'S CONDITION Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SERIAL # (B)(6) (VTICMO 12.6 -12.00/+1 .50/082) WAS SHOWN ON THE PATIENT CHART LABEL INSIDE THE BOX FOR ANOTHER LENS WITH SERIAL# (B)(6) (VTICMO 12.6 -12.00/+1.50/083) INDICATED ON THE EXTERNAL BOX LABEL. REPORTEDLY, "THE LENS WAS NOT IMPLANTED IN 2018 DUE TO PATIENT'S REASONS." THE VTICMO 12.6 LENS WITH A DIOPTER OF -12.00/+1.50/082 (SPHERE/CYLINDER/AXIS), SERIAL (B)(6) WAS IMPLANTED TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULTING. THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. COMPLAINT DESCRIPTIONS PACKAGE LABEL DOES NOT MATCH SERIAL NUMBER ON STICKER. UNABLE TO VERIFY COMPLAINT STATEMENT DUE TO LENS NOT RETURNED IN ORIGINAL PACKAGING. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THE FAILURE IS LIKELY DUE TO FAILURE TO FOLLOW PROCEDURE BY PACKAGING MORE THAN ONE LENS AT A TIME AND THEREFORE INADVERTENTLY PACKAGING THE INCORRECT PATIENT CHART LABEL. A SECOND PROBABLE CAUSE IDENTIFIED IS DUE TO INADEQUATE FINAL VERIFICATION PROCESS OF THE PATIENT CHART LABEL. CORRECTIVE ACTIONS FOR A SIMILAR DISCREPANCY REGARDING THE PACKAGING OF AN INCORRECT PATIENT CHART LABEL HAVE BEEN IDENTIFIED AND HAVE BEEN IMPLEMENTED UNDER CAPA M18-12. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7006984&pc=MTA | 2023826-2017-01692 | 2017-08-29 | Injury | STAAR SURGICAL COMPANY | 2017-11-07 | MTA | IMPLANTABLE COLLAMER LENS | Activation Failure | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: REPORTEDLY, WHILE PREPARING TO INSERT A 13.2MM MICL13.2, DIOPTER -10.0 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE, THE LENS WAS DAMAGED. THE RETURNED BOX STATES THE LENS WAS NOT IMPLANTED AND "LENS REMOVED-WAS NOT OPENING PROPERLY." THE SURGERY DATE IS LISTED AS (B)(6) 2017. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: SURGEON STATES THAT WHEN THE LENS WAS INJECTED INTO THE EYE IT DID NOT UNFOLD PROPERLY. THE LENS WAS IN THE EYE STILL FOLDED "DESPITE WAITING FOR SOME TIME." THE LENS WAS IMPLANTED, REMOVED, AND "ULTIMATELY A BACK-UP LENS WAS IMPLANTED AFTERWARD." (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6810144&pc=MTA | 2023826-2017-01363 | 2015-01-30 | Malfunction | STAAR SURGICAL COMPANY AG | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON DID NOT IMPLANT A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE. THE LENS WAS NOT IMPLANTED DUE TO A DOUBT THAT THE LENS ARE SWAPPED BETWEEN BOTH EYES DURING OPERATION AND THE DR. STOPPED OPERATION AFTER IMPLANTING THE RIGHT EYE. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO DOUBT THAT THE LENS WERE SWAPPED BETWEEN BOTH EYES DURING OPERATION. THE DR. SKIPPED OPERATION AFTER IMPLANTING THE RIGHT EYE. ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON DID NOT IMPLANT A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS NOT IMPLANTED DUE TO A DOUBT THAT THE LENS ARE SWAPPED BETWEEN BOTH EYES DURING OPERATION. THE DOCTOR STOPPED OPERATION AFTER IMPLANTING THE RIGHT EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696923&pc=MTA | 2023826-2018-01032 | 2018-04-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. ON THE SAME DATE ((B)(6) 2018) AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. IT IS UNKNOWN AS TO WHETHER THIS WAS PERFORMED INTRAOPERATIVELY OR IN A 2ND SURGERY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PREVIOUSLY STATED IT WAS UNKNOWN IF THE SURGERY OCCURRED INTRAOPERATIVELY. CORRECTED-THE PROCEDURE WAS COMPLETED WITHIN THE SAME OPERATION. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853015&pc=MTA | 2023826-2018-01343 | 2018-05-28 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE PREPARING TO IMPLANT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/070 (SPHERE/CYLINDER/AXIS) THE LENS TORE/BROKE BEFORE/DURING LOADING. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY FROM THE LEFT (OD) EYE ON (B)(6) 2018. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN; ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS IS ALSO REPORTED. Manufacturer Narrative: CORRECTED DATA: INITIAL REPORT STATES LENS WAS IMPLANTED INTO THE "LEFT (OD) EYE". CORRECTED TO "RIGHT (OD) EYE". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9323796&pc=MTA | 2023826-2019-02147 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM512.6, -03.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. "PATIENT WAS NOT SATISFIED WITH VISUAL ACUITY, REQUESTED MORE THAN 20/16.7." LENS WAS REMOVED AND EXCHANGED FOR A -04.00 DIOPTER LENS ON 01/OCT/2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: H6 - PATIENT CODE (B)(6) CORRECTED TO PATIENT CODE (B)(6) IN INITIAL MDR H6 - DEVICE CODE (B)(4) CORRECTED TO DEVICE CODE (B)(4) IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5638613&pc=MTA | 2023826-2016-00595 | 2016-01-14 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND THERE WAS NO REPORTED INJURY. THE SURGEON DID NOT HAVE ANOTHER 13.2MM LENS AVAILABLE AND AS THE PATIENT WAS FROM (B)(6), A 12.6MM LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THIS IS AN ADVERSE EVENT, NOT A PRODUCT PROBLEM. DISTRIBUTOR NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6042653&pc=MTA | 2023826-2016-01428 | 2016-02-07 | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED. THE EXCHANGED LENS RESOLVED THE PROBLEM AND THERE WAS NO INDICATION OF ANY APPARENT INJURY. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5665329&pc=MTA | 2023826-2016-00737 | 2016-03-15 | Injury | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION. THE LENS WAS REMOVED WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/32. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. UPON VISUAL INSPECTION, THE OPTIC WAS OBSERVED TORN AND THE HAPTIC BROKEN AND BENT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). PREVIOUS CODES NOT APPLICABLE, THERE IS NO BREAK OR TORN MATERIAL ASSOCIATED WITH A MATERIAL INTEGRITY ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668559&pc=MTA | 2023826-2016-00632 | 2016-03-29 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/060 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE LENS TORE/BROKE DURING INJECTION. THE LENS WAS REMOVED WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE. (B)(4): SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CONCLUSION: PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER THE AGE OF 45 YEARS OLD AND WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSURE, LESS THAN 3.0MM. THE PATIENT WAS (B)(6) AND HAD AN ACD OF 2.80MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817533&pc=MTA | 2023826-2016-00918 | 2016-04-29 | Malfunction | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/088 DIOPTER, AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY. ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (NOT SPECIFIED). (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: UNCORRECTED VISUAL ACUITY WAS 20/20, NOT BEST CORRECTED VISUAL ACUITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5985705&pc=MTA | 2023826-2016-01353 | 2016-07-28 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20.Event Description: IMPLANT DATE (B)(6) 2016, EXPLANT DATE (B)(6) 2016, EXCHANGE DATE (B)(6) 2016. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: IMPLANT DATE (B)(6) 2016; EXPLANT DATE (B)(6) 2016; EXCHANGE DATE (B)(6) 2016. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND HAPTIC BROKEN (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923021&pc=MTA | 2023826-2015-00826 | 2015-03-30 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Device Operates Differently Than Expected; Torn Material; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. CORRECTIONAL DATA: SUSPECT MEDICAL DEVICE: DIOPTER -14.0/+2.0/070 IS INCORRECT, CORRECT DATA IS DIOPTER -14.50/1.5/084. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047903&pc=MTA | 2023826-2015-01143 | 2015-02-12 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contamination with Chemical or Other Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND A FOREIGN BODY WAS NOTED. THE LENS WAS REMOVED AND EXCHANGED FOR ANOTHER LENS AND THIS RESOLVED THE PROBLEM. NO PATIENT INJURY WAS REPORTED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND A FOREIGN BODY WAS NOTED. ON THE SAME DAY, THE LENS WAS REMOVED AND EXCHANGED FOR ANOTHER LENS AND THIS RESOLVED THE PROBLEM. NO PATIENT INJURY WAS REPORTED. PATIENT'S LAST POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED, A PIECE OF THE HAPTIC MISSING, FOREIGN MATERIAL ON LENS SURFACE (DEBRIS), AND A SMALL PARTICULATE IN THE OPTIC VISIBLE AT 30X. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5001759&pc=MTA | 2023826-2015-01054 | 2015-01-17 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contamination with Chemical or Other Material; Adverse Event Without Identified Device or Use Problem | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE LENS WAS REMOVED DUE TO FB ADHESION. THE LENS WAS EXCHANGED FOR ANOTHER VICMO12.6 AND THE ISSUE WAS RESOLVED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO FOREIGN BODY ADHESION. THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. ON (B)(6) 2015, THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. NO DAMAGE OBSERVED BUT CANNOT CONFIRM PARTICULATE REPORT IN INCIDENT. NO LOCATION OR ATTRIBUTE OF PARTICULATE REPORTED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - 02 (DEVICE EVALUTION ANTICIPATED BUT NOT YET PERFORMED) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765105&pc=MTA | 2023826-2018-01172 | 2018-07-09 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.5/+5.0/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ACCORDING TO THE PATIENT'S (B)(6) VIDEO LOG, THE ICL IS "SLIGHTLY TURNED AND NEEDS TO BE ADJUSTED BY 10 DEGREES." REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS CURRENTLY AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS REPOSITIONED ON (B)(6) 2018. THE ISSUE IS RESOLVED AND THE PATIENT NO LONGER WEARS GLASSES. (B)(4). Manufacturer Narrative: ADDRESS: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6314690&pc=MTA | 2023826-2017-00140 |  | Injury | STAAR SURGICAL COMPANY AG | 2017-02-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS ROTATED ABOUT THE AXIS AT 84 DEGREES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17/+3.5/90 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS ROTATED ABOUT THE AXIS AT 84 DEGREES. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2016. THE LENS REMAINS IMPLANTED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4)- PREVIOUS CODE NOT APPLICABLE. LENS REPOSITIONING. (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). DEVICE MANUFACTURE DATE IS CORRECTED TO 04/03/2013 FROM 03/27/2013. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6125957&pc=MTA | 2023826-2016-01678 | 2016-09-20 | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Handling Problem | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/2.5/092 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE LENS WAS IMPLANTED INSIDE OUT. THE LENS WAS EXCHANGED FOR A LENS THE SAME MODEL, SIZE BUT A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS EXPLANTED AND EXCHANGED FOR A LENS THE SAME SIZE, MODEL, BUT A DIFFERENT DIOPTER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6124295&pc=MTA | 2023826-2016-01642 | 2016-08-09 | Malfunction | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND THE LENS WAS NOTED TO BE TORN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LENS THE SAME SIZE, MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE OPTIC AND THE HAPTIC WERE TORN. (B)(4). LENS WAS EXCHANGED INTRAOPERATIVELY FOR ANOTHER LENS OF THE SAME MODEL, SIZE AND DIOPTER. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6216174&pc=MTA | 2023826-2016-01851 | 2016-08-17 | Injury | STAAR SURGICAL COMPANY | 2016-12-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND IT WAS NOTED THAT THE LENS WAS TORN DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME SIZE, MODEL, AND DIOPTER AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY PATIENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6226582&pc=MTA | 2023826-2017-00023 | 2016-09-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/093 DIOPTER, TORE/BROKE BEFORE INJECTION INTO THE EYE. THE LENS WAS FOUND CRACKED JUST BEFORE IMPLANTATION. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS BROKEN. (B)(4). Manufacturer Narrative: CORRECTED INFORMATION: FOLLOW-UP #1 - 01/05/2017. (B)(4) PREVIOUS CODE NOT APPLICABLE, SECONDARY INTERVENTION WAS PERFORMED (EXCHANGE). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6260573&pc=MTA | 2023826-2017-00099 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | INTRAOCULAR LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, HAD TORN/BROKEN DURING INJECTION/DELIVERY INTO THE LEFT (OS) EYE. THIS OCCURRED ON (B)(6) 2016. THE LENS WAS NOT IMPLANTED AND THERE IS NO REPORTED PATIENT INJURY. AS OF THE DATE OF MDR SUBMISSION, THIS LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE SURGEON DECIDED TO IMPLANT THE ICL. WITH THE LENS IMPLANTED, THE SURGEON NOTICED THAT THE TORN HAPTIC WAS TOUCHING THE HUMAN LENS. THE SURGEON DECIDED TO REMOVE THE ICL AND CONVERT TO CLEAR LENS EXTRACTION WITH IOL IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6196858&pc=MTA | 2023826-2016-01809 | 2016-10-31 | Injury | STAAR SURGICAL COMPANY | 2016-12-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND IT WAS NOTED THAT THE LENS WAS TORN. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LENS THE SAME SIZE, MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, THE HAPTIC BROKEN, AND LENS IN HALF. CORRECTED DATA: PREVIOUS REPORT AS "PRODUCT PROBLEM" IS INCORRECT. THE CORRECT REPORT IS "ADVERSE EVENT". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PER DFU, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THEN 3.0MM. PATIENT'S ACD WAS REPORTED TO BE 2.85MM. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6196945&pc=MTA | 2023826-2016-01812 | 2016-11-02 | Injury | STAAR SURGICAL COMPANY | 2016-12-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND THE LENS WAS TORN DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS EXPLANTED IN THE SAME SURGERY ON (B)(6) 2016 AND WAS EXCHANGED FOR A LENS OF THE SAME SIZE, MODEL AND DIOPTER. THERE WAS NO REPORT OF ANY PATIENT INJURY. THE PROBLEM WAS RESOLVED AFTER THE LENS WAS EXCHANGED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THIS CASE IS AN ADVERSE EVENT, NOT A PRODUCT PROBLEM. THE REPORTER INDICATED THAT THE SURGEON TRIED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND THE LENS WAS TORN DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS EXPLANTED AND EXCHANGED ON THE SAME DAY WITH A LENS OF THE SAME SIZE, MODEL, AND DIOPTER. THERE WAS NO REPORT OF ANY PATIENT INJURY AND THE PROBLEM WAS RESOLVED AFTER THE LENS EXCHANGE. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191942&pc=MTA | 2023826-2016-01800 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS, -10.5/2.5/126 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 AND IT WAS NOTED THAT THE LENS WAS DAMAGED, THE LENS WAS EXPLANTED ON (B)(6) 2016. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: THE TORN LENS WAS EXCHANGED. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6249963&pc=MTA | 2023826-2017-00067 | 2016-11-23 | Injury | STAAR SURGICAL COMPANY | 2017-01-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO A LENS HAPTIC TEAR DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL AND SIZE, AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6254539&pc=MTA | 2023826-2017-00088 | 2016-12-03 | Injury | STAAR SURGICAL COMPANY | 2017-01-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/075 DIOPTER, TORE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL AND SIZE AND THE PROBLEM WAS RESOLVED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PREVIOUS INFORMATION INCORRECT, IT IS AN ADVERSE EVENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361177&pc=MTA | 2023826-2017-00272 | 2016-12-09 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/1.5/90 DIOPTER, TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A NEW LENS WITH THE SAME MODEL AND SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED AND ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: CORRECTED FROM ADVERSE EVENT TO PRODUCT PROBLEM. (CORRECTED): THE DOCTOR STATED THAT THE ICL DIDN'T MOVE "SMOOTHLY" FORWARD IN THE CARTRIDGE TIP. HE NOTED THE "FOLDED SHAPE OF THE ICL IN THE TIP LOOKED DISTORTED." THE DOCTOR EJECTED THE LENS, PUSHING IT WITH THE PLUNGER, TEARING THE LENS. THERE WAS NO PATIENT CONTACT. THE CARTRIDGE WAS NOT EVER LOADED INTO THE INJECTOR. AS OF THE DATE OF SUPPLEMENTAL MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. (PREVIOUS): STATED THAT THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION-LENS RETURNED IN A CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN, HAPTIC PIECE MISSING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6289977&pc=MTA | 2023826-2017-00162 | 2016-12-17 | Injury | STAAR SURGICAL COMPANY | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, ON (B)(6) 2016. THIS BACKUP LENS WAS IMMEDIATELY PREPARED AND THE SURGERY WAS SUCCESSFULLY COMPLETED WITH THE STATUS OF THE EYE BEING "VERY GOOD" ON (B)(6) 2016. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6332671&pc=MTA | 2023826-2017-00190 | 2017-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/-1.5/75 DIOPTER, WHICH TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS NOT IMPLANTED DUE TO THE ICL LENS BEING STUCK IN THE CARTRIDGE (TRAILING HAPTIC). THE SURGEON "HAD TO REMOVE THE INJECTED ICL AND NOTICED BREAKAGE OF THE ONE EDGE OF THE HAPTIC. THE RIGHT EYE (OD) LENS IMPLANT IS DOING WELL" WHILE THE LEFT EYE (OS) LENS IS BEING "SENT IN THE SAME CONTAINER TO STAAR SURGICAL REPRESENTATIVE" AS OF (B)(6) 2017. THE SURGEON REQUESTED FOR A NEW ICL LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429973&pc=MTA | 2023826-2017-00466 | 2017-01-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/1.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED DUE TO TEAR/BREAK DURING INJECTION INTO THE EYE ON THE SAME DAY. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/1.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS TORE/BROKE DURING INJECTION INTO THE EYE AND WAS EXCHANGED DURING THE SAME SURGERY. THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AS WELL AS DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). RESULT CODE : WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531873&pc=MTA | 2023826-2017-00482 | 2017-02-06 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME SURGERY THE LENS WAS EXCHANGED FOR THE SAME SIZE AND SIMILAR DIOPTER LENS DUE TO TEAR/BREAK DURING INJECTION INTO THE EYE. THE REPORTER STATED THAT "BAD LOADING" WAS THE CAUSE OF THE TEAR. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS BUT DEBRIS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684200&pc=MTA | 2023826-2017-01104 | 2017-05-22 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON THE SAME DAY, THE LENS WAS EXCHANGED DUE TO LENS TEAR/BEAK DURING INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. ATTEMPTS TO CONTACT TO THE CUSTOMER FOR ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE IMPLANT AND EXCHANGE WERE DONE INTRAOPERATIVELY. THE REPORTER STATED THAT THE SURGEON "PULLED THE LEN WITH TOO MUCH POWER, SO THE TEAR WAS BROKEN." (B)(4). Manufacturer Narrative: THE DIOPTER IS -13.0. DATE IS (B)(6) 2017 FOR FOLLOW UP #1. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE LENS MISSING PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF LABEL USE (BELOW 21 YEARS OF AGE AT THE TIME OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320741&pc=MTA | 2023826-2018-00325 | 2017-07-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2017. ON THE SAME DATE THE LENS WAS EXPLANTED. IT IS UNKNOWN WHETHER EXPLANT OCCURRED INTRA-OPERATIVELY OR POST-OPERATIVELY. THE SURGEON PLANS TO EXCHANGE THE LENS ONCE THE REPLACEMENT LENS HAS BEEN RECEIVED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE LENS WAS REMOVED INTRA-OPERATIVELY. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC HAD A TEAR ON IT AND WAS BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). TWENTY-FOUR-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.92MM)]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6817510&pc=MTA | 2023826-2017-01255 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/084 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND NOTED THE LENS WAS TORN. THE LENS WAS REMOVED WITH NO APPARENT INJURY . THE REPORTER INDICATED THE LENS WAS DAMAGED DURING THE LOADING PROCESS. THE LENS WAS NOT EXCHANGED WITH ANOTHER LENS. Manufacturer Narrative: ON (B)(6) 2017 A NEW LENS OF SAME MODEL AND SIMILAR DIOPTER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, HAPTIC BROKEN AND FIBERS PRESENT ON THE LENS SURFACE. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890300&pc=MTA | 2023826-2017-01474 | 2017-08-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS EXCHANGED FOR A SIMILAR LENS ON THE SAME DAY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: PATIENT IDENTIFIER: "UNK" IS CORRECTED TO "(B)(6)". DATE OF BIRTH: (B)(6). SEX: IT IS CORRECTED FROM "NO INFORMATION" TO "FEMALE". THE LENS WAS EXCHANGED FOR A SIMILAR LENS ON THE SAME DAY; INTRAOPERATIVELY. TYPE OF REPORTABLE EVENT: IT IS CORRECTED TO A "MALFUNCTION" FROM" SERIOUS INJURY". (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890301&pc=MTA | 2023826-2017-01473 | 2017-08-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON, DURING IMPLANTATION INTO THE EYE, NOTICED THAT THE LENS TORE. THE LENS WAS EXPLANTED ON THE SAME DAY AND A BACK-UP LENS OF THE SAME MODEL WAS USED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE WITH CLEAR SURGICAL RESIDUE /DEBRIS ON LENS. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC AND CLEAR AND RED RESIDUE ON LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: PATIENT IDENTIFIER: "UNK" IS CORRECTED TO "(B)(6)". DATE OF BIRTH: (B)(6). SEX: "NO INFORMATION" IS CORRECTED TO "FEMALE". THE LENS WAS EXCHANGED FOR A SIMILAR LENS ON THE SAME DAY; INTRAOPERATIVELY. TYPE OF REPORTABLE EVENT: IT IS CORRECTED TO A "MALFUNCTION" FROM "SERIOUS INJURY". (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6908161&pc=MTA | 2023826-2017-01492 | 2017-08-29 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE. DURING INJECTION, THE LENS TORE. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND THEN EXCHANGED FOR A SIMILAR LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THIS EVENT IS CHANGED FROM AN "ADVERSE EVENT" TO " PRODUCT PROBLEM". THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 -6.5 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. DURING INJECTION THE LENS TORE. THE REPORTER STATED THAT "INITIAL IMPLANT AND EXCHANGE OF THE LENS WERE PERFORMED AT SAME TIME, IN ONE SURGERY". THE PROBLEM WAS RESOLVED. IMPLANTED DATE: IMPLANT DATE IS (B)(6) 2017. TYPE OF REPORT: IT IS CHANGED FROM "SERIOUS INJURY" TO "MALFUNCTION". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419953&pc=MTA | 2023826-2018-00529 | 2017-09-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE INJECTING A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. ON THE SAME DATE ((B)(6) 2017) THE LENS WAS EXPLANTED. IT IS UNKNOWN AS TO WHETHER THIS WAS PERFORMED INTRAOPERATIVELY OR IN A 2ND SURGERY. ON A FUTURE DATE (DATE UNKNOWN), AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE TORN PIECE IS MISSING. CLAIM # (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT IT WAS UNKNOWN WHETHER THE LENS WAS REMOVED INTRAOPERATIVELY OR NOT. ALSO, AND ALTERNATE LENS WAS IMPLANTED WITH AN IMPLANT DATE OF UNK. CORRECTED INFORMATION-THE ORIGINAL LENS WAS REMOVED INTRAOPERATIVELY AND AN ALTERNATE LENS WAS IMPLANTED ON (B)(6) 2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050895&pc=MTA | 2023826-2017-01705 | 2017-09-27 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -15.0/+2.5/100 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. DURING THE SAME SURGERY, THE SURGEON EXPLANTED THE LENS AS HE SAW THAT THE LENS TORE BEFORE IT WAS IMPLANTED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED Manufacturer Narrative: A WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. CORRECTED INFORMATION: IT WAS REPORTED AS PRODUCT PROBLEM WITH NO ADVERSE EVENTS. THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A A 12.1MM VTICMO12.1, -15.0/+2.5/100 (SPHERE, CYLINDER, AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) THAT A LENS TEAR WAS NOTICED PRIOR TO IMPLANTATION. THIS OCCURRED ON (B)(6) 2017. ON (B)(6) 2017, A SIMILAR LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FIBERS ON LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183448&pc=MTA | 2023826-2018-00049 | 2017-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) (B)(6) 2017. AFTER INSERTION, THE SURGEON NOTICED THAT THE LENS WAS TORN. AS SOON AS HE NOTICED, HE REMOVED THE LENS AND PERFORMED THE EXCHANGE SURGERY WITH A SIMILAR LENS AFTER AN HOUR. THE SURGERY WAS SUCCESSFUL. THE SURGEON THOUGH THAT THE LENS WAS TORN WHEN IT WAS LOADED INTO THE CARTRIDGE. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT, A HAPTIC PIECE MISSING, AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND WHITE RESIDUE ON LENS). (B)(4) Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050831&pc=MTA | 2023826-2017-01753 | 2017-10-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), DUE TO A LENS/BREAK DURING DELIVERY INTO THE EYE. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED INTRAOPERATIVELY, A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) DUE TO LENS TEAR DURING INJECTION/ DELIVERY INTO THE EYE. THE REPORTER STATED THAT THE "FOAM TIP BENT LEADING TO CUTTING THE LENS DURING DELIVERY". THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LENS OF SAME MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED TO "PRODUCT PROBLEM" FROM "ADVERSE EVENT". "THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LENS OF SAME MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED" IS CORRECTED TO "A NEW LENS OF SAME MODEL AND DIOPTER WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069386&pc=MTA | 2023826-2017-01803 | 2017-10-26 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), AND THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THE LENS WAS REMOVED IMMEDIATELY AND THE SURGERY WAS ABORTED. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40 AND THE PATIENT IS DOING WELL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245223&pc=MTA | 2023826-2018-00195 | 2017-12-25 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/172 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2017. IT IS UNKNOWN WHETHER THE LENS WAS EXPLANTED INTRA-OPERATIVELY OR IN A 2ND SURGERY. Manufacturer Narrative: THE LENS WAS EXPLANTED INTRAOPERATIVELY. B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268475&pc=MTA | 2023826-2018-00180 | 2018-01-16 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS DAMAGED DURING THE LOADING PROCESS AND WAS RECOGNIZED BY THE SURGEON POSTOPERATIVELY AFTER DILATION. THIS WAS DUE TO USER TECHNIQUE ERROR WHILE LOADING THE LENS. Manufacturer Narrative: THE LENS WAS THEN EXCHANGED ON (B)(6) 2018 FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. THE REPORTED STATED THAT "ALL IS FINE NOW". CORRECTED DATA: THE IMPLANTATION DATE IN THE INITIAL MDR IS CORRECTED FROM (B)(6) 2017 TO (B)(6) 2017. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300925&pc=MTA | 2023826-2018-00289 | 2018-01-28 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00/+2.50/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. REPORTEDLY, THE LENS REMAINS IN THE EYE. THE SURGEON PLANS TO EXCHANGE THE LENS ONCE THE REPLACEMENT LENS HAS BEEN RECEIVED. Manufacturer Narrative: THE LENS WAS EXCHANGED AND THE PATIENT IS DOING FINE. (B)(4). IT SHOULD BE NOTED A SUTURE WAS USED DURING INITIAL IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7364232&pc=MTA | 2023826-2018-00445 | 2018-02-19 | Injury | STAAR SURGICAL COMPANY | 2018-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER AND THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. THE REPORTER INDICATED IT WAS UNKNOWN IF THERE WAS ANY PATIENT CONTACT BUT THERE WAS NO INDICATION OF ANY INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2018. IT IS CORRECTED FROM "PRODUCT PROBLEM" TO "ADVERSE EVENT". DATE OF THIS REPORT: IT IS CORRECTED FROM "26-FEB-2018" TO "24-FEB-2018". THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM, VICMO12.6 IMPLANTABLE COLLAMER LENS OF, -13.0 DIOPTER AND THE LENS TORE DURING IMPLANTATION INTO THE PATIENT'S EYE ON (B)(6) 2018. IT IS UNKNOWN IF THERE WAS ANY PATIENT CONTACT BUT THERE WAS NO INDICATION OF ANY INJURY. ANOTHER LENS OF SAME MODEL WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. IT IS UNCLEAR IF THE LENS WAS EXCHANGED INTRAOPERATIVELY OR EXTRAOPERATIVELY. THE CAUSE OF THE EVENT WAS UNKNOWN. IT IS CORRECTED FROM "26-FEB-2018" TO "24-FEB-2018" IN THE INITIAL MDR. TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "MALFUNCTION" TO "SERIOUS INJURY". (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7555374&pc=MTA | 2023826-2018-00806 | 2018-03-01 | Injury | STAAR SURGICAL COMPANY | 2018-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2,-9.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE (DURING INJECTION/DELIVERY INTO THE EYE). THE LENS WAS REMOVED AND EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION- LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND TEARS ON A HAPTIC, A PIECE OF HAPTIC MISSING, AND DEBRIS ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7503535&pc=MTA | 2023826-2018-00680 | 2018-03-08 | Injury | STAAR SURGICAL COMPANY | 2018-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICTED THAT A 12.6MM VICMO12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND EXCHANGED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND PRESENCE OF RESIDUE ON THE LENS SURFACE. THE LENS WAS MISSING THE TORN PIECE OF HAPTIC. DATE OF THIS REPORT: IT IS CORRECTED TO "11-APR-28" IN THE INITIAL MDR. THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER TORE DURING INJECTION/ DELIVERY INTO THE EYE. THE LENS WAS EXPLANTED AND A NEW LENS WAS IMPLANTED ON THE SAME DAY. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557855&pc=MTA | 2023826-2018-00816 | 2018-04-05 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -18.0 INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. REPORTEDLY, THE LENS WAS REPOSITIONED AND EXCHANGED FOR AN ALTERNATE LENS. THE DATE OF LENS IMPLANT AND EXCHANGE IS (B)(6) 2018, HOWEVER IT IS NOT KNOWN IF THE EXCHANGE OCCURED INTRAOPERATIVELY OR NOT. ATTEMPTS TO OBTAIN ADDITIONAL CLARIFICATION HAVE NOT BEEN SUCCESSFUL. THE PROBLEM IS NOTED AS RESOLVED AFTER EXCHANGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7599035&pc=MTA | 2023826-2018-00910 | 2018-04-26 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6, -3.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE, THE LENS WAS EXPLANTED . ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC TORN AND PRESENCE OF FOREIGN MATERIAL (RESIDUE) ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850682&pc=MTA | 2023826-2017-01406 | 2017-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Activation Failure | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0/+3.0/X086 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THE SURGEON THEN EXCHANGED THE LENS FOR THE SAME SIZE, SIMILAR DIOPTER LENS. THIS OCCURRED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: AFTER THE LENS WAS IMPLANTED, THE DOCTOR DISCOVERED IT WAS "FLIPPED INSIDE THE EYE." INTRAOPERATIVELY, THE SURGEON REMOVED THE LENS AND THE LENS BROKE DURING EXPLANT. AN ALTERNATE LENS WAS IMPLANTED. (B)(4). Manufacturer Narrative: PATIENT POST-OP BCVA WAS 20/20 ON (B)(6) 2017. DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.97MM)]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6155425&pc=MTA | 2023826-2016-01743 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-12-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS ATTEMPTING TO IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE AND THE LENS WAS NOTED TO BE DAMAGED. IT WAS INDICATED THERE WAS NO PATIENT CONTACT WITH THE LENS. Manufacturer Narrative: THE LENS WAS DAMAGED DURING INJECTION. PATIENT CONTACT IS UNKNOWN. THE LENS WAS EXCHANGED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. A WORK ORDER SEARCH WAS PERFORMED AND FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). SECONDARY INTERVENTION, LENS EXCHANGED AND PROBLEM RESOLVED. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE OPERATION WAS PERFORMED ON THE PATIENT'S RIGHT EYE (OD), NOT LEFT EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7326990&pc=MTA | 2023826-2018-00335 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS WAS TO BE EXPLANTED AND EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: OUTCOMES ATTACHED TO ADVERSE EVENT: IT IS CORRECTED FROM "OTHER SERIOUS" TO "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/ DAMAGE". THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON A (B)(6) 2018. THE REPORTER INDICATED THAT THE LENS WAS TO BE EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAVE BEEN FORTHCOMING. (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047988&pc=MTA | 2023826-2015-01144 | 2014-12-17 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Scratched Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND SCRATCHES WERE OBSERVED ON THE LENS. THE LENS WAS REMOVED AND EXCHANGED FOR ANOTHER LENS AND THIS RESOLVED THE PROBLEM. NO PATIENT INJURY WAS REPORTED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5985275&pc=MTA | 2023826-2016-01336 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS NOTED TO BE TORN AND WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH THE SAME DIOPTER. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.95MM AT THE TIME OF IMPLANTATION. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5940511&pc=MTA | 2023826-2016-01249 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY | 2016-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/090 DIOPTER, AND THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED AND ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED. THERE WAS NO REPORT OF ANY APPARENT INJURY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/28. Manufacturer Narrative: IMPLANT DATE (B)(6) 2016; EXPLANT DATE (B)(6) 2016; EXCHANGE DATE (B)(6) 2016. DEVICE EVALUATION: THE LENS WAS RETURNED IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN (B)(4). Manufacturer Narrative: EXCHANGE DATE (B)(6) 2016; PREVIOUS DATE INCORRECT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826216&pc=MTA | 2023826-2016-00977 | 2016-06-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS INSERTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016 AND THE LENS TORE. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/ SPLIT, HAPTIC BROKEN AND LENS BROKEN IN TWO PIECES. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONS. THE REPORTER STATED INDICATED THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTICED A LENS TEAR DURING DELIVERY OF ICL INTO THE ANTERIOR CHAMBER. NEXT DAY, THE LENS WAS REMOVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5986996&pc=MTA | 2023826-2016-01340 | 2016-07-05 | Injury | STAAR SURGICAL COMPANY | 2016-09-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS 12.5/+3.5/X096 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A LENS TEAR DURING DELIVERY INTO THE EYE. THE LENS WAS REPLACED AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. THE IMPLANT AND EXPLANT DATE IS (B)(6) 2016. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6001955&pc=MTA | 2023826-2016-01396 | 2016-08-02 | Injury | STAAR SURGICAL COMPANY | 2016-10-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, AND THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY AND WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THIS CASE IS AN ADVERSE EVENT NOT A PRODUCT PROBLEM. THE CORRECT REPORT SOURCE IS DISTRIBUTOR NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5951517&pc=MTA | 2023826-2016-01260 | 2016-08-09 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.0 DIOPTER, AND THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY AND WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS. THE EXCHANGE SOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7508619&pc=MTA | 2023826-2018-00722 | 2010-06-22 | Injury | STAAR SURGICAL COMPANY AG | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2010. THE REPORTER INDICATED THE LENS HAD A LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A LONGER LENS (13.7) BUT DIFFERENT MODEL AND DIOPTER - VTICM5_13.7, -8.00/+1.5/X089 (SPHERE/CYLINDER/AXIS). THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). CORRECTED DATA: PRODUCT PROBLEM TO ADVERSE EVENT. MALFUNCTION TO SERIOUS INJURY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6290476&pc=MTA | 2023826-2017-00159 | 2016-08-05 | Injury | STAAR SURGICAL COMPANY | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 VICMO13.7MM COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074706&pc=MTA | 2023826-2016-01592 | 2016-09-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND THE ANTERIOR CHAMBER WAS SHALLOW. THE LENS REMAINS IMPLANTED BUT WILL BE REMOVED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTED DATA:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024725&pc=MTA | 2023826-2016-01430 | 2016-09-23 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. Manufacturer Narrative: DEVICE REMAINS IMPLANTED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6266899&pc=MTA | 2023826-2017-00106 | 2016-12-08 | Injury | STAAR SURGICAL COMPANY | 2017-01-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON THIS DATE, THE PATIENT EXPERIENCED LOW VAULT. THE LENS REMAINS IMPLANTED. THE PATIENT IS BEING CLOSELY MONITORED BY THE PHYSICIAN TO DETERMINE WHETHER THE LENS WILL BE EXCHANGED AT A LATER DATE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). CONCLUSION CODE: NO LONGER APPLICABLE, PRODUCT RETURNED ON 07 MAR 2018. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL OBSERVATION FOUND THE LENS WITH NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6267285&pc=MTA | 2023826-2017-00107 | 2016-12-08 | Injury | STAAR SURGICAL COMPANY | 2017-01-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON THIS DATE, THE PATIENT EXPERIENCED LOW VAULT. THE LENS REMAINS IMPLANTED. THE PATIENT IS BEING CLOSELY MONITORED BY THE PHYSICIAN TO DETERMINE WHETHER THE LENS WILL BE EXCHANGED AT A LATER DATE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: CHANGED FROM ADVERSE EVENT TO PRODUCT PROBLEM. (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH RESIDUE ON THE LENS SURFACE AND THE LENS HAVING A CURVED SHAPE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6396110&pc=MTA | 2023826-2017-00365 | 2016-12-27 | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. PATIENT DEVELOPED LOW VAULT IN THE RIGHT EYE ONE WEEK AFTER THE SURGERY. AS OF TODAY, (B)(6) 2017, THE LENS HAS YET TO BE EXCHANGED WITH A LONGER LENS. Manufacturer Narrative: UPDATED INFORMATION: LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LONGER ONE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6396126&pc=MTA | 2023826-2017-00373 | 2016-12-27 | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON REPORTED THAT THE "PATIENT DEVELOPED LOW VAULT IN OS ONE WEEK AFTER SURGERY AND WILL NEED TO EXCHANGE LENS WITH BIGGER LENS." THE LENS IS STILL IMPLANTED AS OF (B)(6) 2017. Manufacturer Narrative: UPDATED INFORMATION: THE LENS WAS EXCHANGED FOR A LARGER LENS ON (B)(6) 2017. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A SEALED VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS BOX WAS RETURNED OPEN, LENS VIAL AND CASE SEAL NOT BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). INCORRECT SERIAL NUMBER, CANNOT CHECK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6441301&pc=MTA | 2023826-2017-00517 | 2017-02-08 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/X126 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND WILL BE REMOVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE LENS HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021455&pc=MTA | 2023826-2017-01674 | 2017-03-15 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULTING WAS NOTICED. THE LENS CURRENTLY REMAINS IMPLANTED. THE LENS EXCHANGE SURGERY IS PLANNED FOR THE END OF 2017 OR AT THE BEGINNING OF 2018. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 OF -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULT WAS NOTICED AND THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. EXPLANTED DATE: (B)(6) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050891&pc=MTA | 2023826-2017-01703 | 2017-03-15 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON 14 MAR 2017. EXCESSIVE VAULTING WAS NOTICED. LENS EXCHANGE IS PLANNED EITHER IN THE END OF 2017 OR BEGINNING OF 2018 DUE TO THE PATIENT'S CONVENIENCE. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 AND EXCHANGED FOR A SHORTER LENS. THE EXCHANGE RESOLVED THE PROBLEM. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6513187&pc=MTA | 2023826-2017-00693 | 2017-03-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, DIOPTER -5.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE NEXT DAY, THE PATIENT BEGAN TO EXPERIENCE LOW VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: UPDATED INFORMATION: ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6609995&pc=MTA | 2023826-2017-00920 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, DIOPTER -8.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON IS PLANNING TO EXCHANGE THE LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT. THE PROBLEM IS NOT YET RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. ATTEMPTS TO CONTACT SURGEON HAVE NOT YET BEEN SUCCESSFUL. Manufacturer Narrative: (B)(6) 2017. (B)(4). Manufacturer Narrative: -PREVIOUS: STATED LENS REMAINS IMPLANTED. CURRENT: LENS WAS EXCHANGED ON (B)(6) 2017. EXCHANGE INFORMATION (LENS SIZE, PROBLEM RESOLUTION) IS UNKNOWN. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4).CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.98MM)]. (B)(4). Manufacturer Narrative: UPDATED: THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6756913&pc=MTA | 2023826-2017-01256 | 2017-07-05 | Injury | STAAR SURGICAL COMPANY | 2017-08-01 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED LOW VAULT AND THE LENS IS PLANNED TO BE EXCHANGED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6770823&pc=MTA | 2023826-2017-01288 | 2017-07-13 | Injury | STAAR SURGICAL COMPANY AG | 2017-08-07 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/2.5/75 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULTING. THE LENS IS PLANNED FOR EXCHANGE. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018 FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CENTRIFUGE VIAL WITH PRESENCE OF RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6770898&pc=MTA | 2023826-2017-01289 | 2017-07-13 | Injury | STAAR SURGICAL COMPANY AG | 2017-08-07 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULTING. THE LENS IS PLANNED FOR AN EXCHANGE. Manufacturer Narrative: CORRECTION DATA: DISREGARD MFR REPORT# 2023826-2018-01289 (FOLLOW-UP #:1). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS -11.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THAT THE "UCVA IS 20/20; VAULTING IS GOOD AND PATIENT IS HAPPY". (B)(4). THIS INFORMATION WAS REPORTED ON 25-MAY-2018 UNDER MFR REPORT# 2023826-2018-01289 (FOLLOW-UP#: 1) IN ERROR. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6811835&pc=MTA | 2023826-2017-01364 | 2017-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/169 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND THE LENS IS PLANNED TO BE EXCHANGED AT THE END OF (B)(6). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND RED RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836437&pc=MTA | 2023826-2017-01378 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/032 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS WITHIN THE NEXT COUPLE OF WEEKS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017 FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AT ONE WEEK POST-OP, THE PATIENT'S CONDITION WAS GOOD, WITH UCVA -20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850658&pc=MTA | 2023826-2017-01400 | 2017-08-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2/114 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. INTRAOPERATIVELY, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS FOUND ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: CORRECTED DATA: ON (B)(6) 2017, POST-REMOVAL OBSERVATION OF ANISOCORIA AND UNCORRECTED VISUAL ACUITY (UCVA) OF 20/400 WAS OBSERVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135241&pc=MTA | 2023826-2017-01927 | 2017-09-04 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON REPORTS LOW VAULT, STATES THE "CENTRAL REAR" TOUCHES THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183426&pc=MTA | 2023826-2018-00044 | 2017-09-06 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. AFTER IMPLANTATION, IT WAS NOTICED THAT THERE WAS INSUFFICIENT VAULT. THERE WAS NO PATIENT INJURY. IT IS PLANNED FOR THE LENS TO BE REMOVED AND REPLACED FOR A LONGER LENS. ATTEMPTS AT OBTAINING ADDITION INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LONGER LENS. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. DEVICE EVALUATION - THE PRODUCT WAS RETURNED DRY, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, THE HAPTIC TORN, A MISSING PIECE OF HAPTIC AND WHITE RESIDUE ON LENS. (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183444&pc=MTA | 2023826-2018-00042 | 2017-09-07 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. AFTER IMPLANTATION, IT WAS NOTICED THAT THERE WAS INSUFFICIENT VAULT. THERE WAS NO PATIENT INJURY. IT IS PLANNED FOR THE LENS TO BE REMOVED AND REPLACED FOR A LONGER LENS. ATTEMPTS AT OBTAINING ADDITION INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED ON 21 DEC 2017 FOR A LONGER LENS. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. DEVICE EVALUATION - THE PRODUCT WAS RETURNED DRY, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, THE HAPTIC TORN, A MISSING PIECE OF HAPTIC AND WHITE RESIDUE ON LENS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050893&pc=MTA | 2023826-2017-01708 | 2017-10-19 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED THAT THE SURGEON IMPLANTED A MICL13.7, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. IT WAS NOTICED THAT THE LENS WAS TOO BIG. IT IS UNSURE IF THE LENS IS GOING TO BE EXCHANGED FOR A SHORTER LENS. THERE IS NO PATIENT INJURY. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7 IMPLANTABLE COLLAMER LENS OF -8.0 DIOPTER IN THE PATIENT'S EYE ON (B)(6) 2017. IT WAS NOTICED THAT THE LENS WAS TOO BIG. THERE IS NO PATIENT INJURY. THE LENS WAS EXPLANTED. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. EXPLANTED IN 2017. DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR AND RED RESIDUE ON THE LENS SURFACE. (B)(4). METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: DATE OF EVENT: IT IS CORRECTED FROM (B)(6) 2017 TO (B)(6) 2017. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7 IMPLANTABLE COLLAMER LENS OF -8.0 DIOPTER IN THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON(B)(6) 2017 DUE TO EXCESSIVE VAULT, NARROWING OF ANGLES AND SHALLOWING OF ACD. THERE WAS NO PATIENT INJURY. THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7286008&pc=MTA | 2023826-2018-00247 | 2017-10-22 | Injury | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REMOVED ON THE SAME DAY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: IT IS CHANGED FROM "ADVERSE EVENT" TO "PRODUCT PROBLEM". THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENS DURING THE SAME SURGERY AS IMPLANTATION, DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. TYPE OF REPORTABLE EVENT: IT IS CHANGED FROM "SERIOUS INJURY" TO " MALFUNCTION". (B)(4). CORRECTED DATA: THE SURGERY TOOK PLACE IN THE PATIENT'S LEFT EYE (OS) AND NOT THE RIGHT EYE (OD) AS STATED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. (SECONDARY SURGERY, LENS EXCHANGE). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LATER IN A DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7109937&pc=MTA | 2023826-2017-01882 | 2017-11-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM MICL13.2 LENS, DIOPTER -14.0 WAS TO BE EXPLANTED FROM THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2017, POSSIBLY DUE TO AN OVER VAULT. ATTEMPTS TO OBTAIN ADDITIONAL OR UPDATED INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: PATIENT IDENTIFIER - CORRECTED FROM UNK TO KO. DATE OF BIRTH: CORRECTED TO (B)(6). IMPLANTED DATE: CORRECTED TO (B)(6) 2017. EXPLANTED DATE: CORRECTED FROM (B)(6) 2017 TO NOT APPLICABLE. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 LENS, -14 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 11 NOV 2017 DUE TO "OVER VAULT". THE SYMPTOMS OF "OVER VAULT" WERE NOTED ON (B)(6) 2017 AND TO DATE THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7109939&pc=MTA | 2023826-2017-01881 | 2017-11-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM MICL 13.2 LENS, DIOPTER -15.5 WAS TO BE EXPLANTED FROM THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017, POSSIBLY DUE TO AN OVER VAULT. ATTEMPTS TO OBTAIN ADDITIONAL OR UPDATED INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON PLANS TO EXCHANGE THE LENS DUE TO EXCESSIVE VAULT. THE REPORTER STATES THAT "THE SYMPTOMS BEGAN ON (B)(6) 2017". THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7177066&pc=MTA | 2023826-2018-00007 | 2017-12-01 | Injury | STAAR SURGICAL COMPANY | 2018-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/176 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON 13 MAR 2018 AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7177114&pc=MTA | 2023826-2018-00006 | 2017-12-01 | Injury | STAAR SURGICAL COMPANY | 2018-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/168 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.0/3.0/168 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THAT "ALL IS FINE" AND THAT THE PATIENT IS HAPPY. EXPLANTED DATE: (B)(6) 2018. (B)(4) Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED FROM "ADVERSE EVENT" TO " PRODUCT PROBLEM". TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO " MALFUNCTION". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403688&pc=MTA | 2023826-2018-00513 | 2018-01-09 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULTING. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403692&pc=MTA | 2023826-2018-00515 | 2018-01-09 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+3.0/097 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULTING. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED WITH A SHORTER LENS ON (B)(6) 2018. THE PROBLEM WAS RESOLVED AND "PATIENT IS HAPPY - ALL IS FINE". (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300924&pc=MTA | 2023826-2018-00284 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/010 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS EXCESSIVE VAULT. CURRENTLY, THE LENS REMAINS IMPLANTED, HOWEVER IT IS PLANNED TO BE EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. UPDATED/CORRECTED INFORMATION-THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7749504&pc=MTA | 2023826-2018-01155 | 2018-02-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICMO12.6, -6/1/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS, THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6 VTICMO_12.6 -6/1/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND REMOVED INTRAOPERATIVELY DUE TO EXCESSIVE VAULT ON (B)(6) 2018. A SHORTER REPLACEMENT LENS WAS IMPLANTED POST-OPERATIVELY ON (B)(6) 2018, RESOLVING THE PROBLEM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC DEFORMED. THERE WAS RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, EXCHANGE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7518745&pc=MTA | 2023826-2018-00697 | 2018-02-13 | Injury | STAAR SURGICAL COMPANY | 2018-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENGTH LENS DURING THE SAME SURGERY (INTRAOPERATIVELY) DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED TO (B)(6) 2018 IN THE PREVIOUS SUPPLEMENTAL MDR (FOLLOW-UP #2). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICM5_13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON THE SAME DAY DUE TO OBSERVATION OF EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE ON THE LENS SURFACE, A TEAR ON THE OPTIC AND HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336753&pc=MTA | 2023826-2018-00378 | 2018-02-14 | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED ON (B)(6) 2018 FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RED RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7518671&pc=MTA | 2023826-2018-00696 | 2018-02-15 | Injury | STAAR SURGICAL COMPANY | 2018-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENGTH LENS DURING THE SAME SURGERY (INTRAOPERATIVELY) DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTION: THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7720627&pc=MTA | 2023826-2018-01114 | 2018-04-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS INSERTED, REMOVED AND EXCHANGED DURING THE SAME SURGERY. CORRECTED DATA: ADVERSE EVENT TO PRODUCT PROBLEM, SERIOUS INJURY TO MALFUNCTION (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647922&pc=MTA | 2023826-2018-00975 | 2018-04-26 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM MICL12.6, -14.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS REMOVED FROM THE PATIENT'S RIGHT EYE (OD) AND REPLACED DUE TO WRONG SIZE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7648087&pc=MTA | 2023826-2018-00974 | 2018-04-26 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM MICL12.6, -14.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE PATIENT'S LEFT EYE (OS) AND REPLACED DUE TO WRONG SIZE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM MICL 12.6 IMPLANTABLE COLLAMER LENS -14.5 DIOPTER WAS EXPLANTED FROM THE PATIENT'S LEFT EYE (OS) DUE TO WRONG SIZE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7747114&pc=MTA | 2023826-2018-01161 | 2018-06-17 | Injury | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.50/+3.0/016 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT WAS OBSERVED. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED, IT WAS REPORTED THAT THE SURGEON PLANS TO EXCHANGE THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 -13.50/+3.0/016 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SMALLER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7839690&pc=MTA | 2023826-2018-01324 | 2018-06-19 | Injury | STAAR SURGICAL COMPANY | 2018-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -15.5/+2.0/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LOW VAULT WAS OBSERVED, THE LENS IS PLANNED TO BE EXCHANGED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS USER ERROR . ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION DATA: DEVICE MANUFACTURE DATE CORRECTED TO "07-FEB-2018" IN THE INITIAL MDR. ADDITIONAL DATA: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -15.50/+2.0/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018 DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED, "VAULT IS A BIT LOW, BUT IT'S THE BIGGEST LENS". IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS USER ERROR. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824068&pc=MTA | 2023826-2018-01274 | 2018-07-20 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED UPDATED STATEMENT-ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107291&pc=MTA | 2023826-2018-01721 | 2018-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. ON 20-NOV-2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE LENS WAS RETURNED DRY IN A CONTACT LENS CASE. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7995981&pc=MTA | 2023826-2018-01553 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/178 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM IS RESOLVED. "AFTER EXCHANGE ALL IS FINE/PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021154&pc=MTA | 2023826-2018-01580 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2018-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/055 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. "AFTER EXCHANGE ALL IS FINE/PATIENT IS HAPPY." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073836&pc=MTA | 2023826-2018-01628 | 2018-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -11.50/5.0/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -11.5/5.0/63 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT WAS OBSERVED. ON 03/01/2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT CODE 3191: SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8185176&pc=MTA | 2023826-2018-01879 | 2018-10-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS WAS REMOVED AND A REPLACEMENT LENS (SHORTER LENGTH) WAS IMPLANTED. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT CODE: "3191" NOT INCLUDED IN THE INITIAL MDR. THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 -11.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED AFTER THE REPLACEMENT LENS. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE: CORRECTED FROM 2692 TO 3191, SIGINIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ONLY, PREVIOUS CODE NOT APPLICABLE IN THE SUPPLEMENTAL MDR#1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8184966&pc=MTA | 2023826-2018-01880 | 2018-10-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -06.00DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS WAS REMOVED AND A REPLACEMENT LENS (SHORTER LENGTH) WAS IMPLANTED. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT CODE 3191: SIGNIFICANT REDUCTINO OF IRIDO-CORNEAL ANGLES. PATIENT CODE 2692: PREVIOUS CODE NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND A NEW LENS WAS IMPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE IMPLANT, REMOVAL, AND REPLACEMENT IMPLANT ALL OCCURED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND RESIDUE ON LENS. PATIENT CODE 3191: PREVIOUS CODE NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113346&pc=MTA | 2023826-2018-01754 | 2018-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS. THE LENS TO DATE REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: AGE OR DATE OF BIRTH-CORRECTED TO :(B)(6)1981. ADDITIONAL DATA: EVENT-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -12.50/+2.0/083 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. ON (B)(6)2018, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE REPORTER NOTED, "PATIENT IS HAPPY. NO FURTHER TREATMENT IS NECESSARY. LAST VISIT: (B)(6)2018; BCVA: 20/20; UCVA: 20/20; VAULT: 600UM". PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXCHANGED CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8404606&pc=MTA | 2023826-2019-00353 | 2018-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS REMAINS IMPLANTED. EXPLANTATION OF THE LENS IS NOT PLANNED TO DATE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTERR INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8215185&pc=MTA | 2023826-2019-00030 | 2018-12-11 | Injury | STAAR SURGICAL COMPANY | 2019-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED. CORRECTED DATA: LEFT EYE (OS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8215213&pc=MTA | 2023826-2019-00031 | 2018-12-11 | Injury | STAAR SURGICAL COMPANY | 2019-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A VICMO IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS; RESIDUE ON LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTED TO "06-DEC-2018" IN THE INITIAL MDR. CORRECTED TO "06-DEC-2018" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED. CORRECTED DATA: RIGHT EYE (OD). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298907&pc=MTA | 2023826-2019-00190 | 2018-12-23 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.00/+2.0/079 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD A LOW VAULT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019. THE PATIENT DECIDED NOT TO HAVE ANOTHER LENS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671670&pc=MTA | 2023826-2019-01021 | 2019-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM/ OUTCOMES ATTRIBUTED TO ADVERSE EVENT: CHANGED FROM PRODUCT PROBLEM IN THE INITIAL MDR TO ADVERSE EVENT WITH REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANABLE COLLAMER LENS OF -5.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT ON (B)(6) 2019. THE EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS UNKNOWN. PRODUCT EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656052&pc=MTA | 2023826-2019-00975 | 2019-04-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICMO 13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, OCT SHOWED A VAULT OF MORE THAN 1200 MICRONS, HOWEVER THE IRIDOCORNEAL ANGLES WERE OPEN AND INTRAOCULAR PRESSURE WAS 15 MM HG. ON (B)(6) 2019 ANOTHER OCT WAS PERFORMED AND THE VAULT INCREASED TO 1328 MICRONS. REFRACTION, PRESSURE AND IRIDOCORNEAL ANGLES WERE FINE. ON (B)(6) 2019 THE FOLLOWING INFORMATION WAS RECEIVED "THE PATIENTS (OD) IS CURRENTLY FINE, HER BCVA IS 20/20. ON NEXT DAYS SHE WILL BE INTERVENED AGAIN. WE ARE WAITING FOR THE REPLACEMENT LENS. DOCTOR HAS DECIDED TO EXCHANGE WITH SMALLER LENS (12.6MM)". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.00 DIOPTER,IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8662746&pc=MTA | 2023826-2019-00864 | 2019-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -01.75 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED, REPORTEDLY THE LENS WILL BE REMOVED AND EXCHANGED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM: SHOULD BE CORRECTED TO ADVERSE EVENT IN INITIAL MEDWATCH REPORT. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION SHOULD BE SELECTED. EVENT DESCRIPTION: THE LENS WAS EXCHANGED ON (B)(6) 2019. ONE DAY POST SURGERY PATIENT CONDITION IS FINE. BCVA IS 20/20. EXPLANT DATE: EXPLANTED DATE SHOULD BE (B)(6) 2019. TYPE OF REPORTABLE EVENT: SHOULD BE CORRECTED TO SERIOUS INJURY. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE MANUFACTURE DATE 19-MAR-2010 SHOULD BE CORRECTED TO 19-MAR-2019 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516589&pc=MTA | 2023826-2019-02474 | 2019-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED. ALPHAGAN PRESCRIBED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B3- "(B)(6) 2019" SHOUDL BE CORRECTED TO "UNK" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: "PRODUCT PROBLEM" SHOULD BE CORRECTED TO "ADVERSE EVENT" IN THE INITIAL MDR. B2: "REQUIRED INTERVENTION" SHOULD BE MARKED IN THE INITIAL MDR. H1: "MALFUNCTION" SHOULD BE CORRECTED TO "SERIOUS INJURY" IN THE INITIAL MDR. H6: PATIENT CODE 2692 SHOULD BE CORRECTED TO "CODE 3191- MEDICAL INTERVENTION" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941305&pc=MTA | 2023826-2019-01576 | 2019-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PATIENT CODE UPDATED, 3191: (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941317&pc=MTA | 2023826-2019-01575 | 2019-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE EVENT DESCRIPTION PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9200184&pc=MTA | 2023826-2019-01955 | 2019-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED, SURGEON PLANS TO EXCHANGE LENS WITH A SHORTER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. EXPLANT DATE: (B)(6) 2019. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. TYPE OF REPORTABLE EVENT: SERIOUS. DEVICE EVALUATION: LENS RETURNED DRY WRAPPED IN GAUZE INSIDE A BIOHAZARD BAG WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND GAUZE FIBERS ON LENS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT IS (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312703&pc=MTA | 2023826-2019-02116 | 2019-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING HIGH VAULTING AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM#: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4- EXPIRATION DATE: SHOULD BE CORRECTED TO "31-MAY-2022" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9483492&pc=MTA | 2023826-2019-02397 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE WAS OBSERVED AND THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2019, "PATIENT IS FINE, LENS WILL BE REPLACED (B)(6) 2020." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - "EXCESSIVE WAS OBSERVED" SHOULD BE CORRECTED TO "LOW VAULT WAS OBSERVED". CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. H6: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITIH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559338&pc=MTA | 2023826-2020-00039 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B1, B2-UPDATED. B5-UPDATED INFORMATION: ON 12-FEB-2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM IS RESOLVED. H6-PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869318&pc=MTA | 2023826-2020-00614 | 2020-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERV.ENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- BS THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON 06-FEB-2020. THE SURGEON REPORTS EXCESSIVE VAULT. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869335&pc=MTA | 2023826-2020-00613 | 2020-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE . B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D7- EXPLANT DATE: (B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10224707&pc=MTA | 2023826-2020-01390 | 2020-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D7 - UNKNOWN EXPLANT DATE. B5 - AT AN UNKNOWN DATE THE LENS WAS EXPLANTED DUE TO THE LENS WAS RETURNED BACK WITH NO ADDITIONAL INFORMATION. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271251&pc=MTA | 2023826-2020-01477 | 2020-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, SURGEON WILL EXCHANGE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE REPORT INDICATED A BIOPTICS IS PLANNED FOR THE PATIENT. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6199297&pc=MTA | 2023826-2016-01797 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT REPORTED A LOW VAULT. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION REGARDING LENS EXPLANT. IF ADDITIONAL INFORMATION IS RECEIVED, THEN A SUPPLEMENTAL REPORT WILL BE SENT. Manufacturer Narrative: LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS. PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6281924&pc=MTA | 2023826-2017-00175 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED EVO VISIAN ICL LENSES BILATERALLY IN A PATIENT AND BOTH EYES PRESENT HIGH VAULTING. THE LENSES REMAIN IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE RIGHT EYE AND OVER TIME THE ISSUE OF HIGH VAULTING WAS RESOLVED AND THE LENS IS FINE. THE LENS REMAINS IMPLANTED AND THE RESULT IS GREAT OVERALL. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7184499&pc=MTA | 2023826-2018-00060 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED THE LENS HAS AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED, AND THE SURGEON PLANS TO EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THE DATE OF BIRTH (B)(6) IS INCORRECT, SHOULD BE UNKNOWN. (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7467159&pc=MTA | 2023826-2018-00621 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS, WAS RECEIVED BY THE MANUFACTURER DAMAGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE LENS WAS DAMAGED DURING EXPLANTATION (AS THIS WAS AN EXPLANT CASE). THERE WAS NO PATIENT INJURY. THE PROBLEM WAS SOLVE AND THE PATIENT IS OKAY WITH 0.740 MICRONS OF VAULT. THE SECOND LENS WAS IMPLANTED THE SAME DAY (INTRAOPERATIVELY). (B)(4). CORRECTED INFORMATION COMPLAINT IS AN ADVERSE EVENT DUE TO A SECONDARY SURGICAL INTERVENTION. COMPLAINT IS A SERIOUS INJURY NOT A MALFUNCTION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: LENS EXCHANGE WAS NOT REPORTED. CLAIM # (B)(4). Manufacturer Narrative: EXPIRATION DATE UNK. THE MANUFACTURING SITE IS UNKNOWN BUT THE PRODUCT IS NOT MARKETED IN THE U.S. PRODUCT EVALUATION FOUND TWO LENSES RETURNED IN THE SAME VIAL. BOTH LENSES WERE RETURNED DRY WITH RESIDUE/DEBRIS ON THE LENSES. VISUAL INSPECTION FOUND THE HAPTIC OF THE FIRST LENS BENT AND DEBRIS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC OF THE SECOND LENS BROKEN, A PIECE OF THE HAPTIC MISSING, AND DEBRIS AND RED RESIDUE ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8668923&pc=MTA | 2023826-2019-01030 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER INDICATED THE SURGEON IS PLANNING TO EXPLANT THE LENS DUE TO BEING TOO LONG. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8668925&pc=MTA | 2023826-2019-01029 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER INDICATED THE SURGEON IS PLANNING TO EXPLANT THE LENS DUE TO BEING TOO LONG. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941313&pc=MTA | 2023826-2019-01569 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS HE INTENDS TO EXCHANGE THE LENS WITH A 13.7MM LENS, STATING "THERE IS NOT INADEQUATE VAULT" OF THE LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED INTRAOPERATIVELY WITH A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. THE PATIENT IS HEALING NICELY. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND A PATIENT RELATED FACTOR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID. VISUAL INSPECTION FOUND A TEAR ON THE LENS HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED IN THE RIGHT (OD) EYE. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941318&pc=MTA | 2023826-2019-01568 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS HE INTENDS TO EXCHANGE THE LENS WITH A 13.7MM LENS, STATING "THERE IS NOT INADEQUATE VAULT" OF THE LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: UPDATED INFORMATION: THE LENS WAS IMPLANTED THE LEFT (OS) EYE ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON REPORTING LOW VAULT (VALUE OF 0.5CT). AFTER THE EXCHANGE THE VAULT VALUE IS REPORTED AS 1.0CT AND THE SURGEON NOTES THE PATIENT IS "HEALING WELL." DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9158762&pc=MTA | 2023826-2019-01847 |  | Injury | STAAR SURGICAL COMPANY | 2019-10-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER. THE LENS WAS REPORTED AS HAVING LOW VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS EXPLANTED ON AN UNKNOWN DATE; NO FURTHER INFORMATION HAS BEEN PROVIDED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC SPLIT AND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION SINCE LENS RETURNED SPLIT. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4973899&pc=MTA | 2023826-2015-00940 | 2012-02-03 | Injury | STAAR SURGICAL COMPANY | 2015-08-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. LOW VAULT WAS OBSERVED AND THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2012. THIS RESOLVED THE PROBLEM. PATIENT'S LAST BCVA WAS 20/20. SEE MFR REPORT# (B)(4) FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5673278&pc=MTA | 2023826-2016-00633 | 2016-03-17 | Injury | STAAR SURGICAL COMPANY | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON THE LENS SURFACE (NOT POSSIBLE TO SPECIFY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688626&pc=MTA | 2023826-2016-00750 | 2016-04-12 | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND REMOVED WITHIN THE SAME SURGERY WITH NO APPARENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND FOREIGN MATERIAL - FIBERS PRESENT ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: THE ANTERIOR ENDOTHELIUM CHABMER DEPTH (ACD) IS BELOW 3.0MM PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7434261&pc=MTA | 2023826-2018-00558 | 2018-02-05 | Injury | STAAR SURGICAL COMPANY | 2018-04-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS NOTED TO HAVE LOW VAULTING, WITH LENS ROTATION. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.5 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018 DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7562898&pc=MTA | 2023826-2018-00802 | 2018-05-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/+2.5/174 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD A LOW VAULT, WITH ROTATION. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. A LONGER LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE REPORTER CLARIFIED STATING THAT LENS WAS EXCHANGED IN THE SAME SURGERY AS LENS REMOVAL ON (B)(6) 2018 FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH PRESENCE OF RESIDUE/ DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE LENS TO HAVE A CURVED SHAPE AND PRESENCE OF RESIDUE/ DEBRIS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824433&pc=MTA | 2023826-2018-01255 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/+3.5/X077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2018. THE LENS WAS REPOSITIONED ON (B)(6) 2018 BUT THE ISSUE WAS NOT RESOLVED. THE LENS ROTATED A SECOND TIME ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE LENS IS NOW STABLE. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). CORRECTED DATA: PRODUCT PROBLEM TO ADVERSE EVENT MALFUNCTION TO SERIOUS INJURY CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6213692&pc=MTA | 2023826-2016-01848 | 2016-06-30 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | INTRAOCULAR LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS HAD TORN DURING THE ICL LOADING PROCESS. THIS OCCURRED ON (B)(6) 2016, AS THE SURGEON WAS PREPARING TO IMPLANT THE LENS INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS IMPLANTED AT AN UNKNOWN DATE AND THE PATIENT EXPERIENCED EXCESSIVE VAULT. THE LENS WAS EXPLANTED AT AN UNKNOWN DATE AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074719&pc=MTA | 2023826-2016-01591 | 2016-09-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND THE ANTERIOR CHAMBER WAS SHALLOW. THE LENS REMAINS IMPLANTED BUT WILL BE REMOVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6727761&pc=MTA | 2023826-2017-01209 |  | Injury | STAAR SURGICAL COMPANY | 2017-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Insufficient Information | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENTS EYE ON (B)(6) 2017. THE LENS WILL BE EXPLANTED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXPLANTED TO PREVENT ANGLE CLOSURE (BORDERLINE OCCLUDABLE). ADDITIONAL PERIPHERAL IRIDECTOMY PERFORMED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO MISMATCH SULCUS-SULCUS AND WHITE-TO-WHITE MEASUREMENTS. ANOTHER LENS WILL BE IMPLANTED AT A LATER DATE. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6727763&pc=MTA | 2023826-2017-01210 |  | Injury | STAAR SURGICAL COMPANY | 2017-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Insufficient Information | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WILL BE EXPLANTED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXPLANTED TO PREVENT ANGLE CLOSURE (BORDERLINE OCCLUDABLE). THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO MISMATCH SULCUS-SULCUS AND WHITE-TO-WHITE MEASUREMENTS. ANOTHER LENS WILL BE IMPLANTED AT A LATER DATE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4958187&pc=MTA | 2023826-2015-00981 | 2015-02-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND LOW VAULT, ROTATION AND LENS DISLOCATION/SUBLUXATION WAS NOTED. THE LENS REMAINS IMPLANTED AND THE SURGEON PLANS ON EXCHANGING IT FOR A LONGER LENGTH. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: LENS WAS EXPLANTED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. CORRECTED DATA:(B)(4) PREVIOUS CODE NOT APPLICABLE, LENS EXPLANTED. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN PATIENTS WHO ARE AMBLYOPIC OR BLIND IN FELLOW EYE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6813293&pc=MTA | 2023826-2017-01365 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7327005&pc=MTA | 2023826-2018-00334 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS WAS TO BE EXPLANTED AND EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE LENS WAS EXPLANTED. (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131320&pc=MTA | 2023826-2019-01792 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: A COMPLETED LAF CARD FOR A 13.2MM, MICL13.2, -10.0 DIOPTER, INTRAOCULAR LENS WAS RETURNED BACK TO US. NOTE ON THE LAF CARD INDICATED " LENS DIDNT SET CORRECTLY IN EYE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191 CORRECTED TO PATIENT CODE 2692 IN INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965325&pc=MTA | 2023826-2019-01612 | 2019-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. TREATMENT WITH RINDERON EYE DROPS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM: ADVERSE EVENT. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVANT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). TYPE OF REPORTABLE EVENT: SERIOUS. PATIENT CODE 3191: NO CODE AVAILABLE (MEDICAL TREATMENT WITH RINDERON EYE DROPS). CAUSE OF THE EVENT WAS REPORTED AS DEVICE CAUSALITY. AT A FOLLOW UP UPDATE REPORTER STATED "THE WHITE ACCRETIONS OF THE PATIENTS GOT SMALLER NOW WITHOUT SPECIFIC MEDICAL TREATMENT. DOCTOR THINKS THE RISK OF RECURRENCE WILL BE LOW. REPORTED LENS OPACITY MEANS THAT ICL LENS OPACITY DUE TO WHITE ACCRETIONS, NOT PATIENT'S CATARACT". CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: RESULT CODE: 213 SHOULD BE CORRECTED TO 114. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409176&pc=MTA | 2023826-2019-00414 | 2018-06-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICMO12.1,-9.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. PATIENT CONTACT BUT NO PATIENT INJURY WAS REPORTED. A REPLACEMENT LENS WAS IMPLANTED AND THE REPORTER STATED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT CODE 3191: CODE IN PREVIOUS MDR NOT APPLICABLE. PATIENT CODE 2692: NO KNOWN IMPACT, IS APPLIED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7361884&pc=MTA | 2023826-2018-00463 | 2018-01-31 | Injury | STAAR SURGICAL COMPANY AG | 2018-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -6.5/2.5/89 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2011. THE REPORTER INDICATED THE LENS HAD AN LOW VAULT. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: UPDATED TO ADVERSE EVENT. UPDATED DESCRIPTION OF EVENT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS -6.5/2.5/89 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE REPORTER OBSERVED A LOW VAULT SO THE SURGEON EXPLANTED THE LENS. NO INFORMATION REGARDING THE DETAILS OF THE SECONDARY SURGERY HAS BEEN PROVIDED DESPITE THE REQUESTS FOR DETAILS. PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT WITH RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671662&pc=MTA | 2023826-2019-01016 | 2019-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DANAGE SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MDR#1. PER NEW INFORMATION RECEIVED ON 17-SEP-2019, THE LOW VAULT PREVIOUSLY OBSERVED IN THE REPLACEMENT LENS HAS BEEN RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE SURGEON STATES THE VAULT IS STILL LOW. REFERENCE CLAIM FILE# (B)(4) FOR REPLACEMENT LENS. PATIENT CODE UPDATED TO 3191, NO CODE AVAILABLE: (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. PREVIOUS DEVICE CODE 1494 NO LONGER APPLICABLE. ACD>2.8MM APPROVED FOR ARGENTINA IN FEB 2019. (B)(4). Manufacturer Narrative: PER NEW INFORMATION RECEIVIED ON 17-SEP-2019, THE LOW VAULT PREVIOUSLY OBSERVED IN THE INITIAL LENS HAS BEEN RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: RESULT CODE 114 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186678&pc=MTA | 2023826-2019-01885 | 2019-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETRURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE: ACD< 3.00MM SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: PATIENT CODE: 3191 IS NOT APPLICABLE; SHOULD BE CORRECTED TO CODE 2692 IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203030&pc=MTA | 2023826-2019-01951 | 2019-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. REPORTER INDICATED THAT "IT IS NOT YET CLEAR WHETHER THE LENS WILL BE EXCHANGED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM." D7- EXPLANT DATE: (B)(6) 2019. H1- "MALFUNCTION" SHOULD BE CHANGED TO "SERIOUS INJURY". H6- PATIENT CODE: 3191- SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE - AGE UNDER 21 YEARS OLD SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: "ADVERSE EVENT" SHOULD BE MARKED ON THE SUPPLEMENTAL MDR #1. "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE" SHOULD BE MARKED ON THE SUPPLEMENTAL MDR #1. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936291&pc=MTA | 2023826-2020-00747 | 2020-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. WITHIN THE SAME SURGERY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE 3191 CORRECTED TO PATIENT CODE 2692 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6997275&pc=MTA | 2023826-2017-01652 | 2017-06-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DURING THE IMPLANTATION, THE SURGERY SAW THAT THE LENS WAS UPSIDE DOWN. THE LENS WAS EXPLANTED ON THE SAME DAY WITH THE LENS BEING TORE DUE TO AN UNSKILLED REMOVAL TECHNIQUE. THE LENS WAS THEN EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 -12.5 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DURING THE IMPLANTATION THE SURGEON NOTICED THAT THE LENS WAS UPSIDE DOWN AND REMOVED IT DURING THE SAME SURGERY. DURING LENS REMOVAL THE LENS TORE DUE TO AN UNSKILLED REMOVAL TECHNIQUE. A NEW AND SIMILAR LENS WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS FOUND TO BE 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE PRESENT ON THE LENS. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RED RESIDUE ON THE LENS SURFACE. CORRECTED DATA: DATE OF BIRTH: IT IS CORRECTED FROM (B)(6) TO (B)(6) IN THE INITIAL MDR. THIS EVENT IS CORRECTED FROM ADVERSE EVENT TO PRODUCT PROBLEM. DEVICE MANUFACTURE DATE: IT IS CORRECTED FROM 29-OCT-2015 TO 15-APR-2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6987333&pc=MTA | 2023826-2017-01579 | 2017-09-19 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. THE SURGEON NOTICED THAT THE LENS UNFOLDED UPSIDE DOWN. THE LENS WAS EXCHANGED FOR A BACK-UP LENS. Manufacturer Narrative: METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE SURGEON NOTICED THAT THE LENS UNFOLDED UPSIDE DOWN. THE LENS WAS REMOVED INTRAOPERATIVELY AND THE BACK-UP LENS WAS IMPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(6). CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH CLEAR RESIDUE ON LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077649&pc=MTA | 2023826-2016-01565 | 2016-09-27 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure; Device Operates Differently Than Expected; Torn Material; Device Handling Problem | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO BEING IMPLANTED UPSIDE DOWN AND THEN BROKEN DURING EXPLANT. THE LENS WAS EXCHANGED FOR A LENS THE SAME SIZE, MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN PATIENTS UNDER THE AGE OF 21 YEARS OLD. CORRECTED DATA: IT IS AN ADVERSE EVENT, NOT A PRODUCT PROBLEM. PREVIOUS CODE NOT APPLICABLE, NO TORN MATERIAL ASSOCIATED WITH A MATERIAL INTEGRITY ISSUE. (B)(4). Manufacturer Narrative: PMA/510(K): N/A THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923613&pc=MTA | 2023826-2016-01207 | 2016-04-22 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, AND THE LENS TORE DURING INJECTION INTO THE EYE. THE LENS WAS REMOVED WITH NO APPARENT INJURY. ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: PATIENTS AGE AT TIME OF EVENT (B)(6) YEARS OLD IS INCORRECT, CORRECT PATIENT'S AGE IS (B)(6) YEARS OLD. ADDITIONAL DATA: (B)(4). ONLY SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED WAS REPORTED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND LENS OPTIC TORN/SPLIT, LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS DRIED, DISCOLORED MATERIAL. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6030339&pc=MTA | 2023826-2016-01410 | 2016-08-09 | Injury | STAAR SURGICAL COMPANY | 2016-10-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS - -13.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS TEAR. ANOTHER SAME MODEL AND DIOPTER WAS IMPANTED. THE NEW LENS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/35. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6021626&pc=MTA | 2023826-2016-01440 | 2016-09-06 | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/040 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS TORE/BROKE DURING THE IMPLANTATION AND WAS EXPLANTED ON THE SAME DAY. ANOTHER LENS OF THE SAME SIZE WAS IMPLANTED ON (B)(6) 2016. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. ANOTHER LENS WAS IMPLANTED ON THE SAME DAY. (B)(4). Manufacturer Narrative: (B)(4). PRODUCT EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843646&pc=MTA | 2023826-2019-01332 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS IMPLANTED UPSIDE DOWN. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXCHANGED IN (B)(6) 2019 BUT DID NOT PROVIDE THE DAY. THIS EXCHANGE RESOLVED THE PROBLEM AND THE LENS WAS DISCARDED AND WILL NOT BE RETURNED FOR PRODUCT EVALUATION. EXPLANT DATE: UNKNOWN. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). (B)(6) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059249&pc=MTA | 2023826-2016-01538 | 2015-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material | No Known Impact Or Consequence To Patient; No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, AND THE LENS WAS DAMAGED WHILE INJECTING INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS NOT IMPLANTED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS.Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, AND THE LENS WAS DAMAGED BEFORE INJECTING INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS NOT IMPLANTED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: LENS TEAR HAPPENED BEFORE, NOT WHILE, INJECTION INTO THE EYE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, HAPTIC TORN/BROKEN, AND SPOTS ON LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5937931&pc=MTA | 2023826-2016-01239 | 2016-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Information | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50+1.0/092 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. IT WAS INDICATED THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: OFF-LABEL USE: PATIENT ACD BELOW 3.0 MM. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PRODUCT PROBLEM (MALFUNCTIONS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816262&pc=MTA | 2023826-2016-00971 | 2016-06-10 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Information | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE LENS IS PLANNED TO EXCHANGE WITH ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: THIS IS A PRODUCT PROBLEM, NOT AN ADVERSE EVENT. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/40. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831922&pc=MTA | 2023826-2016-00995 | 2016-06-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS HAD AN EXCESSIVE VAULT AND WILL BE EXPLANTED. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE SURGEON, IN ERROR, IMPLANTED THE WRONG SIZE LENS INSTEAD OF A 12.6MM ICL LENS. Manufacturer Narrative: ADDITIONAL DATA: SURGEON IMPLANTED A LENS OF ANOTHER PATIENT WITH SAME POWER BUT A 13.2MM LENGTH INSTEAD OF THE REQUIRED 12.6MM LENGTH. PATIENT IS HAPPY AND NO EXPLANT OR EXCHANGE IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: DATE OF EXPLANT: UNK. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5957298&pc=MTA | 2023826-2016-01265 | 2016-08-18 | Injury | STAAR SURGICAL COMPANY | 2016-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Information | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND WILL BE EXPLANTED. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE SURGEON DECIDED TO NOT EXCHANGE THE LENS NOW AND CLOSELY MONITOR THE PATIENT. THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. CORRECTED DATA: THE REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4889334&pc=MTA | 2023826-2015-00802 |  | Injury | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE SURGEON IS PLANNING TO EXPLANT THE LENS DUE TO INADEQUATE VAULTING AND EXCHANGE FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER STATED THE PROBLEM OF INADEQUATE VAULTING RESOLVED ITSELF OVER TIME AND THE LENS REMAINS IMPLANTED. THE PATIENT IS DOING WELL. Manufacturer Narrative: METHOD CODE (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS CODE (EVALUATION RESULT): PER MEDICAL REVIEW - ACCURATE DETERMINATION OF ANTERIOR SEGMENT DIMENSIONS IS KEY TO AVOIDANCE OF CENTRAL AND/OR PERIPHERAL CONTACT BETWEEN THE ICL AND THE CRYSTALLINE LENS. THE METHOD OF ICL SIZING FOR THIS PATIENT IS NOT KNOWN. THE REPORT ALSO PROVIDES NO SPECIFIC INFORMATION ON THE DEGREE OF VAULTING, OR THE CAUSE OF INADEQUATE VAULT. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5506284&pc=MTA | 2023826-2016-00273 | 2015-11-24 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Known Impact Or Consequence To Patient; No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/104 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A LOW VAULT. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT THE LENS BUT THE LENS IS CURRENTLY IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: METHOD: WORK ORDER SEARCH PERFORMED, NO SIMILAR COMPLAINTS FOUND. CONCLUSION: PATIENT ACD WAS BELOW 3.0 MM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. PROCESS EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5392219&pc=MTA | 2023826-2016-00105 | 2015-12-02 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER STATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, WAS REMOVED FROM THE LENS VIAL AND THE SURGEON WAS PREPARING TO LOAD THE LENS. A BLACK SPECK WAS NOTED ON THE LENS AND THE SURGEON DECIDED NOT TO USE IT. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE PARTICULATES ON THE LENS OR RELEASE OF NONCONFORMING LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED THE COMPLAINT ISSUE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF WHITE RESIDUE ON THE LENS SURFACE. CONCLUSION - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437901&pc=MTA | 2023826-2019-00448 | 2019-01-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Contamination /Decontamination Problem | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS PREPARING TO IMPLANT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), IT WAS DISCOVERED TO HAVE A FOREIGN BODY ADHESION. THIS WAS DISCOVERED WHEN IT WAS LOADED INTO THE CARTRIDGE. THERE WAS NO PATIENT CONTACT WITH THE LENS. THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437898&pc=MTA | 2023826-2019-00451 | 2019-02-08 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-20 | MTA | IMPLANTABLE COLLAMER LENS | Contamination /Decontamination Problem | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS PREPARING TO IMPLANT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), IT WAS DISCOVERED TO HAVE A FOREIGN BODY ADHESION. THIS WAS DISCOVERED AFTER THE SURGEON REMOVED THE LENS FROM THE VIAL. THERE WAS NO PATIENT CONTACT WITH THE LENS. THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. A PIECE OF FOREIGN BODY WAS ALSO STUCK ON THE LENS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492316&pc=MTA | 2023826-2019-02409 | 2019-11-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Contamination /Decontamination Problem | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON FOUND A BLACK DOT-LIKE FOREIGN MATTER ON A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, WHEN THE LENS WAS REMOVED FROM THE VIAL. THE LENS WAS NOT USED OR IMPLANTED. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5227492&pc=MTA | 2023826-2015-01416 | 2015-06-24 | Malfunction | STAAR SURGICAL COMPANY | 2015-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Defective Device; Scratched Material | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO USE A 12.6MM VICMO12.6, -11.00 DIOPTER IMPLANTABLE COLLAMER LENS. BEFORE OPENING THE LENS VIAL, A SCRATCH WAS FOUND ON THE LENS. LENS WAS RECEIVED WITH SCRATCH. THEN WHEN TAKING OUT THE LENS FROM THE VIAL WITH FORCEPS, IT WAS TORN IN TWO PIECES AND ONE PIECE WAS LOST. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6327406&pc=MTA | 2023826-2017-00121 | 2016-11-18 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contaminated During Manufacture or Shipping | No Patient Involvement | N/A | 1 | Event Description: ON (B)(6) 2016, THE REPORTER INDICATED THAT A FOREIGN BODY ADHESION (BLACK DOT) WAS OBSERVED ON A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, PRIOR TO LOADING AND HENCE HAD NO PATIENT CONTACT. ANOTHER LENS WAS USED FOR IMPLANTATION ON THE SAME DAY WHICH RESOLVED THE PROBLEM. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UCVA 20/20. Manufacturer Narrative: PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6586265&pc=MTA | 2023826-2017-00870 | 2017-04-21 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, FROM ITS VIAL AND WHEN RINSING THE LENS HE NOTED RED DEBRIS IN THE MIDDLE OF THE LENS OPTIC. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: METHOD: (MATERIALS AND CHEMISTRY EVALUATION): THE REDDISH DEBRIS FOUND ON THE LENS WAS SENT FOR ANALYSIS TO BE INVESTIGATED FURTHER USING A FOURIER TRANSFORM INFRARED SPECTROSCOPY AND SCANNING ELECTRON MICROSCOPY WITH ENERGY DISPERSIVE X-RAY SPECTROSCOPY(SEM-EDX); THE PARTICULATE CONSISTED OF OXIDIZED STEEL (COMPOSED OF IRON, CHROMIUM, AND NICKEL), SILICATES (SODIUM, ALUMINUM, POTASSIUM, AND CALCIUM), AND OTHER ORGANIC CARBONATES. REVIEW OF THE LENS MANUFACTURING PROCESS AND EQUIPMENT DID NOT FIND PROBABLE CAUSES FOR THE PARTICULATE REVIEW OF THE LENS MANUFACTURING PROCESS AND EQUIPMENT DID NOT FIND ANY PROBABLE CAUSES FOR THE PARTICULATE. PART OF THE PROCESSING, PRODUCTION LENSES ARE VISUALLY INSPECTED UNDER MAGNIFICATION TWICE TO HELP ENSURE PARTICULATES, SUCH AS THIS REDDISH DEBRIS, ARE DETECTED. METHOD: (DEVICE FROM SAME LOT EVALUATED): A LENS WORK ORDER SEARCH WAS PERFORMED. NO OTHER SIMILAR COMPLAINTS FOUND. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. LENS HAS REDDISH BROWN COLORED DEBRIS/RESIDUE ON OPTIC. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6415613&pc=MTA | 2023826-2017-00409 | 2016-11-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, WAS RETURNED DUE TO A LOADING ERROR/PROBLEM, THEY COULDN'T GET IT TO LOAD. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES. Manufacturer Narrative: LENS WAS RETURNED DRY, IN LENS CASE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DRY RESIDUE AND DEBRIS ON PRODUCT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6552491&pc=MTA | 2023826-2017-00721 | 2017-04-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL12.6, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THE REPORTER CONFIRMED THAT IT WAS NOT DUE TO USER OR TECHNICIAN ERROR AND THAT "THE LOADER TORE THE LENS". THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND PIECE OF HAPTIC MISSING. NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7253167&pc=MTA | 2023826-2018-00168 | 2017-04-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON LOADED A MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS, AND A TEAR WAS NOTED. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. THE REPORTER STATED THAT THE CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED STUCK IN THE CARTRIDGE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO DEVICE AND WHITE RESIDUE ON LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6592833&pc=MTA | 2023826-2017-00895 | 2017-05-08 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE REPORTER DID NOT KNOW IF THE CAUSE WAS DUE TO USER ERROR. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: NO LENS RETURNED FOR EVALUATION. A CARTRIDGE CASE WAS RETURNED WRAPPED WITH GAUZE. A LEN VIAL STOPPER WAS IN THE CASE. A NOTE INDICATING "TORN UPON LOADING" WAS WRITTEN ON THE BOX. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR CLAIM TYPE WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6601312&pc=MTA | 2023826-2017-00886 | 2017-05-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER STATED THAT A MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. THE REPORTER WAS NOT SURE IF THE LENS TEAR WAS CAUSED DUE TO USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND MISSING PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6727139&pc=MTA | 2023826-2017-01184 | 2017-06-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: TYPE OF REPORT: SUPPLEMENTAL #1 WAS NOT SUCCESSFULLY SUBMITTED BEFORE SUPPLEMENTAL #2. USAGE OF DEVICE: INITIAL REPORT SHOULD BE "INITIAL USE OF DEVICE" INSTEAD OF "UNKNOWN". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CONCOMITANT MEDICAL PRODUCTS:: INJECTOR, CARTRIDGE AND INJECTOR SYSTEM MODELS ARE ADDED. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED LOOSE IN LENS TRAY. VISUAL INSPECTION FOUND PORTION OF THE OPTIC AND HAPTIC MISSING. CLEAR RESIDUE ON THE LENS SURFACE WAS OBSERVED NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6727148&pc=MTA | 2023826-2017-01215 | 2017-06-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 TORE DURING THE LOADING PROCESS. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF REPORT; PREVIOUS REPORT SHOULD BE CHECKED "30 DAY" AND "INITIAL". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED STUCK IN CARTRIDGE. THERE IS NO VISIBLE DAMAGE TO THE DELIVERY DEVICE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993821&pc=MTA | 2023826-2017-01658 | 2017-08-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017, AS THE SURGEON ATTEMPTED TO LOAD AND IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -17.0/+0.50/X134 INTO THE PATIENT'S LEFT (OS) EYE, HE DISCOVERED THE LENS BROKE. REPORTEDLY, THIS WAS DISCOVERED PRIOR TO IMPLANT AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7006993&pc=MTA | 2023826-2017-01691 | 2017-10-12 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: REPORTEDLY, WHILE PREPARING TO INSERT A 13.7MM MICL13.7, DIOPTER -8.5 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE, THE LENS TORE. THIS OCCURRED DURING THE LOADING PROCESS. REPORTER STATES THAT THE SURGEON NOTICED "A LOT OF BUBBLES" AND WHEN TRYING TO PULL THE LENS THROUGH "SHE HAD TO GRAB [THE LENS] A COUPLE OF TIMES." THERE WAS NO PATIENT CONTACT WITH THIS LENS. THIS OCCURRED ON (B)(6) 2017. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND A TEAR IN THE OPTIC. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7232868&pc=MTA | 2023826-2018-00161 | 2017-10-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.1MM VICM5_12.1, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. BEFORE IMPLANTATION THE LENS TORE/BROKE. THE LENS WAS NOT IMPLANTED AND NO PATIENT CONTACT IS CONFIRMED. ON THE SAME DATE AN ALTERNATE LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.99MM)] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7155526&pc=MTA | 2023826-2017-01984 | 2017-11-17 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED WITH NO ISSUES. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245214&pc=MTA | 2023826-2018-00200 | 2017-11-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.7MM VICM5_13.7, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. BEFORE IMPLANTATION THE LENS TORE/BROKE. THE LENS WAS NOT IMPLANTED AND NO PATIENT CONTACT IS CONFIRMED. ON THE SAME DATE AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND WHITE RESIDUE AND DEBRIS ON LENS), AND A PIECE OF HAPTIC MISSING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245222&pc=MTA | 2023826-2018-00199 | 2017-11-17 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.7MM VICM5_13.7, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. BEFORE IMPLANTATION THE LENS TORE/BROKE. THE LENS WAS NOT IMPLANTED AND NO PATIENT CONTACT IS CONFIRMED. ON THE SAME DATE AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND WHITE RESIDUE AND DEBRIS ON LENS), LENS IS CURVED, AND PIECE OF OPTIC AND HAPTIC MISSING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261732&pc=MTA | 2023826-2018-00169 | 2017-11-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/90 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017. THE SURGEON EXPERIENCED LENS DAMAGE BEFORE IMPLANTATION OF THE LENS AND A BACK-UP LENS WAS USED INSTEAD ON (B)(6) 2017 TO COMPLETE THE SURGERY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND PORTION OF THE LENS MISSING, A TEAR ON THE HAPTIC, AND RESIDUE ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180342&pc=MTA | 2023826-2018-00058 | 2017-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED WITH NO PROBLEM. THE REPORTER STATED THAT THE CAUSE OF EVENT IS UNKNOWN. REPORTER WAS UNABLE TO PROVIDE ADDITIONAL INFORMATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7242309&pc=MTA | 2023826-2018-00165 | 2017-12-19 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLMER LENS, TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. THE REPORTED STATED THAT THE CAUSE OF EVENT WAS UNKNOW. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT A MICL 12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS, TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC AND A MISSING HAPTIC PIECE. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7262034&pc=MTA | 2023826-2018-00170 | 2017-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE FOOT PLATE OF 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/2.5/105( SPHERE/ CYLINDER/ AXIS), TORE SHARPLY WHEN PUSHING THE LENS BY USING FTP (FOAM TIP PLUNGER . THE SURGEON EXPERIENCED LENS DAMAGE BEFORE INJECTION INTO THE EYE. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND PORTION OF THE LENS MISSING AND RESIDUE ON THE LENS SURFACE. SFC-45 CARTRIDGE WAS RETURNED IN A BIOHAZARD BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE AND CLEAR RESIDUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: IT SHOULD BE NOTED THAT BEFORE LENS TORE, "LENS MOVEMENT WAS NOT SMOOTH WHEN PULLING LENS TOWARDS NOZZLE". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7294653&pc=MTA | 2023826-2018-00265 | 2017-12-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-24 | MTA | MICL | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS DEFECTIVE. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: THE REPORTER STATED, "THE HAPTICS ARE ORIENTED IN THE OPPOSITE DIRECTION AS THE DESIGN OF THE LENS CALLS FOR. IT WAS NOTICED WHEN THE SURGEON (DR. (B)(6)) WAS LOADING THE LENS". A BACK UP LENS WAS LATER IMPLANTED. CORRECTED DATA: CLAIM# (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281443&pc=MTA | 2023826-2018-00251 | 2017-12-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: REPORTEDLY, WHILE PREPARING TO INSERT A 13.2MM MICL13.2, DIOPTER -16.0 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE, THE LENS TORE. THIS WAS DISCOVERED AFTER THE LOADING PROCES. REPORTER STATES THEY ARE UNSURE AS TO WHETHER THE LENS WAS RECEIVED DAMAGED OR IF IT TORE DURING LOADING. THERE WAS NO PATIENT CONTACT WITH THIS LENS. THIS OCCURRED ON (B)(6) 2017. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7294025&pc=MTA | 2023826-2018-00266 | 2018-01-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-23 | MTA | MICL | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT DURING LOADING OF A MICL13.2, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS, IT WAS NOTED THAT A PORTION OF ONE OF THE FOOTPLATES WAS MISSING. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: CORRECTED DATA: EXPIRATION DATE: IT IS CORRECTED FROM "31-JAN-2018" TO "31-JAN-2019". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID INSIDE A LENS VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356139&pc=MTA | 2023826-2018-00432 | 2018-01-30 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VICMO12.6, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS NOT IMPLANTED AND IT IS CONFIRMED THAT THERE WAS NO PATIENT CONTACT WITH THIS LENS. ON (B)(6) 2018 AN ALTERNATE LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7308338&pc=MTA | 2023826-2018-00307 | 2018-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON PARTIALLY LOADED A MICL13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IN THE CARTRIDGE AND DECIDED TO GO WITH THE BACKUP LENS. IT WAS REPORTED THAT THE DOCTOR INSPECTED THE LENS UNDER A MICROSCOPE AND DID NOT LIKE HOW IT LOOKED, POSSIBLE FLAW IN THE LENS. NO PATIENT CONTACT. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339633&pc=MTA | 2023826-2018-00396 | 2018-02-08 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTED INDICATED THAT A 13.2MM MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE BACK UP LENS WAS USED. THE REPORTER STATED THAT THE CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7395923&pc=MTA | 2023826-2018-00504 | 2018-02-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON NOTICED A TEAR ON A 12.6 VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, DURING LOADING . ACCORDING TO THE SURGEON THE LENS WAS TORN WHILE PUSHING THE LENS WITH THE FOAM TIP PLUNGER DUE TO ITS RESISTANCE. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2018 THE SURGEON IMPLANTED A SAME MODEL AND DIOPTER LENS INTO THE PATIENT'S RIGHT EYE (OD) WITH NO ISSUES. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS MISSING TWO PIECES OF HAPTIC WITH RESIDUE AND DEBRIS ON LENS SURFACE. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7366323&pc=MTA | 2023826-2018-00460 | 2018-02-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON LOADED A 13.2MM MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS AND A TEAR WAS NOTICED. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. EVENT CAUSE WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7539104&pc=MTA | 2023826-2018-00765 | 2018-02-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO LOAD A 12.6 MM VTICM5_12.6, 14.0/+1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018 AFTER ATTEMPTING TO LOADING ANOTHER LENS. THE SURGEON HAS ENCOUNTERED SIMILAR DIFFICULTIES LOADING THE LENS AS THE FIRST LENS SHE TRIED TO LOAD. THERE WAS NO PATIENT CONTACT. THE SURGEON WENT BACK TO THE FIRST LENS SHE ATTEMPTED TO LOAD AND HE WAS SUCCESSFUL IN LOADING IT. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE TORN PIECE IS MISSING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7539110&pc=MTA | 2023826-2018-00766 | 2018-02-23 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO LOAD A 12.6 MM VTICM5_12.6, 14.0/+1.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. SHE HAD DIFFICULTY LOADING THE LENS IN THE CARTRIDGE. THE SURGEON THEN ATTEMPTED TO LOAD THE BACK-UP LENS AFTERWARDS. THE SURGEON ENCOUNTERED SIMILAR DIFFICULTIES LOADING THE BACK-UP LENS AS THE FIRST LENS. THE SURGEON WENT BACK TO THE FIRST LENS SHE ATTEMPTED TO LOAD AND SHE WAS SUCCESSFUL IN LOADING IT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7364249&pc=MTA | 2023826-2018-00440 | 2018-02-26 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, BUT THE PATIENT "SQUEEZED". THE SURGEON PLACED THE LENS ON THE TRAY TO RELOAD AND NOTED THERE WAS A MARK ON THE LENS. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) BUT THE PATIENT "SQUEEZED". DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7382965&pc=MTA | 2023826-2018-00488 | 2018-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON LOADED A 12.6MM MICL12.6, IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, AND THE LENS RIPPED. THERE WAS NO PATIENT CONTACT. THE BACK UP LENS WAS USED WITH NO ISSUES. THE REPORTER STATED THAT THE CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA CONCOMITANT MEDICAL PRODUCTS: CARTRIDGE MODEL NO.: SFC-25 AND CARTRIDGE LOT NO.: 1380994. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED STUCK IN SFC-25 CARTRIDGE. CARTRIDGE WAS RETURNED IN DEVICE TRAY WITH CLEAR RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE, RESIDUE ON LENS AND THE LENS STUCK IN THE CARTRIDGE. A LENS WORK ORDER SEARCH WAS PERFORMED. NO OTHER SIMILAR COMPLAINT TYPE EVENT FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7430699&pc=MTA | 2023826-2018-00551 | 2018-03-14 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM MICL12.6, -15.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS TORN DURING LOADING. THE REPORTER STATED THAT ALTERNATE FORCEPS WERE USED TO LOAD THE LENS SINCE THEY ONLY HAD ONE FORCEP FOR VISIAN ICLS AND THE FORCEP WAS BROKEN WHEN LOADING THE PRIMARY LENS. THE SURGEON DECIDED TO USE A SIMILAR FORCEP AND THE LENS WAS TORN. THERE WAS NO PATIENT CONTACT. THE REPORTER STATED THAT THE CAUSE OF EVENT WAS A MECHANICAL ERROR, COULD NOT LOAD THE LENS CORRECTLY. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7449090&pc=MTA | 2023826-2018-00573 | 2018-03-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2, 10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WOULD NOT FOLD IN THE CARTRIDGE CORRECTLY WHEN LOADING. NO PATIENT CONTACT. BACK UP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE REPORTER STATED THAT WHEN THE LENS WAS REMOVED FROM THE VIAL TO LOAD IT THE LENS BECAME STIFF WITHIN SECONDS AND THEY COULDN'T GET IT IN THE CARTRIDGE. THE LENS DIDN'T STAY MOIST. THE BACKUP LENS WAS USED WITH NO ISSUES. CAUSE OF EVENT WAS UNKNOWN. (B)(4) Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7508767&pc=MTA | 2023826-2018-00684 | 2018-03-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO13.2, -14.5/+4.0/85 (SPHERE/CYLINDER/AXIS). IMPLANTABLE COLLAMER LENS TORE WHILE ADJUSTING THE LENS POSITION WITHOUT EXCESSIVE OPERATION. THE LENS WAS NOT IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS SURFACE. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7469543&pc=MTA | 2023826-2018-00582 | 2018-03-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON LOADED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, AND THE LENS TORE. THE LENS WAS NOT IMPLANTED AND THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED WITH ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS COILED INTO A TUBE SHAPE. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488387&pc=MTA | 2023826-2018-00661 | 2018-03-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTED INDICATED THAT THE HAPTICS OF A 13.7MM MICL13.7, IMPLANTABLE COLLAMER LENS WERE BENT. THE SURGEON DID NOT WANT TO USE THE LENS SO THE BACKUP LENS WAS USED. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: DATE OF THIS REPORT: IT IS CORRECTED TO "(B)(6) 2018". CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7487378&pc=MTA | 2023826-2018-00659 | 2018-04-05 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6, -8.0, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. IT WAS REPORTED THAT THERE WAS NO PATIENT CONTACT, THE BACKUP LENS WAS USED AND THE SURGERY WAS A SUCCESS. THE REPORTER STATED THAT THEY HAVE HAD NO ISSUES WITH A LENS TEAR PREVIOUSLY SO THIS WAS CERTAINLY AN ISOLATED INCIDENT. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7520478&pc=MTA | 2023826-2018-00754 | 2018-04-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2, -12.5 DIOPTER TORE DURING THE LOADING PROCESS. THERE WAS NO PATIENT CONTACT AND THE BACK UP LENS WAS USED WITH NO PROBLEM. THE REPORTER STATED THAT THE CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF A HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7520052&pc=MTA | 2023826-2018-00694 | 2018-04-19 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING THE LOADING PROCESS. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. REPORTER STATED THAT THE CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF A HAPTIC. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7582671&pc=MTA | 2023826-2018-00867 | 2018-04-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A 13.7MM VICM5_13.7, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE CAUSE OF THE EVENT WAS REPORTED TO BE UNKNOWN. ON (B)(6) 2018 A SAME MODEL AND DIOPTER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7580442&pc=MTA | 2023826-2018-00868 | 2018-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON OBSERVED A CREASE IN THE CENTRAL OPTIC OF A 13.2MM MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS UPON LOADING INTO THE INJECTOR. THERE WAS NO PATIENT CONTACT WITH THE LENS. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7648911&pc=MTA | 2023826-2018-00989 | 2018-05-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON LOADED A 13.2MM MICL13.2, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WRONG INTO THE CARTRIDGE. THE REPORTER STATED THAT IT WAS A LOADING ISSUE AND THAT CAUSE OF EVENT IS UNKNOWN. THERE WAS NO PATIENT CONTACT WITH THE DEVICE AND THE BACKUP LENS WAS USED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6727022&pc=MTA | 2023826-2017-01212 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER WAS RETURNED TORN IN TWO PIECES. THE REPORTER STATED THAT THE CAUSE OF THE DAMAGE TO THE LENS IS UNKNOWN , COULD HAVE BEEN DAMAGE DURING LOADING. THE REPORTER STATED THAT THERE WAS NO PATIENT CONTACT, THE LENS WAS NEVER INSERTED IN THE EYE. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS TORN IN TWO PIECES. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6740904&pc=MTA | 2023826-2017-01238 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM MICL12.6, -8.0 DIOPTER, WAS TORN DURING THE LOADING PROCESS. THERE WAS NO PATIENT CONTACT AND THE BAKUP LENS WAS USED. THE REPORTER STATED THAT THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING FROM THE LENS. CORRECTED DATA: TYPE OF REPORT: INITIAL REPORT SHOULD BE "30-DAY" AND "INITIAL". (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7361609&pc=MTA | 2023826-2018-00457 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER WOULD NOT LOAD PROPERLY IN THE CARTRIDGE. NO HARM TO THE PATIENT, NO PATIENT CONTACT. THE BACK UP LENS WAS USED WITH NO ISSUES. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER TORE DURING LOADING AS IT WOULD NOT LOAD PROPERLY IN THE CARTRIDGE. THERE WAS NO HARM TO THE PATIENT AND NO PATIENT CONTACT. THE BACK-UP LENS WAS USED WITH NO ISSUES. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7663994&pc=MTA | 2023826-2018-01005 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-07-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM MICL13.7, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND PORTION OF THE LENS MISSING. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8383527&pc=MTA | 2023826-2019-00346 | 2019-01-30 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Ejection Problem | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT WHEN PUSHING THE ROD OF A KS-1 AQ2017V, 23.50 DIOPTER, PRELOADED INJECTOR THE ROD FELT HEAVIER THAN USUAL SO STOPPED TO USE AND THE SURGERY WAS SUCCESSFULLY COMPLETED WITHIN THE SAME SURGERY BY USING THE BACKUP. THERE WAS NO PATIENT CONTACT WITH THE DEVICE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. IOL WAS TWISTED AND REMAINED INSIDE THE NOZZLE. VOLUME OF USED VISCOELASTIC MATERIAL APPEARED TO BE ENOUGH. TOP FLANGE WAS PUSHED DOWN CORRECTLY. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5365485&pc=MTA | 2023826-2016-00037 | 2015-12-07 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Improper or Incorrect Procedure or Method; Torn Material | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS , -12.5 DIOPTER, TORE DURING THE LOADING PROCESS AND THERE WAS NO PATIENT CONTACT. THE REPORTER INDICATED THE EVENT WAS NOT DUE TO USER ERROR. THIS WAS ONE OF TWO LENSES USED FOR THIS PATIENT - SEE MFR. REPORT # 2023826-2016-00036. ANOTHER SAME MODEL LENS, BUT DIFFERENT DIOPTER POWER WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE DAMAGE TO THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, INJECTOR LOT SEARCH, CARTRIDGE LOT SEARCH, FOAM TIP PLUNGER LOT SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH, INJECTOR LOT SEARCH, CARTRIDGE LOT SEARCH, FOAM TIP PLUNGER LOT SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. AN INJECTOR LOT SEARCH FOUND ANOTHER SIMILAR COMPLAINT. A CARTRIDGE LOT SEARCH FOUND ANOTHER SIMILAR COMPLAINT. A FOAM TIP PLUNGER LOT SEARCH FOUND ANOTHER SIMILAR COMPLAINT. BASED ON THE COMPLAINT HISTORY, LENS WORK ORDER SEARCH, INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER LOT SEARCHES, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120425&pc=MTA | 2023826-2019-01738 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inaccurate Information | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A PACKAGED 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5, WAS POSSIBLY LABELED WITH A DIFFERENT SERIAL NUMBER THAN THE PATIENT LABELS INCLUDED INSIDE THE BOX (REPORTER DID NOT OPEN THE BOX). REPORTEDLY, THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN AN UNOPENED BOX. THE BOX AND LABEL SERIAL NUMBERS ARE THE SAME. (B)(4). Manufacturer Narrative: PATIENT INFORMATION: NO PATIENT INVOLVEMENT. DATE OF EVENT: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: TWO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8358682&pc=MTA | 2023826-2019-00292 | 2019-01-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Discolored | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AFTER OPENING THE VIAL OF A 12.1MM, VTICM5_12.1, -12.5/2.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, THE DOCTOR FOUND A FOREIGN BODY,IT LOOKED LIKE A BLACK LINE. THE LENS HAD NO PATIENT CONTACT, NOT IMPLANTED. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TEAR ON THE OPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. THE LENS WAS SOAKED IN BSS SOLUTION IN ORDER TO NOTE ANY FOREIGN BODIES, NONE WHERE NOTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8350795&pc=MTA | 2023826-2019-00278 | 2019-01-25 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AFTER OPENING THE LENS VIAL IN ORDER TO IMPLANT A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE, IT WAS NOTICED THAT THE LENS WAS "CLOUDY." THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THIS LENS. THE REPORTER STATES, "ICL HAD OPACIFICATION." Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND CLOUDY RESIDUE ON THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254836&pc=MTA | 2023826-2019-00060 | 2017-10-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE PREPARING TO LOAD A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+4.0/084 (SPHERE/CYLINDAR/AXIS) THEN INJECT INTO THE PATIENT'S LEFT EYE (OS), HE NOTICED THE LENS WAS TORE/BROKE. THIS OCCURRED ON (B)(6) 2017. REPORTEDLY, THIS WAS DISCOVERED PRIOR TO OPENING THE VIAL, HOWEVER IT IS ALSO NOTED THAT THIS COULD HAVE BEEN USER ERROR. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND A PORTION OF THE LENS MISSING. METHOD CODE ADDED, RESULT CODE UPDATED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7898026&pc=MTA | 2023826-2018-01419 | 2018-02-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -9.50/+2/028 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS RECEIVED DAMAGE AND BROKEN ON ONE SIDE. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY WITH RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND TEARS IN HAPTICS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7682694&pc=MTA | 2023826-2018-01021 | 2018-06-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING THE LOADING PROCESS. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED WITH NO ISSUES. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC (TORN PIECE IS MISSING). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824055&pc=MTA | 2023826-2018-01280 | 2018-07-03 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON PREPARED TO IMPLANT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), HE NOTICED IT HAD TORN. REPORTEDLY, THE LENS GOT JAMMED IN THE INJECTOR PRIOR TO INJECTION. THERE WAS NO PATIENT CONTACT WITH THE LENS. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND A TORN HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7974480&pc=MTA | 2023826-2018-01491 | 2018-07-05 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING ON (B)(6) 2018. THERE WAS NO PATIENT CONTACT. BACKUP LENS WAS IMPLANTED DURING SAME SURGERY. THE REPORTER STATED THAT THEY THOUGHT THE LENS FOLDED CORRECTLY, HOWEVER, IT WAS NOT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8073847&pc=MTA | 2023826-2018-01677 | 2018-08-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_13.7, -8.00/0.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS DID NOT SLIDE SMOOTHLY INTO THE INJECTOR. LENS WAS INJECTED OUR OF INJECTOR ONTO TABLE WHERE IT APPERED TORN. NO PAITEN CONTACT AND NO PATIENT INJURY WAS REPORTED. BACK UP LENS WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: RETURNED TO MANUFACTURER CORRECTED TO: "25-OCT-2018" IN THE INITIAL MDR. RESULT CODE CORRECTED TO: "114" IN THE INITIAL MDR, PREVIOUS CODE NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8072137&pc=MTA | 2023826-2018-01640 | 2018-08-13 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), BUT THE LENS TORE DURING INJECTION INTO THE EYE. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/DIOPTER LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7947525&pc=MTA | 2023826-2018-01507 | 2018-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE ANTERIOR LEAFLET OF A MICL13.2, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING ON (B)(6) 2018, WITH NO PATIENT CONTACT. THE BACKUP LENS IMPLANTED DURING INITIAL SURGERY WITH NO PATIENT INJURY. PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: "VISUAL INSPECTION FOUND THE OPTIC TORN" SHOULD BE CORRECTED TO "VISUAL INSPECTION FOUND THE HAPTIC TORN" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906907&pc=MTA | 2023826-2018-01387 | 2018-08-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT VTICMO13.2, -18.00/2.0/083 (SPHERE/CYLINDER/AXIS), INTRAOCULAR LENS TORE DURING LOADING. IT WAS REPORTED THAT THE LENS STUCK TO THE CARTRIDGE MORE THAN USUAL AND THAT THERE WAS NO PATIENT CONTACT. REPLACEMENT LENS WAS IMPLANTED ONE WEEK LATER. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: DIOPTER CORRECTED TO " -18.00/ 2.0/ 093 " IN THE INITIAL MDR. RETURNED TO MANUFACTURER DATE: 14-SEP-2018 IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8071101&pc=MTA | 2023826-2018-01679 | 2018-08-21 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. THERE WAS NO PATIENT CONTACT AND THE CAUSE OF THE EVENT WAS UNKNOWN. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN DRY; IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7898353&pc=MTA | 2023826-2018-01431 | 2018-08-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL13.2, CC4204A, -14.5 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. IT WAS REPORTED THAT THERE WAS NO PATIENT CONTACT. CAUSE OF EVENT CARTRIDGE, MODEL AND LOT NUMBERS NOT AVAILABLE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. LENS TORN INTO TWO PIECES. (B)(4). Manufacturer Narrative: CORRECTED DATA: INITIAL REPORT STATES DEVICE IS "A MICL 13.2, CC4204A, -14.5 DIOPTER"; CORRECT DESCRIPTION IS "A MICL 13.2, -14.5 DIOPTER". (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7927070&pc=MTA | 2023826-2018-01470 | 2018-09-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT, AFTER OPENING THE LENS BOTTLE OF A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0, THE SURGEON NOTICED A FOREIGN OBJECT INSIDE THE BOTTLE. THE OBJECT IS "YELLOWISH" IN COLOR AND THE LENS WAS NOT USED OR IMPLANTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THE LENS VIAL WAS NOT RETURNED. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND ANOTHER HAPTIC BENT. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8208778&pc=MTA | 2023826-2019-00002 | 2018-10-11 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IT TORE. THIS OCCURRED ON 11-OCT-2018. THIS WAS REPORTEDLY DUE TO A LOADING ERROR. THERE WAS NO PATIENT CONTACT AND A BACK-UP LENS WAS USED. ADDITIONAL INFORMATION IS NOT AVAILABLE. Manufacturer Narrative: CORRECTED DATA: CORRECTED TO: "UNKNOWN" THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS -10.5 DIOPTER, IT TORE DURING LOADING. THE CAUSE OF THE EVENT WAS DUE TO LOADING ERROR. THERE WAS NO PATIENT CONTACT. A BACK-UP LENS WAS USED. NO OTHER INFORMATION WAS PROVIDED. DEVICE EVALUATION CORRECTED TO: " NO LENS IN BOX. THE LENS VIAL WAS RETURNED EMPTY." (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6-LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8158123&pc=MTA | 2023826-2018-01788 | 2018-11-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, WAS TORN WHILE BEING PULLED INTO THE SUPER FUNNEL CARTRIDGE. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8159596&pc=MTA | 2023826-2018-01821 | 2018-11-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL12.1, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WOULD NOT LOAD PROPERLY. IT WAS REPORTED THAT THERE WAS NO PATIENT CONTACT AND THE BACK-UP LENS (SERIAL NUMBER (B)(4) WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN VIAL, IN LIQUID, IN A PLASTIC BAG. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8270845&pc=MTA | 2023826-2019-00069 | 2018-11-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VICM5 12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT CONTACT AND NO PATIENT INJURY REPORTED. IT WAS REPORTED THAT DURING INJECTION, THE USER REALIZED LENSES ABNORMITY SO THE LENS WAS NOT INJECTED. PATIENTS CURRENT CONDITION IS FINE AND A REPLACEMENT LENS IS PLANNED TO BE IMPLANTED AT A LATER DATE. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND PRESENCE OF RESIDUE ON LENS SURFACE. MODEL NUMBER CORRECTED TO VICM5 12.6. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8371356&pc=MTA | 2023826-2019-00337 | 2018-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/2.00/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING, WHEN POSITIONING IN THE INJECTOR. THERE WAS NO PATIENT CONTACT WITH THE DEVICE. A NEW LENS WAS ORDERED AND IMPLANTED ON (B)(6) 2018, PROBLEM RESOLVED. IT WAS REPORTED THAT THE FORCEPS WERE CHECKED FOR BURRS AND THEY HAD NO DEVIATIONS. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8192158&pc=MTA | 2023826-2018-01908 | 2018-11-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MIL13.7, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING WITH NO PATIENT CONTACT. THE BACK UP LENS WAS IMPLANTED WITHIN THE SAME SURGERY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. 2 SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8204888&pc=MTA | 2023826-2018-01927 | 2018-11-27 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, WAS TORN DURING THE LOADING PROCESS, AS THE LENS WAS BEING PULLED THROUGH THE CARTRIDGE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A CARTRIDGE CASE. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8270848&pc=MTA | 2023826-2019-00070 | 2018-12-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VICM5 13.2, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT CONTACT AND NO PATIENT INJURY REPORTED. IT WAS REPORTED THAT DURING INJECTION, THE USER REALIZED LENSES ABNORMITY SO THE LENS WAS NOT INJECTED. PATIENTS CURRENT CONDITION IS FINE AND A REPLACEMENT LENS WAS SUCCESSFULLY IMPLANTED ON THE SAME DATE. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372268&pc=MTA | 2023826-2019-00305 | 2018-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, - 11.0/+1.0/087 (SPHERE/CYLINDER/AXIS) TORE WHILE BEING LOADED AND THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE BECAUSE THE LENS HOLDER FORCEP DID NOT GRAB THE LENS PROPERLY SO THERE WAS A LENS BITE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8302662&pc=MTA | 2023826-2019-00227 | 2019-01-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE HAPTICS OF A 13.2MM, VTICM5_13.2, -13.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WERE BROKEN WHEN THE DOCTOR OPENED THE LENS. THERE WAS NO PATIENT CONTACT WITH THE LENS. A REPLACEMENT LENS WAS IMPLANTED AND THE PATIENT IS DOING WELL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372540&pc=MTA | 2023826-2019-00326 | 2019-01-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT, AFTER OPENING THE VIAL AND BEFORE LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/094 (SPHERE/CYLINDAR/AXIS), IT TORE/BROKE. THIS WAS MEANT TO BE IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT CONTACT WITH THE LENS AND THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. ON THE SAME DATE AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8588755&pc=MTA | 2023826-2019-00737 | 2019-03-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL LENS WAS IMPLANTED DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8620025&pc=MTA | 2023826-2019-00821 | 2019-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON FOUND THAT A13.2MM, VICM5_13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT CONTACT WITH THE DEVICE. A REPLACEMENT LENS WAS IMPLANTED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; LENS HAPTIC IS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8569669&pc=MTA | 2023826-2019-00724 | 2019-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON FOUND A 12.6MM, VICM5_12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS DAMAGE AFTER LOADING. THE SURGEON FOUND THE HATICS WAS BROKEN WHEN IT WAS LOADED INTO THE NOZZLE. THERE WAS NO PATIENT CONTACT. A BACKUP LENS WAS IMPLANTED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A CARTRIDGE CASE. THERE WAS CLEAR RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THAT THE LENS WAS CURVED, A HAPTIC WAS TORN AND RESIDUE ON LENS SURFACE. THE CARTRIDGE WAS RETURNED IN DEVICE TRAY AND THERE WAS CLEAR RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMGE TO DEVICE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595422&pc=MTA | 2023826-2019-00761 | 2019-04-03 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS REMOVING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, FROM THE VIAL, HE NOTICED IT WAS TORN. THIS OCCURRED ON (B)(6) 2019. "DEFECTIVE" IS RECORDED ON THE BOX. THERE WAS NO PATIENT CONTACT WITH THE LENS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8625107&pc=MTA | 2023826-2019-00784 | 2019-04-09 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 THE SURGEON FELT THAT THE INJECTOR WAS "BINDING" WHEN HE TRIED TO ADVANCE THE FOAM TIP PLUNGER AND WAS CONCERNED THE 13.2MM, MICL 13.2, IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER WAS "STUCK." SURGEON EXTRUDED THE LENS, "THOUGHT HE SAW A SMALL TEAR IN THE LENS AND DECIDED TO USE THE BACK UP LENS." THERE WAS NO PATIENT CONTACT WITH THE LENS. THE PATIENT IS DOING WELL POSTOPERATIVELY WITH UNCORRECTED VISION OF 20/20. Manufacturer Narrative: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TEARS/TORN. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609209&pc=MTA | 2023826-2019-00826 | 2019-04-17 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IT TORE. THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS AND THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876526&pc=MTA | 2023826-2019-01399 | 2019-06-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, TORE WHILE BEING LOADED INTO THE CARTRIDGE. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895278&pc=MTA | 2023826-2019-01423 | 2019-06-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, WAS TORN DURING LOADING INTO THE CARTRIDGE. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANTED: NA. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912813&pc=MTA | 2023826-2019-01512 | 2019-07-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER TORE/BROKE DURING THE LOADING PROCESS. THERE WAS NO PATIENT CONTACT WITH THE LENS. THIS OCCURRED ON (B)(6) 2019. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER: "26-JUL-2019" SHOULD BE CORRECTED TO "29-JUL-2019" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THAT THE LENS OPTIC AND HAPTIC HAD TEARS AND THERE WAS FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9130177&pc=MTA | 2023826-2019-01788 | 2019-07-18 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.00 DIOPTER, IMPLANTABLE COLANDER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION : LENS WAS RETURNED IN TWO PIECES IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THAT THE OPTIC AND HAPTIC WERE TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941319&pc=MTA | 2023826-2019-01570 | 2019-07-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO LOAD A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IT TORE. THE SURGEON REPORTS THAT HE FELT LIKE HE HAD APPROPRIATELY GRASPED THE LENS BUT THE EDGE RIPPED WHEN PULLING IT THROUGH. THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL, INSIDE OF A BIOHAZARD BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO PATIENT INVOLVEMENT. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8957661&pc=MTA | 2023826-2019-01560 | 2019-07-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS NOTICED TO BE TORN BEFORE VIAL WAS OPENED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT. BACK-UP LENS WAS IMPLANTED AND PROBLEM RESOLVED. PER THE REPORTER ON (B)(6) 2019, "UPON OPENING THE PACKAGE, THE SURGEON NOTICED LOOKING THROUGH THE VIAL THAT THE LENS WAS TORN." Manufacturer Narrative: ADDITIONAL INFORMATION; DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. LENS VIAL RETURNED WITH VIAL SEAL REMOVED. (B)(4). Manufacturer Narrative: RESULT CODE 114 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9076365&pc=MTA | 2023826-2019-01715 | 2019-08-26 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO LOAD A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IT TORE. THE SURGEON REPORTS THAT THE LENS WAS LOADED IN THE SAME MANNER AS ALL OTHER ICLS USING THE RECOMMENDED INSTRUMENTATION. BACKUP LENS AND NEW INJECTOR USED SUCCESSFULLY. NOTHING UNUSUAL OCCURRED DURING THE LOADING PROCESS. IT'S POSSIBLE THE SURGEON MAY HAVE PULLED TOO HARD/FAST ON THE LENS HOWEVER THE SAME LOADING PROCESS IS FOLLOWED IN ALL CASES. THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS. Manufacturer Narrative: NO PATIENT INVOLVEMENT. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9221386&pc=MTA | 2023826-2019-01970 | 2019-09-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER TORE/BROKE DURING INJECTION INTO THE LEFT (OS) EYE. THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: EVENT: THE LENS WAS REPLACED WITH A SAME LENGTH/MODEL LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9372551&pc=MTA | 2023826-2019-02255 | 2019-10-29 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE PLUNGER KEPT DRAGGING AND PULLING A 12.6MM, MICL12.6, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS. THE LENS WOULD NOT FULLY INSERT AND THEN IT TORE. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TEAR ON THE OPTIC AND A TORN HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9372698&pc=MTA | 2023826-2019-02254 | 2019-10-29 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE PLUNGER KEPT DRAGGING AND PULLING A 12.6MM, MICL12.6, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS. THE LENS WOULD NOT FULLY INSERT AND THEN IT TORE. THERE WAS NO PATIENT CONTACT. LENS WAS SHREDDED INTO LITTLE PIECES AND DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441178&pc=MTA | 2023826-2019-02349 | 2019-11-12 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, HE HAD TROUBLE "FOLDING THE LENS WITH THE CARTRIDGE." THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS. REPORTER STATES THE LENS "WOULDN'T FOLD LIKE USUAL" AND THE SURGEON "DIDN'T WANT TO RUN INTO ANY ISSUES." Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS.-LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9684506&pc=MTA | 2023826-2020-00257 | 2019-12-10 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, - 10.0 DIOPTER, GOT STUCK DURING LOADING AND WAS DAMAGED. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PATIENTS CURRENT STATUS IS GOOD. THE CAUSE OF THE EVENT WAS DUE TO A LOADING ERROR. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559348&pc=MTA | 2023826-2020-00034 | 2019-12-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, THE LENS TORE. THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS. REPORTER STATES THE CAUSE OF THE EVENT IS UNKNOWN AS NORMAL LOADING PROCEDURE WAS FOLLOWED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, STUCK ON TAPE. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THE LENS WAS RETURNED COILED AND UNABLED TO BE EVALUATED. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. AN INJECTOR, CARTRIDGE, AND INJECTOR SYSTEM WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781694&pc=MTA | 2023826-2020-00401 | 2019-12-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER TORE/BROKE DURING LOADING. REPORTEDLY THE "LENS WAS DAMAGED DURING LOADING THE LENS INTO THE NOZZLE." THIS OCCURRED ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON THE SAME DATE AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE AND DEBRIS ON LENS. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721640&pc=MTA | 2023826-2020-00335 | 2020-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT WHEN PREPARING A 13.2MM, VICM5_13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IT WAS NOTED THAT IT WAS RIPPED. THE BACKUP ICL WAS USED AND THE SURGERY WENT WELL AND ALL WAS FINE WITH THE PATIENT POST-OP.CAUSE OF EVENT "DEVICE". Manufacturer Narrative: H3- LENS WAS RETURNED WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN DEBRIS AND A SCRATCH ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869330&pc=MTA | 2023826-2020-00610 | 2020-02-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER TORE/BROKE BEFORE/DURING LOADING. THE DAMAGE WAS NOTED AFTER OPENING THE LENS VIAL. THIS OCCURRED ON (B)(6) 2020. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETUNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988288&pc=MTA | 2023826-2020-00940 | 2020-03-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM, VICM5_12.6, -6.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS FOUND DAMAGE DURING INJECTION. THE SURGEON DISCONTINUED USING THE LENS, NO PATIENT CONTACT. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965152&pc=MTA | 2023826-2020-00895 | 2020-03-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS ATTEMPTING TO LOAD A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, INTO THE CARTRIDGE AND IT WAS NOTED THE LENS WAS DAMAGED. THE SURGEON FELT THE LENS WAS TORN DURING MANIPULATING USING THE FORCEP. THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED FOR THE BACKUP LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988436&pc=MTA | 2023826-2020-00951 | 2020-03-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, AND FOUND THE LENS DAMAGED. THERE WAS NO PATIENT CONTACT. THE SURGEON IMPLANTED ANOTHER SAME MODEL/LENGTH/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA : H1- TYPE OF REPORTABLE EVENT IN PREVIOUS MDR IS CORRECTED FROM SERIOUS TO MALFUNCTION. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135965&pc=MTA | 2023826-2020-01169 | 2020-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, WAS TORN WHILE BEING PULLED THROUGH THE CARTRIDGE DURING THE LOADING PROCESS. THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271245&pc=MTA | 2023826-2020-01492 | 2020-05-15 04:00 | Malfunction | Unknown Manufacturer | 2020-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS NO PATIENT CONTACT. DURING INITIAL SURGERY ON (B)(6) 2020, A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189418&pc=MTA | 2023826-2020-01279 | 2020-05-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, FROM THE VIAL AND NOTED THE LENS WAS PARTICULARLY SOFT AND COULD NOT SPREAD IN THE BSS. THE LENS WAS NOT IMPLANTED IN THE PATIENTS EYE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10463781&pc=MTA | 2023826-2020-01915 | 2020-06-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS GOT STUCK IN THE CARTRIDGE OR INJECTION SYSTEM, NO PATIENT CONTACT. THE REPORTER STATED " LENS WAS SET AS USUAL. WHEN INJECTING THE LENS, THE LENS SEEMED TO BE A LITTLE STUCK BUT ENFORCED TO INJECT. THEN LENS WAS DAMAGED." A REPLACEMENT LENS SUCCESSFULLY COMPLETED THE SURGERY. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7680214&pc=MTA | 2023826-2018-01020 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 13.7MM MICL13.7, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN CARTRIDGE (SFC-45). VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE AND RESIDUE ON THE DEVICE. WORK ORDER SEARCH: (B)(4) SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8067019&pc=MTA | 2023826-2018-01613 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-11-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: IT WAS REPORTD THAT A 12.6MM, VICM5_12.6, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. THERE WAS NO PATIENT CONTACT WITH THE DEVICE AND THE SURGEON PLANS TO IMPLANT ANOTHER LENS. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372297&pc=MTA | 2023826-2019-00307 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, WAS TORN WHILE BEING LOADED AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9416981&pc=MTA | 2023826-2019-02236 | 2019-03-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Ejection Problem | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VICM5 13.7, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS "STUCK IN THE INJECTOR" DURING INJECTION INTO PATIENTS LEFT EYE (OS). NO PATIENT CONTACT. REPLACEMENT LENS IMPLANTED ON (B)(6) 2019. REPORTEDLY "THE LENS WAS PLACED IN THE INJECTOR CORRECTLY." THE LENS "DID NOT MOVE" WHEN THE ROD WAS PRESSED, "THE LENS WAS PUT BACK INTO POSITION." DURING INJECTION OF THE LENS INTO PATIENT'S EYE, THE ROD PASSED OVER THE LENS. LENS DID NOT "MOVE FORWARD OR TOUCH THE EYE." Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS - INJECTOR MODEL: MSI-PF, LOT# UNK. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AS WELL AS THE HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8710344&pc=MTA | 2023826-2019-01098 | 2019-04-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; Contamination /Decontamination Problem | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS PREPARING TO IMPLANT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), IT WAS DISCOVERED TO HAVE A FOREIGN BODY ADHESION. AS THE SURGEON CHECKED THE LENS BEFORE THE SURGERY, HE FOUND A TINY "PLASTIC DEBRIS" ON THE LENS. HE WASHED THE LENS OUT WITH BSS AND IMPLANTED IT. THE LENS REMAINS IMPLANTED. THIS OCCURRED ON (B)(6) 2019. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307846&pc=MTA | 2023826-2019-00217 | 2018-07-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE PREPARING TO IMPLANT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2018. THIS LENS WAS NOT IMPLANTED, NOR WAS THERE ANY PATIENT CONTACT WITH THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. Manufacturer Narrative: H3: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CORRECTED DATA H10: COMMENT "H6-PATIENT CODE 3191: NO CODE AVAILABLE (SEONDARY SURGERY, LENS EXCHANGE)" IS NOT APPLICABLE FOR THIS CLAIM.(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10356711&pc=MTA | 2023826-2020-01607 | 2020-07-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -16.00 DIOPTER, AND NOTED A NICK IN THE LENS. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED AND THE PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # : (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7449228&pc=MTA | 2023826-2018-00601 | 2018-03-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Scratched Material | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON NOTICED A SCRATCH ON A 13.2 MM MICL13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WHILE LOADING IT UNDER A MICROSCOPE. THERE WAS NO PATIENT CONTACT WITH THE LENS. A BACK UP LENS WAS USED, BUT THE PROCEDURE WAS NOT INTRAOPERATIVE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LENS VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TEAR. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8471403&pc=MTA | 2023826-2019-00507 | 2019-02-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Ejection; No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, DEPLOYED PREMATURELY AND WAS DEFECTIVE. THERE WAS NO PATIENT CONTACT OR INJURY AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON LOADED AND WAS PRIMING THE LENS BEFORE INJECTING AND THE LENS PREMATURELY DEPLOYED. THERE WAS NO PATIENT CONTACT. THE SURGEON WAS NOT SURE IF THE LENS WAS DAMAGED SO HE LOADED AND IMPLANTED THE BACKUP LENS, WHICH WAS IMPLANTED WITH NO PROBLEM. THE PATIENT IS DOING WELL. THE REPORTER INDICATED THERE WAS NO ISSUE WITH THE LENS OR INJECTION SYSTEM AND WAS MOST LIKELY DUE TO A USER ERROR. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A CARTRIDGE CASE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON THE LENS. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7206051&pc=MTA | 2023826-2018-00102 | 2017-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Cut In Material; Device Operates Differently Than Expected | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON NOTICED A TEAR ON A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS,-8.5/+3.0/94 (SPHERE/CYLINDER/AXIS), WHEN THE LENS WAS REMOVED FROM THE VIAL. THE LENS WAS NOT USED, NO PATIENT CONTACT. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE LENS MISSING A PIECE OF ITS HAPTIC. THERE WAS PRESENCE OF FIBER- LIKE DEBRIS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7188509&pc=MTA | 2023826-2018-00066 | 2017-12-15 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | No Patient Involvement; No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON NOTICED MICROSCOPIC BLACK FLAKES ON A MICL13.2, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS WHEN THEY OPENED THE LENS VIAL AND WERE PREPARING TO LOAD THE LENS IN THE CARTRIDGE FOR THE INJECTION ON (B)(6) 2017. THE SURGEON DESCRIBED OPENING THE LENS VIAL WITH HEMOSTATS, REMOVING THE METAL RING, AND REMOVING THE GREY STOPPER WITH A STERILE GLOVED HAND. THE LENS WAS POURED OUT OF THE VIAL ONTO THE INJECTOR CARTRIDGE TRAY WITH THE ASSISTANCE OF A WEK-CELL TO GUIDE THE LENS OUT OF THE VIAL. WHEN THE LENS WAS PICKED UP WITH ICL LOADING FORCEPS, THE BLACK FLAKES WERE NOTICED. THERE WAS NO PATIENT CONTACT. THEY USED A NEW CARTRIDGE/INJECTOR AND A NEW ICL WAS LOADED AND USED TO PERFORM THE SURGERY. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6200475&pc=MTA | 2023826-2016-01808 | 2016-05-26 | Malfunction | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS ATTEMPTING TO IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/179 DIOPTER IN THE PATIENT'S LEFT EYE (OS) AND IT WAS NOTED THAT THE LENS WAS TORN DURING LOADING. BACKUP LENS OF THE SAME SIZE, MODEL AND DIFFERENT DIOPTER WAS USED. THERE WAS NO REPORT OF ANY APPARENT INJURY. Manufacturer Narrative: (B)(4). BASED ON INFORMATION FROM REPORTER, IT CANNOT BE CONFIRMED THAT DEVICE MALFUNCTION OR USER ERROR CONTRIBUTED TO THIS EVENT. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD CODE(S): CLAIM EVALUATED BASED ON REPORTER RESPONSE ON COMPLAINT QUESTIONNAIRE. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050906&pc=MTA | 2023826-2017-01715 | 2017-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6 MM VICMO12.6, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). BEFORE THE LENS WAS INJECTED, IT WAS NOTICED THAT THE LENS BROKE. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL . VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7072216&pc=MTA | 2023826-2017-01809 | 2017-10-26 | Malfunction | STAAR SURGICAL COMPANY | 2017-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/145 DIOPTER, AND THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE. THE REPORTER INDICATED THE LENS WAS TORN WHILE LOADING INTO THE CARTRIDGE AND THERE WAS NO PATIENT CONTACT. THE SURGERY WAS ABORTED. ANOTHER SAME MODEL LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL WITH VISON OF 20/20. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: PATIENT IDENTIFIER: IT IS CORRECTED FROM (B)(4). (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND DEBRIS ON LENS, THE LENS MISSING A PIECE OF HAPTIC AND THE LENS SHAPE DISTORTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9158754&pc=MTA | 2023826-2019-01833 | 2019-09-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Ejection Problem | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, WAS DAMAGED DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS), DUE TO THE FOAM TIP PLUNGER OVERRODE THE LENS UPON INJECTION. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THERE WAS NO PATIENT CONTACT. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339933&pc=MTA | 2023826-2018-00415 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Operates Differently Than Expected | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS UNABLE TO INJECT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, - 10.0/+2.0/091 (SPHERE/CYLINDER/AXIS), DUE TO WRONG VISCOELASTIC WAS USED. THE LENS WAS DAMAGED DURING THE LOADING PROCESS. THE LENS WAS NOT IMPLANTED AND THERE WAS NO REPORT OF PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY ON A MICRO CENTRIFUGE VIAL WITH PRESENCE OF CLEAR RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE LENS WAS MISSING A PART AND THE PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7975171&pc=MTA | 2023826-2018-01530 | 2018-04-08 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON FOUND A VICM5_12.6, -07.50 DIOPTER, INTRAOCULAR LENS DAMAGE DURING LENS SETTING, WHILE REMOVING FROM VIAL. THE REPLACEMENT LENS (SAME MODEL AND DIOPTER) WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172173&pc=MTA | 2023826-2018-01853 | 2018-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER WAS REPORTED AS A COMPLAINT. RECEIVED INFORMATION THAT THE LENS WAS NOT IMPLANTED AND "CORNEA IS CLEAR." ATTEMPTS TO OBTAIN CLARIFICATION/ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTED INDICATED THAT, BEFORE OPENING THE VIAL CONTAINING A 13.2MM VICMO 13.2 -12.0 DIOPTER IMPLANTABLE COLLAMER LENS, IT WAS DISCOVERED TO HAVE A TEAR/BREAK. THIS OCCURRED ON (B)(6) 2018. THE CAUSE OF THE EVENT IS REPORTED AS "TRANSPORTATION TO DISTRIBUTOR." NOTING THAT 2 LENSES WERE SHIPPED AND 1 WAS DAMAGED. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PLEASE DISREGARD PREVIOUS STATEMENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8271267&pc=MTA | 2023826-2019-00020 | 2018-11-02 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 BEFORE IMPLANT, A 13.2 MM, MICL13.2, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS LOOKED AT UNDER THE MICROSCOPE, THE SHAPE LOOKED DIFFERENT THAN ANY OTHER ICL DOCTOR HAD IMPLANTED. THE BACK UP LENS WAS IMPLANTED, THIS RESOLVED THE ISSUE. CAUSE OF THE EVENT IS REPORTED AS DEVICE AND OTHER; POSSIBLE DEFECTIVE LENS. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8218424&pc=MTA | 2023826-2019-00014 | 2018-11-12 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -16.50 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING. THERE WAS NO PATIENT CONTACT, CAUSE OF THIS EVENT WAS REPORTED AS USER ERROR. HOWEVER, IT WAS ALSO REPORTED THAT THE LENS WAS LOADED CORRECTLY/ PATIENT IS OKAY. ADDITIONAL INFORMATION/CLARIFICATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: DATE SHOULD BE 05-DEC-2018 IN ORIGINAL MDR. DATE SHOULD BE 05-DEC-2018 IN ORIGINAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER CLARIFIED THAT THE REPLACEMENT LENS WAS LOADED AND IMPLANTED CORRECTLY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9612927&pc=MTA | 2023826-2020-00151 | 2019-12-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER TORE BEFORE LOADING THE LENS INTO THE CARTRIDGE. THIS OCCURRED ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON THE SAME DATE AND THE PROBLEM IS RESOLVED. THE SURGEON SUSPECTS THE LENS WAS TORN BEFORE OPENING THE VIAL AS THE TEAR WAS DISCOVERED PRIOR TO LOADING. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL , DRY WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CORRECTED DATA: H6 - PATIENT CODE: 2692 SHOULD BE CORRECTED TO 2645. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782558&pc=MTA | 2023826-2020-00428 | 2020-01-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING. THERE WAS NO PATIENT CONTACT AND THE CAUSE OF THE EVENT WAS UNKNOWN. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE REPORTER STATED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5: THE IMPLANTABLE COLLAMER LENS TORE/BROKE BEFORE/DURING LOADING. DAMAGE WAS NOTICED DURING LOADING. B5: "THERE WAS NO PATIENT CONTACT" SHOULD BE REMOVED AS IT WAS NOT STATED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC AND OPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7627372&pc=MTA | 2023826-2018-00927 | 2018-05-22 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Issue; No Apparent Adverse Event | No Patient Involvement; Patient Problem/Medical Problem | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS NOT HAPPY WITH THE WAY A 13.2MM MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS FOLDED AT LOADING. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN A ZIPLOC BAG. VISUAL INSPECTION FOUND RESIDUE ON THE LENS SURFACE AND PORTION OF THE LENS MISSING. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816882&pc=MTA | 2023826-2019-01223 | 2019-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Not Applicable; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM, VICMO12.1, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816881&pc=MTA | 2023826-2019-01220 | 2019-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Not Applicable; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICMO13.2, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). PATIENT CODE 3189 CORRECTED TO 3191. PATIENT CODE 3189 NOT APPLICABLE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9484640&pc=MTA | 2023826-2019-02417 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Not Applicable; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, MICL12.6, - 16.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. ENLARGEMENT OF PI (YAG) WAS ALSO REPORTED AS A SECONDARY SURGICAL INTERVENTION. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 3189 SHOULD BE CORRECTED TO 3191. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT:UNK. ETHNICITY:UNK. RACE:UNK. DATE OF EVENT:UNK. DEVICE EVALUATION : LENS WAS RETURNED DRY, IN LENS CASE. VISUAL INSPECTION FOUND A HAPTIC TORN AND BENT. RESIDUE ON THE LENS WAS ALSO OBSERVED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816891&pc=MTA | 2023826-2019-01224 | 2018-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Not Applicable; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM, VICMO12.1, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. DEVICE CODE 1494: OFF-LABEL USE- ACD < 3.00MM. CLAIM #: (B)(4). Manufacturer Narrative: PATIENT CODE "3189" SHOULD BE CORRECTED TO PATIENT CODE "3191" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816885&pc=MTA | 2023826-2019-01225 | 2018-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Not Applicable; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM, VICMO12.1, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: DATE OF REPORT: AWARE DATE IN INITIAL MDR CORRECTED TO 24-JUN-2019. DATE RECEIVED BY MANUFACTURER SHOULD BE CORRECTED TO 24-JUN-2019 IN INITIAL MDR. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816889&pc=MTA | 2023826-2019-01216 | 2018-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Not Applicable; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM, VICMO12.1, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: AWARE DATE IN INITIAL MDR CORRECTED TO 24-JUN-2019. DATE RECEIVED BY MANUFACTURER SHOULD BE CORRECTED TO 24-JUN-2019 IN INITIAL MDR. DATE RECEIVED BY MANUFACTURER SHOULD BE CORRECTED TO 31-JUL-2019 IN FOLLOW-UP#1. PATIENT CODE 3191: NO CODE AVAILABLE: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE PATIENT CODE 3189 IN INITIAL MDR NO LONGER APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCETRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DEVICE CODE 1494: OFF LABEL USE - ACD<3.0MM. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5984621&pc=MTA | 2023826-2016-01334 | 2016-01-08 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Obstruction/Occlusion; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND TORN. (B)(4). REPLACE "THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20" TO "THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5984622&pc=MTA | 2023826-2016-01335 | 2016-01-31 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Obstruction/Occlusion; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: THE PATIENT IS UNDER THE AGE OF 21 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8574197&pc=MTA | 2023826-2019-00672 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | N/A | 1 | Event Description: THE REPORTER INDICATED THAT "PATIENT COMPLAINTS ABOUT PAIN BINOCULAIR WHILE LOOKING UP CLOSE TO A COMPUTER, PHONE OR IPAD (ONLY DEVICES THAT ARE GIVING LIGHT). PATIENT DOES NOT HAVE ANY REFRACTIVE ERROR AND ALSO ALL THE OTHER PARAMETERS ARE FINE." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8575590&pc=MTA | 2023826-2019-00673 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | N/A | 1 | Event Description: THE REPORTER INDICATED THAT "PATIENT COMPLAINTS ABOUT PAIN BINOCULAIR WHILE LOOKING UP CLOSE TO A COMPUTER, PHONE OR IPAD (ONLY DEVICES THAT ARE GIVING LIGHT). PATIENT DOES NOT HAVE ANY REFRACTIVE ERROR AND ALSO ALL THE OTHER PARAMETERS ARE FINE." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10496543&pc=MTA | 2023826-2020-01974 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VISIAN ICL LENS, -17.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. THE SURGEON REPORTS PAIN IN THE CILIARY ARCH. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. REPORTEDLY, THE DOCTOR THINKS POSSIBLY THIS PAIN IS NOT DUE TO THE ICL. Manufacturer Narrative: B5-UPDATED INFORMATION: REPORTEDLY, THE PATIENT IS "PERFECT" NOW AND WITHOUT SYMPTOMS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11547918&pc=MTA | 2023826-2021-00739 | 2021-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Defective Device; Unintended Movement | Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT WAS REPORTED WITH HYPEROPIA AND ASTIGMATISM, WHICH CAUSED BAD AND BLURRED VISION. ALSO REPORTED PERMANENT LOSS OF BCVA. THE LENS WAS ALSO REPORTED AS HAVING FLIPPED. THE PATIENT HAD A SMALL EYE, THERE WAS PUPILAR CONSTRICTION AND THE PUPIL BEGAN TO CLOSE DURING THE SURGERY. THE IRIS GOT EXTRUDED OUT THE WOUND DURING THE SURGERY. THE LENS WAS EXPLANTED AND RE-IMPLANTED ON (B)(6) 2021. PERIBULBAR BLOCK WAS ADMINISTERED DURING 2ND SURGERY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY WITH THE RESULTS OF THE SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS ALSO INDICATED THE PATIENT EXPERIENCED PAIN; NON CONFORMITY THAT WASN'T SPECIFIED WHAT TYPE; LENS WAS IMPLANTED UPSIDE DOWN/FLIPPED, LENS WAS EXPLANTED ON 03-MARCH-2021 AND THE DAY AFTER WAS RE-IMPLANTED. H6: 4581- LOSS OF BCVA, IRIS PROLAPSE. CLAIM#: (B)(4). Manufacturer Narrative: IT WAS LATER CLARIFIED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND PROPERLY REIMPLANTED, INTRAOPERATIVELY AND NOT IN A SEPARATE SURGERY. THIS HAD RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11557439&pc=MTA | 2023826-2021-00740 | 2021-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Defective Device; Unintended Movement | Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT WAS REPORTED WITH HYPEROPIA AND ASTIGMATISM, WHICH CAUSED BAD AND BLURRED VISION. ALSO REPORTED PERMANENET LOSS OF BCVA. THE LENS WAS ALSO REPORTED AS HAVING FLIPPED. THE PATIENT HAD A SMALL EYE, THERE WAS PUPILAR CONSTRICTION AND THE PUPIL BEGAN TO CLOSE DURING THE SURGERY. THE LENS WAS EXPLANTED AND RE-IMPLANTED ON (B)(6) 2021. PERIBULBAR BLOCK WAS ADMINISTERED DURING THE 1ST SURGERY (INITIAL) AND 2ND SURGERY (RE-IMPLANT). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. TWO HOURS AFTER THE RE-IMPLANT SURGERY, ANTERIOR CHAMBER A LITTLE BIT NARROW AND PUPIL DILATED. THE PATIENT IS HAPPY WITH THE RESULTS OF THE SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS ALSO INDICATED THAT THE PATIENT EXPERIENCED PAIN. REPORTER INDICATED THERE WAS PRODUCT NONCONFORMITY BUT THIS WAS NOT SPECIFIED ON. LENS WAS IMPLANTED UPSIDE DOWN/FLIPPED. IT WAS LATER CLARIFIED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021, AND THE NEXT DAY RE-IMPLANTED. H6 - 4581 - LOSS OF BCVA. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5/H10- IT WAS LATER CLARIFIED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND PROPERLY REIMPLANTED, INTRAOPERATIVELY AND NOT IN A SEPARATE SURGERY. THIS HAD RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7676124&pc=MTA | 2023826-2018-01028 | 2018-06-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Pain; Blurred Vision; Discomfort | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6, +4.0/+5.5/035 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE AND REPORTEDLY THE PATIENT IS "REALLY UNCOMFORTABLE." THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721248&pc=MTA | 2023826-2018-01078 | 2017-12-07 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Pain; Blurred Vision; Halo; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES, BLURRED VISION, PAIN, AND "NO COMFORT FOR NEAR AND MIDDLE VISION." THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917237&pc=MTA | 2023826-2015-00799 | 2015-06-17 | Injury | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO NO LENS VAULT. THERE WAS NO LENS OPACITY. THE PATIENT EXPERIENCED BLURRY VISION AND EYE PAIN. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY INDICATED THE DATE OF INITIAL SURGERY WAS (B)(6) 2015. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. USE OF THE WRONG LENS WAS MOST LIKELY THE CAUSE OF LACK OF VAULT. (B)(4). Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: CORRECTION - IMPLANTED DATE - (B)(6) 2015. METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - THE CONSEQUENCE OF INADEQUATE VAULTING COULD BE CATARACT THAT WARRANT LENS EXCHANGE TO PREVENT FURTHER DAMAGE. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074717&pc=MTA | 2023826-2016-01575 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS REPOSITIONED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS BUT THAT LENS ALSO IS VAULTED. THE PATIENT'S POST-OP BSCVA WAS 20/30 +1, WITH BLURRY VISION AND PAIN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403798&pc=MTA | 2023826-2019-00367 | 2019-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; No Apparent Adverse Event | Pain; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -9.50/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2017. THE REPORTER INDICATED THAT AN ACCIDENT TO THE PATIENTS EYE HAS RESULTED IN BLUR VISION WHEN READING OR CONCENTRATING, AND PAIN WITH LIGHT CHANGES AND WHILE LOOKING AT A MOVING OBJECT. REPORTEDLY, THE PATIENT HAD A VISUAL CHECK WITH THE SURGEON AND THE VISUAL CHECK WAS "PERFECT." THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2017. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6143452&pc=MTA | 2023826-2016-01723 | 2016-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-12-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Misfocusing; Device Operates Differently Than Expected | Pain; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE PATIENT HAD AN ICL WITH AN UNKNOWN SERIAL NUMBER AND DIOPTER IMPLANTED IN BANGALORE 28 MONTHS PRIOR TO (B)(6) 2016, AND FOR THE LAST TWO MONTH THE PATIENT HAS EXPERIENCED SOME DISCOMFORT SEEING OBJECTS IN THE RIGHT EYE, WHICH NEED MORE FOCUS. THE PATIENT FEELS SLIGHT PAIN IN THE LEFT EYE. PATIENT VISITED SAME SURGEON WHO IMPLANTED ICL. EXAM INDICATES THAT EYE POWER AT -0.50 AND -0.25 "IN BOTH EYES", SPECIFIC EYE/POWER WAS NOT CLEAR. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT CODE: (B)(4) - REFRACTIVE CHANGE OVER TIME. CORRECTED DATA: THE REPORTER INDICATED THAT THE PATIENT HAD AN ICL WITH AN UNKNOWN SERIAL NUMBER AND DIOPTER IMPLANTED IN PATIENT'S LEFT EYE (OS) IN (B)(4) 28 MONTHS PRIOR TO (B)(6) 2016. THE PATIENT FEELS SLIGHT PAIN IN THE LEFT EYE. PATIENT VISITED SAME SURGEON WHO IMPLANTED ICL. EXAM INDICATES THAT EYE POWER AT -0.50 AND -0.25 "IN BOTH EYES," SPECIFIC EYE/POWER WAS NOT CLEAR. (B)(4). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): ATTEMPTED TO CONTACT PATIENT FOR ADDITIONAL INFORMATION, NONE FORTH COMING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14107495&pc=MTA | 2023826-2022-01095 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Pain; Discomfort | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 FOR RELIEF OF PAIN AND DISCOMFORT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12226996&pc=MTA | 2023826-2021-02915 | 2021-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Pain; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -9 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTED PIGMENT DISPERSION AND PATIENT EXPRESSED A BIT OF PAIN/DISCOMFORT. AT A FOLLOW UP DATE AN OCT SCAN WAS PERFORMED AND IT WAS INDICATED THAT IOP WAS NORMAL AND PATIENT IS OKAY, THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "THIS IRIS SHOULD BE FRAGILE BECAUSE OD IS OK, 13.2MM ICL WAS IMPLANTED AS WELL". Manufacturer Narrative: (B)(4). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8628687&pc=MTA | 2023826-2019-00889 | 2018-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain; Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICM5_12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED DISCOMFORT AND ACHE, ON (B)(6) 2018 THE LENS WAS EXPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER EXPLANTING THE LENS. THERE IS NO PROBLEM WITH THE PATIENT. AT PATIENTS LAST VISIT, BCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: EVENT DESCRIPTION (EXPLANT DATE) "(B)(6) 2018" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186507&pc=MTA | 2023826-2019-01860 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain; Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5 12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY "THE PATIENT HAD A PAIN OF THE EYE AND THERE WAS NO TREATMENT FOR THE PAIN." ON (B)(6) 2019 THE LENS WAS REMOVED. PER THE REPORTER ON (B)(6) 2019, "AFTER REMOVING THE LENS", THE PATIENT IS IN GOOD CONDITION. LENS REMOVAL RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION; LENS EXPLANT. RESULT CODE 213 CORRECTED TO 114. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE ON LENS, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186511&pc=MTA | 2023826-2019-01861 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain; Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5 12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON 20/FEB/2018. REPORTEDLY "THE PATIENT HAD A PAIN OF THE EYE AND THERE WAS NO TREATMENT FOR THE PAIN." ON 10/AUG/2019 THE LENS WAS REMOVED. PER THE REPORTER ON 22/SEP/2019, "AFTER REMOVING THE LENS", THE PATIENT IS IN GOOD CONDITION. LENS REMOVAL RESOLVED THE PROBLEM. Manufacturer Narrative: RESULT CODE 213 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO RESULT CODE 114 IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE ON LENS, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE 3191: 'EXCHANGE" SHOULD BE CORRECTED TO "EXPLANT" IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5922831&pc=MTA | 2023826-2016-01210 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Pain; Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT PAIN AND DISCOMFORT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (NOT POSSIBLE TO SPECIFY) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN A PATIENT UNDER THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT WAS (B)(6) AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917994&pc=MTA | 2023826-2016-01169 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Pain; Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT DISCOMFORT AND PAIN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL WAS OBSERVED TO BE ON THE LENS (NOT POSSIBLE TO SPECIFY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4) DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094953&pc=MTA | 2023826-2016-01636 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO PATIENT WAS IN EXTREME PAIN. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721911&pc=MTA | 2023826-2020-00312 | 2020-01-13 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-18 | MTA | IMPLANTABLE COLLAMER LENS | Contamination /Decontamination Problem | Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. APPROXIMATELY 1-MONTH AFTER SURGERY THE SURGEON REPORTS WHITE DEPOSITS FOUND ON THE LENS AND THE PATIENT HAS PAIN AND CONGESTION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248831&pc=MTA | 2023826-2015-01539 | 2015-05-30 | Injury | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -04.00/+4.50/X114 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT STARTED TO HAVE SEVERE PAIN 15 MINUTES POST-OP. THE LENS WAS EXPLANTED THAT SAME DAY. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PEN MARKS ON LENS SURFACE. (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. PATIENT'S ACD IS 2.97 MM. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -04.00/+4.50/X114. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6296898&pc=MTA | 2023826-2017-00165 | 2017-01-02 | Injury | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON PLANS TO EXCHANGE LENS FOR A LARGER SIZE DUE TO LOW VAULTING. LENS REMAINS IMPLANTED. THE PATIENT HAS PAIN IN THEIR RIGHT EYE (OD). Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON 01/25/2017 AND EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: FOLLOWUP #1 WAS INADVERTENTLY SKIPPED AND FOLLOWUP #2 AND #3 WERE SUBMITTED INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261493&pc=MTA | 2023826-2018-00173 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.11/+0.33/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE PATIENT BEGAN TO COMPLAIN OF OCULAR PAIN. REPORTEDLY, ALSO NOTED IS EXCESSIVE VAULT AND THAT THE LENS IS "COMPROMISING THE ANGLES." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: LENS WAS EXCHANGED AS THE VAULT WAS EXCESSIVE. (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261496&pc=MTA | 2023826-2018-00174 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.50, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE PATIENT BEGAN TO COMPLAIN OF OCULAR PAIN. REPORTEDLY, ALSO NOTED IS EXCESSIVE VAULT AND THAT THE LENS IS "COMPROMISING THE ANGLES." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: LENS WAS EXCHANGED WITH A DIFFERENT POWER, VAULT WAS FINE (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6775540&pc=MTA | 2023826-2017-01298 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, PUPILLARY BLOCK (ELEVATED IOP), BLURRED VISION, AND PAIN. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A CASE/VIAL. THE LENS VIAL WAS RETURNED TAPED TO LAF CARD. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7146624&pc=MTA | 2023826-2017-01945 | 2016-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Pain; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -7 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE, GLARES, HALOES, INDUCED ASTIGMATISM, AND CILIARY GROOVE PAIN. THE LENS WAS EXCHANGED FOR A DIFFERENT LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285658&pc=MTA | 2023826-2020-01486 | 2020-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Pain; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 LENS, -17.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED, THE SURGEON REPORTING LENS ROTATION, LENS DISLOCATION/SUBLUXATION, CILIARY PAIN, LENS DECENTRATION, TYNDALL 2+, AND DOUBLE IMAGE. TOPICAL STERIODS WERE ALSO ADMINISTERED. REPOSITIONING DID NOT RESOLVE THE PROBLEM. REPORTEDLY, ON (B)(6) 2020 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER 1 BROKEN HAPTIC WAS DISCOVERED UPON EXPLANT WHICH COULD BE A CONTRIBUTING CAUSE. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY, A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE. PAIN PERSISTS. REFERENCE FILE# (B)(4) FOR REPLACEMENT LENS INFORMATION. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. THERE WAS DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9385347&pc=MTA | 2023826-2019-02271 | 2019-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Pain; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PATIENT REPORTED, "ISSUES WITH BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILIARY BODY." LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - PATIENT CODE: 3191 - NO CODE AVAILABLE (CONTRACTION OF THE CILIARY BODY) WAS REPORTED IN MANUFACTURER NARRATIVE, BUT CODE SHOULD HAVE ALSO BEEN SELECTED IN H6. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9385361&pc=MTA | 2023826-2019-02272 | 2019-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Pain; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -08-50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PATIENT REPORTED, "ISSUES WITH BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILIARY BODY." LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (CONTRACTION OF THE CILIARY BODY) WAS REPORTED IN THE MANUFACTURER NARRATIVE, BUT SHOULD HAVE ALSO BEEN SELECTED IN H6 SECTION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7448696&pc=MTA | 2023826-2018-00587 |  | Injury | STAAR SURGICAL COMPANY | 2018-04-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Patient Problem/Medical Problem; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT HE IMPLANTED A 12.6 MM MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO CHRONIC IRITIS. IT WAS REPORTED THAT THE PATIENT WAS TREATED WITH DUREZOL AND PRED FORTE FOR OVER THREE MONTHS. THE REPORTED STATED THAT IN HIS OPINION THE CAUSE OF EVENT WAS DUE TO PATIENT SULCUS ANATOMY. ON (B)(6) 2018 THE SURGEON REPORTED THAT THE PATIENT IS CONTINUING A STEROID TAPER AND THAT THE PATIENT CONTINUES TO RECOVER. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A CARTRIDGE CASE WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14704535&pc=MTA | 2023826-2022-01861 | 2022-05-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pupillary Block | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT HIS PATIENT EXPERIENCED PUPILLARY BLOCK POST ICL IMPLANTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15321217&pc=MTA | 2023826-2022-02873 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pupillary Block | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. AT ONE DAY POST-OP, THE PATIENT RETURNED WITH PUPILLARY BLOCK. WITH SOME DIFFICULTY, PIS WERE PERFORMED (WRONG MEDICATION USED FOR DILATION). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPIRATION DATE UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15331790&pc=MTA | 2023826-2022-02874 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pupillary Block | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. AT ONE DAY POST-OP, THE PATIENT RETURNED WITH PUPILLARY BLOCK. WITH SOME DIFFICULTY, PIS WERE PERFORMED (WRONG MEDICATION USED FOR DILATION). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14885125&pc=MTA | 2023826-2022-02094 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pupillary Block | P030016 | 1 | Event Description: A PLANNED LENS EXPLANT FOR A PATIENT WHO EXPERIENCED ADVERSE EVENT (PUPILLARY BLOCK) POST ICL IMPLANTATION WAS REPORTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5751024&pc=MTA | 2023826-2016-00865 | 2014-06-11 | Injury | STAAR SURGICAL COMPANY | 2016-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Pupillary Block | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AN ICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S EYE. THE REPORTER INDICATED THE PATIENT HAD A PUPIL BLOCK DESPITE A GOOD IRIDOTOMY. IT WAS UNKNOWN IF THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2014. THE REPORTER INDICATED THE PATIENT EXPERIENCED A PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND HAD A UNREACTIVE PUPIL (FIXED). A SECONDARY YAG WAS PERFORMED TO ENLARGE OR ADDITION OF PIS. THE LENS REMAINS IMPLANTED AND AT LAST VISIT DATE ON (B)(6) 2015 THE PUPIL WAS COMING DOWN. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487527&pc=MTA | 2023826-2017-00618 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pupillary Block; Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, AND THE LENS WAS NOTED AS HAVING AN EXCESSIVE VAULT. THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED ON (B)(6) 2017 IN THE PATIENT'S RIGHT EYE (OD), -11.0 DIOPTER. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO HYPEROPIC SURPRISE, EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND ANGLE CLOSURE. A YAG PERIPHERAL IRIDOTOMY WAS PERFORMED TO RELIEVE PUPIL BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS POST-OP BSCVA WAS 20/25. (B)(4). CORRECTED DATA: THE IMPLANTED DATE IS (B)(6) 2017 NOT (B)(6) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL, IN A BIOHAZARD BAG. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: CONCLUSION CODE: (VISIAN® IMPLANTABLE COLLAMER® LENSES ARE INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. THIS PATIENT WAS (B)(6) YEARS OF AGE). (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6629958&pc=MTA | 2023826-2017-00949 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Pupillary Block; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON COMMENTED THAT THE PATIENT ALSO EXPERIENCED PUPIL BLOCK, AND SUSPICION OF PIGMENT DISPERSION. THE DR. PLANS TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. MULTIPLE ATTEMPTS TO CONTACT TO CUSTOMER HAVE BEEN UNSUSCCESSFUL. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403689&pc=MTA | 2023826-2018-00510 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pupillary Block; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED. THE SURGEON REPORTS THE FOLLOWING: EXCESSIVE VAULT WITH "NEAR CORNEAL TOUCH IN PERIPHERY," "INTRAOPERATIVE PUPILLARY BLOCK WITH EXTRUSION OF IRIS THROUGH WOUNDS," LENS REPOSITIONING AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. REPORTEDLY, THE PROBLEM WAS RESOLVED WITH EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL PLASTIC BAG AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND FOUND THE PRESENCE OF FIBERS AND BLACK SPOTS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, LENS REPOSITIONING, NEAR CORNEAL TOUCH IN PERIPHERY, EXTRUSION OF IRIS THROUGH WOUNDS, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16765383&pc=MTA | 2023826-2023-01438 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Retinal Detachment | P030016 | 1 | Event Description: ON 20-MAR-2023 WE RECEIVED NOTIFICATION OF A JOURNAL ARTICLE TITLED, 'TEN-YEAR PREVALENCE OF RRD AFTER ICL IMPLANT IN MYOPIC EYE'. THE PAPER REPORTS 7 (OUT OF 409) EYES WITH RETINAL DETACHMENT POST-ICL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924973&pc=MTA | 2023826-2023-01824 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Retinal Detachment | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). POST-OP, THE PATIENT HAD A RETINAL DETACHMENT AND A VITREORETINAL PROCEDURE WAS PERFORMED. THE LENS WAS REMOVED. SILICONE OIL WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - " THE LENS WAS REMOVED" SHOULD BE CORRECTED TO "THE LENS REMAIN IMPLANTED" IN THE INITIAL MDR. H6 - DEVICE CODE 4627 IS NOT APPLICABLE. CLAIM #: (B)(4). Manufacturer Narrative: A2: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - VITREORETINAL PROCEDURE (VITRECTOMY), SILICONE OIL. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299039&pc=MTA | 2023826-2021-00350 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN FRONTIERS IN MEDICINE WAS RECEIVED ON 19-JAN-2021, ENTITLED 'LONG-TERM OUTCOMES OF RETINAL DETACHMENT IN PHAKIC EYES AFTER IMPLANTATION OF IMPLANTABLE COLLAMER LENS V4C FOR HIGH MYOPIA CORRECTION.' THE ARTICLE STATES THAT 14 OUT OF 704 EYES EXPERIENCED RETINAL DETACHMENT. 11 EYES HAD CAUSATIVE BREAKS LOCATED AT PERIPHERAL RETINA AND 3 EYES HAD A CAUSATIVE BREAK LOCATED AT THE POSTERIOR POLE. SCLERAL BUCKLING (SB) OR EXTRASCLERAL COMPRESSION (EC) WAS PERFORMED IN 10 OF THE EYES. PARS PLANA VITRECTOMY ALONE WAS PERFORMED IN 4 EYES. ONE EYE UNDERWENT ADDITIONAL VITREORETINAL SURGERY AFTER THE INITIAL SB PROCEDURE BECAUSE OF PROLIFERATIVE VITREORETINOPATHY. THE SURGEON CONCLUDES V4C ICL IMPLANTATION IS A SAFE METHOD FOR HIGH MYOPIA CORRECTION. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15654635&pc=MTA | 2023826-2022-03767 | 2022-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -18.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS REPORTED THAT THE PATIENT HAD RETINAL DETACHMENT AND HIS VISION WAS LP. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6527426&pc=MTA | 2023826-2017-00698 | 2015-03-14 | Injury | STAAR SURGICAL COMPANY | 2017-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Retinal Detachment; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -21.50/+4.5/100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED RETINAL DETACHMENT, AND LOST OF BCVA. THE PATIENT WAS TREATED WITH "EXTERNAL CERCLAGE" AND RETINAL REPAIR WAS PERFORMED ON (B)(6) 2015. THE CUSTOMER STATED HE THOUGH THE RETINAL DETACHMENT IS INDEPENDENT OF THE LENS. ON (B)(6) 2015 THE LENS WAS EXPLANTED. AT PRESENT, THE PATIENT REFUSES TO IMPLANT ANOTHER LENS. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535077&pc=MTA | 2023826-2018-00789 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Retinal Detachment; Blurred Vision; No Code Available | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS THAT AT 25 MONTHS AFTER INITIAL SURGERY, RETINAL DETACHMENT OCCURRED, LEADING TO A LOSS OF 7 LINES IN BCVA COMPARED TO PREOPERATIVE VALUE. REPORTEDLY, THE PATIENT WAS GIVEN A FUNDUS OPERATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT INFORMATION: UNK. PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9031445&pc=MTA | 2023826-2019-01696 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Retinal Detachment; No Code Available | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF CATARACT & REFRACTIVE SURGERY IN JUNE 2019 TITLED, '10-YEAR CLINICAL OUTCOMES AFTER IMPLANTATION OF A POSTERIOR CHAMBER PHAKIC LENS FOR MYOPIA.' THE ARTICLE STATES THAT PATIENT UNDERWENT SCLERAL BUCKLING WITH CRYOTHERAPY. RESOLUTION OF RETINAL DETACHMENT WAS OBSERVED AND THE CDVA RECOVERED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INITIAL MDR SHOULD HAVE ALSO STATED: 'THERE WAS ONE REPORTED CASE OF RHEMATOGENOUS RETINAL DETACHMENT.' DATE RECEIVED BY MANUFACTURER LITERATURE SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION - SCLERAL BUCKLING, CRYOTHERAPY) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275449&pc=MTA | 2023826-2018-00226 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Retinal Detachment; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A V4 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED AFTER BEING IMPLANTED IN THE PATIENTS EYE 7 YEARS AGO, DUE TO A RHEGMATOGENOUS RETINAL DETACHMENT. PHACOEMULSIFICATION SURGERY WAS PERFORMED AND AN IOL (INTRAOCULAR LENS) WAS IMPLANTED. THE REPORTER ATTRIBUTED THE RETINAL DETACHMENT TO HIGH MYOPIA. Manufacturer Narrative: THE LENS WAS EXPLANTED AFTER BEING IMPLANTED IN THE PATIENT'S EYE FOR 7 YEARS, DUE TO RHEGMATOGENOUS RETINAL DETACHMENT. (B)(4). Manufacturer Narrative: THE SURGEON ALSO PERFORMED A VITREORETINAL SURGERY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED, PHACOEMULSIFICATION SURGERY AND IOL IMPLANTED (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7277477&pc=MTA | 2023826-2018-00227 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Retinal Detachment; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A V3 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED AFTER 10 YEARS OF BEING IMPLANTED IN THE PATIENTS EYE, DUE TO A RHEGMATOGENOUS RETINAL DETACHMENT. PHACOEMULSIFICATION SURGERY WITH IOL (INTRAOCULAR LENS) IMPLANTATION WAS PERFORMED. THE REPORTER ATTRIBUTED THE RETINAL DETACHMENT TO HIGH MYOPIA. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED, PHACOEMULSIFICATION SURGERY WITH IOL (INTRAOCULAR LENS) IMPLANTATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536949&pc=MTA | 2023826-2016-00301 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Retinal Detachment; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HER RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT REPORTED HAD A SECOND SURGICAL PROCEDURE BECAUSE OF A RETINAL DETACHMENT AND THE LENS WAS EITHER REMOVED, REPLACED OR REPOSITIONED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE FACILITY WAS UNABLE TO CONFIRM THE PATIENT'S REPORT. THE DATE OF THE LAST VISIT WAS (B)(6) 2013 AND THEY WERE UNAWARE OF ANY ADVERSE EVENT WITH THE LENS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5449342&pc=MTA | 2023826-2016-00156 |  | Injury | STAAR SURGICAL COMPANY | 2016-02-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Retinal Tear; Vitreous Detachment | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HIS RIGHT EYE (OD). THE PATIENT REPORTED RETINAL TEAR/VITREOUS DETACHMENT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PER MEDICAL REVIEW: THE PATIENT COMPLETED 48-MONTH FOLLOW-UP FORM FOR ICL ON (B)(6) 2016 AND REPORTED THAT HE HAD EXPERIENCED "RETINAL TEAR/VITREOUS DETACHMENT". NO ADDITIONAL INFORMATION IS OBTAINED. IT SHOULD BE NOTED THAT THERE HAS BEEN NO DEFINITE PROOF OF A CAUSAL LINK BETWEEN ICL PROCEDURES AND VITREORETINAL COMPLICATIONS. THE RARITY OF VITREORETINAL EVENTS AFTER ICL IMPLANTATION MAKES IT DIFFICULT TO EXAMINE THE EXACT INCIDENCE AND RISK FACTORS OF VITREORETINAL COMPLICATIONS. HOWEVER, IT IS WELL KNOWN THAT MYOPIA IS PREDISPOSING FACTOR FOR VITREORETINAL COMPLICATIONS REGARDLESS OF ANY SURGICAL PROCEDURE. THE REASSESSMENT WILL BE PERFORMED WHEN (IF) ADDITIONAL INFORMATION IS RECEIVED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE FACILITY REPORTED THEY WERE UNABLE TO AGREE WITH THE PATIENT REPORT. NO MEDICATIONS OR TREATMENT WERE REQUIRED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215077&pc=MTA | 2023826-2022-01277 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER INDICATED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED AFTER LENS IMPLANTATION. DIAGNOSTICS WERE NOT PERFORMED. ORAL PREDNISOLONE AND BETAMETHASONE SODIUM PHOSPHATE EYE DROPS WERE PRESCRIBED. THE PROBLEM WAS RESOLVED AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (ALLERGIES TO LENSES). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN, EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15521040&pc=MTA | 2023826-2022-03398 | 2022-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -8.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED TASS - TOXIC ANTERIOR SEGMENT SYNDROME. MEDICATION OF RINDERON WAS PRESCRIBED TO THE PATIENT. LENS REMAINS IMPLANTED. PROBLEM IS NOT RESOLVED. ADDITIONALLY IT WAS NOTED ACCUJECT DEVICE WAS USED TO IMPLANT THE LENS AND "I WOULD LIKE TO KNOW IF IT IS RELATED TO ACCUJECT; LIKE TASS WHICH IS HAPPENING IN OTHER HOSPITALS". Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17051932&pc=MTA | 2023826-2023-02079 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED BILATERAL ICL IMPLANTABLE COLLAMER LENSES. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED IN THE LEFT EYE (OS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: REPORT # MW5117323. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15170613&pc=MTA | 2023826-2022-02617 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) |  | 1 | Event Description: THE REPORTER INDICATED THAT AN UNKNOWN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE (OD) ON AN UNKNOWN DATE. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. LINDERONE EYE DROPS/INJECTION WAS PRESCRIBED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE SURGERY WAS PERFORMED AROUND THE BEGINNING OF MAY.THE PATIENT IN QUESTION HAD ALWAYS HAD TEARY EYES. LINDELON EYE DROPS AND INJECTIONS WERE PRESCRIBED. SINCE THE PATIENT HAS NOT VISITED THE HOSPITAL SINCE THEN, WE ASSUME THAT HIS SYMPTOMS HAVE RESOLVED. EXPLAINED THAT SINCE THE PATIENT HAD ALWAYS HAD TEARY EYE SYMPTOMS, I WOULD ASK HIM TO USE MORE EYE DROPS THAN USUAL. HOWEVER, THE PATIENT DID NOT TAKE THE DROPS AS INSTRUCTED, WHICH WAS PROBABLY THE CASUE OF THE TASS. BASED ON THE ABOVE, IN THE PHYSICIANS OPINION, HE DOES NOT BELIEVE THAT THIS CASE WAS CAUSED BY THE LENS.Event Description: THE REPORTER INDICATED THAT AN UNKNOWN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE (OD) ON AN UNKNOWN DATE. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. LINDERONE EYE DROPS/INJECTION WAS PRESCRIBED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE SURGERY WAS PERFORMED AROUND THE BEGINNING OF MAY. THE PATIENT IN QUESTION HAD ALWAYS HAD TEARY EYES. LINDELON EYE DROPS AND INJECTIONS WERE PRESCRIBED. SINCE THE PATIENT HAS NOT VISITED THE HOSPITAL SINCE THEN, WE ASSUME THAT HIS SYMPTOMS HAVE RESOLVED. EXPLAINED THAT SINCE THE PATIENT HAD ALWAYS HAD TEARY EYE SYMPTOMS, I WOULD ASK HIM TO USE MORE EYE DROPS THAN USUAL. HOWEVER, THE PATIENT DID NOT TAKE THE DROPS AS INSTRUCTED, WHICH WAS PROBABLY THE CAUSE OF THE TASS. BASED ON THE ABOVE, IN THE PHYSICIANS OPINION, HE DOES NOT BELIEVE THAT THIS CASE WAS CAUSED BY THE LENS. Manufacturer Narrative: A4-A6:UNK. D2 IS UNKNOWN BUT FIELD CANNOT BE LEFT BLANK. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. D2 IS UNKNOWN BUT FIELD CANNOT BE LEFT BLANK. SUPP1: DUPLICATE ERROR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15218205&pc=MTA | 2023826-2022-02750 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) |  | 1 | Event Description: THE REPORTER INDICATED THAT AN UNKNOWN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "I DON'T THINK IT IS DERIVED FROM ICL PRODUCTS.WITH AN IOL LENS (CATARACT), THE INTRAOCULAR CAN BE THOROUGHLY CLEANED BECAUSE THERE IS NO LENS, BUT AN ICL LENS CANNOT BE THOROUGHLY CLEANED BECAUSE OF THE LENS. SINCE ALL OF THE VISCOELASTIC MATERIAL CANNOT BE REMOVED FROM WITHIN THE EYE, WE SUSPECT THAT THE REMAINING VISCOELASTIC MATERIAL IS THE CAUSE OF TASS. Manufacturer Narrative: CORRECTED DATA: H4: DEVICE MANUFACTURE DATE: 17-MAR-2021 SHOULD BE OMITTED FOR CORRECTION. Manufacturer Narrative: COMMON DEVICE NAME IS UNKNOWN BUT FIELD CANNOT BE LEFT BLANK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15556900&pc=MTA | 2023826-2022-03445 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A SURGEON OBSERVED A PATIENT DEVELOP AN INFECTION WHICH MAY BE POSSIBLE TASS. SECONDARY INTERVENTION WAS PERFORMED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CLAIM#(B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15557127&pc=MTA | 2023826-2022-03444 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A SURGEON OBSERVED A PATIENT DEVELOP AN INFECTION WHICH MAY BE POSSIBLE TASS. SECONDARY INTERVENTION WAS PERFORMED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15562518&pc=MTA | 2023826-2022-03520 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. TREATMENT PERFORMED OF UNKNOWN KIND AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: INVESTIGATION TYPE : LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15562543&pc=MTA | 2023826-2022-03521 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. TREATMENT PERFORMED OF UNKNOWN KIND AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17786414&pc=MTA | 2023826-2023-04091 | 2023-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. DIAGNOSTICS WAS PERFORMED: IOP, VISUAL ACUITY, AND CASIA. RINDERON WAS USED FOR TREATMENT. THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS STATED, "ANTERIOR CHAMBER ABSCESS 3 WEEKS AFTER SURGERY (1 AUG) RECOVERED (FELT A LITTLE KP REMAING ON THE LENS) ON 21 AUG." CAUSE OF THE EVENT WAS THE DEVICE WITH THE LENS NOT PERFORMING AS INTENDED. REPORTEDLY, "FOUND FOREIGN OBJECTS ON THE LENS WHEN INSERTING" Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H10 - PRIOR SUPPLEMENTAL MDR #1 SHOULD BE N/A AS IT WAS SUBMITTED IN ERROR - WAS INTENDED FOR ANOTHER COMPLAINT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND WITH FIBRE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18421049&pc=MTA | 2023826-2023-05808 | 2023-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -4.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 16-SEP-2023, THE LENS WAS REMOVED DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS). THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18421051&pc=MTA | 2023826-2023-05807 | 2023-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -5.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON 16-SEP-2023, THE LENS WAS REMOVED DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS). THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322213&pc=MTA | 2023826-2023-02770 | 2023-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. DIAGNOSTICS WERE PERFORMED. ANTERIOR CHAMBER LAVAGE SUBCONJUNCTIVAL STEROID INJECTION, ORAL MEDICINE, AND OPHTHALMIC DROPS. DECADRON SUBCONJUNCTIVAL INJECTION, PREDNISONE 10MG, EVERY HOUR (VEGAMOX, TOBRACIN, RINDERON. THE PROBLEM WAS RESOLVED. REPORTEDLY, "TASS IS OCCURING ON A REGULAR BASIS." CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.00 INTO THE LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED TOXIC ANTERIOR SEGMENT SYNDROME (TASS). TREATMENT AND DIAGNOSTICS WERE PERFORMED. ANTERIOR CHAMBER LAVAGE WAS PERFORMED AND SUBCONJUNCTIVAL STEROID INJECTION, ORAL STEROIDS AND TOPICAL DROPS (ANTIBIOTICS AND STEROIDS) WERE PRESCRIBED. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "TASS IS OCCURRING ON A REGULAR BASIS." H6-HEALTH EFFECT- IMPACT CODE: "4625" AND "4614" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "1401" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190876&pc=MTA | 2023826-2022-01294 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6, -16.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS TASS-DIAGNOSTICS WERE NOT PERFORMED AND MEDICATION WAS ADMINISTERED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR: THE DEVICE DID NOT FAIL TO PERFORM AS EXPECTED. THE SURGEON ALSO REPORTS, "ALLERGIES TO LENSES." Manufacturer Narrative: B5 - CAUSE WAS REPORTED AS PATIENT-RELATED FACTOR, NOT USER ERROR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN, RACE:UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16887509&pc=MTA | MW5117323 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Toxic Anterior Segment Syndrome (TASS) |  | 1 | Event Description: BILATERAL ICL (IMPLANTABLE COLLAMER LENS) IMPLANTATION WITH TOXIC ANTERIOR SEGMENT SYNDROME OF LEFT EYE ONLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14222057&pc=MTA | 2023826-2022-01278 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS, -03.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE REPORTER INDICATED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED AFTER LENS IMPLANTATION. DIAGNOSTICS WERE NOT PERFORMED. INTRAOCULAR RINSING AND REMOVAL OF FIBRIN WAS PERFORMED. THE PROBLEM WAS RESOLVED AND THE LENS REMAINED IMPLANTED. THERE WAS IRIS CONTRACTION. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (ALLERGIES TO LENSES). Manufacturer Narrative: UNK. NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH IMPACT- CLINICAL CODE:4581- IRIS CONTRACTION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16535465&pc=MTA | 2023826-2023-00880 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: ON (B)(6) 2023 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN BMC OPHTHALMOLOGY TITLED, 'LATE ONSET TOXIC ANTERIOR SEGMENT SYNDROME AFTER ICL IMPLANTATION: TWO CASE REPORTS'. THE ARTICLE REPORTS TWO CASES OF TASS POST ICL (IMPLANTABLE COLLAMER LENS) IMPLANTATION THAT RESOLVED WITH MEDICAL MANAGEMENT. AT 1-WEEK POSTOPERATIVE BOTH CASES PRESENTED FIBRIN FORMATION IN THE ANTERIOR CHAMBER SUCH AS KERATIC PRECIPITATES AND WHITE PUS ON THE SURFACE OF THE ICL. FUNDUS EXAMINATION WAS NORMAL. AFTER 4 TO 5 WEEKS OF TOPICAL AND ORAL STEROID TREATMENT, VISUAL ACUITY AND FIBRIN FORMATION IN THE ANTERIOR CHAMBER IMPROVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED INFORMATION: H6- HEALTH EFFECT- CLINICAL CODE: "4581- MONOCULAR VISION LOSS" SHOULD BE REMOVED AND "4581- ICL SURFACE PIGMENTS" SHOULD BE ADDED. B5- AN ARTICLE WAS PUBLISHED IN BMC OPHTHALMOLOGY ENTITLED, 'LATE- ONSET TOXIC ANTERIOR SEGMENT SYNDROME AFTER ICL IMPLANTATION: TWO CASE REPORTS'. THE ARTICLE REPORTS ADVERSE EVENTS FOLLOWING ICL IMPLANTATION. THE PATIENTS EXPERIENCED TASS, ICL SURFACE PIGMENTS AND MEDICATION WAS PRESCRIBED. THE FIRST CASE REPORTS "IN HER RIGHT EYE, KERATIC PRECIPITATES (KP) WERE OBSERVED ON THE POSTERIOR SURFACE OF THE CORNEA, WHILE CELLS AND FLARES WERE OBSERVED IN THE ANTERIOR CHAMBER". ADDITIONALLY, "NUMEROUS PIGMENTS AND WHITE PUS WERE OBSERVED ON THE SURFACE OF THE ICL. AT ONE WEEK OF TREATMENT, THE ANTERIOR CHAMBER WAS CLEAR AND THE WHITE PUS ON THE SURFACE OF ICL WAS SIGNIFICANTLY DECREASED". THE LENS REMAINS IMPLANTED. THE SECOND CASE REPORTS "ONE WEEK AFTER SURGERY, THE UDVA OF THE LEFT EYE WAS 20/25" AND "THERE WAS A SMALL KP ON THE POSTERIOR SURFACE OF THE CORNEA AND AMOUNTS OF CELLS AND FLARES IN THE ANTERIOR CHAMBER. A LARGE NUMBER OF PIGMENTS AND WHITE PUS WERE OBSERVED ON THE SURFACE OF THE ICL LENS. AFTER TREATMENT, THE NUMBER OF CELLS IN THE ANTERIOR CHAMBER DECREASED WITHIN THE NEXT 5 DAYS. AT 1 WEEK OF TREATMENT, THE ANTERIOR CHAMBER WAS CLEAR, AND THE WHITE PUS ON THE SURFACE OF THE LENS WAS ALSO SIGNIFICANTLY DECREASED". THE LENS REMAINS IMPLANTED. THE SERIAL NUMBERS WERE NOT PROVIDED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16032529&pc=MTA | 2023826-2022-04638 | 2020-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 -10.00 DIOPTER IMPLATABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO UVEITIS( PRED FORTE. PRESCRIBED). IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16032568&pc=MTA | 2023826-2022-04639 | 2020-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 -9.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO UVEITIS( PRED FORTE. PRESCRIBED). IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17877870&pc=MTA | 2023826-2023-04418 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis | P030016 | 1 | Event Description: THE SURGEON REPORTS A CASE OF UVEITIS THREE WEEKS AFTER SURGERY. PER UPDATED INFORMATION THE PATIENT IS DOING WELL AND THE TREATMENT ( EYE DROPS) HAS SOLVED THE ISSUE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17877885&pc=MTA | 2023826-2023-04419 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis | P030016 | 1 | Event Description: THE SURGEON REPORTS A CASE OF UVEITIS THREE WEEKS AFTER SURGERY. PER UPDATED INFORMATION THE PATIENT IS DOING WELL AND THE TREATMENT ( EYE DROPS) HAS SOLVED THE ISSUE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6314691&pc=MTA | 2023826-2017-00142 |  | Injury | Unknown Manufacturer | 2017-02-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Uveitis | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE PATIENT DEVELOPED UVEITIS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SERIAL NUMBER REPORTED. THIS PRODUCT WAS NOT MARKETED IN THE U.S. CLAIM #(B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6325653&pc=MTA | 2023826-2017-00206 |  | Injury | Unknown Manufacturer | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Uveitis | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE PATIENT DEVELOPED UVEITIS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816245&pc=MTA | 2023826-2016-00976 | 2016-02-21 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Vertigo; Blurred Vision; Loss of Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOSS OF VISION AND VERTIGO. THE LENS WAS EXCHANGED WITH ANOTHER SAME MODEL LENS WITH A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. Manufacturer Narrative: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ACD BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS DRIED, DISCOLORED MATERIAL. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO UNEXPECTED REFRACTIVE OUTCOME AND VERTIGO. THE LENS WAS EXCHANGED WITH ANOTHER LENS OF THE SAME MODEL BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/30. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073304&pc=MTA | 2023826-2020-03183 | 2000-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -17.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO " PATIENTS VISION WAS YELLOW". IT WAS REPORTED THAT THE PROBLEM RESOLVED, PATIENTS VISION RETURNED TO NORMAL AFTER EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12644941&pc=MTA | 2023826-2021-04020 | 2021-07-07 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM6_13.2,-6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT HAS PROBLEMS WITH NEAR VISION, LENS REMAINS IMPLANTED. THE REPORTER STATED "THE PATIENT DOESN'T WANT ANY FURTHER TREATMENT FOR THE TIME BEING." FOR PATIENTS CURRENT CONDITION " PROBLEMS WITH READING" WAS REPORTED. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7835543&pc=MTA | 2023826-2018-01307 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | P030016 | 1 | Event Description: RECEIVED A REPORT WHERE THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS AND THE PATIENT COMPLAINED OF BAD NIGHT TIME GLARE. THE PATIENT WAS BEING TREATED WITH ALPHAGAN, WHICH WAS HELPING THE GLARE BUT WAS STILL VERY BOTHERSOME. THERE WERE NO ISSUES WITH THE SURGERY ITSELF AND LENS PLACEMENT LOOKED PERFECT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182716&pc=MTA | 2023826-2022-00027 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS THE LENS IS STILL NOT CENTERED AND THE PATIENT IS UNCOMFORTABLE WITH THE GLARE/HALOS. SEE MFR REP# 2023826-2021-04895 FOR INITIAL LENS. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923026&pc=MTA | 2023826-2015-00829 | 2015-06-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND WILL BE EXCHANGED FOR ANOTHER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. METHOD - WORK ORDER SEARCH RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE MANUFACTURE DATE: 03/19/2015. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VILE. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. CONCLUSION: PATIENT IS UNDER THE AGE OF 21 AND HAS AN ACD OF 2.9. AS PER DFU FOR THIS MODEL FOR THE PATIENTS UNDER THE AGE OF 21 YEARS AND IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923027&pc=MTA | 2023826-2015-00828 | 2015-06-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND WILL BE EXCHANGED FOR ANOTHER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO -(OTHER): LENS REMAINS IMPLANTED. EVALUATION CODES: METHOD (OTHER): WORK ORDER SEARCH RESULTS -(OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/1/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. EXPLANT DATE: - UNK. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CONCLUSION CODE: THE PATIENT IS UNDER 21 YEARS OLD AND HAS AN ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD AND ACD BELOW 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6158574&pc=MTA | 2023826-2016-01737 | 2016-11-11 | Injury | STAAR SURGICAL COMPANY | 2016-12-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT IS EXPERIENCING ISSUE WITH VISION WHEN IT IS DARK. THE LENS REMAINS IMPLANTED AT THE TIME OF THIS REPORT. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6158460&pc=MTA | 2023826-2016-01736 | 2016-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2016-12-09 05:00 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.00/40.5/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT IS EXPERIENCING ISSUE WITH VISION WHEN IT IS DARK. THE LENS REMAINS IMPLANTED AT THE TIME OF THIS REPORT. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS -7.00/0.5/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2015. THE PATIENT IS EXPERIENCING ISSUE WITH VISION WHEN IT IS DARK. ON (B)(4) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, DIFFERENT LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6623193&pc=MTA | 2023826-2017-00938 | 2017-04-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances | P030016 | 1 | Event Description: RECEIVED A VOLUNTARY EVENT REPORT FROM THE FDA- # MW5069251. THE PATIENT REPORTED TO HAVE HAD VISIAN ICL SURGERY IN THE LEFT EYE ON (B)(6) 2016. THE PATIENT REPORTED TO HAVE FLASHES AND BAD SHADOW ON THE VISION FIELD. RETINA EXAMS WERE NORMAL. DOCTOR COULD NOT EXPLAIN CAUSE OF EVENT. RX MEDS: OFLOXACIN EYE DROPS, PREDNISOLONE EYE DROPS, KETOROLAC EYE DROPS. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6623706&pc=MTA | 2023826-2017-00939 | 2017-04-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances | P030016 | 1 | Event Description: RECEIVED A VOLUNTARY EVENT REPORT FROM THE FDA- # MW5069251. THE PATIENT REPORTED TO HAVE HAD VISIAN ICL SURGERY ON THE RIGHT EYE ON (B)(6) 2017. THE PATIENT REPORTED TO HAVE FLASHES AND BAD SHADOW ON THE VISION FIELD POST -OP #3. RETINA EXAMS WERE NORMAL. DOCTOR COULD NOT EXPLAIN CAUSE OF EVENT. RX MEDS: OFLOXACIN EYE DROPS, PREDNISOLONE EYE DROPS, KETOROLAC EYE DROPS. Manufacturer Narrative: THIS COMPLAINT FALLS UNDER THE CATEGORY OF "PRODUCT PROBLEM" AND NOT " ADVERSE EVENT" AS DETERMINED FROM THE MEDICAL REVIEW PERFORMED.(B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6451166&pc=MTA | 2023826-2017-00537 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-31 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A V4B IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND PATIENT EXPERIENCES " 'REFLECTIONS' FROM THE CENTRAL FLOW HOLE." LENS PLANS TO BE EXCHANGED FOR ANOTHER. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME, MULTIPLE ATTEMPTS MADE TO CONTACT CUSTOMER. Manufacturer Narrative: MANUFACTURED IN U.S. BUT NOT DISTRIBUTED IN U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7329352&pc=MTA | 2023826-2018-00356 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-09 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.0 DIOPTER ON (B)(6) 2017 IN THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTED THAT "THE PATIENT IS SUFFERING FROM SEVERE GLARE AND HE BELIEVES THAT IT IS FROM THE CENTRAL PORTS OF THE LENSES IN HER EYES". THE REPORTER STATED THAT THE PATIENT STILL HAS "GLARES/ DECENTRATION". THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE REPORTER BUT NONE HAVE BEEN FORTHCOMING. A SUPPLEMENTAL MDR WILL BE SUBMITTED IF NEW INFORMATION IS RECEIVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7495936&pc=MTA | 2023826-2018-00711 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-08 | MTA | UNK | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE PATIENT CANNOT "GET USED TO THE GLARES." REPORTEDLY, AN EXPLANT IS PLANNED, HOWEVER ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535056&pc=MTA | 2023826-2018-00782 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS GLARE AFTER IMPLANTATION; RESOLVED AFTER 12 MONTHS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535057&pc=MTA | 2023826-2018-00783 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS GLARE AFTER IMPLANTATION; RESOLVED AFTER 12 MONTHS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535058&pc=MTA | 2023826-2018-00778 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS GLARE AFTER IMPLANTATION; RESOLVED AFTER 12 MONTHS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535061&pc=MTA | 2023826-2018-00780 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS GLARE AFTER IMPLANTATION; RESOLVED AFTER 12 MONTHS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535062&pc=MTA | 2023826-2018-00781 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS GLARE AFTER IMPLANTATION; RESOLVED AFTER 12 MONTHS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535064&pc=MTA | 2023826-2018-00776 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: PLOS ONE ARTICLE WAS RECEIVED ON 27/04/2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS GLARE AFTER IMPLANTATION; RESOLVED AFTER 12 MONTHS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535090&pc=MTA | 2023826-2018-00785 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS GLARE AFTER IMPLANTATION; RESOLVED AFTER 12 MONTHS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY) PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535109&pc=MTA | 2023826-2018-00779 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Visual Disturbances | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS GLARE AFTER IMPLANTATION; RESOLVED AFTER 12 MONTHS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571619&pc=MTA | 2023826-2018-00856 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Visual Disturbances | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN BMC OPHTHALMOLOGY TITLED, 'ANALYSIS OF INTRAOCULAR POSITIONS OF POSTERIOR IMPLANTABLE COLLAMER LENS BY FULL-SCALE ULTRASOUND BIOMICROSCOPY.' THE ARTICLE REFERENCES 1 EYE WITH IDEAL VAULT, BUT WITH ICL DECENTRALIZATION AND SERIOUS GLARE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7304830&pc=MTA | 2023826-2018-00300 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE PATIENT COMPLAINED OF GLARE AND FOUND IT HARD TO DRIVE. THE SURGEON PLANS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023228&pc=MTA | 2023826-2021-04895 | 2021-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT, LENS ROTATION, AND GLARE/HALOS. SEE MFR FILE# (B)(4) FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "LENGTH OF THE ICL." Manufacturer Narrative: B5- CORRECTED TO : "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULTS, LENS ROTATION, AND GLARE/HALOS. SEE MFR FILE# (B)(4) FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "LENGTH OF THE ICL." CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16044661&pc=MTA | 2023826-2022-04707 | 2022-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS WITH A DIOPTER OF -6.5 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND GLARE/HALOES. CAUSE OF THE EVENT IS REPORTED AS THE DEVICE. REPORTEDLY, "LENS TOO BIG SO HIGH VAULT BUT ALSO GLARE AT NIGHT SO NOT GOING TO EXCHANGE. PATIENT HAD DRY EYE BUT THIS HAS CLEARED SO HAVE DECIDED TO LASER INSTEAD." STATUS OF THE EYE IS SAID, "LENS HAS BEEN REMOVED AND WILL LASER INSTEAD." IT WAS ALSO STATED, "PATIENT HAD DRY EYE WHEN FIRST SEEN SO DECIDED UPON ICL, HOWEVER THE PATIENT HAS A HIGH VAULT AND THE LENSES NEED REMOVING. SINCE THE PATIENT NO LONGER HAS DRY EYE AND IS GETTING GLARE AT NIGHT WHERE SHE CANNOT GO OUT AFTER 4PM, HAVE DECIDED TO LASER RATHER THAN EXCHANGE." PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC BROKEN. RECEIVED TWO LENSES IN ONE BAG AND ONE VIAL,DUE TO SIGNIFICANT LENS DAMAGE, UNABLE TO PERFORM DIMENSIONAL INSPECTION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16044665&pc=MTA | 2023826-2022-04705 | 2022-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS WITH A DIOPTER OF -7.0 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND GLARE/HALOES. CAUSE OF THE EVENT IS REPORTED AS THE DEVICE. REPORTEDLY, "LENS TOO BIG SO HIGH VAULT BUT ALSO GLARE AT NIGHT SO NOT GOING TO EXCHANGE. PATIENT HAD DRY EYE BUT THIS HAS CLEARED SO HAVE DECIDED TO LASER INSTEAD." STATUS OF THE EYE IS SAID, "LENS HAS BEEN REMOVED AND WILL LASER INSTEAD." IT WAS ALSO STATED, "PATIENT HAD DRY EYE WHEN FIRST SEEN SO DECIDED UPON ICL, HOWEVER THE PATIENT HAS A HIGH VAULT AND THE LENSES NEED REMOVING. SINCE THE PATIENT NO LONGER HAS DRY EYE AND IS GETTING GLARE AT NIGHT WHERE SHE CANNOT GO OUT AFTER 4PM, HAVE DECIDED TO LASER RATHER THAN EXCHANGE." PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. RECEIVED TWO LENSES IN ONE BAG AND ONE VIAL, THUS UNABLE TO IDENTIFY WHICH LENS PERTAINS TO WHICH CLAIM. CLAIM# (B)(4). Manufacturer Narrative: PT INFO:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16742913&pc=MTA | 2023826-2023-01257 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) -2023 DUE TO LOW VAULTING AND THE PATIENT EXPERIENCED DOUBLE VISION AND GLARE AT NIGHT. THE PATIENT WAS PRESCRIBED MEDICATION BUT THE SYMPTOMS DID NOT IMPROVE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC SPLIT. THE HAPTIC WAS BROKEN AND BENT. UNABLE TO EVALUATE THE LENS DUE TO LENS WAS RETURNED IN PIECES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595988&pc=MTA | 2023826-2017-00841 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HER VISION HAD WORSENED. THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE DATE OF THE LAST VISIT WAS (B)(6) 2014. THE REPORTER STATED THE CAUSE OF THE EVENT WAS RESIDUAL ASTIGMATISM, THE PRE-OP ASTIGMATISM COULD NOT BE FULLY CORRECTED WITH THE ICL. THE VA AT PATIENT'S 6 MONTH VISIT WAS 20/20 AND AT 12 MONTH VISIT IT WAS 20/25. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL EVALUATION OF COMPLAINT EVENTS FOUND NO SERIOUS INJURY TO PATIENT REQUIRING SURGICAL OR MEDICAL INTERVENTION. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE DATE OF THE LAST VISIT WAS (B)(6) 2014. THE REPORTER STATED THE CAUSE OF THE EVENT WAS RESIDUAL ASTIGMATISM, THE PRE-OP ASTIGMATISM COULD NOT BE FULLY CORRECTED WITH THE ICL. THE VA AT PATIENT'S 6 MONTH VISIT WAS 20/20 AND AT 12 MONTH VISIT IT WAS 20/25. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429111&pc=MTA | 2023826-2020-01845 | 2020-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) OF 2015. THE SURGEON REPORTS DOUBLE VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: DATE IS APPROXIMATE. PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11070813&pc=MTA | 2023826-2020-03179 | 2020-08-15 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1,-7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER EXPLANTED, THE REPORTER STATED "THE PATIENT WAS DISSATISFIED WITH THE QUALITY OF VISION AND REQUESTED TO BE REMOVED THE LENS." IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17599756&pc=MTA | 2023826-2023-03628 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF A 5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND LATER REPLACED FOR A SAME SIZE LENS DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: DIOPTER SHOULD BE CORRECTED TO -5.0. H6: 4627: EXPLANT SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980359&pc=MTA | 2023826-2020-02957 |  | Injury | STAAR SUGICAL COMPANY | 2020-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | NA | 1 | Event Description: FACEBOOK COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENTS RIGHT EYE (OD). THE PATIENT REPORTS DIPLOPIA VISION AFTER IMPLANTATION, ASTIGMATISM. LENS WAS EXPLANTED. PATIENT EXPERIENCED NEWBORN FLOATERS ONE MONTH AFTER EXPLANT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: H10 - CLAIM WAS DISCOVERED TO BE A DUPLICATE OF CLAIM# (B)(4), - MFR# 2019-02196. DISREGARD CURRENT INFOMATION AND SEE NOTED CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE U.S. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980374&pc=MTA | 2023826-2020-02958 |  | Injury | STAAR SUGICAL COMPANY | 2020-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | NA | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENTS LEFT EYE (OS). THE PATIENT REPORTS DIPLOPIA VISION AFTER IMPLANTATION, ASTIGMATISM. LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM WAS DISCOVERED TO BE A DUPLICATE OF CLAIM# 714996 - MFR# 2019-02197. DISREGARD CURRENT INFOMATION AND SEE NOTED CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # :(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634490&pc=MTA | 2023826-2020-02296 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, APPROXIMATELY 10 YEARS AGO AND NOW THE PATIENT IS EXHIBITING "MYOPTIC AND ASTIGMATIC REGRESSION" (POSSIBLE CHANGE IN REFRACTIVE ERROR). THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11412923&pc=MTA | 2023826-2021-00570 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-03-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | NA | 1 | Event Description: THE REPORTER STATED " I RECEIVED A COMPLAINT ABOUT ACCOMMODATION POST ICL IMPLANTATION IN THE CONTEXT OF GOOD VISION AND REFRACTION. DOCTOR WILL MONITOR FIRST AND LET US KNOW IF SHE NEEDS ASSISTANCE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH-RACE: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK. EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14962865&pc=MTA | 2023826-2022-02196 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | P030016 | 1 | Event Description: AN ARTICLE PUBLISHED TITLED "A NATIONWIDE MULTICENTER STUDY ON 1-YEAR OUTCOMES OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS IMPLANTATION FOR LOW MYOPIA"; WAS TO ASSESS THE OUTCOME OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS IMPLANTATION WITH A CENTER HOLE (EVO-ICL; STAAR SURGICAL) FOR PATIENTS WITH LOW MYOPIA. THE STUDY COVERS 172 EYES OF 111 PATIENTS. A 12.1MM VICM5_12.1; -3.0 DIOPTER IMPLANTABLE COLLAMER LENS THAT WAS IMPLANTED INTO THE PATIENT'S EYE; NEEDED TO BE EXCHANGED WITH A LONGER LENGTH/DIFFERENT POWER LENS DUE TO AN UNDERCORRECTION. LENS REMAINS IMPLANTED. SEE MFR# 2023826-2022-02195 FOR AN ADDITIONAL EYES RELATED TO THIS ARTICLE. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15411920&pc=MTA | 2023826-2022-03050 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "MACULAR HOLE 3 DAYS AFTER ICL." Manufacturer Narrative: B5 - VITRECTOMY WAS PERFORMED ON (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: B5: LENS WAS IMPLANTED ON (B)(6) 2022 AND REMOVED ON (B)(6) 2022, DUE TO A MACULAR HOLE 3 DAYS AFTER ICL. CAUSE OF THE EVENT IS UNKNOWN. PROBLEM WAS NOT RESOLVED. REPORTEDLY, "PATIENT TRANSFERRED TO GENERAL HOSPITAL, RECENT DETAILS ARE UNAVAILABLE." CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135982&pc=MTA | 2023826-2020-01176 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN HIS RIGHT EYE (OD), IN (B)(6) 2005 TO CORRECT STRONG NEARVISION PROBLEMS. RECENTLY, THE PATIENT HAS BEEN HAVING ISSUES WITH VISION ABILITY (NIGHT VISION, READING, COMPUTER ISSUES). THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO PRODUCT PROBLEM IN INITIAL MDR. B2 - NOT APPLICABLE IN INITIAL MDR. H1 - CORRECTED TO MALFUNCTION IN INITIAL MDR. H6 - DEVICE CODE 2993 IS NOT APPLICABLE IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135995&pc=MTA | 2023826-2020-01177 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN HIS LEFT EYE (OS), IN (B)(6) 2005 TO CORRECT STRONG NEARVISION PROBLEMS. RECENTLY, THE PATIENT HAS BEEN HAVING ISSUES WITH VISION ABILITY (NIGHT VISION, READING, COMPUTER ISSUES). THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO PRODUCT PROBLEM IN INITIAL MDR. B2 - NOT APPLICABLE IN INITIAL MDR. H1 - CORRECTED TO MALFUNCTION IN INITIAL MDR. H6 - DEVICE CODE 2993 IS NOT APPLICABLE IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073301&pc=MTA | 2023826-2020-03182 | 2020-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749981&pc=MTA | 2023826-2020-02523 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Patient-Device Incompatibility | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS A 'REFRACTIVE MISS.' ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE SURGEON STATES THERE IS NO ISSUE WITH THE IMPLANTED LENS. THIS IS A CORNEAL ISSUE. CLAIM# (B)(4). Manufacturer Narrative: UNK. DATE OF EVENT-UNK. PRODUCT CODE: UNK; SUBMITTER SOFTWARE DOES NOT ALLOW FOR UNK OR BLANK ENTRY. UNK. DEVICE MANUFACTURE DATE -UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9825902&pc=MTA | 2023826-2020-00443 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances | N/A | 1 | Event Description: RECEIVED FINAL ANALYSIS OF OUR PRODUCT SATISFACTORY SURVEY EU 2019, " SOMETIMES SEVERE ISSUES WITH NIGHT VISION," WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBERS REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359504&pc=MTA | 2023826-2021-03266 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL), IN THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2021. THE PATIENT COMPLAINED OF POOR VISION IN LOW-LIGHT CONDITIONS DUE TO HER VERY LARGE PUPILS EXPANDING BEYOND THE OPTICAL ZONE OF THE VISIAN ICL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12365649&pc=MTA | 2023826-2021-03267 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL), IN THE PATIENTS LEFT EYE (OS) IN (B)(6) 2021. THE PATIENT COMPLAINED OF POOR VISION IN LOW-LIGHT CONDITIONS DUE TO HER VERY LARGE PUPILS EXPANDING BEYOND THE OPTICAL ZONE OF THE VISIAN ICL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. EXPLANT DATE: NA. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14940366&pc=MTA | 2023826-2022-02076 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED A FEMALE PATIENT HAD AN IMPLANTABLE COLLAMER LENS (POSSIBLE ICL LENS) IMPLANTED IN THEIR EYE IN MEXICO ONE YEAR AGO. THE PATIENT IS NOW COMPLAINING OF VISUAL DISTURBANCES (THE PATIENT IS SEEING CIRCLES WHEN MOVING HER EYES). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE DOCTOR INDICATED SHE FELT THE PATIENT DID NOT HAVE AN ICL IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16870253&pc=MTA | 2023826-2023-01636 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances | UNK | 1 | Event Description: A PATIENT INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED; AND REPORTEDLY, "I HAD ICL SURGERY IN (B)(6) 2020. SINCE THEN I'VE SEEN CLOUDY WHITE VISION WHEN I'M IN A DARK/LOW IT ROOM AND LOOKING OUT THE WINDOW WHEN IT IS SUNNY OUTSIDE.". Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPIRATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511816&pc=MTA | 2023826-2020-01994 | 2020-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1M VICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT REQUESTED TO CHANGE THE LENS TO A LOWER POWER BECAUSE HE FELT THE VISUAL ACUITY WAS TOO GOOD. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL BUT DIFFERENT DIOPTER LENS (-06.00) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL,DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524373&pc=MTA | 2023826-2021-03721 | 2021-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6,-8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND DIFFERENT DIOPTER LENS DUE TO LOW VAULT AND DISCOMFORT WITH VISION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (OVER 45 YRS OF AGE AT DATE OF IMPLANT) ADDED TO INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT - RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911484&pc=MTA | 2023826-2022-00962 | 2020-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Distortion | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS DIOPTER -17.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME SIZE SAME DIOPTER LENS DUE TO POSTOPERATIVE VISUAL DISTORTION AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. AFTER EXCHANGE IT IS REPORTED THAT, "SELF-COMPLAINT VISUAL DISTORTION IMPROVED SLIGHTLY." Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). . CLAIM# (B)(4). Manufacturer Narrative: "THE REPORTER MARKED THE CAUSE OF THE EVENT AS PATIENT RELATED FACTOR AND STATED 'THE PATIENT HAS HIGH MYOPIA, POSTERIOR SCLERAL STPHYLOMA' AS THEIR CAUSE ANALYSIS." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082456&pc=MTA | 2023826-2021-04965 | 2021-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -8.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS GLARE/HALOS. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -8.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS GLARE/HALOS. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558556&pc=MTA | 2023826-2021-03809 | 2021-05-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS GLARE/HALOS AND PILOCARPINE ADMINISTRATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16856855&pc=MTA | 2023826-2023-01586 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022 . THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE PATIENT EXPERIENCED GLARE AND THE PROBLEM WAS RESOLVED. THE PATIENT COULD NOT ACCEPT THE GLARE AND REQUESTED THE LENS BE REMOVED. Manufacturer Narrative: H6 WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18297786&pc=MTA | 2023826-2023-05508 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18297821&pc=MTA | 2023826-2023-05509 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17916944&pc=MTA | 2023826-2023-04390 | 2023-08-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS EYE. AT ONE MONTH POST-OP, THE PATIENT COMPLAINED OF GLARE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK . D6A: UNK. D6B: UNK. H4: UNK. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. AT ONE MONTH POST-OP THE PATIENT COMPLAINED OF PERSISTENT GLARE AND THE LENS REMAINED IMPLANTED. POST-OP, THE PROBLEM IS GETTTING BETTER. THE CAUSE OF THE EVENT IS PATIENT-RELATED FACTOR, THE PATEINT IS A CAREFUL OBSERVER. CLAIM # 433379.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786657&pc=MTA | 2023826-2020-02581 |  | Injury | STAAR SUGICAL COMPANY | 2020-11-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | NA | 1 | Event Description: A LENS EXPLANT IN KERATACONUS PATIENT DUE TO GLARE WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: G3 - DATE RECEIVED BY MANUFACTURER CORRECTED TO 07-OCT-2020 IN THE INITAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (DEVICE PRODUCT CODE), UNK, UNABLE TO LEAVE FIELD BLANK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786720&pc=MTA | 2023826-2020-02580 |  | Injury | STAAR SUGICAL COMPANY | 2020-11-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | NA | 1 | Event Description: A LENS EXPLANT IN KERATACONUS PATIENT DUE TO GLARE WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: G3 - DATE RECEIVED BY MANUFACTURER CORRECTED TO 07-OCT-2020, IN THE INITAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (DEVICE PRODUCT CODE), UNK, UNABLE TO LEAVE FIELD BLANK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498032&pc=MTA | 2023826-2019-02401 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER INDICATED THE PATIENT WAS EXPERIENCING GLARE AT NIGHT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: NA. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498117&pc=MTA | 2023826-2019-02400 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER INDICATED THE PATIENT WAS EXPERIENCING GLARE AT NIGHT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9634639&pc=MTA | 2023826-2020-00204 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-01-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018 AS A REPLACEMENT LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE LENS RESOLVED THE PROBLEM BUT THE PATIENT HAD SOME ISSUES OF GLARE AT NIGHT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. REFERENCE MFR# 2023826-2020-00202 FOR INITIAL LENS. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9634784&pc=MTA | 2023826-2020-00205 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-01-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AS A REPLACEMENT LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE LENS RESOLVED THE PROBLEM BUT THE PATIENT HAD SOME ISSUES OF GLARE AT NIGHT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. REFERENCE MFR# 2023826-2020-00203 FOR INITIAL LENS. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11270326&pc=MTA | 2023826-2021-00264 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S RIGHT (OD) EYE. THE PATIENT REPORTS CLOSURE OF PIS; PI REPERFORMED. ALSO REPORTED IS MILD FLARES IN BRIGHT LIGHT. VISUAL ACUITY IS EXCELLENT AND THE PATIENT IS SATISFIED WITH THE ICL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11270275&pc=MTA | 2023826-2021-00265 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S LEFT (OS) EYE. THE PATIENT REPORTS MODERATE HORIZONTAL FLARES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: H6: TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH. ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: LENS IMPLANTED IN EITHER (B)(6) 2013. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599727&pc=MTA | 2023826-2021-03909 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -6.5 INTO THE PATIENT'S EYE ON (B)(6) 2021. THE SURGEON REPORTS HALOS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT IDENTIFIER, WEIGHT, ETHNICITY, RACE- UNK. DATE OF EVENT - UNK. INITIAL REPORTER -ESTABLISHMENT NAME/ADDRESS: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14254418&pc=MTA | 2023826-2022-01387 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE PATIENT REPORTS GLARE/HALOS AND STARBURSTS AROUND LIGHTS AT NIGHT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: B5 - PATIENT REPORTS "I WANTED TO GIVE AN UPDATE ON MY EYES SURGERY BECAUSE I HAD MY SECOND EYE SURGERY WITH DR BERG AND THE RESULTS ARE BEYOND EXCELLENT! AFTER DR BERG DID MY OTHER EYE ICL SURGERY, BOTH MY EYES GAINED 20 20 OR BETTER VISION. I NEVER COULD SEE THIS SHARP WITH GLASSES OR CONTACT LENS EVER. I HAVE TO APOLOGIZE TO DR BERG, I WAS WRONG ABOUT MY NIGHT VISION CAUSES. NOW WITH BOTH EYES CORRECTED, MY NIGHT VISION IS GREAT WITH VERY LITTLE HALOS, NOT AS MUCH AT ALL." SUPPLEMENTAL MDR#1 IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE -UNK. DATE OF EVENT-UNK. COMMON DEVICE NAME-UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK/UNK ENTRY MODEL #, IMPLANT DATE-UNK. INITIAL REPORTER OCCUPATION-UNK. DEVICE MANUFACTURE DATE-UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14492552&pc=MTA | 2023826-2022-01382 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY GLARES/HALOES OBSERVED. TATTOO INK WAS INJECTED INTO CORNEAL POCKETS AND ANTERIOR CHAMBER, STAINING THE LENS. THE LENS WAS REMOVED ON (B)(6) 2022 AND REPLACED IN A SEPARATE SURGERY WITH AN IDENTICAL LENS. CAUSE IS DUE TO THE DEVICE - ICL WAS PERFECT. PATIENT HAD GLARE FROM PI, TATTOO WAS INJECTED INTO CORNEAL POCKETS. Manufacturer Narrative: PT INFO: UNK. EVENT DATE: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15415025&pc=MTA | 2023826-2022-03052 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: A YOUTUBE CONTENT CREATOR SHARED HIS EXPERIENCE AFTER IMPLANTATION WITH IMPLANTABLE COLLAMER LENSES. 2 MONTHS AFTER IMPLANTATION, HE STATED THAT HE WAS EXPERIENCING GLARES AND DOUBLE VISION, AND BRIMONIDINE WAS PRESCRIBED WHICH SOLVED THE ISSUE. 2 MONTHS LATER IN ANOTHER VIDEO, THE CREATOR STOPPED USING THE EYE DROPS, BUT STILL EXPERIENCES GLARES. IN BOTH VIDEOS, HE ACKNOWLEDGES THAT HE HAS OVERSIZE PUPILS ARE NOT PERFECTLY FITTED TO THE LENS. HE STATED THAT HE IS HAPPY WITH THE PROCEDURE. LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY:UNKNOWN, RACE: UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15934261&pc=MTA | 2023826-2022-04459 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. THE SURGEON STATES THAT SINCE DAY 1 THE PATIENT HAS BEEN COMPLAINING OF GLARE. PER UPDATED INFORMATION THE PATIENT IS GETTING BETTER AND THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16567705&pc=MTA | 2023826-2023-00938 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. AT ONE MONTH POST-OP, THE PATIENT WAS COMPLAINING OF GLARE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587314&pc=MTA | 2023826-2023-00939 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. AT ONE MONTH POST-OP, THE PATIENT WAS COMPLAINING OF GLARE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16636051&pc=MTA | 2023826-2023-01065 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REMOVED ON (B)(6), 2023 DUE TO THE PATIENT COMPLAINED OF GLARE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK RACE: UNK DATE OF EVENT: UNK. IMPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16638118&pc=MTA | 2023826-2023-01066 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT COMPLAINED OF GLARE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18311165&pc=MTA | 2023826-2023-05091 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF A -15.0 DIOPTER INTO THE PATIENT'S EYE. PATIENT SUBJECTIVE DISTURBANCE OF "GLARE AND HALOS" WAS REPORTED. THERE IS NO VISUAL ACUITY LOSS OR LOSS OF BCVA REPORTED AND THE LENS REMAINS IMPLANTED. THE ICL DIRECITONS FOR USE IDENTIFIES GLARE AND HALO VISUAL DISTURBANCE AS A KNOWN POTENTIAL ADVERSE EVENT. PATIENT RISK FACTORS INCLUDE PRE-EXISTING/CURRENT OCULAR STATUS AND MESOPIC PUPIL SIZE THAT IS GREATER THAN THE ICL OPTIC DIAMETER. BASED ON THE INFORMATION RECEIVED, THE ROOT CAUSE OF THE REPORTED EVENT CANNOT BE DETERMINED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18121663&pc=MTA | 2023826-2023-04973 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; No Apparent Adverse Event | Visual Disturbances | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). GLARES ARE REPORTED. Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075964&pc=MTA | 2023826-2020-01085 | 2020-04-24 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-20 | MTA | IMPLANTABLE COLLAMER LENS | Patient-Device Incompatibility; Optical Problem | Visual Disturbances | N/A | 1 | Event Description: THE PATIENT REPORTED A VISIAN ICL WAS IMPLANTED INTO HIS LEFT (OS) EYE. THE PATIENT REPORTS SEEING A CIRCLE OF COLORS. IT IS UNKNOWN AS TO WHETHER THE LENS WAS REMOVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ESUBMITTER SOFTWARE DOES NOT ALLOW FOR UNK ENTRY. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323156&pc=MTA | 2023826-2015-01725 |  | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WILL BE EXPLANTED DUE TO THE LENS HAS BEEN SHIFTING AND THE PATIENT HAS BEEN EXPERIENCING VISUAL DISTURBANCE. THE SURGEON PLANS TO EXCHANGE THE LENS FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE SHIFTING OF THE IMPLANTED LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. PER MEDICAL REVIEW: REPORTEDLY, MICL (13.2 MM) CHANGED THE POSITION ("SHIFTED") FOLLOWING IMPLANTATION IN (B)(6) 2015 CAUSING VISUAL DISTURBANCES TO PATIENT. THE SURGEON WAS PLANNING TO EXCHANGE THE LENS, IN (B)(6) 2015, FOR A LONGER LENGTH ICL BUT HE PERFORMED ONLY MANUAL ROTATION INSTEAD. ACCORDING TO EMAIL REPORT DATED ON (B)(6) 2015, THE MICL SHIFTED AGAIN SO THE SURGEON WAS PLANNING LENS EXCHANGE FOR MICL (13.7 MM) TO PREVENT FUTURE COMPLICATIONS. DESPITE MULTIPLE ATTEMPTS NO CONFIRMATION ABOUT LENS EXCHANGE WAS RECEIVED TO DATE. THE SERIOUSNESS IS BASED ON INFORMATION ABOUT SECONDARY SURGICALLY INTERVENTION PERFORMED (ROTATION).DEVICE CAUSALITY IS BASED ON INFORMATION THAT AT THE TIME OF THE SURGERY PATIENT WAS (B)(6) AND PER DFU :"INDICATIONS: THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18187453&pc=MTA | 2023826-2023-05231 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. REFRACTION IS PLANO -0.25 BUT ISSUES WITH NIGHT VISION ARE REPORTED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE DESCRIPTION: THE EVO ICL LENS FEATURES AN OPTIC DIAMETER THAT VARIES WITH THE DIOPTRIC POWER; THE SMALLEST OPTIC DIAMETER BEING 4.9MM AND THE LARGEST 6.1MM. THE EVO ICL LENS IS CAPABLE OF BEING FOLDED AND INSERTED INTO THE POSTERIOR CHAMBER THROUGH AN INCISION OF 3.5MM OR LESS. THE EVO ICL LENS IS INTENDED TO BE PLACED ENTIRELY WITHIN THE POSTERIOR CHAMBER DIRECTLY BEHIND THE IRIS AND IN FRONT OF THE ANTERIOR CAPSULE OF THE HUMAN CRYSTALLINE LENS. WHEN CORRECTLY POSITIONED, THE EVO ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. MODE OF ACTION: THE ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. WARNINGS: A PATIENT WITH MESOPIC PUPIL SIZE THAT IS GREATER THAN THE OPTIC DIAMETER OF THE ICL LENS MAY EXPERIENCE SYMPTOMS OF GLARE AND/OR HALOS. PATIENTS SHOULD BE ADVISED ABOUT THIS POTENTIAL RISK PRIOR TO ICL LENS IMPLANTATION. PRECAUTIONS: PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE. PATIENTS WITH HIGHER DEGREES OF MYOPIA AND/OR MYOPIC ASTIGMATISM EXPERIENCE LOWER EFFICACY AND HIGHER RATES OF ADVERSE EVENTS (AES) AND COMPLICATIONS. IF A METHOD OF POWER CALCULATION DIFFERENT FROM THAT USED IN THE ICL LENS CLINICAL STUDY (I.E., LENS POWER CALCULATED BY STAAR SURGICAL USING STAAR'S PROPRIETARY SOFTWARE) IS USED, THE EFFECTIVENESS OF THE ICL LENS FOR MYOPIA WITH OR WITHOUT ASTIGMATISM MAY NOT BE CONSISTENT WITH THE RESULTS REPORTED IN THE ICL LENS CLINICAL STUDY RESULTS SECTION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16177977&pc=MTA | 2023826-2023-00153 | 2022-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED NIGHT TIME GLARE CAUSED BY LIGHTS. THE VA HAS BEEN GOOD AND PATIENT IS DOING WELL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE PATIENT ALSO EXPERIENCED KERATOCONUS. H6: HEALTH EFFECT- CLINICAL CODE: 4581-KERATOCONUS. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178186&pc=MTA | 2023826-2023-00154 | 2022-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED NIGHT TIME GLARE CAUSED BY LIGHTS. THE VA HAS BEEN GOOD AND PATIENT IS DOING WELL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5- "THE PATIENT ALSO EXPERIENCED KERATOCONUS." H6- HEALTH EFFECT- CLINICAL CODE: 4581- KERATOCONUS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313440&pc=MTA | 2023826-2023-02767 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5 INTO THE PATIENT'S LEFT EYE (OS) ON(B)(6) 2020. IRIS TRANSILLUMINATION DEFECTS ARE REPORTED. CAUSE OF THE EVENT WAS THE DEVICE. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "ADVERSE EVENT STARTED 23JUNE23, NOTED AS MILD AND RELATED TO DEVICE. (B)(6) 2023 WAS THE LAST PROTOCOL REQUIRED STUDY VISIT, SUBJECT EXITED SAME DAY AND WILL BE FOLLOWED PER STANDARD OF CARE." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581- IRIS TRANSILLUMINATION DEFECTS" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED IRIS TRANSILLUMINATION DEFECTS. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "ADVERSE EVENT STARTED (B)(6) 23, NOTED AS MILD AND RELATED TO DEVICE. (B)(6) 2023 WAS THE LAST PROTOCOL REQUIRED STUDY VISIT, SUBJECT EXITED SAME DAY AND WILL BE FOLLOWED PER STANDARD OF CARE". ADDITIONAL INFORMATION PROVIDED: 31AUG23, THE TID SEEN AT THE 36-MONTH VISIT WAS VERY MILD. IT WAS DIFFICULT TO EVEN PHOTOGRAPH. CLINICIAN DID NOT NOTICE IT AT THE 24-MONTH VISIT OR OTHER PREVIOUS VISITS. NO OTHER CHANGES WERE NOTED IN SUBJECT'S HEALTH, MEDS, VISION AND REFRACTION WERE STABLE." THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313441&pc=MTA | 2023826-2023-02766 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. IRIS TRANSILLUMINATION DEFECTS ARE REPORTED. CAUSE OF THE EVENT WAS THE DEVICE. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "ADVERSE EVENT STARTED (B)(6) 2023, NOTED AS MILD AND RELATED TO DEVICE. (B)(6) 2023 WAS THE LAST PROTOCOL REQUIRED STUDY VISIT, SUBJECT EXITED SAME DAY AND WILL BE FOLLOWED PER STANDARD OF CARE." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581- IRIS TRANSILLUMINATION DEFECTS" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED IRIS TRANSILLUMINATION DEFECTS. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "ADVERSE EVENT STARTED (B)(6) 2023, NOTED AS MILD AND RELATED TO DEVICE. (B)(6) 2023 WAS THE LAST PROTOCOL REQUIRED STUDY VISIT, SUBJECT EXITED SAME DAY AND WILL BE FOLLOWED PER STANDARD OF CARE". ADDITIONAL INFORMATION PROVIDED: "(B)(6) 2023 THE TID SEEN AT THE 36-MONTH VISIT WAS VERY MILD. IT WAS DIFFICULT TO EVEN PHOTOGRAPH. CLINICIAN DID NOT NOTICE IT AT THE 24-MONTH VISIT OR OTHER PREVIOUS VISITS. NO OTHER CHANGES WERE NOTED IN SUBJECT'S HEALTH, MEDS, VISION AND REFRACTION WERE STABLE." THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772254&pc=MTA | 2023826-2019-01177 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; Discomfort | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 REPORTER INDICATED THAT LOW VAULT WITH TEMPORAL/INFERIOR DECENTRATION AND WITH "DOUBLING IMAGE AND DISCOMFORT WAS REPORTED. DOCTOR AND PATIENT DO NOT WANT MORE INTERVENTIONS AT THIS TIME." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT : LENS ROTATION (30 DEGREES) "AGAIN", LENS LOCATED OUTSIDE THE PUPIL, OFF-CENTER AND PATIENT "FEELS IMAGE UNFOLDING AGAIN." FOUND NO ZONULAR DEFECT AND "SURGEON WANTS TO EXPLANT FOR FEAR OF MAJOR PROBLEMS." CLAIM# (B)(4). Manufacturer Narrative: LENS ROTATION WAS ALSO REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852586&pc=MTA | 2023826-2020-02687 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Discomfort | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT WAS UNCOMFORTABLE AND EXPERIENCED VISION FATIGUE. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279016&pc=MTA | 2023826-2023-02617 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1, -01.25 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO PATIENT DISCOMFORT AND GLARE ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2330" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1IMPLANTABLE COLLAMER LENS OF DIOPTER -1.25 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED DISCOMFORT/ GLARE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "NO. THERE IS NO PLAN TO IMPLANT ANOTHER LENS." CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6521049&pc=MTA | 2023826-2017-00696 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT WAS EXPERIENCING GLARE AND HALOS, WORSE IN AM AND AT NIGHT. THE PATIENTS PUPILS ARE 6.5MM IN DIM ILLUMINATION, ARE ESSENTIALLY EQUAL IN SIZE AND THERE IS NO EVIDENCE OF TRAUMATIC MYDRIASIS. THE PATIENT IS BEING TREATED WITH DROPS AND THE LENS REMAINS IMPLANTED. THE PATIENTS VISION IS 20/20 UNCORRECTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6521056&pc=MTA | 2023826-2017-00695 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT WAS EXPERIENCING GLARE AND HALOS, WORSE IN AM AND AT NIGHT. THE PATIENTS PUPILS ARE 6.5MM IN DIM ILLUMINATION, ARE ESSENTIALLY EQUAL IN SIZE AND THERE IS NO EVIDENCE OF TRAUMATIC MYDRIASIS. THE PATIENT IS BEING TREATED WITH DROPS AND THE LENS REMAINS IMPLANTED. THE PATIENTS VISION IS 20/20 UNCORRECTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020633&pc=MTA | 2023826-2017-01701 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT HAS BEEN BOTHERED BY PERSISTENT GLARE AND HALOS TWO MONTHS AFTER THE SURGERY. THE PATIENT HAS TRIED GLASSES AND MIOTICS BUT WITH NO IMPROVEMENT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 6 MONTHS. Manufacturer Narrative: CORRECTED DATA: THE EVENT IS CORRECTED TO "PRODUCT PROBLEM" FROM "ADVERSE EVENT". TYPE OF REPORTABLE EVENT: IT IS CORRECTED TO "MALFUNCTION" FROM "SERIOUS INJURY". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020681&pc=MTA | 2023826-2017-01702 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT HAS BEEN BOTHERED BY PERSISTENT GLARE AND HALOS TWO MONTHS AFTER THE SURGERY. THE PATIENT HAS TRIED GLASSES AND MIOTICS BUT WITH NO IMPROVEMENT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 6 MONTHS. Manufacturer Narrative: CORRECTED DATA: THE EVENT IS CORRECTED TO "PRODUCT PROBLEM" FROM "ADVERSE EVENT". TYPE OF REPORTABLE EVENT: IT IS CORRECTED TO "MALFUNCTION" FROM "SERIOUS INJURY". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930935&pc=MTA | 2023826-2020-02813 | 2020-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930952&pc=MTA | 2023826-2020-02812 | 2020-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11738945&pc=MTA | 2023826-2021-01199 | 2020-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OS) ON 29/SEP/2020. REPORTEDLY, LENS TOO LONG, SURGEON PREFERENCE TO CHANGE. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENGTH, HIGHER POWER LENS. PER PATIENT'S (B)(6) 2021 VISIT, GLARE AND HALOS ARE BEING EXPERIENCED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS EXCHANGE RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7945358&pc=MTA | 2023826-2018-01501 | 2018-08-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS), THE HAPTICS TORE DURING INJECTION. THE SURGEON STATES THE LENS IS "TILTED AND DISLOCATED" AND "THE LENS WENT OUT WITH A BROKEN HAPTIC." THE LENS WAS IMPLANTED ON (B)(6) 2018 AND REPORTEDLY REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS LISTED AS "LENS TORN." ADDITIONALLY, THE SURGEON REPORTS LENS DISLOCATION/SUBLUXATION AND GLARE/HALOS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: SERIOUS INJURY CORRECTED TO MALFUNCTION IN SUPPLEMENTAL #1. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7287284&pc=MTA | 2023826-2018-00243 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS TORN IN THE HAPTIC AREA, DURING DELIVERY INTO THE EYE AND REMAINS IMPLANTED. THE PATIENT COMPLAINS OF GLARE AND HALOS, ESPECIALLY AT NIGHT. THE SURGEON PLANS TO EXPLANT THE LENS. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7578068&pc=MTA | 2023826-2018-00864 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLMER LENS INTO THE RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO HALOS AND GLARE. IT WAS REPORTED THAT THE HALOS AND GLARES WERE NOTICED ON THE SAME DAY AS THE IMPLANT SURGERY. THERE IS NO PLANS TO IMPLANT ANOTHER LENS. PATIENT IS STABLE BUT STILL HAS GLARE FROM PI. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A TORN HAPTIC AND RESIDUE ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7954143&pc=MTA | 2023826-2018-01500 | 2018-04-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717807&pc=MTA | 2023826-2020-00333 | 2020-01-18 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SEVERE GLARE AND HALOES. THE REPORTER STATED THAT THE SEVERE HALOES MAY BE DUE TO VAULT PUSHING PUPIL WIDER. THE LENS REMAINS IMPLANTED. THE DOCTOR PLANS ON EXCHANGING THE LENS FOR A SHORTER LENGTH LENS BUT THE PATIENT IS CURRENTLY PREGNANT, SO HE WILL HAVE TO WAIT UNTIL SHE DELIVERS LATER THIS YEAR. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717808&pc=MTA | 2023826-2020-00332 | 2020-01-18 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SEVERE GLARE AND HALOES. THE REPORTER STATED THAT THE SEVERE HALOES MAY BE DUE TO VAULT PUSHING PUPIL WIDER. THE LENS REMAINS IMPLANTED. THE DOCTOR PLANS ON EXCHANGING THE LENS FOR A SHORTER LENGTH LENS BUT THE PATIENT IS CURRENTLY PREGNANT, SO HE WILL HAVE TO WAIT UNTIL SHE DELIVERS LATER THIS YEAR. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403835&pc=MTA | 2023826-2019-00347 | 2018-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM, MICL12.6, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS EYE. THE SURGEON REPORTS EXCESSIVE VAULT, GLARE, AND HALOS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091302&pc=MTA | 2023826-2020-03208 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO THE LENS HAD A HIGHER VAULT AND WAS CAUSING BAD HALO AND GLARE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS IMPLANTED ON 30-OCT-2020. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CORRECTED DATA: B3: DATE OF EVENT - (B)(6) 2020. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6072194&pc=MTA | MW5065752 | 2016-10-10 | Injury | VISIAN STARR SURGICAL AW | 2016-10-31 | MTA | VISIAN ICL LENS | Insufficient Information | Visual Disturbances; Halo |  | 1 | Event Description: I HAD VISIAN ICL VISION CORRECTION PERFORMED BY DR. (B)(6) WITH (B)(6). THEY DID NOT CHECK MY PUPIL SIZE, AND NOW WITH THE ICL LENSES IN MY EYES WHICH MY PUPILS IN LOW LIGHT DILATE LARGER THAN THE USABLE AREA OF, I CANNOT DRIVE SAFELY AT NIGHT DUE TO EXTREME HALOS, STAR BURST AND DOUBLE/GHOSTING VISION. I ALSO HAVE GHOSTING DURING OFFICE ENVIRONMENT LIGHTING SO IT IS NOT ONLY AT NIGHT AND IN LOW LIGHT. THEY ALSO PERFORMED LASER IRIDOTOMY TO BOTH MY EYES PRIOR TO INSERTING THESE LENS. I HAVE PERMANENT WHITE LINE GLARES AND GHOSTING AFFECTS FROM THIS ALSO. I HAVE CRIED OVER THIS AS A (B)(6) MAN THAT MY EYES ARE RUINED. I WAS ONLY TOLD ABOUT POTENTIAL INFECTIONS, AND PRESSURE INCREASE AFTER THE SURGERY. NEVER WAS I TOLD ABOUT THESE SIDE EFFECTS. I WANT SOMETHING DONE ABOUT THIS! ALSO, DR. (B)(6) TOLD ME FACE TO FACE WHEN I ASKED HOW MANY OF THESE PROCEDURES HAD HE DONE HE SAID "20 EYES" HOWEVER ON THE DAY OF THE LENS IMPLANTS THE NURSE SAID I WAS HIS 4TH PATIENT, THAT SHE HAD BEEN IN THERE WITH HIM WITH ALL OF THEM SO FAR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7636085&pc=MTA | 2023826-2018-00952 | 2018-04-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE PATIENT COMPLAINED OF GLARE AND HALOS AND COULD SEE THE BORDER OF THE LENS. THE LENS REMAINS IMPLANTED. SEE MFR REPORT # 2023826-2018-00951 FOR INITIAL LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574193&pc=MTA | 2023826-2017-00831 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS; -16.00/2.5/085 DIOPTER; IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THIS WAS A 2ND LENS IMPLANTED IN THIS PATIENT (EXCHANGED TO A SHORTER LENS) AND THE PATIENT REPORTED THE EXCHANGE DID NOT RESOLVE THE PROBLEM. THE PATIENT REPORTED VISION WAS POOR IN BAD LIGHTING AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574194&pc=MTA | 2023826-2017-00830 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS; -16.50/2.5/097 DIOPTER; IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THIS WAS A 2ND LENS IMPLANTED IN THIS PATIENT (EXCHANGED FOR A SHORTER LENS) AND THE PATIENT REPORTED THE EXCHANGE DID NOT RESOLVE THE PROBLEM. THE PATIENT REPORTED VISION WAS POOR IN BAD LIGHTING AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6938912&pc=MTA | 2023826-2017-01557 |  | Injury | STAAR SURGICAL COMPANY | 2017-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO VISIAN TORIC IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S LEFT EYE (OS) IN EARLY (B)(6) 2017, FOR CORRECTION OF SEVERE MYOPIA OF -12.5 , WITH ASTIGMATISM. THE REPORTER INDICATED THE PATIENT HAD NOTED CUTTING OFF OF THE PERIPHERAL VISION IN LEFT EYE. THE SYMPTOM IS WORSE WHEN ACCOMMODATING (READING). THE PERIPHERAL VISION ALMOST APPEARS LIKE SEEING THE EDGE OF THE LENS WITH A CRESCENT SHAPE CUTOFF. THE SYMPTOM IS BOTHERSOME SINCE THE PATIENT DOES A LOT OF READING FOR THEIR JOB AND THE VISION FIELD FEELS ASYMMETRICAL. THE PATIENT IS ALSO EXPERIENCING HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18223261&pc=MTA | 2023826-2023-05166 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18223321&pc=MTA | 2023826-2023-05167 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10936807&pc=MTA | 2023826-2020-02841 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM A PATIENT'S EYE. THE PATIENT REPORTS GLARES, HALOS, DOUBLE VISION, AND STARBURSTS. THE PROBLEM WAS RESOLVED WITH EXPLANT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR UNK ENTRY. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10936888&pc=MTA | 2023826-2020-02842 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: FACEBOOK COMPLAINT: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM A PATIENT'S EYE. THE PATIENT REPORTS GLARES, HALOS, DOUBLE VISION, AND STARBURSTS. THE PROBLEM WAS RESOLVED WITH EXPLANT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK; PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195876&pc=MTA | 2023826-2022-02607 | 2021-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE PATIENT DID NOT WANT THE LENS REPLACED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195903&pc=MTA | 2023826-2022-02608 | 2021-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE PATIENT DID NOT WANT THE LENS REPLACED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY:UNKNOWN, RACE: UNKNOWN. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921079&pc=MTA | 2023826-2022-04454 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921081&pc=MTA | 2023826-2022-04455 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295992&pc=MTA | 2023826-2023-02835 | 2023-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295993&pc=MTA | 2023826-2023-02836 | 2023-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17722680&pc=MTA | 2023826-2023-03959 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715576&pc=MTA | 2023826-2023-03874 | 2023-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715689&pc=MTA | 2023826-2023-03875 | 2023-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18115810&pc=MTA | 2023826-2023-04768 | 2023-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK . A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167101&pc=MTA | 2023826-2021-02514 | 2021-03-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE RECEIVED BY MANUFACTURER: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167100&pc=MTA | 2023826-2021-02513 | 2021-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: G3: CORRECTED TO 21-JUN-2021 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16579641&pc=MTA | 2023826-2023-00988 | 2023-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS GLARE AND HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS NOTED "AFTER IMPLANTATION THE PATIENT SHOWED SYMPTOMS OF GLARE; AND THE CLINICIAN TRIED TO COMMUNICATE WITH HIM FOR A MONTH; BUT STILL COULD NOT ADAPT; THE PATIENT STRONGLY REQUESTED REMOVAL". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16580026&pc=MTA | 2023826-2023-00989 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS GLARE AND HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS NOTED "AFTER IMPLANTATION THE PATIENT SHOWED SYMPTOMS OF GLARE; AND THE CLINICIAN TRIED TO COMMUNICATE WITH HIM FOR A MONTH; BUT STILL COULD NOT ADAPT; THE PATIENT STRONGLY REQUESTED REMOVAL". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM) . CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18508980&pc=MTA | 2023826-2024-00216 | 2023-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Distortion; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -4.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON 30-NOV-2023 THE LENS WAS EXCHANGED FOR A LONGER LENS OF THE SAME MODEL AND POWER DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11750667&pc=MTA | 2023826-2021-01280 | 2018-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 14 AUG 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11751109&pc=MTA | 2023826-2021-01279 | 2018-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307835&pc=MTA | 2023826-2019-00226 | 2019-01-10 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS GLARE/HALOS. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273785&pc=MTA | 2023826-2021-00245 | 2019-09-11 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO _13.2, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. PER THE REPORTER ON (B)(6) 2021, ANOTHER LENS WILL NOT BE IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13385986&pc=MTA | 2023826-2022-00343 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13386000&pc=MTA | 2023826-2022-00342 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10819682&pc=MTA | 2023826-2020-02645 | 2020-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE SURGEON PLANS ON IMPLANTING A REPLACEMENT LENS BUT THE DATE IS NOT FIXED YET. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. RESIDUE ON LENS SURFACE WAS OBSERVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10819835&pc=MTA | 2023826-2020-02644 | 2020-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE SURGEON PLANS ON IMPLANTING A REPLACEMENT LENS BUT THE DATE IS NOT FIXED YET. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. RESIDUE ON LENS SURFACE WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12112700&pc=MTA | 2023826-2021-02210 | 2020-08-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF EVENT. IT WAS UNKNOWN IF THE DEVICE FAIL TO PERFORM AS INTENDED. THE REPORTER STATED "AFTER THE OPERATION, THE PATIENTS VISION IS NOT ADAPTED AND CANNOT BE RELIEVED." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507468&pc=MTA | 2023826-2021-00669 | 2020-08-31 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY IMMEDIATELY AFTER THE ICL SURGERY THE PATIENT HAD GOOD VISION. HOWEVER, HE COMPLAINED OF POSTOPERATIVE HALOS ESPECIALLY IN THE DARK. ONE MONTH AFTER SURGERY THE SITUATION WAS WORSE. THE PATIENT HAD DIFFICULTY GOING OUTSIDE AND DRIVING AT NIGHT. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507480&pc=MTA | 2023826-2021-00668 | 2020-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY IMMEDIATELY AFTER THE ICL SURGERY THE PATIENT HAD GOOD VISION. HOWEVER, HE COMPLAINED OF POSTOPERATIVE HALOS ESPECIALLY IN THE DARK. ONE MONTH AFTER SURGERY THE SITUATION WAS WORSE. THE PATIENT HAD DIFFICULTY GOING OUTSIDE AND DRIVING AT NIGHT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507478&pc=MTA | 2023826-2021-00673 | 2020-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE, RESIDUE, AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).