https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587682&pc=QCB | 2023826-2020-00122 | 2019-11-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608210&pc=QCB | 2023826-2020-00163 | 2019-12-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -08.50/1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2019, "THE DOCTOR WANTS TO EXCHANGE LENSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608211&pc=QCB | 2023826-2020-00164 | 2019-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -09.00/1.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2019, "THE DOCTOR WANTS TO EXCHANGE LENSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9614042&pc=QCB | 2023826-2020-00133 | 2019-12-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+3.0/165 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9618777&pc=QCB | 2023826-2020-00134 | 2019-12-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+4.0/073 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9656158&pc=QCB | 2023826-2020-00211 | 2019-10-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 7.50/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE. THE REPORTER INDICATED THE LENS WILL BE EXPLANTED DUE TO TOO MUCH VAULTING. THE LENS WILL BE EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9696241&pc=QCB | 2023826-2020-00224 | 2020-01-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient |  | 1 | Event Description: THE REPORTER INDICATED THAT A TORIC ICL 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -14.5/2.0/094 (SPHERE/CYLINDER/AXIS) TORE IN THE INJECTOR DURING SURGERY ON THE PATIENT'S RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2020. ON A LATER DATE AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN AS WELL AS RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703351&pc=QCB | 2023826-2020-00303 |  | Injury | STAAR SURGICAL COMPANY | 2020-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A12.1MM, VTICH12.1, 2.50/4.0/085, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION AND "SOMEWHAT" OF A LOW VAULT WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- "LENS REMAINS IMPLANTED." SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9705873&pc=QCB | 2023826-2020-00226 | 2018-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/1.5/087 (SPHERE/CYLINDER/AXIS) TORE/BROKE BEFORE/DURING LOADING AFTER OPENING THE VIAL FOR THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2018. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717757&pc=QCB | 2023826-2020-00293 | 2020-01-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.00/0.5/031 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY "PATIENT IS HAPPY." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - "FOREIGN MATERIAL ON THE LENS SURFACE" SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721912&pc=QCB | 2023826-2020-00328 | 2019-12-09 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-18 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/089 (SPHERE/CYLINDER/AXIS) WAS RETURNED TO STAAR SURGICAL. THE WORDS "UPSIDE DOWN" AND "DID NOT USE-OD" WERE WRITTEN ON THE BOX. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: ON (B)(6) 2019 THE TMICL LENS WAS IMPLANTED UPSIDE DOWN. IT WAS REMOVED AND REPLACED SUCCESSFULLY WITH A BACK-UP LENS. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED ALL WITHIN THE SAME SURGERY. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782978&pc=QCB | 2023826-2020-00410 | 2020-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -1.0/+1.0/005 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781525&pc=QCB | 2023826-2020-00339 | 2020-02-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -09.50/0.5/022 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. DAMAGE WAS NOTED IN CARTRIDGE. "SURGERY RESCHEDULED." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781714&pc=QCB | 2023826-2020-00420 | 2019-10-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/078 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# : (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510K - THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789273&pc=QCB | 2023826-2020-00412 | 2020-01-28 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020 AS A REPLACEMENT LENS - SEE MFR. REPORT 2023826-2020-00411 FOR INITIAL LENS. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789280&pc=QCB | 2023826-2020-00413 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824533&pc=QCB | 2023826-2020-00358 | 2019-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9..50/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H10: ADDITIONAL NARRATIVES/DATA: "H6: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE)" SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824542&pc=QCB | 2023826-2020-00360 | 2019-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824550&pc=QCB | 2023826-2020-00362 | 2019-11-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE DENTAL ARCH HEIGHT WAS SIGNIFICANTLY HIGHER AFTER IMPLANTATION, AND THE ANTERIOR CHAMBER WAS SHALLOW. TO BE ON THE SAFE SIDE, REMOVE THE 13.2 SIZE LENS AND IMPLANTED THE 12.6 SIZE LENS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824551&pc=QCB | 2023826-2020-00364 | 2019-07-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.50/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824555&pc=QCB | 2023826-2020-00363 | 2019-09-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824568&pc=QCB | 2023826-2020-00486 | 2019-11-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9836557&pc=QCB | 2023826-2020-00541 | 2019-04-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS LATER REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1-IN INITIAL MDR SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2-IN INITIAL MDR REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT /DAMAGE IS NOT APPLICABLE. B3- DATE OF EVENT IS UNK. B5-THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS IMPLANTED/REMOVED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. H1-IN INITIAL MDR AND SUPPLEMENTAL #1; SERIOUS INJURY SHOULD BE CORRECTED TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837917&pc=QCB | 2023826-2020-00529 | 2019-10-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/102 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9845342&pc=QCB | 2023826-2020-00575 | 2019-11-27 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM, VTICMO13.2, -16.5/4.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION WAS NOT NECESSARY. " VAULT IS 2+ ICL THICKNESS. VA HAS IMPROVED ENOUGH TI WARRANT NO FURTHER TX AT THIS TIME." CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9841763&pc=QCB | 2023826-2020-00551 | 2019-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS LATER REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B5 - THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842793&pc=QCB | 2023826-2020-00511 | 2019-11-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/059 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9855298&pc=QCB | 2023826-2020-00533 | 2020-02-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICH13.2, 8.50/3.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VICH13.2 OF +9.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. REPORTER STATES THAT "VAULTING GETS BACK TO NORMAL RANGE." LENS REMAINS IMPLANTED. D2-PHAKIC INTRAOCULAR LENS, PRODUCT CODE: MTA. D4- MODEL#: VICH 13.2. SERIAL #: (B)(6). EXPIRATION DATE: 31-OCT-2020. D6-IMPLANT DATE (B)(6) 2020. G5- THIS PRODUCT IS NOT MARKETED IN THE US. H6-METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT.(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9874569&pc=QCB | 2023826-2020-00637 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-03-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A TORIC ICL PATIENT HAS A -1.00 MYOPIC SPH OUTCOME (OS). LENS REMAINS IMPLANTED. PER NEW INFORMATION BOTH THE AMOUNT OF RESIDUAL MYOPIA AND THE UCDVA HAVE IMPROVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9914135&pc=QCB | 2023826-2020-00776 | 2020-02-28 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM, VTICM5_12.6, -7.50/1.00/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. THE REPORTER STATED " THE VAULTS ARE VERY LOW. THE ACUITIES ARE STABLE AND GOOD AND THE LENS ARE OK." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913329&pc=QCB | 2023826-2020-00738 | 2020-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.00/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2020 AND A LONGER LENGTH NON-STAAR IPCL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. REPORTEDLY, "LASER TOUCH UP IS PLANNED AND PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916653&pc=QCB | 2023826-2020-00732 | 2019-11-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.5/085 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916672&pc=QCB | 2023826-2020-00733 | 2019-08-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/080 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928353&pc=QCB | 2023826-2020-00711 | 2020-01-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/3.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928355&pc=QCB | 2023826-2020-00712 | 2019-11-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.50/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9940832&pc=QCB | 2023826-2020-00844 | 2020-02-14 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WILL BE REPLACED WITH A SHORTER LENGTH LENS. THE REPORTER STATED THAT THE LENS IS TOO BIG FOR THE PATIENTS EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9940866&pc=QCB | 2023826-2020-00845 | 2020-02-14 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.5/1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WILL BE REPLACED WITH A SHORTER LENGTH LENS. THE REPORTER STATED THAT THE LENS IS TOO BIG FOR THE PATIENTS EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938080&pc=QCB | 2023826-2020-00710 | 2019-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.00/2.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938130&pc=QCB | 2023826-2020-00709 | 2019-11-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9940468&pc=QCB | 2023826-2020-00834 | 2019-09-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/4.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED AND THE LENS WAS REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9955260&pc=QCB | 2023826-2020-00888 | 2019-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-14.5/2.5/007 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CORRECTED DATA: B7-CATARACT IN THE OPERATIVE EYE. H6-DEVICE CODE 1494 OFF-LABEL; PRE-EXISTING CATARACT IN THE OPERATIVE EYE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9955312&pc=QCB | 2023826-2020-00889 | 2019-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-13.5/2.0/008 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED AND RESIDUE ON LENS. CLAIM# :(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957590&pc=QCB | 2023826-2020-00778 | 2019-01-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -9.50/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H1 - CORRECTED TO MALFUNCTION IN INITIAL MDR. H6 - PATIENT CODE 3191 IS NOT APPLICABLE IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9977665&pc=QCB | 2023826-2020-00915 | 2020-03-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS IMPLANTED IN THE PATIENT WAS NOT THE PLANNED LENS(DIFFERENT LENS). LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: "THE SURGEON "MADE THE MISTAKE OF IMPLANTING THE LENS OF ANOTHER PATIENT." " SHOULD BE ADDED TO INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9986805&pc=QCB | 2023826-2020-00942 | 2020-01-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-12.5/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9991678&pc=QCB | 2023826-2020-00974 | 2019-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.00/+6.0/096 (SPHERE/CYLINDER/AXIS), FROM THE PATIENTS LEFT EYE (OS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9991680&pc=QCB | 2023826-2020-00975 | 2019-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.00/+6.0/099 (SPHERE/CYLINDER/AXIS), FROM THE PATIENTS RIGHT EYE (OD), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10048954&pc=QCB | 2023826-2020-01041 | 2020-04-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS "OTHER" WITH NO FURTHER INFORMATION. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116220&pc=QCB | 2023826-2020-01162 | 2020-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116222&pc=QCB | 2023826-2020-01163 | 2020-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THAT THE HAPTIC WAS BENT AND DEFORMED. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116227&pc=QCB | 2023826-2020-01161 | 2019-09-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. WITHIN THE SAME SURGERY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE VAULT VALUE WENT FROM 230UM PRIOR TO EXCHANGE, TO 440UM AFTER EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: PROBLEM WAS RESOLVED WITH EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135946&pc=QCB | 2023826-2020-01170 | 2020-03-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00/+1.5/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), AND NOTED A SMALL DEFORMITY AT THE TOP EDGE OF THE LENS. THE LENS WAS PARTIALLY INSERTED INTO THE EYE AND WAS REMOVED WITHIN THE SAME SURGERY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138549&pc=QCB | 2023826-2020-01180 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM, VTICMO13.2, -11.00/2.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. REPORTEDLY, "INITIAL LENS WAS FOUND BROKEN AFTER IMPLANTATION AND REMOVED DURING INITIAL SURGERY. REPLACEMENT LENS NOT YET IMPLANTED DUE TO PATIENT'S ISSUES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN AND HAPTIC BENT WITH DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138558&pc=QCB | 2023826-2020-01138 | 2020-01-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.50/2.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189314&pc=QCB | 2023826-2020-01275 | 2020-05-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 15.50/+1.0/087 (SPHERE/CYLINDER/AXIS), WAS TORN DURING LOADING. THE SURGEON IMPLANTED ANOTHER SAME MODEL/LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197347&pc=QCB | 2023826-2020-01330 | 2019-08-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. WITHIN THE SAME SURGERY, THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON THE SAME DATE A LONGER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197350&pc=QCB | 2023826-2020-01325 | 2020-05-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/113 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B3- "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197357&pc=QCB | 2023826-2020-01328 | 2019-09-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222172&pc=QCB | 2023826-2020-01411 | 2020-03-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222181&pc=QCB | 2023826-2020-01413 | 2020-02-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10224705&pc=QCB | 2023826-2020-01391 | 2020-06-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.50/1.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.50/1.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020. REPORTEDLY, THE "PATIENT IS OK" AFTER LENS REPLACEMENT. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242498&pc=QCB | 2023826-2020-01421 | 2020-06-03 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS ATTEMPTING TO INSERT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/040 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) AND NOTED THE LENS WAS TORN BEFORE IMPLANTING. THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271257&pc=QCB | 2023826-2020-01479 | 2020-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -4.00/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS OPACITY WAS OBSERVED ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS REMOVED AND CATARACT SURGERY WAS PERFORMED. REPORTEDLY, "PATIENT IS HAPPY." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: B4: DATE OF THIS REPORT IN INITIAL MDR 26-JUN-2020 IS CORRECTED TO 25-JUN-2020. G4 IN INITIAL MDR 26-JUN-2020 IS CORRECTED TO 25-JUN-2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). SECONDARY SURGICAL INTERVENTION, LENS EXPLANT, CATARACT SURGERY. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308836&pc=QCB | 2023826-2020-01529 | 2020-07-01 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/1.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. AN "OIL-LIKE ATTACHMENTS WERE OBSERVED ON THE SURFACE OF THE LENS." LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, "PATIENT'S ACUITY WAS 1.0 ON (B)(6) 2020. LITTLE ASTIGMATISM, SURGEON BELIEVES CAUSED BY INCISION. EVERYTHING ELSE OK, AC WAS QUIET. SURGEON WILL CONTINUE TO OBSERVE CLOSELY. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308790&pc=QCB | 2023826-2020-01519 | 2020-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308807&pc=QCB | 2023826-2020-01520 | 2020-06-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/3.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK, PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308810&pc=QCB | 2023826-2020-01518 | 2020-03-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.50/0.5/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308829&pc=QCB | 2023826-2020-01523 | 2019-10-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.1MM VTICMO12.1, -10.00/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2019, A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335345&pc=QCB | 2023826-2020-01531 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/088 (SPHERE/CYLINDER/AXIS) IT TORE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332986&pc=QCB | 2023826-2020-01583 | 2020-06-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -13.00/2.0/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS "REMOVED FROM VIAL AND FOUND THAT THERE WERE CRACKS IN THE LENS. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335336&pc=QCB | 2023826-2020-01545 | 2020-03-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE, RESIDUE, AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10356719&pc=QCB | 2023826-2020-01606 | 2020-07-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/090 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO A BLEMISH/SCRATCH ON THE OPTIC, NOTED ONCE THE LENS WAS INSERTED. THE BACKUP LENS WAS IMPLANTED AND THE EXCHANGE WAS UNEVENTFUL WITH NO SEQUELAE. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS LOST. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361982&pc=QCB | 2023826-2020-01623 | 2020-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.50/0.5/146 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UNK. UNK. UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361986&pc=QCB | 2023826-2020-01622 | 2020-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.50/0.5/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361989&pc=QCB | 2023826-2020-01626 | 2020-06-24 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.00/2.5/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, "THE LENS IS IN THE RIGHT AXIS (18") BUT THE ASTIGMATISM IS NOT CRRECTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398653&pc=QCB | 2023826-2020-01747 | 2020-01-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AND REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. ON (B)(6) 2020 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-DEVICE EVALUATION: LENS RETURNED IN A MIRCO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403868&pc=QCB | 2023826-2020-01723 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTEDLY IS CONSIDERING TO REMOVE THE LENS BECAUSE THE LENS IS "TOO BIG." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: INITIAL MFR REPORT SUBMITTED IN ERROR. PLEASE DISREGARD THE INFORMATION IN THIS REPORT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 - DISREGARD INITIAL MEDWATCH REPORT DUE TO CLAIM IS A DUPLICATE OF ANOTHER CLAIM. MFR# 2023826-2020-01707 FOR CORRECT CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10405558&pc=QCB | 2023826-2020-01753 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: A DAMAGED 12.6MM, TMICL12.6, -14.0/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS RETUNED TO US. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, INSIDE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10406783&pc=QCB | 2023826-2020-01751 | 2020-07-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -10.5/5.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: "PATIENT'S LAST VISIT WAS ON (B)(6) 2020. BCVA 20/20; VAULT 900. PATIENT IS HAPPY. LENS REMAINS IMPLANTED." (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429647&pc=QCB | 2023826-2020-01803 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICMO13.2, -9.5/2/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT LEFT EYE (OS) AS A REPLACEMENT LENS DUE TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM (VAULT IS STILL NOT IDEAL). ON (B)(6) 2020 THE REPORTER STATED " THE PATIENT CONDITION IS GOOD NOW EXCEPT THE VAULT IS STILL LOW: ABOUT 200 MICROMETER. THE CAUSE OF THE UNIDEAL VAULT AFTER THE CHANGED ICL SURGERY MIGHT BE THE STRUCTURE OF THE PATIENT'S EYE IS ABNORMAL. THE DOCTOR HASN'T DONE THE INTERVENTION FOR THE LOW VAULT AND DECIDES TO DO THE FOLLOW-UP CLOSELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429503&pc=QCB | 2023826-2020-01793 | 2019-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440595&pc=QCB | 2023826-2020-01808 | 2020-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/102 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476241&pc=QCB | 2023826-2020-01932 | 2020-06-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.50/+1.5/077(SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY ON (B)(6) 2020, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10480709&pc=QCB | 2023826-2020-01998 | 2020-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -12.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491990&pc=QCB | 2023826-2020-01960 | 2019-08-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/175 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10495978&pc=QCB | 2023826-2020-02014 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-09-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.00/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AND EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. THE REPORTER STATED " PLEASE NOTE THAT WE ARE STILL MONITORING THE PATIENT. NO FURTHER INTERVENTION DONE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10495983&pc=QCB | 2023826-2020-02015 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-09-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020 AND EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. THE REPORTER STATED " PLEASE NOTE THAT WE ARE STILL MONITORING THE PATIENT. NO FURTHER INTERVENTION DONE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912816&pc=QCB | 2023826-2019-01509 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -14.00/2.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- EXPLANT DATE: (B)(6) 2019. H6- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10. IN INITAL MDR CLAIM # (B)(4) IS NOT APPLICABLE, INITIAL MDR IS FOR CLAIM# (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271249&pc=QCB | 2023826-2020-01473 | 2020-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED ON (B)(6) 2020. LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, LENS TO BE EXCHANGED WITH A LONGER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION. B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE . B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTER STATES LOW VAULT; AND THE LENS WAS ROTATED 40 DEGREES 1 AFTER 1 WEEK. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- EXPLANT DATE: (B)(6) 2020. H6- CLINICAL CODE 4582: VAULTING. CORRECTED DATA: H6- RESULT FINDINGS IN INITIAL MDR 114 SHOULD BE CORRECTED TO 213. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 STATEMENT IN INITIAL MDR SHOULD INCLUDE THAT THE PATIENT ALSO EXPERIENCED LOW VISION WITH LOW VAULT. H6-CLINICAL CODE 4581- LOW VISION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652115&pc=QCB | 2023826-2019-00928 | 2018-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THIS DID NOT RESOLVE THE PROBLEM OF LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. REFERENCE MFR REPORT# 2023826-2019-00927 FOR INITIAL IMPLANTED LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON 22/FEB/2019 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608594&pc=QCB | 2023826-2020-00174 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -16.5/5.0/138 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -16.50/5.0/138 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLAMER LENS INTO THE PATIENT RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. D6B- EXPLANT DATE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT WITH DEBRIS ON LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. D6B- EXPLANT DATE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608630&pc=QCB | 2023826-2020-00175 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -14.5/5.0/51 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE BS-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -14.50/5.0/051 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLAMER LENS INTO THE PATIENT LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- EXPLANT DATE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT WITH RESIDUE ON LENS AND LENS IS CURVED. (B)(4). Manufacturer Narrative: CORRECTED DATA: D6B- EXPLANT DATE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782975&pc=QCB | 2023826-2020-00409 | 2020-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -1.0/+1.0/012 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXPLANTED AND ON THE SAME DATE, THE LENS WAS REPLACED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361997&pc=QCB | 2023826-2020-01631 | 2020-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: B5 - LENS BEING REPLACED DURING INITIAL SURGERY SHOULD BE CORRECTED TO LENS WAS EXCHANGED LATER ON THE SAME DATE ON (B)(6) 2020 IN A SECOND SURGERY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392937&pc=QCB | 2023826-2020-01670 | 2020-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+0.5/076 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS TORE DURING IMPLANTATION AND A LARGE PART OF THE LENS REMAINED IN THE EYE. THE PLAN IS TO EXPLANT THE LENS AT A LATER DATE. THE REPORTER INDICATED PART OF THE LENS WAS STUCK IN THE PLUNGER AND THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED ON (B)(6) 2020 AND EXPLANTED ON (B)(6) 2020 DUE TO LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416196&pc=QCB | 2023826-2020-01721 | 2020-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS HAD LOW VAULTING AND REMAINED IMPLANTED. THE PLAN IS TO EXPLANT/EXCHANGE THE LENS AT A LATER DATE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020 AND WAS EXCHANGED FOR A LONGER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8433521&pc=QCB | 2023826-2019-00443 | 2019-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.50/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REPORTEDLY, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM UPDATE - FROM PRODUCT PROBLEM TO ADVERSE EVENT UPDATE TO INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) NOTED AS "REMAINS IMPLANTED" - UPDATE - THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. PREVIOUSLY REPORTED DISTRIBUTOR. CORRECTION, DISTRIBUTOR SHOULD NOT BE REPORTED. PREVIOUSLY NOT REPORTED - CORRECTION, THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475385&pc=QCB | 2023826-2019-00524 | 2019-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/020 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULTING. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED. PATIENT CODE UPDATED TO 3191, NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8479491&pc=QCB | 2023826-2019-00534 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -11.5/3/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT HAS BEEN OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THIS EVENT IS UNKNOWN. IT WAS REPORTED THAT A LENS EXCHANGE IS PLANNED IN ABOUT 2 MONTHS. THE PATIENT IS CURRENTLY FEELING WELL (NO PUPIL OR HIGH IOP ISSUES), CURRENT BCVA IS 20/20. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICTED THAT A 13.7MM, VTICM5_13.7, -11.5/3/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PATIENT IS HAPPY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION,EXCHANGE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8544997&pc=QCB | 2023826-2019-00635 | 2019-03-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2, 5.50/2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED DURING SAME SURGERY DUE TO EXCESSIVE VAULT. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. A LENS REPLACEMENT IS PLANNED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- CAUSE OF THE EVENT HAS BEEN CHANGED TO UNKNOWN. IT HAS BEEN REPORTED THAT "PATIENT DOES NOT WANT TO HAVE ANOTHER SURGERY AT THE MOMENT." CORRECTED DATA: H6- PATIENT CODE 3191: IN ORIGINAL MDR IS NOT APPLICABLE. PATIENT CODE 2692: NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8547366&pc=QCB | 2023826-2019-00643 |  | Injury | STAAR SURGICAL COMPANY | 2019-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Material Split, Cut or Torn | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +7.00/+2.0/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS REMOVED DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE AND A DIFFERENT MODEL BUT SAME LENGTH LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. EXPLANT DATE: (B)(6) 2018. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED DEVICE CODE: 3189 TO 4008 - PREVIOUS CODE NOT APPLICABLE, AN ADVERSE EVENT DID OCCUR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8591938&pc=QCB | 2023826-2019-00779 | 2019-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -9.5/2.5/067 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. AFTER INSERTION, THE LENS "FLIPPED IN THE EYE." THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPLANT DATE: (B)(6) 20019. PATIENT CODE 2692: NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT. PATIENT CODE 3191 IN ORIGINAL MDR IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICTED THAT THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. AFTER INSERTION THE LENS "FLIPPED IN THE EYE". "THE LENS WAS TAKEN OUT AND REPLACED WITH A DIFFERENT ICL." (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8592238&pc=QCB | 2023826-2019-00638 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -13.00/1.5/090 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. REPORTER INDICATED LENS EXCHANGE PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE LENS WAS REMOVED AND EXCHANGED ON (B)(6)B 2019, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. "PATIENT AND DOCTOR ARE HAPPY." ADDED EXPLANT DATE OF (B)(6) 2019. PATIENT CODE - 2692 NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT TO 3191 NO CODE AVAILABLE - SECONDARY SURGERY, LENS EXCHANGE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631512&pc=QCB | 2023826-2019-00848 | 2019-03-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANT SHOULD BE CORRECTED TO PATIENT CODE: 2692 - NO KNOWN IMPACT OR CONSEQUENCE TO PATIENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631532&pc=QCB | 2023826-2019-00856 | 2019-03-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/089 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633183&pc=QCB | 2023826-2019-00887 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/146 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LENS DECENTRATION INFERIORLY AND TEMPORALLY AND LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: "REPORTER INDICATED THAT THE LENS WAS EXPLANTED" SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: EXPLANT DATE: UNK. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: DATE RECEIVED BY MANUFACTURER: 13-SEP-2019 SHOULD BE CORRECTED TO 10-SEP-2019 IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8648029&pc=QCB | 2023826-2019-00966 | 2019-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13. 2 IMPLANTABLE COLLAMER LENS, - 11.00/+1.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE THE LENS FOR A SMALLER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 11-SEP-2019. REPORTEDLY, THE LENS WAS TOO LARGE. ON (B)(6) 2019 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. THE PROBLEM WAS RESOLVED. PATIENT CODE 2692 NO LONGER APPLICABLE DUE TO ADDITIONAL INFO. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736582&pc=QCB | 2023826-2019-01124 | 2019-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/163 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772293&pc=QCB | 2023826-2019-01198 | 2019-06-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.00/1.5/080 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019.THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT CODE: 3191 SHOULD BE CORRECTED TO 2692 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8819695&pc=QCB | 2023826-2019-01302 | 2019-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/101 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD A LOW VAULT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC BENT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8819703&pc=QCB | 2023826-2019-01305 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/055 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED, THE VAULT WAS STILL LOW. THE CAUSE OF THE EVENT WAS UNKNOWN. SEE MFR REPORT # 2023826-2019-01546 FOR THE EXCHANGED LENS. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8883523&pc=QCB | 2023826-2019-01357 | 2019-03-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DURING INITIAL SURGERY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894305&pc=QCB | 2023826-2019-01431 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7 LENS, -12.5/+2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. THE SURGEON STATES THE LENS IS TOO LARGE AND THE VAULT IS TOO HIGH. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION OR UPDATES HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: THE LENS WAS RETURNED TO THE MANUFACTURER AND EVALUATED, THEREFORE LENS HAS BEEN EXPLANTED. NO ADDITIONAL INFORMATION AVAILABLE. IS THIS A SINGLE-USE DEVICE THAT WAS REPROCESSED AND REUSED ON A PATIENT?: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL, INSIDE OF A STERILIZATION PACK. VISUAL INSPECTION FOUND THAT THE THERE WAS A TEAR IN THE HAPTIC. PATIENT CODE UPDATED: 3191 NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894309&pc=QCB | 2023826-2019-01432 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7 LENS, -6.5/+1.5/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. THE SURGEON STATES THE LENS IS TOO LARGE AND THE VAULT IS TOO HIGH. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: THE LENS HAS BEEN RETURNED AND EVALUATED, THEREFORE THE LENS HAS BEEN EXPLANTED. NO ADDITIONAL INFORMATION IS KNOWN. EXPLANT DATE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN A LENS VIAL, IN A STERILIZATION POUCH. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. PATIENT CODE UPDATED: 3191 NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8910141&pc=QCB | 2023826-2019-01447 | 2019-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.50/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTER INDICATED ON (B)(6) 2019 THAT A LENS REMOVAL AND EXCHANGE ARE PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965327&pc=QCB | 2023826-2019-01614 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.00/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. PER REPORTER, "THE PATIENT HAD NO SUBJECTIVE SYMTOMS AND THE LENS OPACITY WAS MILD AND IMPROVED WITH FLUOROMETHOLONE EYE DROPS IN 1 MONTH. CURRENTLY, THE LENS OPACITY REMAINS SLIGHTLY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM: SERIOUS. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT. IMPAIREMENT/DAMAGE (DEVICE). PATIENT CODE 3191: EYE TREATED WITH FLUOROMETHOLONE EYE DROPS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 3189 IS CORRECTED TO 1426 FOR OPACITY ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965332&pc=QCB | 2023826-2019-01613 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. PER THE REPORTER, "THE PATIENT HAD NO SUBJECTIVE SYMPTOMS AND THE LENS OPACITY WAS MILD AND IMPROVED WITH FLUOROMETHOLONE EYE DROPS IN 1 MONTH. CURRENTLY, THE LENS OPACITY REMAINS SLIGHTLY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE: 2692 SHOULD BE CORRECTED TO 3191: NO CODE AVAILABLE (MEDICAL INTERVENTION, TOPICAL STEROIDS). DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1426. PATIENT IDENTIFIER: (B)(6) SHOULD BE CORRECTED TO (B)(6). PRODUCT PROBLEM SHOULD BE CORRECTED TO ADVERSE EVENT. REQUIRED INTERVENTION SHOULD BE MARKED. AT A FOLLOW UP VISIT IT WAS REPORTED THAT THE WHITE ACCRETIONS OF THE PATIENTS GOT SMALLER WITHOUT SPECIFIC MEDICAL TREATMENT. THE DOCTOR THINKS THE RISK OF RECURRENCE WILL BE LOW. REPORTED LENS OPACITY MEANS THAT ICL LENS OPACITY DUE TO WHITE ACCRETIONS AND NOT THE PATIENT'S CATARACT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112033&pc=QCB | 2023826-2019-01777 | 2019-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED BUT DUE TO STABLE CONDITION THE SURGEON DECIDED TO LEAVE THE LENS IN THE EYE AND MONITOR THE PATIENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: "(B)(4)" SHOULD BE CORRECTED TO "(B)(4)" IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#:(B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115525&pc=QCB | 2023826-2019-01730 | 2019-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -15.50/4.5/152 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -15.50/4.5/152 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131806&pc=QCB | 2023826-2019-01759 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+5.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PLAN IS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINED IMPLANTED. Manufacturer Narrative: REC'D BY MFR: CORRECTED TO 26-OCT-2019 IN SUPPLEMENTAL MEDWATCH REPORT #1. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED TO SERIOUS INJURY IN SUPPLEMENTAL MEDWATCH REPORT #2. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131823&pc=QCB | 2023826-2019-01760 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.00/+5.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PLAN IS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINED IMPLANTED. Manufacturer Narrative: TYPE OF REPORTABLE EVENT: CORRECTED TO SERIOUS INJURY IN SUPPLEMENTAL MEDWATCH REPORT #2. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6)2019. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED (B)(4). Manufacturer Narrative: PMA\510K: CORRECTED TO (B)(6) 2019 IN SUPPLEMENTAL MEDWATCH REPORT #1 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9139622&pc=QCB | 2023826-2019-01818 | 2019-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -07.50/3.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS REMAINS IMPLANTED. AN ADDITION OF A NEW PI WAS PERFORMED. REPORTEDLY THE PATIENT IS "ASYMPTOMATIC." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CASE/ VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AS WELL AS THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION ON 19/SEP/2019 THE LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT PROBLEM CORRECTED TO ADVERSE EVENT IN INITIAL MDR. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. EXPLANT DATE: (B)(6) 2019 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. PATIENT CODE 2692 CORRECTED TO PATIENT CODE 3191 IN INITIAL MDR. PATIENT CODE 3191 CORRECTED TO 'SECONDARY SURGICAL INTERVENTION: SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, SECONDARY YAG' IN INITIAL MDR. 'LENS EXCHANGE' PREVIOUSLY STATED IS NOT APPLICABLE. PER NEW INFORMATION RECEIVED, THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS REPORTEDLY HAPPY AND SATISFIED WITH THE FINAL RESULT. PATIENT CODE 3191 - SECONDARY SURGICAL INTERVENTION: LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147370&pc=QCB | 2023826-2019-01800 | 2017-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/3.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT DESCRIPTION: "BUT DID NOT RESOLVE THE PROBLEM" SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT, AS IT IS NOT APPLICABLE. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS WAS IMPLANTED ON (B)(6)2017. LENS WAS REPOSITIONED ON (B)(6)2017 AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE REPORT STATED "PATIENT IS HAPPY". CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147383&pc=QCB | 2023826-2019-01801 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.00/2.0/057 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED . LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: "PRODUCT PROBLEM" SHOULD BE CORRECTED TO "ADVERSE EVENT" IN THE INITIAL MDR. "REQUIRED INTERVENTION" SHOULD BE MARKED IN THE INITIAL MDR. "LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED." SHOULD BE REMOVED FROM THE INITIAL MDR. PATIENT CODE 2692 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION; REPOSITIONING IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147400&pc=QCB | 2023826-2019-01802 | 2018-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT: LENS REPOSITIONING RESOLVED THE PROBLEM. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS WAS IMPLANTED ON (B)(6) 2018. THE LENS WAS REPOSITIONED ON (B)(6) 2018. POSTOPERATIVE REPORT STATED "PATIENT IS HAPPY". CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147401&pc=QCB | 2023826-2019-01803 | 2017-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -7.00/2.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: "PRODUCT PROBLEM" SHOULD BE CORRECTED TO "ADVERSE EVENT" IN INITIAL MDR. "REQUIRED INTERVENTION" SHOULD BE CHECKED IN THE INITIAL MDR. "BUT DID NOT RESOLVE THE PROBLEM" SHOULD BE REMOVED FROM THE INITIAL MDR. PATIENT CODE: 2692 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO "3191: SECONDARY SURGICAL INTERVENTION; REPOSITIONING" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B3- (B)(6) 2017. B5- " LENS WAS REPOSITIONED ON (B)(6) 2017 AND PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY.". D6- (B)(6) 2017. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION: LENS REPOSITIONING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9226358&pc=QCB | 2023826-2019-02019 | 2019-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -7.50/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019 AND LOW VAULT WAS OBSERVED. PATIENTS CURRENT BCVA WAS REPORTED TO BE 20/20 AND PATIENT IS DOING WELL. LENS REMAINS IMPLANTED. THE SURGEON PLANS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: "LENS REMAINS IMPLANTED. THE SURGEON PLANS TO EXCHANGE THE LENS WITH A LONGER LENS." SHOULD BE REMOVED AND CORRECTED TO "THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM." IN THE INITIAL MDR. TYPE OF REPORTABLE EVENT: "MALFUNCTION" SHOULD BE CORRECTED TO "SERIOUS". (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: FOLLOW-UP TYPE: "CORRECTION" SHOULD BE CORRECTED TO "ADDITIONAL INFORMATION" IN SUPPLEMENTAL MEDWATCH REPORT #1. "CORRECTED DATA" SHOULD BE CORRECTED TO "ADDITIONAL MANUFACTURER NARRATIVE" IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9227249&pc=QCB | 2023826-2019-02022 | 2019-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -10.5/3/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) AND LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED, LENS EXCHANGE IS PLANNED FOR NOVEMBER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PER UPDATED INFORMATION, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE). (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANTED:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9246813&pc=QCB | 2023826-2019-02029 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+0.5/163 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT. THE LENS CURRENTLY REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM IS RESOLVED. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). (B)(4). Manufacturer Narrative: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; REPOSITIONING IS NOT APPLICABLE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9271546&pc=QCB | 2023826-2019-02040 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/5.0/091 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY THE SURGEON PLANS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/5.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE REPORTER STATED "ALL IS FINE- PATIENT IS HAPPY!" H6- PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9272143&pc=QCB | 2023826-2019-02041 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -7.50/6.0/086 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY THE SURGEON PLANS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL DEBRIS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -7.50/6.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. REPORTER STATES "PATIENT IS HAPPY." H6: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9278970&pc=QCB | 2023826-2019-01991 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A LENS INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. SURGEON NOTICED "TRANS ILLUMINATION DEFECTS, DEPIGMENTATION, AND EYE BETWEEN 3-9 O'CLOCK." REPORTEDLY, IOP WAS NOTED AS REASONABLE BUT SURGEON IS NOTICING MID DILATED, POORLY REACTIVE PUPILS. SURGEON STATES "THE VAULT IS PERHAPS AROUND 1000 MICRONS. REPORTER INDICATED "PATIENT SAYS VISION IS VERY GOOD." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: G4- "07-OCT-2019" SHOULD BE CORRECTED TO "06-OCT-2019" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B4- "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H1- "SERIOUS INJURY" SHOULD BE CORRECTED TO "MALFUNCTION" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK. MODEL #: UNK. DEVICE MANUFACTURER DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9278981&pc=QCB | 2023826-2019-01992 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A LENS INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2019. SURGEON NOTICED "TRANS ILLUMINATION DEFECTS AND DEPIGMENTATION, EYE BETWEEN 8-12 O'CLOCK." IOP WAS NOTED AS REASONABLE BUT SURGEON IS NOTICING MID DILATED AND POORLY REACTIVE PUPILS. SURGEON STATES "THE VAULT IS PERHAPS AROUND 1000 MICRONS. REPORTEDLY "PATIENT SAYS VISION IS VERY GOOD." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE: 2692 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B4 - DATE OF THIS REPORT: SHOULD BE CORRECTED TO (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK. MODEL #: UNK. DEVICE MANUFACTURER DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441174&pc=QCB | 2023826-2019-02353 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+6.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY THE LENS IS "SMALL AND TURNING." IT IS UNKNOWN AS TO WHETHER THIS LENS REMAINS IMPLANTED OR NOT; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+6.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY THE LENS IS "SMALL AND TURNING." IT IS UNKNOWN AS TO WHETHER THIS LENS HAS BEEN EXPLANTED OR NOT; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. H6: PATIENT CODE 2692 CORRECTED TO PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION (LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441177&pc=QCB | 2023826-2019-02354 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+6.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY THE LENS IS "SMALL AND TURNING." IT IS UNKNOWN AS TO WHETHER THIS LENS REMAINS IMPLANTED OR NOT; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5 - CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+6.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY THE LENS IS "SMALL AND TURNING. IT IS UNKNOWN AS TO WHETHER THIS LENS HAS BEEN EXPLANTED OR NOT; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. H6 - PATIENT CODE 2692 CORRECTED TO PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION (LENS EXPLANT). CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9483557&pc=QCB | 2023826-2019-02395 | 2019-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.50/5.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY "PATIENT NEEDS EXPLANT DUE TO EXCESS VAULTING AT AROUND 1400UM." PER THE REPORTER ON (B)(6) 2019 PATIENT "IS CURRENTLY EMMETROPIC." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B4 - CORRECTED TO 17-NOV-2019 IN INITIAL MDR. H6 - RESULT CODE 114 CORRECTED TO RESULT CODE 213 IN INITIAL MDR. (B)(4). Manufacturer Narrative: G4 - CORRECTED TO 17-NOV-2019 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5: UPDATED TO: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.50/5.5/089 (SPHERE / CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT OF 1500UM, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PERIPHERAL ANTERIOR SYNECHIAE, AND IRIDOCORNEAL CONTACT WERE REPORTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE EYE IS "HEALTHY." THE CAUSE OF THIS EVENT IS REPORTED AS THE DEVICE. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGICAL INTERVENTION - LENS EXCHANGE). H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES; PERIPHERAL ANTERIOR SYNECHIAE, IRIDOCORNEAL CONTACT). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9483562&pc=QCB | 2023826-2019-02396 | 2019-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -3.00/5.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY "PATIENT NEEDS EXPLANT DUE TO EXCESS VAULTING AT AROUND 1400UM." PER THE REPORTER ON (B)(6) 2019 PATIENT "IS CURRENTLY EMMETROPIC." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B4 - CORRECTED TO 17-NOV-2019 IN INITIAL MDR. H4 - CORRECTED TO 23-SEP-2016 IN INITIAL MDR. H6 - RESULT CODE 114 CORRECTED TO RESULT CODE 213 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: G4 - CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: UPDATED TO: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -3.00/5.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULT OF 1500UM, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PERIPHERAL ANTERIOR SYNECHIAE, AND IRIDOCORNEAL CONTACT WERE REPORTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE EYE IS "HEALTHY." THE CAUSE OF THIS EVENT IS REPORTED AS THE DEVICE. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGICAL INTERVENTION - LENS EXCHANGE) PATIENT CODE 3191: NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES; PERIPHERAL ANTERIOR SYNECHIAE, IRIDOCORNEAL CONTACT). CLAIM# (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528009&pc=QCB | 2023826-2019-02496 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 12.50/+1.5/029 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING HIGH VAULTING AND WILL BE EXPLANTED AT A LATER DATE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO UNKNOWN/PATIENT RELATED FACTOR. POST-OP PATIENT CONDITION - IMPROVED VA/IMPROVE PHOTOPHOBIA. THE LENS WAS DISCARDED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED H6 - PATIENT CODE: 3191 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528016&pc=QCB | 2023826-2019-02497 | 2019-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 13.50/+4.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING HIGH VAULTING AND WILL BE EXPLANTED/EXCHANGED FOR A SHORTER LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6: DEVICE CODE: 1583 SHOULD BE CORRECTED TO 3189. B5: LENS IS REPORTED TO BE EXCHANGED WITH A SHORTER LENGTH LENS, NO REASON FOR EXCHANGE WAS REPORTED. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE REPORT INDICATED "PATIENT IS HAPPY" AND "PERFECT, IMPROVED VISION OVER PRE-OP". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE LENS WAS REPORTED AS HAVING HIGH VAULTING SHOULD BE CORRECTED TO THE LENS WAS REPORTED TO BE TOO LONG. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9595583&pc=QCB | 2023826-2020-00123 | 2019-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/1.5/033 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: DATE OF EVENT IS (B)(6) 2020. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/1.5/033 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO LOW VUALT. THE EXCHANGE RESOLVED THE PROBLEM. DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HATIC TORN AND DEBRIS ON THE LENS. PATIENT CODE:3191 SECONDARY SURGERY (EXCHANGE). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9664304&pc=QCB | 2023826-2020-00232 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). PER THE REPORTER ON (B)(6) 2020, "LENS NEEDS TO BE EXCHANGED FOR A LARGER SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1, -14.50/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2019, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "TOO SMALL LENS SIZE.". H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE, REPOSITIONING. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - DATE RECEIVED BY MANUFACTURER: 31/JAN/2020 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. H6 - DEVICE CODE: 2923 SHOULD HAVE BEEN INCLUDED. D10 - RETURNED TO MANUFACTURER: 12/FEB/2020 SHOULD BE CORRECTED TO 13/FEB/2020. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B5/H10 - INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019 IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9684571&pc=QCB | 2023826-2020-00264 | 2019-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT IS HAPPY. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED CORRECTED DATA: B3: DATE OF EVENT: (B)(6) 2019 CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689269&pc=QCB | 2023826-2020-00283 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -7.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS EXCHANGE TO OCCUR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON 26-FEB-2020, THE SURGEON OPERATED ON THE LEFT EYE TO PERFORM ICL EXCHANGE." H6- PATIENT CODE 3191: LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4- "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020" IN THE INITIAL MDR. G4-"16-JAN-2020" SHOULD BE CORRECTED TO "15-JAN-2020" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- ADVERSE EVENT SHOULD BE MARKED IN MDR SUPPLEMENTAL #2. B2- REQUIRED INTERVENTION SHOULD BE MARKED IN MDR SUPPLEMENTAL #2. D7- EXPLANT DATE: (B)(6) 2020 SHOULD BE IN MDR SUPPLEMENTAL #2. H1- SERIOUS INJURY SHOULD BE IN MDR SUPPLEMENTAL #2. CLAIM#(B)(4). Manufacturer Narrative: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717739&pc=QCB | 2023826-2020-00329 | 2020-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -1.00/2.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT THE INITIAL LENS WAS IMPLANTED ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SMALLER LENGTH LENS AND THE PROBLEM RESOLVED. "PATIENT IS HAPPY". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - RESULT CODE: 114 SHOULD BE CORRECTED TO 213, IN INITIAL MEDWATCH REPORT. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717740&pc=QCB | 2023826-2020-00330 | 2020-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -1.50/1.50/149 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 - RESULT CODE: 114 SHOULD BE CORRECTED TO 213, IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE INITIAL LENS WAS IMPLANTED ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, AND A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S EYE BUT WAS IMPLANTED UPSIDE DOWN. DURING REMOVAL OF THE LENS, THE LENS TORE. THIS OCCURRED INTRAOPERATIVELY. ANOTHER SHORTER LENGTH REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9811296&pc=QCB | 2023826-2020-00447 | 2020-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/+3.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/+3.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9811419&pc=QCB | 2023826-2020-00446 | 2020-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/+3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/+3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814229&pc=QCB | 2023826-2020-00471 | 2020-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.50/+0.5/042 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814241&pc=QCB | 2023826-2020-00470 | 2020-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.00/+0.5/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9855436&pc=QCB | 2023826-2020-00536 | 2020-02-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.50/1.0/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REPORTEDLY "PATIENT IS UNHAPPY WITH HER VISION." PATIENT WAS GIVEN AN RX OF ALPHAGAN P OU QHS. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA:) B5- IN INITIAL MDR STATEMENT "THE PATIENT WAS GIVEN AN RX FOR ALPHAGAN" IS NOT APPLICABLE. THE "RX" IS REFERENCE TO EYEGLASS PRESCRIPTION PROVIDED BY THE OPTOMETRIST. PATIENT IS CONCURRENTLY TAKING ALPHAGAN OU QHS. H6: PATIENT CODE 3191: RESIDUAL REFRACTION ERROR. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880942&pc=QCB | 2023826-2020-00665 | 2020-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+4.0/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND REMAINS IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020. THE LENS EXCHANGE RESOLVED THE PROBLEM. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - "-13.5/+4.0/102 (SPHERE/CYLINDER/AXIS)" CORRECTED TO "-13.5/+4.0/107 (SPHERE/CYLINDER/AXIS)". H6 - DEVICE CODE 4117 IS NO LONGER APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880973&pc=QCB | 2023826-2020-00666 | 2020-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/+1.5/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE REPORTER INDICATED THE LENS WAS IMPLANTED INTO THE WRONG EYE, THE LENS FOR THE LEFT EYE (OS) WAS IMPLANTED INTO THE RIGHT EYE (OD). THE LENS WILL BE EXPLANTED AND TO DATE REMAINS IMPLANTED. CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE 2692 UPDATED TO PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION - LENS EXCHANGE). H6 - DEVICE CODE 3189 UPDATED TO DEVICE CODE 2993. CLAIM# (B)(4). Manufacturer Narrative: B5 - 'THE LENS WAS EXCHANGED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED' UPDATED TO: 'THE LENS WAS EXPLANTED FROM THE OD ON (B)(6) 2020 AND THE CORRECT LENS WAS IMPLANTED INTO THE OS. THE PROBLEM WAS RESOLVED.' CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889401&pc=QCB | 2023826-2020-00585 | 2019-11-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -08.50/0.5/144 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC DEFORMED . CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA : H6-PATIENT CODE 3191 IN INITIAL MDR IS NOT APPLICABLE AND IS CORRECTED TO 2692. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9937637&pc=QCB | 2023826-2020-00707 | 2019-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - 'SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS ALSO REPORTED.' SHOULD BE ADDED TO INITIAL MDR. H6 - PATIENT CODE 2692 CORRECT TO PATIENT CODE 3191 (NO CODE AVAILABLE - SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES) IN INITIAL MDR. H6 - DEVICE CODE 1494 (OFF-LABEL USE - AGE < 21 YEARS OLD) SHOULD BE ADDED TO INITIAL MDR. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938269&pc=QCB | 2023826-2020-00815 | 2019-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS HAD LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/2.5/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D7- (B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. METHOD CODE 4117 IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10057316&pc=QCB | 2023826-2020-01071 | 2020-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.00/1.0/044 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, "SURGEON IMPLANTED LENS INTO THE WRONG EYE" OF PATIENT. "EXCHANGE IS PLANNED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE). B5- ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D7-EXPLANTED DATE: (B)(6) 2020. H6-PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10112279&pc=QCB | 2023826-2020-01151 | 2020-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, WAS RECEIVED DAMAGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/1.5/089 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. NAME AND ADDRESS: (B)(6). TELEPHONE: (B)(6). CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - CORRECTION TO PREVIOUS LENS EVALUATION- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8452984&pc=QCB | 2023826-2019-00483 | 2019-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-26 | QCB | IMPLANTABLE COLLAMER LENS | Ejection Problem | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE FOOTPLATES OF THE LENS GOT STUCK IN THE INJECTOR. THERE WAS NO PATIENT CONTACT WITH THE LENS. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED; "ALL WENT WELL WITH PATIENT INTERACTIONS". Manufacturer Narrative: LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8453013&pc=QCB | 2023826-2019-00487 | 2019-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-26 | QCB | IMPLANTABLE COLLAMER LENS | Ejection Problem | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE FOOTPLATES OF THE LENS GOT STUCK IN THE INJECTOR. THERE WAS NO PATIENT CONTACT WITH THE LENS. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED; "ALL WENT WELL WITH PATIENT INTERACTIONS". Manufacturer Narrative: LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8480532&pc=QCB | 2023826-2019-00511 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-04-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.50/+2.5/090 (SPHERE/CYLINDER/AXIS), WAS DAMAGED DURING LOADING AND THERE WAS NO PATIENT CONTACT OR INJURY. THE BACKUP LENS WAS USED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE LENS. THE REPORTER DID NOT PROVIDE ANY ADDITIONAL INFORMATION. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8569319&pc=QCB | 2023826-2019-00723 | 2019-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT 13.2MM, VTICM5_13.2, -13.00/1.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTIFUGE VIAL. CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692893&pc=QCB | 2023826-2019-01062 | 2019-05-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.0/+4.0/070 (SPHERE/CYLINDER/AXIS) IT TORE. THIS OCCURRED ON (B)(6) 2019 AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. A BACK UP LENS WAS THEN USED. THE SURGEON DOES NOT KNOW THE CAUSE OF THE EVENT. Manufacturer Narrative: CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CORRECTED TO DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN THE CARTRIDGE. VISUAL INSPECTION OF THE LENS FOUND THE LENS RETURNED IN THE CARTRIDGE. THE CARTRIDGE WAS RETURNED IN THE LENS CASE WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION OF THE CARTRIDGE FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE CARTRIDGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN THE CARTRIDGE AND IN THE LENS CASE. THERE WAS CLEAR SURGICAL RESIDUE ON THE CARTRIDGE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692895&pc=QCB | 2023826-2019-01061 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/108 (SPHERE/CYLINDER/AXIS) IT BROKE. THIS OCCURRED ON (B)(6) 2019 AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. THE SURGEON DOES NOT KNOW THE CAUSE OF THE EVENT. Manufacturer Narrative: DEVICE EVALUATION: "RESIDUE ON THE LENS" CORRECTED TO "RESIDUE ON THE CARTRIDGE". Manufacturer Narrative: DATE OF REPORT CORRECTED TO 15-MAY-2019 IN INITIAL MDR. DATE RECEIVED BY MANUFACTURER CORRECTED TO 15-MAY-2019 IN INITIAL MDR. CORRECTED TO - DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN THE CARTRIDGE. VISUAL INSPECTION OF THE LENS FOUND THE LENS STUCK IN THE CARTRIDGE. THE CARTRIDGE WAS RETURNED IN THE LENS CASE WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION OF THE CARTRIDGE FOUND NO VISIBLE DAMAGE TO THE CARTRIDGE AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN THE CARTRIDGE AND IN THE LENS CASE. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8841796&pc=QCB | 2023826-2019-01371 | 2019-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.0/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE WHILE LOADING. THERE WAS NO PATIENT CONTACT. ANOTHER LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TEAR IN ONE OF THE HAPTIC. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895274&pc=QCB | 2023826-2019-01422 | 2019-07-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED HE NOTICED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/091 (SPHERE/CYLINDER/AXIS) WAS TORN/BROKEN WHEN OPENING THE VIAL AND GRABBING THE LENS WITH FORCEPS. THERE WAS NO PATIENT CONTACT WITH THE LENS. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE LENS OPTIC AND ONE HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: UDI # CORRECTED FROM UNK TO N/A. INITIAL REPORTER NAME AND ADDRESS: INCORRECT DOCTOR NAME REPORTED IN INITIAL MDR CORRECTED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANTED: NA. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8902740&pc=QCB | 2023826-2019-01460 | 2019-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS NOTED TO BE TORN IN THE INJECTOR. THERE WAS NO PATIENT CONTACT OF INJURY AND A BACKUP LENS WAS IMPLANTED. A STAAR INJECTOR SYSTEM WAS USED BUT MODEL AND LOT NUMBERS ARE UNKNOWN. THE REPORTER STATED " I THINK THE LENS TEAR WAS DUE TO THE FOAM TIP PLUNGER NOT SEATING TIGHTLY INTO THE INJECTOR". Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912823&pc=QCB | 2023826-2019-01502 | 2019-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AFTER OPENING THE VIAL AND PLACING ONTO THE CARTRIDGE, THE SURGEON NOTED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/069 (SPHERE/CYLINDER/AXIS) HAD A "SCRATCH" ON IT. THIS OCCURRED ON (B)(6) 2019 AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. A BACK UP LENS WAS THEN USED. Manufacturer Narrative: LENS WORK ORDER SEARCH - NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CARTRIDGE CASE. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO PATIENT INVOLVEMENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962472&pc=QCB | 2023826-2019-01586 | 2019-08-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 8.50/+2.5/087 (SPHERE/CYLINDER/AXIS), TORE DURING THE LOADING PROCESS, WHILE BEING INSERTED INTO THE CARTRIDGE. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE STATEMENT IS CORRECTED TO; " LENS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CARTRIDGE RETURNED IN LENS BOX WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO DEVICE AND RESIDUE ON CARTRIDGE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. A CARTRIDGE WAS RETURNED AND THERE WAS NO VISIBLE DAMAGE TO THE DEVICE. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9176558&pc=QCB | 2023826-2019-01931 | 2019-07-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE DOCTOR LOADED A 12.6MM VTICMO12.6, -15.00/1.0/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS ACCORDING TO THE INSTRUCTION MANUAL, BUT FOUND LENS DAMAGED WHEN DELIVERING THE LENS INTO THE NOZZLE. THERE WAS NO PATIENT CONTACT. SURGEON DISCONTINUED THE SURGERY BECAUSE HE HAD NO ALTERNATIVE LENS AND PLANS TO IMPLANT A NEW LENS IN THE FUTURE. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS AND TWO PIECES OF THE LENS WERE RETURNED. MSI-PF INJECTOR WAS RETURNED AND DEVICE EVALUATION FOUND CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO DEVICE WITH RESIDUE ON THE INJECTOR, CARTRIDGE WAS NOT RETURNED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9257733&pc=QCB | 2023826-2019-01963 | 2019-09-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/081 (SPHERE/CYLINDER/AXIS), WAS DAMAGED DURING LOADING. THERE WAS NO PATIENT CONTACT. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK. IMPLANTED: NA. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9338152&pc=QCB | 2023826-2019-02180 | 2019-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/2.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. THERE WAS NO PATIENT CONTACT. THE REPORTER STATED THAT THE DAMAGE WAS NOTED DURING PULLING TO THE TIP OF THE CARTRIDGE. RUSHED LOADING WAS REPORTED FOR CAUSE OF EVENT, THE REPORTER STATED, "THE DOCTOR IS EXPERIENCED AND ITS NOT COMMON FOR HER TO MESS UP THE LOADING OF LENS". ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED, AND IT RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9386207&pc=QCB | 2023826-2019-02243 | 2019-07-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON OPENED THE VIAL OF A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/097 (SPHERE/CYLINDER/AXIS), AND NOTED THE LENS WAS DAMAGED BEFORE REMOVING THE LENS FROM THE VIAL. THERE WAS NO PATIENT CONTACT AND THE SURGERY WAS NOT STARTED. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANTED: NA. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498043&pc=QCB | 2023826-2019-02406 | 2019-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE NOZZLE AND THE LENS WAS DAMAGED. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9577753&pc=QCB | 2023826-2020-00084 | 2019-12-19 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 10.00/+4.0/091 (SPHERE/CYLINDER/AXIS), WAS TORN WHILE LOADING AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9612926&pc=QCB | 2023826-2020-00149 | 2019-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS LOADING A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/092 (SPHERE/CYLINDER/AXIS) IT TORE. THIS OCCURRED ON (B)(6) 2019 AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND A BACK-UP LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9684576&pc=QCB | 2023826-2020-00260 | 2019-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; Manufacturing, Packaging or Shipping Problem | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.00/+1.5/082 (SPHERE/CYLINDER/AXIS), WAS REMOVED FROM THE VIAL AND WAS NOTED TO BE OPAQUE. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE SURGEON IMPLANTED A LONGER LENS THAT WAS AVAILABLE, AND THE PROBLEM WAS RESOLVED. THE PATIENTS EYE IS STABLE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: "THE REPORTER STATED THAT WHEN THE SURGEON OPENED THE LENS HE SAW SOMETHING LIKE BUBBLE ON THE LENS SURFACE." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: LENS WAS RETURNED IN LIQUID WITH CLEAR SURGICAL RESIDUE AND MOISTURE IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- DEVICE HISTORY REVIEW- (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H6- DEVICE CODE: 2975 IS NOT APPLICABLE IN PREVIOUS MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706459&pc=QCB | 2023826-2020-00302 | 2019-12-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/1.5/068 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING LOADING PROCEDURE. THERE WAS NO PATIENT CONTACT WITH THE DEVICE. THE REPORTER STATED " DOCTOR ALWAYS PREPARE THE LOADING OF THE LENS BEFORE SHE OPENS THE EYE." IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER STATED THAT THE LENS WAS DAMAGED DURING LOADING. THE SURGEON DAMAGED THE LENS WITH THE FORCEPS. H6- CONCLUSION CODE: 4315 TO 61. CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706706&pc=QCB | 2023826-2020-00301 | 2020-01-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE NURSE NOTICED THAT A 12.1MM, VTICMO12.1,-12.5/+1.5/132 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS TORN WHEN SHE OPENED THE LENS VIAL. THE LENS WAS NOT USED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALAUTION: LENS WAS IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9722922&pc=QCB | 2023826-2020-00336 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.0/2.5/064 (SPHERE/CYLUNDER/AXIS), IMPLANTABLE COLLAMER LENS TORE WHILE LOADING. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781716&pc=QCB | 2023826-2020-00395 | 2019-10-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE ATTEMPTING TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), IT WAS DISCOVERED TO BE TORN. REPORTEDLY, THE LENS TORE DURING LOADING. THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS. ON (B)(6) 2019 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4).. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781725&pc=QCB | 2023826-2020-00402 | 2019-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/072 (SPHERE/CYLINDER/AXIS) FROM THE VIAL, IT WAS DISCOVERED THAT THE LENS WAS TORN/BROKEN. THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE LENS DID NOT TOUCH THE PATIENT'S EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. RETURNED PRODUCT EVALUATION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819167&pc=QCB | 2023826-2020-00450 | 2020-02-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO INJECT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+0.5/125 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE, IT WAS DISCOVERED THAT THE LENS WAS TORN/BROKEN. THIS OCCURRED ON (B)(6) 2020. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE??" THE LENS DID NOT TOUCH THE PATIENT'S EYE AND THERE WAS NO PATIENT INJURY. REPORTEDLY, THE LENS WAS STUCK TO THE PLUNGER. THE SURGEON TRIED TO REMOVE AND RELOAD BUT IT WAS "CUT." AN ALTERNATE LENS WAS SUCCESSFULLY PLANTED ON (B)(6) 2020. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: CORRECTED DATA : D10- LENS RETURNED TO MANUFACTURER IN PREVIOUS MDR SHOULD BE 02-APR-2020 CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880944&pc=QCB | 2023826-2020-00654 | 2020-03-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Deformation | No Patient Involvement | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON NOTED THE EDGE OF A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 7.00/+1.5/092 (SPHERE/CYLINDER/AXIS), WAS NOT SMOOTH AND DID NOT USE. THERE WAS NO PATIENT CONTACT AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THE LENS WAS VIEWED UNDER 10X, THE TORN HAPTIC WAS NOTED BUT NO OTHER OBSERVATIONS. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910965&pc=QCB | 2023826-2020-00680 | 2019-12-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE ATTEMPTING TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), IT WAS TORN. THIS OCCURRED ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT WITH THE LENS. ON (B)(6) 2019 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9924474&pc=QCB | 2023826-2020-00835 | 2020-03-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.0/1.0/062 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS FOUND DAMAGE BEFORE/DURING LOADING THE LENS. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965164&pc=QCB | 2023826-2020-00896 | 2020-02-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO LOAD A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/087 (SPHERE/CYLINDER/AXIS), INTO THE CARTRIDGE AND IT WAS NOTED THE LENS WAS DAMAGED. THE SURGEON FELT THE LENS WAS TORN DURING MANIPULATING USING THE FORCEP. THERE WAS NO PATIENT CONTACT. THE LENS WAS EXCHANGED FOR THE BACKUP LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988428&pc=QCB | 2023826-2020-00952 | 2020-03-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/006 (SPHERE/CYLINDER/AXIS), AND FOUND THE LENS DAMAGED. THERE WAS NO PATIENT CONTACT. THE SURGEON IMPLANTED A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5. IT WAS REPORTED CAUSE OF THE EVENT WAS USER ERROR. SHOULD HAVE BEEN INCLUDED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116211&pc=QCB | 2023826-2020-01158 | 2020-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+1.5/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS BECAME STUCK IN THE CARTRIDGE/INJECTION SYSTEM. THIS OCCURRED ON (B)(6) 2020. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED INTRAOPERATIVELY AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT CONTACT WITH THE LENS AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135940&pc=QCB | 2023826-2020-01168 | 2020-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Integrity Problem; No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON FOUND A "LINE" ON THE LENS SURFACE OF A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/081 (SPHERE/CYLINDER/AXIS). THE SURGEON LOADED THE LENS BUT WAS WORRIED ABOUT THE SURGERY EFFECT AND VISUAL QUALITY OF THE LENS, SO PUT LENS BACK IN THE VIAL. THE SURGEON NOTED THE LENS WAS CRACKED IN CENTER OF LENS. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE REPLACEMENT LENS WAS IMPLANTED THE NEXT DAY AND THE PROBLEM WAS RESOLVED. THE PATIENTS CURRENT CONDITION IS GREAT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189346&pc=QCB | 2023826-2020-01280 | 2019-06-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/098 (SPHERE/CYLINDER/AXIS), IN THE CARTRIDGE AND THE LENS TORE. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242559&pc=QCB | 2023826-2020-01433 | 2020-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED WHILE REMOVING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/110 (SPHERE/CYLINDER/AXIS), FROM THE VIAL, IT WAS NOTED THE LENS WAS DAMAGED/SCRATCHES. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE REPORTER INDICATED THE LENS WAS RECEIVED IN GOOD CONDITION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10112184&pc=QCB | 2023826-2020-01149 | 2018-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/089 (SPHERE/CYLINDER/AXIS), WAS TORN DURING INJECTION INTO THE PATIENT'S LEFT EYE (OS) AND THERE WAS NO PATIENT CONTACT OR INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS, ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/089 (SPHERE/CYLINDER/AXIS) GOT STUCK IN THE CARTRIDGE OR INJECTION SYSTEM ON 26/04/2018. THERE WAS NO PATIENT CONTACT OR INJURY. REPORTEDLY "THE LENS WAS DAMAGED DURING THE INJECTION." ON (B)(6) 2018 A REPLACEMENT LENS OF THE SAME MODEL AND LENGTH WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8574021&pc=QCB | 2023826-2019-00722 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.0/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS FOUND TO BE BROKEN AND WITHOUT ASTIGMATISM LINES. NO PATIENT CONTACT. Manufacturer Narrative: THE DAMAGE (BREAK) WAS NOTICED WHEN THE LENS WAS REMOVED FROM THE LENS VIAL AND LOADED. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A TRANSPARENT BAG. VISUAL INSPECTION FOUND THAT A HAPTIC WAS BROKEN AND CONFIRMED THAT THE LENS HAS ASTIGMATISM MARKERS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966509&pc=QCB | 2023826-2019-01628 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: ON (B)(6) 2019 A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.0/+1.0/074 (SPHERE/CYLINDER/AXIS) WAS RETURNED TO STAAR SURGICAL. IT WAS RETURNED STILL INSIDE OF THE LENS CARTRIDGE AND THE BOX HAD "RESIDUE ON LENS" WRITTEN ON IT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: EVENT UPDATED INFORMATION: ON 06-SEP-2019 IT WAS DISCOVERED THAT THIS EVENT IS NOT MDR-REPORTABLE. HOWEVER, THE EVENT HAD ALREADY BEEN REPORTED. REPORTER STATES THAT THE SURGEON BELIEVES THE FORCEPS USED TO GRASP THE LENS HAD RUST ON THEM. THE SURGEON THEN DID NOT WANT TO RISK USING THE LENS. NO PATIENT CONTACT OCCURRED. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED STILL INSIDE THE CARTRIDGE, INSIDE THE LENS BOX. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE ON THE LENS. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9597959&pc=QCB | 2023826-2020-00129 | 2019-11-14 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -08.50/2.0/105 (SPHERE/CYLINDER/AXIS), INTRAOCULAR LENS TORE DURING LOADING ON (B)(6) 2019. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10225159&pc=QCB | 2023826-2020-01385 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Patient Involvement; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: A LENS TEAR/BREAK BEFORE LOADING ON A 12.6MM, VTICMO12.6, -13.5/3.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS REPORTED. IT WAS REPORTED THAT THE DAMAGE WAS NOTED WHILE REMOVING FROM VIAL. THE LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REMOVED WITHIN THE SAME SURGERY. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT IS UNKNOWN . ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DAT: B4-DATE OF THIS REPORT IN INITIAL MDR SHOULD BE CORRECTED TO (B)(6) 2020 G4- DATE RECEIVED BY MANUFACTURER IN INITIAL MDR SHOULD BE CORRECTED TO (B)(6) 2020 CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : B5- THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.5/3.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE WHILE REMOVING FROM VIAL ON (B)(6) 2020. THE LENS WAS NOT IMPLANTED. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. D6-IMPLANT DATE (B)(6) 2020 IN INITIAL MDR IS NOT APPLICABLE. D7-EXPLANTED DATE (B)(6) 2020 IN INITIAL MDR IS NOT APPLICABLE. CORRECTED DATA H6-PATIENT CODE 2692 IN INITIAL MDR IS CORRECTED TO 2645. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816879&pc=QCB | 2023826-2019-01218 | 2018-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Not Applicable | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -12.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816886&pc=QCB | 2023826-2019-01217 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Not Applicable | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -14.0/2.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816888&pc=QCB | 2023826-2019-01222 | 2019-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Not Applicable | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -9.0/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816887&pc=QCB | 2023826-2019-01219 | 2019-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Not Applicable; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -8.0/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT CODE:3189 SHOULD BE CORRECTED TO 3191 IN INITAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: FOAM TIP PLUNGER MODEL-FTP; LOT#-UNK SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INLCUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8816890&pc=QCB | 2023826-2019-01221 | 2019-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Not Applicable; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -7.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT CODE: 3189 SHOULD BE CORRECTED TO 3191 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: FOAM TIP PLUNGER MODEL-FTP; LOT#-UNK SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. PATIENT CODE 1494: OFF LABEL-USE (PATIENT'S ACD <3.00MM) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073178&pc=QCB | 2023826-2021-02032 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/+2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT IS EXPERIENCING PAIN OVER THE INNER EYE FOR NEARLY 10 MONTHS. " UNEVENTFUL SURGERY, GOOD VISION, GOOD POSITIONING AND GOOD VAULTING, CLEAR CORNEA, CLEAN LENS, NO PIGMENT DISPERSION" WAS REPORTED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073179&pc=QCB | 2023826-2021-02031 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/+2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT IS EXPERIENCING PAIN OVER THE INNER EYE FOR NEARLY 10 MONTHS. "UNEVENTFUL SURGERY, GOOD VISION, GOOD POSITIONING AND GOOD VAULTING, CLEAR CORNEA, CLEAN LENS, NO PIGMENT DISPERSION" WAS REPORTED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14620821&pc=QCB | 2023826-2022-01794 | 2022-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00/+1.0/029 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT WAS COMPLAINING OF PAIN, POOR ABILITY TO READ AND LOW VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2022 AND PROBLEM RESOLVED. REPORTEDLY, "ICL HAS BEEN EXPLANTED, PATIENT GOES BACK TO USING REFRACTIVE APPLIANCES. CAUSE OF EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639450&pc=QCB | 2023826-2022-01795 | 2022-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/161 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT WAS COMPLAINING OF PAIN, POOR ABILITY TO READ AND LOW VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2022 AND PROBLEM RESOLVED. REPORTEDLY, "ICL HAS BEEN EXPLANTED, PATIENT GOES BACK TO USING REFRACTIVE APPLIANCES. CAUSE OF EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704775&pc=QCB | 2023826-2023-03953 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.0/+2.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON EXPLANTED THE LENS ON (B)(6) 2023, DUE TO THE PATIENT HAVING DISSATISFACTION AND FOREIGN BODY SENSATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PATIENT DISSATISACTION, FOREIGN SENSATION. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- PATIENT ALSO HAD PAIN. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532320&pc=QCB | 2023826-2019-02455 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-12-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Pain; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A TMICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. REPORTEDLY, 10-DAYS POSTOP THE PATIENT FELT PAIN, EXPERIENCED BLURRED VISION, AND THE LENS ROTATED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652119&pc=QCB | 2023826-2019-00915 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Pain; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME, BLURRED VISION, PAIN, AND "MINOR TEMPORALLY ICL SHIFT." CORNEA IS CLEAR, NO INJUNCTIVAL INJECTION, IOP AT 14MMHG AT LAST VISIT. PATIENT CLAIMS NO HISTORY OF EYE INJURY. Manufacturer Narrative: PRODUCT PROBLEM: SHOULD BE CORRECTED TO ADVERSE EVENT IN INITIAL MEDWATCH REPORT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD BE MARKED IN INITIAL MEDWATCH REPORT. DATE OF THIS REPORT: SHOULD BE CORRECTED TO 02/MAY/2019 IN INITIAL MEDWATCH REPORT. DATE RECEIVED BY MANUFACTURER: SHOULD BE CORRECTED TO 02/MAY/2019 IN INITIAL MEDWATCH REPORT. POST-OP PATIENT UCVA WAS RECORDED AS 20/25 AND THE VAULT WAS 458UM. REPOSITIONING WAS CONDUCTED ON A UNSPECIFIED DATE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10454797&pc=QCB | 2023826-2020-01864 | 2020-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.00/1.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PUPIL BLOCK WITH ELEVATED IOP (37MMHG) OBSERVED ON (B)(6) 2020. ADDITIONAL TREATMENT OF ENLARGEMENT OF PI AND A YAG WAS NECESSARY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS NON PATENT PI. REPORTEDLY, PATIENT CONDITION "IMPROVED - NO COMPLAINTS." Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED BLURRY VISION WITH EXTREME PAIN. H6 - DEVICE CODE: 2993 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335335&pc=QCB | 2023826-2020-01547 | 2019-11-28 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; Blurred Vision; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +8.5/+1.0/119 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WASEXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), LOSS OF VISION, DIM/MISTY VISION, AND ACHING PAIN. ADDITION OF NEW PI WAS ALSO REPORTED. THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "SIZE TOO BIG." Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525151&pc=QCB | 2023826-2019-00599 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2019 AND PROBLEM WAS RESOLVED. THE PATIENT EXPERIENCED PAIN AND THE PAIN WAS RESOLVED WITH ICL EXCHANGE. THE CAUSE OF THE EVENT WAS DUE TO SIZING NOMOGRAM. THE PATIENT HAD LASIK OU ON (B)(6) 2019 TO CORRECT RESIDUAL ASTIGMATISM. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525163&pc=QCB | 2023826-2019-00598 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS IMPLANTED ON 25-FEB-2019 AND EXPLANTED ON 27-MAR-2019. THE PATIENT EXPERIENCED PAIN AND THE PAIN WAS RESOLVED AFTER ICL EXCHANGE.. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT HAD LASIK OU ON 14-MAY-2019 TO CORRECT RESIDUAL ASTIGMATISM. THE CAUSE OF THE EVENT WAS DUE TO SIZING NOMOGRAM. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135938&pc=QCB | 2023826-2020-01166 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.00/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) IN 2019. THE REPORTER HAS INDICATED THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS DUE TO THE LENS HAS VAULTED FORWARD AND POSSIBLY RUBBING AGAINST THE IRIS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5:THE LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION AND INTERMITTENT PAIN IN THE DARK. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. H6 - PATIENT CODE: 3191 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PIGMENT DISPERSION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478865&pc=QCB | 2023826-2021-00636 | 2021-02-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE PUPIL WAS UNREACTIVE (FIXED) (BLOCKED PUPIL). THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE PATIENT IS NOW 8/10 IN THE LEFT EYE AND THE PUPIL IS ALMOST PERFECT BUT STILL NOT 100%. THE PATIENT IS BEING MONITORED. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT DATE: UNK THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311393&pc=QCB | 2023826-2022-00176 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Pupillary Block; Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM TMICL12.6; -7.00/+3.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021; AS A REPLACEMENT LENS. IT WAS NOTED THE PATIENT EXPERIENCED PUPIL BLOCK ON THE REPLACEMENT LENS WHICH WAS CAUSED FROM THE INITIAL IMPLANTED LENSES ANGLE CLOSURE. THE PATIENT THEN FELT VISION WASN'T GOOD AND STILL HAD A BLOCK PUPIL WITH RESIDUAL RX RIGHT EYE -0.25; +1.25; 145. THE SURGEON DECIDED TO PERFORM PRK SURGERY. IT WAS INDICATED THE PATIENT'S VISION IS STILL BLURRY. SEE MDR# 2023826-2022-00174 FOR INITIAL IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT 12.6MM TMICL12.6; -7.00/+3.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021; AS A REPLACEMENT LENS. IT WAS NOTED THE PATIENT EXPERIENCED PUPIL BLOCK ON THE REPLACEMENT LENS WHICH WAS CAUSED FROM THE INITIAL IMPLANTED LENSES ANGLE CLOSURE. THE PATIENT THEN FELT VISION WASN'T GOOD AND STILL HAD A BLOCK PUPIL WITH RESIDUAL RX RIGHT EYE -0.25; +1.25; 145. THE SURGEON DECIDED TO PERFORM PRK SURGERY. IT WAS INDICATED THE PATIENT'S VISION IS STILL BLURRY. SEE MDR# 2023826-2022-00174 FOR INITIAL IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16739723&pc=QCB | 2023826-2023-01370 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Red Eye(s); Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/1.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. IT WAS INDICATED THAT EVERYONE WAS NORMAL FOR THE 1ST FOUR DAYS AFTER SURGERY BUT ON THE 5TH DAY THE EYE WAS DILATED; WAS PRESCRIBED MEDICATION (PILOCARPINE); BUT PIGMENT DISPERSION STARTED AND RED EYE AND SERVER EYE PAIN. A DIFFERENT SURGEON HAD REPOSITIONED THE LENS BUT THIS DID NOT RESOLVE THE PROBLEM. A 3RD DOCTOR HAD PRESCRIBED ALPHAGAN P. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - DILATED PUPIL, PIGMENT DISPERSION. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502410&pc=QCB | 2023826-2022-00410 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/123 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER INDICATED THE REFRACTION WAS NOT STABLE AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02786. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/123 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER INDICATED THE REFRACTION WAS NOT STABLE AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02786. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE CAUSE OF THE REPORTED POSTOPERATIVE UNSTABLE REFRACTION WAS DUE TO A RETINAL DETACHMENT WHICH WILL BE TREATED AT THE MOMENT. THE RETINAL DETACHMENT WAS NOT CAUSED BY THE EVO VISIAN ICL. IT SEEMED TO BE CAUSED BY THE FURTHER HIGH MYOPIA OF THE PATIENT EYE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. GMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. GMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586209&pc=QCB | 2023826-2019-00734 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Retinal Detachment; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS, IN (B)(6) 2019. THE REPORTER INDICATED THERE WAS A RETINAL DETACHMENT AND THE SURGERY WAS UNEVENTFUL. THE SURGEON INDICATED THE PATIENT WAS A VERY HIGH MYOPIC, WHICH WAS HER RISK FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/169 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENTS LEFT EYE (OS) ON 03/26/2019. THE PATIENT HAD A RETINAL DETACHMENT AND THE RETINA WAS REPAIRED. A SCLERAL BUCKLE WAS PLACED, WITH 20% GAS FILL AT THAT VISIT. THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. THE PATIENT IS DOING WELL. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, RETINAL REPAIR, SCLERAL BUCKLE IN PLACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120437&pc=QCB | 2023826-2019-01742 | 2016-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Retinal Detachment; Vitrectomy; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +0.5/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS RETINAL DETACHMENT. RETINAL REPAIR WAS PERFORMED, INCLUDING PARS PLANA VITRECTOMY WITH C3F8 GAS TAMPONADE AND EYE DROPS. THE PROBLEM IS RESOLVED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. INFORMATION INITIALLY RECEIVED VIA AN ARTICLE PUBLISHED IN CLINICAL OPHTHALMOLOGY ENTITLED 'A PROSPECTIVE PILOT STUDY USING A LOW POWER PIGGY-BACK TORIC IMPLANTABLE COLLAMER LENS TO CORRECT RESIDUAL REFRACTIVE ERROR AFTER MULTIFOCAL IOL IMPLANTATION.' Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND WITHIN ASSOCIATED LOTS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15099175&pc=QCB | 2023826-2022-02430 | 2022-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Swelling/ Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.0/1.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SWELLING; IT WAS NOTED "THE PATIENT CONTINUED TO FEEL THE WHOLE SWELLING; WHICH WAS NOT TOLERATED AFTER SEVERAL CONSUTATION". ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. LENS RETURNED IS NOT MODEL STATED IN THE CLAIM, THEREFORE DIMENSIONAL INSPECTION IS NOT REQUIRED. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11448805&pc=QCB | 2023826-2021-00610 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Swelling/ Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS EDEMA, PACHYMETRY AT 600 MICRONS, WITH LIMITED VISION 20/25 AND REFRACTION -0.5. THE DOCTOR BELIEVES THIS SHOULD IMPROVE WITH TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13085403&pc=QCB | 2023826-2021-04924 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. STEROIDS WERE PRESCRIBED. AT THE POST-OP VISIT ON (B)(6) 2021, THE PROBLEM WAS RESOLVED AND STEROIDS WERE NO LONGER PRESCRIBED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. STEROIDS WERE PRESCRIBED. AT THE POST-OP VISIT ON (B)(6) 2021, THE PROBLEM WAS RESOLVED AND STEROIDS WERE NO LONGER PRESCRIBED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510/K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510/K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182687&pc=QCB | 2023826-2022-00035 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS TASS WAS OBSERVED WHICH SUBSIDED WITH MEDICAL TREATMENT. THE LENS IS STABLE IN THE EYE. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13958873&pc=QCB | 2023826-2022-00953 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Toxic Anterior Segment Syndrome (TASS) |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.5/2.5/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS). DIAGNOSTIC PERFORMED: INTENSE INFLAMMATION WITHIN THE ANTERIOR CHAMBER. TREATMENT PERFORMED: ANTIMICROBIAL TREATMENT AND ANTI-INFLAMMATORY TREATMENT WERE GIVEN. THE REPLACEMENT LENS RESOLVED THE PROBLEM. REPORTER STATES "IT IS DIFFICULT TO DETERMINE THE CAUSE. Manufacturer Narrative: WEIGHT - RACE:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190878&pc=QCB | 2023826-2022-01293 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -18.0/+1.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS TASS-DIAGNOSTICS WERE NOT PERFORMED AND MEDICATION WAS ADMINISTERED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR: THE DEVICE DID NOT FAIL TO PERFORM AS EXPECTED. THE SURGEON ALSO REPORTS, "ALLERGIES TO LENSES." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H6-INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5 - CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR, NOT USER ERROR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153793&pc=QCB | 2023826-2022-02512 | 2022-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+1.0/009 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE REPORTER INDICATED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS FOUND AFTER SURGERY. TREATMENT WAS PERFORMED - ANTERIOR CHAMBER LAVAGE AND KENACORT CONJUNCTIVAL INJECTION. RECOVERY TREND WITH STEROIDS TREATMENT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15681653&pc=QCB | 2023826-2022-03832 | 2022-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030017 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -6.0/2.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. NO DIAGNOSTICS WERE PERFORMED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED. LINDELON WAS PRESCRIBED. REPORTEDLY, "AFTER 2 MONTHS PRESCRIPTION OF LINDELON, SHE IS ON THE MEND." PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. STATEMENT SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15681668&pc=QCB | 2023826-2022-03833 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030017 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. NO DIAGNOSTICS WERE PERFORMED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED. LINDELON WAS PRESCRIBED. REPORTEDLY, "AFTER 1 MONTHS PRESCRIPTION OF LINDELON, SHE IS ON THE MEND." PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. STATEMENT SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT#: (B)(4) IS ASSOCIATED WITH THE PATIENT'S FELLOW EYE (OS). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15681720&pc=QCB | 2023826-2022-03834 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Toxic Anterior Segment Syndrome (TASS) | P030017 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. NO DIAGNOSTICS WERE PERFORMED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED. LINDELON WAS PRESCRIBED. REPORTEDLY, "AFTER 1 MONTHS PRESCRIPTION OF LINDELON, SHE IS ON THE MEND." PROBLEM WAS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT #(B)(4) IS ASSOCIATED WITH THE PATIENT'S FELLOW EYE (OD). CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. - STATEMENT SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15765123&pc=QCB | 2023826-2022-03998 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.50/1.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. THE LENS REMAINS IMPLANTED. STEROID EYE DROPS AND ORAL MEDICATION WERE PRESCRIBED. REPORTER STATES THE PROBLEM HAS NOT RESOLVED AND PATIENT OUTCOME STATUS IS UNKNOWN. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15765127&pc=QCB | 2023826-2022-03999 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.00/3.0/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) (B)(6) 2022. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS REPORTED. THE LENS REMAINS IMPLANTED. STEROID EYE DROPS AND ORAL MEDICATION WERE PRESCRIBED. REPORTER STATES THE PROBLEM HAS NOT RESOLVED AND PATIENT OUTCOME STATUS IS UNKNOWN. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515441&pc=QCB | 2023826-2023-03338 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | Uveitis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REPORTEDLY THE SURGERY WAS UNEVENTFUL, GOOD VISUAL RESULT ACHIEVED. POST OPERATIVELY MAXIDEX (STEROID DROPS) UNABLE TO STOP USE AS ONCE CEASED UVEITIS FLARE UP OCCURS. LENS REMAIN IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-HEALTH IMPACT 4644: MAXIDEX DROPS. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515558&pc=QCB | 2023826-2023-03339 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | Uveitis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY THE SURGERY WAS UNEVENTFUL, GOOD VISUAL RESULT ACHIEVED. POST OPERATIVELY MAXIDEX (STEROID DROPS) UNABLE TO STOP USE AS ONCE CEASED UVEITIS FLARE UP OCCURS. LENS REMAIN IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. . Manufacturer Narrative: ADDITIONAL INFORMATION: H6-HEALTH IMPACT 4644: MAXIDEX DROPS. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777436&pc=QCB | 2023826-2023-04124 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.00/+1.00/133 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS THE PATIENT HAS UVEITIS WHICH WAS NOTED (B)(6) 2023. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18164007&pc=QCB | 2023826-2023-05151 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis | P030016 | 1 | Event Description: INFLAMMATORY REACTION 9 DAYS AFTER IMPLANTABLE COLLAMER LENS IMPLANTATION WAS REPORTED. RPORTER IS UNCERTAIN OF THE ETIOLOGY BUT IS TREATING IT AS UVEITITS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15849035&pc=QCB | 2023826-2022-04266 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ENDOPHTHALMITIS WAS REPORTED. DIAGNOSTICS WA PERFORME 'UVEITIS'. TREATMENT WAS PERFORMED: MINOCIN CAP 50MG; PREDNISOLONE 5MG; AXADIN CAP 150MG; VIGAMOX EYE DROPS 0.5%. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS BUT DIFFERENT AXIS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. "IMPROVEMENT IN SYMPTOMS AFTER REPLACEMENT." Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. IN THE COMPLAINT QUESTIONNAIRE, THE BOX MARKED ENDOPHTHALMITIS WAS TICKED AND THE DIAGNOSTICS PERFORMED WAS DESCRIBED AS 'UVEITIS'. LATER CLARIFICATION FROM THE REPORTER STATED THAT ENDOPHTHALMITIS WAS NOT DIAGNOSED NOR WAS THERE ANY EVIDENCE OF THIS. AT THIS TIME, MEDICAL MANAGEMENT IN THE FORM OF ANTIBIOTICS (TOPICAL AND SYSTEMIC) WITH SYSTEMIC STEROIDS WITH GASTRIC PROTECTION WAS PRESCRIBED. FOLLOWING THIS, THE REPORTER NOTED A "SIGNIFICANT IMPROVEMENT IN SYMPTOMS". HOWEVER SUBSEQUENT COMMUNICATION STATED THAT "INFLAMMATION OCCURRED A WEEK AFTER WHEN STEROIDS DISCONTINUED". THREE WEEKS POST-IMPLANTATION THE LENS WAS EXCHANGED FOR A NEW LENS OF SAME PARAMETERS WHICH RESOLVED THE ISSUE. H65: CLINICAL CODE: 4581:ENDOPHTHALMITIS IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLINICAL CODE: (B)(4):ENDOPHTHALMITIS HEALTH IMPACT : 4644:MINOCIN CAP 50MG; PREDNISOLONE 5MG; AXADIN CAP 150MG; VIGAMOX EYE DROPS 0.5%. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930479&pc=QCB | 2023826-2020-02833 | 2020-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED, DIAGNOSTICS AND TREATMENT PERFORMED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. UVEITIS; PUPIL MEMBRANE WAS CLOSED, IRIS SYNECHIA WAS REPORTED. TREATMENT OF HORMONES; NON STEROIDAL ANTI-INFLAMMATORY DRUGS, MYDRIASIS TREATMENT, TREATMENT IMPROVED, REPEATED ATTACKS, AFTER CONSULTATION DECIDED TO TAKE OUT THE LENS. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. H6 CLINICAL CODE: 4581: PUPIL MEMBRANE CLOSED; IRIS SYNECHIA. HEALTH IMPACT: 4625: MYDRIASIS TREATMENT. HEALTH IMPACT : 4644: HORMONES; NON-STEROIDAL ANTI- INFLAMMATORY DRUGS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010278&pc=QCB | 2023826-2022-04648 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Uveitis; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -1.50/+3.5/078 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE SURGEON REPORTS UVEITIS AFTER SURGERY WITHOUT PREVIOUS HISTORY; DEPOSITS IN ICL THAT DECREASE VISUAL ACUITY; PATIENT IS TREATED WITH CORTICOSTEROIDS; AND THERE IS PIGMENT DISPERSION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN; AND DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL DEVICE PROBLEM CODE-1494 SHOULD BE ADDED FOR CORRECTION. OFF-LABEL USAGE. CLAIM#(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706691&pc=QCB | 2023826-2020-00304 | 2020-01-10 | Injury | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Uveitis; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.00/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PATIENT WAS DIAGNOSED WITH UVEITIS ON (B)(6) 2020 (VISUAL ACUITY 6/4.5-2) AND WAS PRESCRIBED PRED FORTE. AT PATIENT'S (B)(6) 2020 FOLLOW-UP, BLURRY VISION WAS OBSERVED (VISUAL ACUITY: 6/7.5-1, PINHOLE: 6/6-1). PATIENT WAS PRESCRIBED CIPROFLOXACIN, DEXAMETHASONE, CYCLOPENTOLATE, CRAVIT. PER REPORTER ON (B)(6) 2020, "PATIENT HAVE FULLY RECOVERED FROM THE UVEITIS." LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930493&pc=QCB | 2023826-2020-02834 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Uveitis; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED, DIAGNOSTICS AND TREATMENT PERFORMED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. UVEITIS WAS REPORTED. TREATMENT OF HORMONES; NON STEROIDAL ANTI-INFLAMMATORY DRUGS, MYDRIASIS TREATMENT, TREATMENT IMPROVED, REPEATED ATTACKS, AFTER CONSULTATION DECIDED TO TAKE OUT THE LENS. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. D9- (B)(6) 2020. HEALTH IMPACT: 4625: MYDRIASIS TREATMENT. HEALTH IMPACT : 4644: HORMONES; NON-STEROIDAL ANTI- INFLAMMATORY DRUGS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312775&pc=QCB | 2023826-2019-02160 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT THAT HAD ICL SURGERY IN THE LEFT EYE (OD) HAD A 14 DEGREE ROTATION. PATIENT IS COMPLAINING THAT WHEN LOOKING AT LINES, LIKE CROSSINGS THERE IS A DISTORTION OF THE VERTICAL AXIS. IT WAS REPORTED THAT THE PATIENT WANTS TO CONSULT WITH OTHER DOCTORS FOR ANOTHER OPINION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941512&pc=QCB | 2023826-2020-00870 | 2020-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | 1 | Event Description: THE RE4PORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/1.0/90 9 (SPHERE/CYLINDER/AXIS),IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT IS REPORTED TO BE EXPERIENCING "DOUBLE VISION THAT CAN BE SEEN UP AND DOWN IN THE DARK PLACES". THE SYMPTOMS DO NOT OCCUR IN BRIGHT PLACES. THE SURGEON REPORTED THAT THE "PATIENT HAS NORMAL IOP, NORMAL VAULT, NORMAL OCULAR BODY AND NO COMPLICATIONS." THE REPORTER STATED " THE DOCTOR DOES NOT INTEND TO POINT OUT ANY PRODUCT DEFECTS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941534&pc=QCB | 2023826-2020-00871 | 2020-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2,-11.0/1.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT IS REPORTED TO BE EXPERIENCING "DOUBLE VISION THAT CAN BE SEEN UP AND DOWN IN THE DARK PLACES". THE SYMPTOMS DO NOT OCCUR IN BRIGHT PLACES. THE SURGEON REPORTED THAT THE "PATIENT HAS NORMAL IOP, NORMAL VAULT, NORMAL OCULAR BODY AND NO COMPLICATIONS." THE REPORTER STATED " THE DOCTOR DOES NOT INTEND TO POINT OUT ANY PRODUCT DEFECTS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740194&pc=QCB | 2023826-2020-02482 | 2020-09-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+2.0/058 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE PATIENT WAS UNHAPPY WITH VISION (RESIDUAL MYOPIA). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740224&pc=QCB | 2023826-2020-02481 | 2020-09-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+2.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE PATIENT WAS UNHAPPY WITH VISION (RESIDUAL MYOPIA). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098621&pc=QCB | 2023826-2020-03250 | 2020-10-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Optical Problem | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER STATES THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. REPORTEDLY, THE PATIENT IS EXPERIENCING LIGHTNING/FLASHES WHILE HE LIES IN BED AND TURNS HIS HEAD; STANDING UP THE VISION IS FINE. RETINAL DIAGNOSTIC SHOWS NO CHANGES AND THE SURGEON THINKS THIS COULD BE A "STRESS SYMPTOMATIC." Manufacturer Narrative: B5-UPDATED INFORMATION: THE HALOS HAVE BECOME LESS AND LESS. THE PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238074&pc=QCB | 2023826-2021-00173 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -11.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND THE LENS WAS REPLACED IN A SECOND SURGERY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. REPORTEDLY, "WTW MEASUREMENT CALLED FOR 12.6MM LENS(STAAR OPTIMIZED). UNCLEAR IF THERE WAS A AN ABNORMAL SULCUS ANATOMY CAUSING EXCESS VAULT. PER REPORTER, "VAULT IS NOW EXCELLENT, PATIENT NOTICES SIGNIFICANTLY REDUCED GLARE." Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238075&pc=QCB | 2023826-2021-00213 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -11.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "WE ARE PLANNING TO EXCHANGE LENS IN MARCH." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: REPORTEDLY, "WE EXCHANGED HER LEFT EYE A COUPLE OF MONTHS AGO. SIGNIFICANTLY IMPROVED HER OVERALL VISION, BUT ESPECIALLY HER NIGHT VISION. SHE HAD A WTW OF 11.5 WTW OU AND I PLACED A 12.6 LENS IN EACH EYE. 300% CT VAULT OU." CLAIM (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11411793&pc=QCB | 2023826-2021-00568 | 2021-01-15 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.00/5.0/045 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WITHOUT ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED DIFFICULTY WITH ACCOMMODATION FOR NEAR OBJECTS, THE REPORTER STATED " SX REPORTS DIFFICULTY WITH ACCOMMODATION POST ICL IMPLANTATION IN THE CONTEXT OF GOOD VISION AND REFRACTION DURING 1 DAY AND 1 WEEK POST OP REVIEW. BETTER WHEN OBJECTS ARE HELD FURTHER AWAY." LENS REMAINS IMPLANTED FOR NOW THE DOCTOR WILL MONITOR THE EYE. THE REPORTER STATED " PATIENTS CURRENT CONDITION IS FINE, ONLY ISSUE IS DIFFICULTY FOR READING NEAR SOMETIMES. VA/ EYE PRESSURE IS GOOD. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT-RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558804&pc=QCB | 2023826-2021-00706 | 2020-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+1.0/052 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER IMPLANT, THE PATIENT FELT UNCOMFORTABLE WITH THE GLARE WHEN DRIVING AT NIGHT AND SO THE ICL WAS REMOVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563271&pc=QCB | 2023826-2021-00759 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/4.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. DIPLOPIA AND ASTIGMATISM WAS OBSERVED, WITH POSSIBLE ROTATION. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, LENS ROTATION IS PLANNED. WILL PROVIDE FURTHER INFORMATION ONCE A RESOLUTION IS REACHED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.0/4.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, ORIGINAL LENS TOO LARGE, SMALLER REPLACED LENS ROTATED. MFR # 2023826-2021-00652. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563284&pc=QCB | 2023826-2021-00757 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/4.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 201. DIPLOPIA AND ASTIGMATISM WAS OBSERVED, WITH POSSIBLE ROTATION. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, LENS ROTATION IS PLANNED. WILL PROVIDE FURTHER INFORMATION ONCE A RESOLUTION IS REACHED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/4.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 201. DIPLOPIA AND ASTIGMATISM WAS OBSERVED, WITH POSSIBLE ROTATION. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, "DIPLOPIA AND ASTIGMATISM HAVE RESOLVED." CORRECTED INFORMATION: B1 - REPORTED AS ADVERSE EVENT - CORRECT TO PRODUCTION PROBLEM. B2 - REPORTED AS REQUIRED INTERVENT TO PREVENT PERMANENT IMPAIRMENT/DAMAGE - NO OUTCOMES ATTRIBUTED TO ADVERSE EVENT. H1 - REPORTED AS SERIOUS INJURY - CORRECT TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820848&pc=QCB | 2023826-2021-01410 | 2021-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER/PATIENT INDICATED THAT A 13.2MM, TMICL13.2, -11.00/1.5/053 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, "GLARE, VISUAL QUALITY IS BAD AND NOT IMPROVING." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER/PATIENT INDICATED THAT A 13.2MM TMICL13.2 -11.00/1.5/053 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY "GLARE VISUAL QUALITY IS BAD AND NOT IMPROVING." GLARE/HALOES REPORTED. LENS REMAINS IMPLANTED. PROBLEM NOT RESOLVED. STATUS OF THE EYE IS QUIET - WELL HEALED. REPORTEDLY THE CAUSE OF THIS EVENT IS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820914&pc=QCB | 2023826-2021-01411 | 2021-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER/PATIENT INDICATED THAT A 13.2MM, TMICL13.2, -10.50/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, "GLARE, VISUAL QUALITY IS BAD AND NOT IMPROVING." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER/PATIENT INDICATED THAT A 13.2MM TMICL13.2 -10.50/1.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, "GLARE, VISUAL QUALITY IS BAD AND NOT IMPROVING." LENS REMAINS IMPLANTED. PROBLEM NOT RESOLVED. STATUS OF THE EYE IS QUIET - WELL HEALED. REPORTEDLY THE CAUSE OF THIS EVENT IS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281705&pc=QCB | 2023826-2021-03069 | 2021-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Dislodged or Dislocated | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS WAS MISALIGNED BY 15 DEGREES. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, MOST LIKELY CAUSE IS THAT INK AXIS MARKING WERE OFF TARGET FOR THE PROCEDURE. REPORTEDLY, LENS REPOSITIONING IS PLANNED. WILL ADVISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS WAS MISALIGNED BY 15 DEGREES. LENS REPOSITIONED ON (B)(6) 2021 DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, MOST LIKELY CAUSE IS THAT INK AXIS MARKING WERE OFF TARGET FOR THE PROCEDURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478533&pc=QCB | 2023826-2021-03579 | 2021-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND GLARE (THE PATIENTS PUPIL WAS 5MM). THE PUPIL SIZE WAS STILL 5MM FOR HALF A MONTH AFTER SURGERY, SO THE SURGEON DECIDED NOT TO IMPLANT ANOTHER ICL LENS FOR THE TIME BEING. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479095&pc=QCB | 2023826-2021-03592 | 2021-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER STATED THAT THE PATIENT DEVELOPED DIPLOPIA WITH VERY LOW VAULT AND LENS ROTATED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479109&pc=QCB | 2023826-2021-03593 | 2021-06-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE REPORTER STATED THAT THE PATIENT DEVELOPED DIPLOPIA WITH VERY LOW VAULT AND LENS ROTATED. LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558495&pc=QCB | 2023826-2021-03810 | 2021-05-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE SURGEON REPORTS GLARE/HALOS AND PILOCARPINE ADMINISTRATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725161&pc=QCB | 2023826-2021-04177 | 2020-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -15.5/3.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO GLARE/HALOES. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ABERATIONS IMPROVED BUT STILL PRESENT. PLAN ON DOING LASIK FOR RESIDUAL REFRACTIVE ERROR." IN THE REPORTERS OPINION " DEVICE AND PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED " DEVICE CAUSES MILD ABERATIONS- PT VERY SENSITIVE TO ABERATIONS". Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725173&pc=QCB | 2023826-2021-04176 | 2020-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -15.0/4.0/089 (SPHERE / CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO GLARE/ HALOES. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ ABERRATIONS IMPROVED BUT STILL PRESENT. PLAN ON DOING LASIK FOR RESIDUAL REFRACTIVE ERROR." IN THE REPORTERS OPINION "DEVICE AND PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED "DEVICE CAUSES GLARE/ABERRATIONS- PT VERY SENSITIVE TO GLARE/ABERRATIONS". Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872872&pc=QCB | 2023826-2021-04514 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | NA | 1 | Event Description: ON (B)(6) 2021 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN SPRINGER NATURE, ENTITLED, "A CRITICAL EVALUATION OF LONGITUDINAL CHANGES OF ASTIGMATISM FOLLOWING IMPLANTATION OF TORIC IMPLANTABLE COLLAMER LENSES (TICL): A COMPARISON BETWEEN TREATED AND UNTREATED CASES OVER 4 YEARS". THE ARTICLE REPORTS RESIDUAL ASTIGMATISM FOLLOWING TORIC IMPLANTATION IN 86 CASES. IT WAS REPORTED THAT IN 50% OF TORIC IMPLANTABLE COLLAMER LENSES PRESENTING WITH ASTIGMATISM LESS THAN OR EQUAL TO -0.75 DC, THE RESIDUAL ASTIGMATISM CAN BE NATURALIZED BY REALIGNING THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH-RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057182&pc=QCB | 2023826-2021-04829 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF NIGHT GLARE, WHICH AFFECTED HIS NORMAL LIFE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF NIGHT GLARE, WHICH AFFECTED HIS NORMAL LIFE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(6). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057258&pc=QCB | 2023826-2021-04828 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF NIGHT GLARE, WHICH AFFECTED HIS NORMAL LIFE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082455&pc=QCB | 2023826-2021-04974 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO "UNACCEPTABLE GLARE AFTER IMPLANTATION OF LENS." THE GLARE DISAPPEARED AFTER THE REMOVAL OF THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO "UNACCEPTABLE GLARE AFTER IMPLANTATION OF LENS." THE GLARE DISAPPEARED AFTER THE REMOVAL OF THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13362150&pc=QCB | 2023826-2022-00347 | 2021-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4/+6/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "PATIENT SAW GLARE".Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4/+6/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "PATIENT SAW GLARE". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14093007&pc=QCB | 2023826-2022-01195 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL_12.1; -9.00/1.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE PATIENT WANTED TO TRY FOR MONOVISION; IN WHICH SHE WAS UNHAPPY WITH THE DEGREE OF MONOVISION AND WOULD PREFER TO HAVE AN EXCHANGE FOR A POWER CLOSER TO HER DISTANCE REQUIREMENT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14706970&pc=QCB | 2023826-2022-01749 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 -13.5/3.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REPORTEDLY GLARES/HALOES OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2022 WHICH RESOLVED THE PROBLEM. CAUSE IS DUE TO THE DEVICE - THE PATIENT FELT THAT THE POSTOPERATIVE HALO WAS BECOMING MORE SERIOUS AND COULD NOT BEAR IT AND ASKED TO REMOVE IT. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14706973&pc=QCB | 2023826-2022-01748 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 -14.5/5.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS OF WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY GLARES/HALOES OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2022 WHICH RESOLVED THE PROBLEM. CAUSE IS DUE TO THE DEVICE - THE PATIENT FELT THAT THE POSTOPERATIVE HALO WAS BECOMING MORE SERIOUS AND COULD NOT BEAR IT AND ASKED TO REMOVE IT. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15003548&pc=QCB | 2023826-2022-02198 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/154 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6). POST-OP, THE UCVA WAS GOOD BUT THERE WAS GLARE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK .EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567406&pc=QCB | 2023826-2022-03564 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICM5_13.2, -9.50/4.0/91. (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AS A REPLACEMENT LENS AFTER THE INITIAL LENS WAS FOUND DAMAGE BEFORE LOADING. IT WAS REPORTED THAT POST ICL IMPLANTATION THE PATIENT EXPERIENCED GLARE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE GLARE RESOLVED. LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000576&pc=QCB | 2023826-2022-04596 | 2022-11-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE REPORTER STATES THAT THE PATIENT HAS AN USTABLE DISTANCE VISION. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, PROBLEM IS RESOLVED. "THE PATIENT HAS GOTTEN USED TO THE SITUATION AND DOES NOT WANT ANY FURTHER INTERVENTIONS." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040405&pc=QCB | 2023826-2022-04702 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -3.5/6.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS REMOVED. REPORTEDLY, "EXCESSIVE GLARE AND HALO, PATIENT CAN'T TOLERATE. CAUSE IS REPORTED AS THE DEVICE AND PATIENT RELATED FACTOR. REPORTEDLY, STATUS OF THE EYE IS "RETURNED TO PRE-OP STATUS." Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296735&pc=QCB | 2023826-2023-00431 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -15.0/2.00/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED DECREASED VISION AND HALO GHOST IMAGE. THE REPORTER STATES THAT ICL ROTATION WAS PERFORMED IN JUNE 2020 AND THAT HE NOW PLANS TO EXCHANGE FOR LONGER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). IMPLANT DATE: UNKNOWN. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS. REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296738&pc=QCB | 2023826-2023-00429 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.00/5.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES THAT A LENS WITH WRONG LENGTH WAS ORDERED AND THAT THE CAUSE OF ROTATION IS LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(6). IMPLANT DATE: (B)(6) 2019. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296741&pc=QCB | 2023826-2023-00433 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.5/1.0/64 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANT DATE: (B)(6) 2022. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296747&pc=QCB | 2023826-2023-00430 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.00/4.5/77(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES THAT A LENS WITH WRONG LENGTH WAS ORDERED AND THAT THE CAUSE OF ROTATION IS LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). IMPLANT DATE: (B)(6) 2019. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296766&pc=QCB | 2023826-2023-00428 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/3.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2021. LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES "CAUSE OF ROTATION IS LOW VAULT". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). IMPLANT DATE: (B)(6) 2021. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: PRODUCT PROBLEM SHOULD BE CORRECTED TO ADVERSE EVENT. B2: OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS SHOULD BE ADDED FOR CORRECTION. H3: SERIOUS INJURY SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296833&pc=QCB | 2023826-2023-00432 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN SEPTEMBER 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: PATIENT IDENTIFIER:(B)(6). IMPLANT DATE: SEPTEMBER 2022 TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7, -7.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". H6: MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE: AGE>45 YEARS OLD AT THE TIME OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7, -7.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16742900&pc=QCB | 2023826-2023-01256 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/177 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023. DUE TO LOW VAULTING AND THE PATIENT EXPERIENCED DOUBLE VISION AND GLARE AT NIGHT. THE PATIENT WAS PRESCRIBED MEDICATION BUT THE SYMPTOMS DID NOT IMPROVE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES AND UNABLE TO PERFORM THE DIMENSIONAL INSPECTION DUE TO THE DAMAGED LENS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16877245&pc=QCB | 2023826-2023-01587 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6)2022. THE LENS WAS EXPLANTED ON (B)(6)2023 DUE TO THE PATIENT EXPERIENCED GLARE AND THE PROBLEM WAS RESOLVED. THE PATIENT COULD NOT ACCEPT THE GLARE AND REQUESTED THE LENS BE REMOVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17273274&pc=QCB | 2023826-2023-02719 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -18.0/2.0/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED SEVERE DOUBLE VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS, AND LIMBAL RELAXING INCISION WAS PERFORMED TO CORRECT ITS ASTIGMATISM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302685&pc=QCB | 2023826-2023-02829 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+3.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322117&pc=QCB | 2023826-2023-02830 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+2.5/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18210927&pc=QCB | 2023826-2023-05249 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2023 ELEVATED IOP (34MMHG), POOR VISUAL ACCUITY WAS OBSERVED. TAPERING TOPICAL STEROID AND ALPHAGAN IS BEING TAKEN. LENS REMAINS IMPLANTED. PROBLEM NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. G4 - UNK. H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18279670&pc=QCB | 2023826-2023-05479 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -2.0/0.5/1 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18279671&pc=QCB | 2023826-2023-05480 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -2.0/0.5/69 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6). LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# 732815.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18456573&pc=QCB | 2023826-2024-00038 | 2023-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND GLARE/HALOS WERE OBSERVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18456592&pc=QCB | 2023826-2024-00039 | 2023-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND GLARE/HALOS WERE OBSERVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684401&pc=QCB | 2023826-2022-00551 | 2021-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -9.50/1.0/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE, AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, GLARE REDUCED A LOT, VISION BETTER PUPIL BACK TO NORMAL. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852592&pc=QCB | 2023826-2020-02686 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Discomfort | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT WAS UNCOMFORTABLE AND EXPERIENCED VISION FATIGUE. THE SURGEON WILL NOT IMPLANT ANOTHER ICL LENS. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10938663&pc=QCB | 2023826-2020-02843 |  | Injury | STAAR SUGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.50/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2019. THE PATIENT CONTINUED TO STRUGGLE WITH GLARE AND HALO SINCE INITIAL SURGERY, POSSIBLE LENS EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.50/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOES. THE REPORTER STATED " PUPIL SIZE GREATER OD > OS DUE TO EXCESSIVE VAULT CAUSING GLARE/HALO OD. " THIS EXCHANGE RESOLVED THE PROBLEM, "GLARE/HALO COMPLAINT HAS RESOLVED WITH 12.6MM ICL IMPROVING VAULT." CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322676&pc=QCB | 2023826-2019-00229 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, VTICMO13.7, -12.0/4.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018 AND THE PATIENT IS COMPLAINING OF GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IT WAS REPORTED THAT REFRACTIVE TREATMENT WAS PERFORMED. ON PATIENTS LAST VISIT DATE: (B)(6) 2019, STATUS OF THE EYE WAS REPORTED AS "NORMAL", BCVA: 20/25-2, UCVA: 20/30 AND IOP: 12 MMHG. THE FOLLOWING WAS REPORTED AS ADDITIONAL INFORMATION: " PATIENT EXPERIENCING RESIDUAL HYPEROPIA AFTER SURGERY, ALSO COMPLAINING OF GLARE & HALOES". ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, VTICMO13.7, -12.0/4.0/075 (SPHERE/CYLINDE/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018 AND THE PATIENT IS COMPLAINING OF GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. ON PATIENT'S LAST VISIT DATE: (B)(6) 2019, STATUS OF THE EYE WAS REPORTED AS "NORMAL", BCVA 20/25-2, UCVA:20/30 AND IOP: 12MMHG. THE FOLLOWING WAS REPORTED AS ADDITIONAL INFORMATION: "PATIENT EXPERIENCING RESIDUAL HYPEROPIA AFTER SURGERY, ALSO COMPLAINING OF GLARE AND HALOES". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8544980&pc=QCB | 2023826-2019-00594 | 2018-03-07 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -1.00/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THAT GLARE/HALOES WERE OBSERVED. CAUSE OF THE EVENT IS REPORTED AS CENTRAL FLOW HOLE. LENS REMAINS IMPLANTED. REPORTEDLY, LENS EXCHANGE IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9323816&pc=QCB | 2023826-2019-02164 | 2018-09-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; No Apparent Adverse Event | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -0.50/2.0/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. ON (B)(6) 2019 REPORTER INDICATED "THIS PATIENT HAS ALREADY EXCHANGED HIS LENS ON OS AND MIGHT NEED TO DO IT ON OD."ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - DEVICE CODE 3189 CORRECTED TO DEVICE CODE 3001. (B)(4). Manufacturer Narrative: H6 - RESULT CODE 213 CORRECTED TO RESULT CODE 114. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427016&pc=QCB | 2023826-2019-02296 | 2019-11-14 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO GLARE/HALOES. PER REPORTER, "PATIENT EXPERIENCED CONCENTRIC CIRCLES OF LIGHT THRU LENS HOLD - INTOLERABLE." PATIENT WILL NOT BE PROCEEDING WITH ANOTHER LENS." Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633749&pc=QCB | 2023826-2020-00161 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/4.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. GLARES/HALOES WERE OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY AFTER EXPLANT (REMOVAL)." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT IMPACT CODE 4629 SHOULD BE IN THE INITIAL MDR. H10- DEVICE EVALUATION: "RETURNED IN CARTRIDGE" SHOULD BE REMOVED FROM PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: THE LENS WAS RETURNED WITH CLEAR SURGICAL RESIDUE AND MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS RETURNED IN THE CARTRIDGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135948&pc=QCB | 2023826-2020-01171 | 2020-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT COMPLAINED OF GLARE/HALOS AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1 - SHOULD BE CORRECTED TO MALFUNCTION. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135999&pc=QCB | 2023826-2020-01172 | 2020-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT COMPLAINED OF GLARE/HALOS AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1 - CORRECTED TO MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740529&pc=QCB | 2023826-2020-02488 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.50/+3.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, WITH GLARE AND HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY" CORRECTED INFORMATION: B5 - (B)(6) 2020 IMPLANT DATE SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740575&pc=QCB | 2023826-2020-02487 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.00/+4.5/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, WITH GLARE AND HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - (B))(6) 2020 IMPLANT DATE SHOULD BE CORRECTED TO (B)(6) 2020. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984748&pc=QCB | 2023826-2020-02858 | 2020-11-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. GLARE/HALOES HAVE BEEN OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: B1 - REPORTED AS PRODUCT PROBLEM - UPDATE TO ADVERSE EVENT. B2 - UPDATE TO REQUIRD INTERVENTION TO PREVENT IMPAIREMENT/DAMAGE. B5 - NOTED LENS REMAINS IMPLANTED. LENS REMOVED ON (B)(6) 2020 AND THE PROBLEM RESOLVED. D6B = (B)(6) 2020. H6- HEALTH EFFECT IMPACT CODE - REPORTED AS 2199, UPDATED TO 4629. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN WITH FOREIGN MATERIAL ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206966&pc=QCB | 2023826-2021-00106 | 2020-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -10.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO LOW VAULT AND GLARE/HALOES. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE IS NO PLAN TO IMPLANT ANOTHER LENS." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238056&pc=QCB | 2023826-2021-00172 | 2020-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10/1.0/131 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "IRIS IS NOT FULLY JUMPING OVER OPTICAL ZONE. MESOPIC PUPIL MID-DILATED POST OP (5MM). EDGE OF THE PUPIL IS NOT IN THE WAY OF THE LENS AND THE PUPIL DOES CONSTRICT AND RELAX BUT DOES NOT FULLY CONSTRICT. PATIENT TO BE MONITORED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238066&pc=QCB | 2023826-2021-00171 | 2020-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "IRIS IS NOT FULLY JUMPING OVER OPTICAL ZONE. MESOPIC PUPIL MID-DILATED POST OP (5MM). EDGE OF THE PUPIL IS NOT IN THE WAY OF THE LENS AND THE PUPIL DOES CONSTRICT AND RELAX BUT DOES NOT FULLY CONSTRICT. PATIENT TO BE MONITORED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368705&pc=QCB | 2023826-2021-00458 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368840&pc=QCB | 2023826-2021-00457 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+4.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11642581&pc=QCB | 2023826-2021-00943 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.5/1.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2020. THE REPORTER STATED " ONE YEAR AGO I GOT THE EVO VISIAN ICL IMPLANTED IN MY EYES (I'M MYOPIC WITH ASTIGMATISM). THE SURGERY WENT WELL AND MY EYES HEALED WELL. I DON'T HAVE ANY HEALTH PROBLEMS, BUT DESPITE HAVING 100% EYESIGHT, I HAVE STRONG SIDE EFFECTS: I SEE LARGE CIRCLES OF LIGHT AT ANY TIME OF THE DAY WHEN I AM ILLUMINATED BY LAMPS OR THE SUN. THE LIGHT CIRCLES ARE EXTREMELY BRIGHT, LARGE AND DEPEND ON HOW THE LIGHT FALLS INTO THE EYE. SOME TUNNELS OR AVENUES OF STREET LAMPS TRIGGER SEVERAL CIRCLES OF LIGHT ONE AFTER THE OTHER, WHICH ARE GETTING BIGGER AND BIGGER. FROM THE SUNLIGHT I ADDITIONALLY SOMETIMES SEE A SMALL CIRCLE OF LIGHT IN THE MIDDLE, WHICH APPEARS IN LIGHT RAINBOW COLORS. IN DAWN AND IN POORLY LIT ROOMS I SEE HIGH-CONTRAST EDGES / CONTOURS TWICE, WHEREAT THEY BLUR INTO EACH OTHER. THE DOUBLE EDGES ARE NOTICEABLY LIGHTER AND BLURRY. EVEN WHEN I WATCH TV IN THE EVENING, THE IMAGES KEEP BLURRING. DUE TO THE RISKS OF A SURGERY, I DO NOT WANT TO HAVE THE LENSES REMOVED PREMATURELY, BUT I AM VERY UNHAPPY WITH MY CURRENT CONDITION AND I REGRET THE SURGERY VERY MUCH." ON (B)(6) 2021 THE REPORTER STATED " WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WILL STAY IN THE. THE REPORTER STATED " THE LENS WILL STAY IN THE EYE, BECAUSE THE UCVA 20/16 IS PERFECT , PATIENT IS TRYING TO GET USED TO HALOS." CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -6.5/1.0/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY THE PATIENT EXPERIENCED GLARES; HALOES; BLURRY VISION AND LARGE CIRCLES OF LIGHT. ON (B)(6) 2021 THE REPORTER STATED "WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." AFTER SOME TIME REPORTS THAT "THE LENS WILL STAY IN THE EYE BECAUSE THE UCVA 20/16 IS PERFECT; PATIENT IS TRYING TO GET USED TO HALOS." LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11642627&pc=QCB | 2023826-2021-00944 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.5/1.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2020. THE REPORTER STATED " ONE YEAR AGO I GOT THE EVO VISIAN ICL IMPLANTED IN MY EYES (I'M MYOPIC WITH ASTIGMATISM). THE SURGERY WENT WELL AND MY EYES HEALED WELL. I DON'T HAVE ANY HEALTH PROBLEMS, BUT DESPITE HAVING 100% EYESIGHT, I HAVE STRONG SIDE EFFECTS: I SEE LARGE CIRCLES OF LIGHT AT ANY TIME OF THE DAY WHEN I AM ILLUMINATED BY LAMPS OR THE SUN. THE LIGHT CIRCLES ARE EXTREMELY BRIGHT, LARGE AND DEPEND ON HOW THE LIGHT FALLS INTO THE EYE. SOME TUNNELS OR AVENUES OF STREET LAMPS TRIGGER SEVERAL CIRCLES OF LIGHT ONE AFTER THE OTHER, WHICH ARE GETTING BIGGER AND BIGGER. FROM THE SUNLIGHT I ADDITIONALLY SOMETIMES SEE A SMALL CIRCLE OF LIGHT IN THE MIDDLE, WHICH APPEARS IN LIGHT RAINBOW COLORS. IN DAWN AND IN POORLY LIT ROOMS I SEE HIGH-CONTRAST EDGES / CONTOURS TWICE, WHEREAT THEY BLUR INTO EACH OTHER. THE DOUBLE EDGES ARE NOTICEABLY LIGHTER AND BLURRY. EVEN WHEN I WATCH TV IN THE EVENING, THE IMAGES KEEP BLURRING. DUE TO THE RISKS OF A SURGERY, I DO NOT WANT TO HAVE THE LENSES REMOVED PREMATURELY, BUT I AM VERY UNHAPPY WITH MY CURRENT CONDITION AND I REGRET THE SURGERY VERY MUCH." ON 06 APR 2021 THE REPORTER STATED " WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WILL STAY IN THE. THE REPORTER STATED " THE LENS WILL STAY IN THE EYE, BECAUSE THE UCVA 20/16 IS PERFECT , PATIENT IS TRYING TO GET USED TO HALOS." CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -6.5/1.0/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY THE PATIENT EXPERIENCED GLARES; HALOES; BLURRY VISION AND LARGE CIRCLES OF LIGHT. ON 06/04/2021 THE REPORTER STATED "WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." AFTER SOME TIME REPORTS THAT "THE LENS WILL STAY IN THE EYE BECAUSE THE UCVA 20/16 IS PERFECT; PATIENT IS TRYING TO GET USED TO HALOS." LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: SEX-RACE:UNK. DATE OF EVENT:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651557&pc=QCB | 2023826-2021-00883 | 2020-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -4.50/1.5/143 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. PER THE REPORTER ON (B)(6) 2021, NO PLANS TO IMPLANT ANOTHER LENS, UCVA OF THIS PATIENT WAS GREAT AFTER THE SURGERY, BUT THE PATIENT COULDN'T STAND THE HALO AND GLARE, PATIENT INSISTED TO TAKE OUT THE LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4), ACD <3.0MM, OVER 45 YRS OF AGE AT DATE OF IMPLANT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721921&pc=QCB | 2023826-2021-01040 | 2019-10-01 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778986&pc=QCB | 2023826-2021-01325 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/1.0/139 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT HAD SERIOUS NIGHTTIME DAZZLES AFTER SURGERY AND REQUESTED LENS BE REMOVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779037&pc=QCB | 2023826-2021-01326 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/1.00/051 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT HAD SERIOUS NIGHTTIME DAZZLES AFTER SURGERY AND REQUESTED LENS BE REMOVED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820836&pc=QCB | 2023826-2021-01405 | 2018-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.50/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS,INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. GLARE/HALOS WERE OBSERVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -8.50/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, LENS NATURALLY SITS VERY SLIGHTLY SUPERIOR/TEMPORAL. EYE DROPS (ALPHAGAN) DO NOT WORK. LENS REMAINS IMPLANTED. STATUS OF EYE; "ALL WITHIN NORMAL LIMITS." H6: HEALTH IMPACT CODE- 2199 IN INITIAL MDR IS CORRECTED TO 4644. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222777&pc=QCB | 2023826-2021-02889 | 2021-02-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.0/6.0/104, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR AND IT WAS UNKNOWN IF THE DEVICE FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222778&pc=QCB | 2023826-2021-02888 | 2020-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.0/1.5/115, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED LOW VAULT, LENS/DISLOCATION/SUBLUXATION AND GLARE/HALOES. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER STATED " LENS INITIALLY HAD A SMALL VAULT AND IS OFF-CENTERED AS A RESULT. IT IS CURRENTLY OFF-CENTERED TEMPORALLY INFERIORLY DISLOCATION. AXIAL POSITION AND OVER-OVER REFRACTION THEREFORE ARE NOT RELIABLE. PATIENT IS DETECTING HALOS AND GLARE DUE TO DISLOCATION." LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H6: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: B5: 12.6MM VTICM5 -13.0/1.5/104 (SPHERE/CYLINDER/AXIS). CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 - "THE LENS NO LONGER REMAINS IMPLANTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/ LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION, LENS DISLOCATION/SUBLUXATION, AND GLARE HALOES. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587386&pc=QCB | 2023826-2021-03898 | 2021-08-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.0/090, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587379&pc=QCB | 2023826-2021-03897 | 2021-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/1.0/075, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5- CORRECTION: INITIAL LENS WAS IMPLANTED (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K : THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12593044&pc=QCB | 2023826-2021-03919 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.0/+1.0/012, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED " THE PATIENT HAS AN UCVA OF 20/20 WITH ICLS BUT MASSIVE ISSUES WITH HALOS AND GLARE AT NIGHT. SO THE SURGEON EXPLANTED THE ICL AND NOW ALL IS FINE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE : UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12593615&pc=QCB | 2023826-2021-03920 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.0/+0.5/033, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED " THE PATIENT HAS AN UCVA OF 20/20 WITH ICLS BUT MASSIVE ISSUES WITH HALOS AND GLARE AT NIGHT. SO THE SURGEON EXPLANTED THE ICL AND NOW ALL IS FINE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12791947&pc=QCB | 2023826-2021-04321 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER STATED " 2 WEEKS POST-OP ICL PATIENT HAS COMPLAINTS OF HALO". PER UPDATED INFORMATION THE REPORTER STATED THAT THE PATIENT IS DOING QUITE WELL, WITH GLARE THAT CAN BE MANAGED WITH BRIMONIDINE. LENS REMIANS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DATE IMPLANTED: UNK. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -6.0/1.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. PATIENT COMPLAINTS OF HALO/GLARE POST ICL IMPLANTATION. PER (B)(6) 2021 EMAIL FROM REPORTED THE REPORTER STATES "PATIENT IS ACTUALLY DOING QUITE WELL, WITH GLARE THAT HE CAN MANAGE WITH BRIMONIDINE". PATIENT-RELATED FACTOR AND DEVICE WAS REPORTED FOR CAUSE OF EVENT. HOWEVER, IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. D4: (B)(6) 2023. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872823&pc=QCB | 2023826-2021-04493 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/011 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED (THE LENS WAS IMPLANTED VERTICALLY). THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHINICITY , RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872831&pc=QCB | 2023826-2021-04494 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED GLARE/ HALOES AND THE PROBLEM WAS RESOLVED. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14169183&pc=QCB | 2023826-2022-01222 | 2021-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE AND HALOES AND THE PATIENT FEELING UNCOMFORTABLE AND COULD NOT TOLERATE. AFTER REPEATED COMMUNICATION, THE PATIENT STILL INSISTED ON REMOVING THE LENS. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14169248&pc=QCB | 2023826-2022-01221 | 2021-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/015 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PATIENT FEELING UNCOMFORTABLE AND COULD NOT TOLERATE. AFTER REPEATED COMMUNICATION, THE PATIENT STILL INSISTED ON REMOVING THE LENS. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14734140&pc=QCB | 2023826-2022-01850 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO GLARE/HALOES. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15003801&pc=QCB | 2023826-2022-02199 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT COMPLAINED OF GLARE/HALOS. THE LENS WAS REPLACED WITH AN IOL LENS AND THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591507&pc=QCB | 2023826-2022-03594 | 2022-08-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM13.2, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, GALRE AND HALOES, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT A 13.2MM VTICM13.2" SHOULD BE CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2." CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 -8.50/1.0/130 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, GLARE AND HALOES. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030530&pc=QCB | 2023826-2022-04434 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS -9.5/1.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTD TO BE PATIENT-RELATED FACTOR. HOWEVER, THE REPORTER STATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16130089&pc=QCB | 2023826-2023-00094 | 2022-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.00/+1.5/087, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT COMPLAINED OF GLARE/HALOES AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301533&pc=QCB | 2023826-2023-00382 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00/5.5/047 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. "JUST IN CASE OF KERATOCONUS HIGH VAULT IS COMMON EVENT AND THIS CASES AS HISTORY OF KERRARING (INTRA CORNEAL RING)." Manufacturer Narrative: IMPLANT DATE: (B)(6) 2021. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301607&pc=QCB | 2023826-2023-00383 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -75/5.0/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. "JUST IN CASE OF KERATOCONUS HIGH VAULT IS COMMON EVENT AND THIS CASES AS HISTORY OF KERRARING (INTRA CORNEAL RING)." Manufacturer Narrative: IMPLANT DATE: (B)(6) 2021. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301634&pc=QCB | 2023826-2023-00384 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO : "THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/068 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN." IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2020. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301672&pc=QCB | 2023826-2023-00385 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: IMPLANT DATE: (B)(6) 2021. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). (B)(4). Manufacturer Narrative: B1 - CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/ DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO : SERIOUS INJURY IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: B1 - CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/ DAMAGE SHOULD BE REMOVED FROM MFR REPORT #: 2023826-2023-00385-SUPP1. H1 - CORRECTED TO : MALUNCTION IN THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16506881&pc=QCB | 2023826-2023-00625 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+2.5/168 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595253&pc=QCB | 2023826-2023-00996 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/051 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 , THE LENS WAS REMOVED AND REPLACED FOR A DIFFERENT MODEL, LONGER LENGTH LENS DUE TO LOW VAULT, LENS OPACITY, AND GLARE HALOES. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER. REPORTEDLY, "PATIENT STILL WITH LENS OPACITY SUPRISINGLY, BUT BETTER VAULT" Manufacturer Narrative: CORRECTED DATA: H6- 1494: OFF-LABEL USAGE- PRE-EXISTING CONDITION OF PROGRESSIVE MYOPIA. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434410&pc=QCB | 2023826-2023-03186 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -14.5/4.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 2(B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT STRONGLY REQUESTED REMOVAL AND NO REPLACEMENT. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17440152&pc=QCB | 2023826-2023-03185 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -13.0/3.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT STRONGLY REQUESTED REMOVAL AND NO REPLACEMENT. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495490&pc=QCB | 2023826-2023-03323 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0/0.5/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SPHERICAL LENS DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2140" SHOULD BE REMOVED AND "2227" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE REMOVED AND "3001" SHOULD BE ADDED. (B)(4) Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733767&pc=QCB | 2023826-2023-04036 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.00/ 3.0/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733799&pc=QCB | 2023826-2023-04037 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/145 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882968&pc=QCB | 2023826-2023-04447 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -16.00/1.0/120 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS AN EXCHANGE LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. HALOS/GLARE WITH NIGHTTIME DRIVING ARE REPORTED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17961829&pc=QCB | 2023826-2023-04557 | 2023-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067593&pc=QCB | 2023826-2023-04969 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075196&pc=QCB | 2023826-2023-04725 | 2023-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6. -11.5/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO NIGHTTIME GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: NIGHTTIME GLARE. REPORTEDLY, PATIENT REQUEST REMOVAL OF LENS. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075197&pc=QCB | 2023826-2023-04726 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6. -12.0/2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO NIGHTTIME GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: NIGHTTIME GLARE. REPORTEDLY, PATIENT REQUEST REMOVAL OF LENS. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18080758&pc=QCB | 2023826-2023-04992 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13/2.5/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18117390&pc=QCB | 2023826-2023-04761 | 2021-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+4.0/077 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND WAS NOT SATISFIED WITH RESULTS. THE PROBLEM WAS RESOLVED. POST-OP THE EYE WAS NORMAL AND WAS WITHOUT SIGNS AND SYMPTOMS. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136579&pc=QCB | 2023826-2023-04762 | 2021-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+5.5/103 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND WAS NOT SATISFIED WITH RESULTS. THE PROBLEM WAS RESOLVED. POST-OP THE EYE WAS NORMAL AND WAS WITHOUT SIGNS AND SYMPTOMS. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440045&pc=QCB | 2023826-2024-00048 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.00/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WITH GLARE/HALOES WAS REPORTED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440457&pc=QCB | 2023826-2024-00049 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0//055 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WITH GLARE/HALOES WAS REPORTED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129006&pc=QCB | 2023826-2021-00086 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.50/1.0/010(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129010&pc=QCB | 2023826-2021-00085 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.00/1.0/170(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTFIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368769&pc=QCB | 2023826-2021-00471 | 2020-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11386930&pc=QCB | 2023826-2021-00544 | 2020-12-22 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.00/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ICA, AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. "NORMAL" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. FOR ADDITIONAL INFORMATION "OS- VAULT 1000 OBSERVED" WAS REPORTED. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. DATE OF EVENT: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697117&pc=QCB | 2023826-2021-01089 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.0/+1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "UNREACTIVE PUPIL AND GLARE HALO DUE TO HIGH VAULT WERE CONFIRMED. THE LENS SIZE DID NOT FIT THE PATIENT." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697124&pc=QCB | 2023826-2021-01088 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.0/+1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "UNREACTIVE PUPIL AND GLARE HALO DUE TO HIGH VAULT WERE CONFIRMED. THE LENS SIZE DID NOT FIT THE PATIENT." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119230&pc=QCB | 2023826-2021-02331 | 2021-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS THEN EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOS, BLURRED VISION, AND SPARKLES. THE EYE "LOOKS GOOD" NOW. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: SPHERE/CYLINDER/AXIS OF THE IMPLANTED LENS IS -0.5/+1.5/049. THE LENS WAS IMPLANTED ON (B)(6) 2021. THE PATIENT CAN SEE CENTRAL RINGS. ON (B)(6) 2021 THE LENS WAS EXCHANGED. THE PROBLEM WAS NOT RESOLVED. SEE MFR REP# 2023826-2021-02969 FOR THE REPLACEMENT LENS. H6-HEALTH EFFECT CLINICAL CODE 4581 (SEES CENTRAL RINGS). H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. SERIAL NUMBER, EXPIRATION DATE: UNK. IMPLANT, EXPLANT DATE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MFG DATE - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669520&pc=QCB | 2023826-2021-04044 | 2021-09-24 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED DYSPHOTOPSIAS AT NIGHT, WITH MESOPIC PUPIL (6MM). THE PATIENT REPORTED GLARE/HALOS. THE PATIENT WAS TREATED WITH ALPHAGAN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15296711&pc=QCB | 2023826-2022-02793 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00/+0.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS QUIET. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: EXPIRATION DATE: 31-OCT-2024 (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307873&pc=QCB | 2023826-2022-02794 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00/+0.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS QUIET. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4), SHOULD BE ADDED TO INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPIRATION DATE: 30-APR-2025. HEALTH IMPACT- CLINICAL CODE: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15476840&pc=QCB | 2023826-2022-03279 | 2022-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18/1.00/106 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON . THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE AND GLARE/HALOES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "ERROR ON LENS DIAMETER CALCULATION. ANGLE CLOSURE ON ONE CUADRANT (NASAL)". Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939032&pc=QCB | 2023826-2023-01654 | 2023-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.0/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE/HALOES, AND EXCESSIVE VAULT HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. PATIENT REMAINS UNDER OBSERVATION. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, PATIENTS ANATOMY MAY BE ABNORMAL OR A MEASUREMENT ERROR. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -7.00/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARES/HALOES. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "NARROW ANGLE, HIGH VAULT, NO OTHER SYMPTOMS, MONITORING FOR NOW". ADDITIONAL INFORMATION PROVIDED: "WOULD LIKE TO REQUEST A CONSULTATION WITH AN MD TO DISCUSS THE SCENARIO AND TO RECEIVE MEDICAL PERSPECTIVE ON HIS PLANT TO OBSERVE." THE CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "HIGH VAULT OD DESPITE ACCURATE PREOP MEASUREMENTS, PATIENTS ANATOMY MAY BE ABNORMAL; OR A MEASUREMENT ERROR. " CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939036&pc=QCB | 2023826-2023-01655 | 2023-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.5/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE/HALOES, AND EXCESSIVE VAULT HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. PATIENT REMAINS UNDER OBSERVATION. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, PATIENTS ANATOMY MAY BE ABNORMAL OR A MEASUREMENT ERROR. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -12.0/+1.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARES/HALOES. THE LENS REMAINS IMPLANTED, AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "NARROW ANGLE, HIGH VAULT, NO OTHER SYMPTOMS, MONITORING FOR NOW". ADDITIONAL INFORMATION PROVIDED: "WOULD LIKE TO REQUEST A CONSULTATION WITH AN MD TO DISCUSS THE SCENARIO AND TO RECEIVE MEDICAL PERSPECTIVE ON HIS PLANT TO OBSERVE." THE CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "HIGH VAULT OS DESPITE ACCURATE PREOP MEASUREMENTS, PATIENTS ANATOMY MAY BE ABNORMAL; OR A MEASUREMENT ERROR." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -10.50/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) HE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARES/HALOES. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "NARROW ANGLE, HIGH VAULT, NO OTHER SYMPTOMS, MONITORING FOR NOW". ADDITIONAL INFORMATION PROVIDED: "WOULD LIKE TO REQUEST A CONSULTATION WITH AN MD TO DISCUSS THE SCENARIO AND TO RECEIVE MEDICAL PERSPECTIVE ON HIS PLANT TO OBSERVE." THE CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "HIGH VAULT OS DESPITE ACCURATE PREOP MEASUREMENTS, PATIENTS ANATOMY MAY BE ABNORMAL; OR A MEASUREMENT ERROR. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17138291&pc=QCB | 2023826-2023-02315 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/046 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, UNREACTIVE (FIXED) PUPIL AND GLARE/HALOES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WERE NO COMPLAINTS, EYE WAS CALM AND THE PUPIL WAS REACTIVE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - UNREACTIVE (FIXED) PUPIL. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: A PI WAS PERFORMED ON (B)(6) 2023 AND WAS THE FIRST PI (NO NEED OF BEFORE PI - MYOPIC LENS). H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ADDITION OF PI. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581- WIDE PUPIL " SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER-18.0/+1.5/46 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, 'WIDE' PUPIL, AND GLARES/HALOES. ON (B)(6) 2023 PI WAS PERFORMED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER, SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE: "NO COMPLAINTS, CALM EYE, REACTIVE PUPIL (STATUS VERY GOOD)". THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "TO BIG SIZE OF THE LENS UNUSUAL STRUCTURE OF THE EYE". CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11887128&pc=QCB | 2023826-2021-01641 | 2021-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING ROTATED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED TOGETHER WITH THE PATIENT, THE DOCTOR DECIDED NOT TO PERFORM A SURGICAL INTERVENTION. ALL IS GOOD AND THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: IN INITIAL MDR THE WORK ORDER SEARCH RESULTS SHOULD BE CORRECTED TO: NO SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6 MEDICAL DEVICE CODE- 2923 CODE IN INITIAL MDR IS NOT APPLICABLE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: AT A POST-OP EXAMINATION ON (B)(6) 2021, THE SURGEON DETECTED AN EXCESSIVE VAULTING. THERE WAS PIGMENT DISPERSION AND PATIENT COMPLAINED OF GLARE/HALOS. H6: HEALTH IMPACT CLINICAL CODE: 4581 PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PIGMENT DISPERSION, AND GLARE/HALOS. THE PROBLEM WAS RESOLVED AND THE "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12756627&pc=QCB | 2023826-2021-04227 | 2021-09-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Discomfort | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7,-12.5/2.5/101, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PATIENT EXPERIENCED EXCESSIVE VAULT AND GLARE/HALOES, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN, CAUSE DETAILS " LENS IS TO BIG FOR PATIENT" WAS REPORTED. PER UPDATED INFORMATION THE REPORTER STATED " PATIENT IS JUST DISCOMFORT WITH THE VISUAL ACUITY IOP IS FINE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "PATIENT ALSO EXPERIENCED DISCOMFORT". CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17057307&pc=QCB | 2023826-2023-02178 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/4.5/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/4.5/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT, GLARES, AND HALOES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17091344&pc=QCB | 2023826-2023-02233 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/044 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR DIFFERENT MODEL, SHORTER LENGTH LENS, DUE TO EXCESSIVE VAULT AND GLARE/HALOES. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "2227" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519001&pc=QCB | 2023826-2019-00612 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WASEXPLANTED DUE TO GLARE/HALOS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519013&pc=QCB | 2023826-2019-00608 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552494&pc=QCB | 2023826-2019-00659 | 2019-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8746386&pc=QCB | 2023826-2019-01163 | 2019-05-04 | Injury | STAAR SURGICAL COMPANY | 2019-06-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/096 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT COULD NOT TOLERATE GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895522&pc=QCB | 2023826-2019-01453 | 2019-03-02 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM WAS RESOLVED, PATIENT CONDITION WAS STABLE. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966508&pc=QCB | 2023826-2019-01647 | 2019-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS EXCESSIVE VAULT, GLARE/HALOS, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) AND IRIS PROLAPSE. THE CAUSE OF THE EVENT IS REPORTED AS "NOMOGRAM ERROR." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DATE OF EVENT CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: EVENT DESCRIPTION CORRECTED TO STATE THAT A 13.7MM VTICM5 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.00/6.0/031 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN ASSOCIATED LOTS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966519&pc=QCB | 2023826-2019-01646 | 2019-05-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS EXCESSIVE VAULT, GLARE/HALOS, AND NEGATIVE DYSPHOTOPSIA. THE CAUSE OF THE EVENT IS REPORTED AS "NOMOGRAM ERROR." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPES WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# : (B)(4). Manufacturer Narrative: CORRECTED TO STATE THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50/1.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). CLAIM# : (B)(4). Manufacturer Narrative: DATE OF EVENT: CORRECTED TO (B)(6) 2019 IN INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: INITIAL MDR SHOULD HAVE INCLUDED PATIENT CODE 3191: NEGATIVE DYSPHOTOPSIA. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK. MODEL #, IMPLANTED, EXPLANTED: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8969999&pc=QCB | 2023826-2019-01662 | 2018-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -0.50/2.0/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. HALOS AND "LIGHTSPRED" WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTER INDICATED THAT "THE SURGEON DELAYED POSSIBLE LENS EXCHANGE DUE TO VERY GOOD REFRACTIVE RESULT. PATIENT WOULD LIKE TO EXCHANGE LENS, STILL HAVING DAILY PROBLEMS WITH GLARE." PER THE REPORTER ON (B)(6) 2019 EXCHANGE WILL NOT HAPPEN UNTIL END OF (B)(6). Manufacturer Narrative: PATIENT GENDER: UNK SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT. ADVERSE EVENT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL BUT SAME POWER/SIZE LENS AND THE PROBLEM WAS RESOLVED. PATIENT STATUS WAS REPORTED AS "GOOD". EXPLANT DATE: (B)(6) 2019. TYPE OF REPORTABLE EVENT: SERIOUS. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115540&pc=QCB | 2023826-2019-01765 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -9.00/1.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. GLARE/HALOES WERE OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "LENS WAS REMOVED DUE TO STRONG HALOES. PATIENT WEARS GLASSES AGAIN." Manufacturer Narrative: CORRECTED DATA: PATIENT CODE 2227 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: DATE OF EVENT UPDATED TO (B)(6) 2019 IN INITIAL MDR. PER UPDATED INFORMATION, THE CAUSE OF THE EVENT IS DUE TO A PATIENT RELATED FACTOR. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL STUCK TO A PIECE OF PAPER WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, BROWN RESIDUE ON THE LENS AND THE LENS RETURNED STUCK TO A PIECE OF PAPER WITH INK. THE LENS HAS BLUISH/ PURPLE TINT POSSIBLY FROM THE INK ON THE PAPER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115541&pc=QCB | 2023826-2019-01764 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARE/HALOES WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "LENS WAS REMOVED DUE TO STRONG HALOES. PATIENT WEARS GLASSES AGAIN." Manufacturer Narrative: PATIENT CODE 2227 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. DATE OF EVENT: UPDATED TO (B)(6) 2019 IN INITIAL MDR. EVENT DESCRIPTION: PER UPDATED INFORMATION, THE CAUSE OF THE EVENT IS DUE TO A PATIENT RELATED FACTOR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RED RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203032&pc=QCB | 2023826-2019-01953 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -08.00/1.0/133 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARE/HALOES WERE OBSERVED. THE LENS REMAINS IMPLANTED. REPORTEDLY SURGEON NOTED "HALOS/GLARE PRESENT BOTH WITH SMALL AND DILATED PUPIL. RELATED IN MY OPINION TO CENTRAL HOLE." CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE, SMALL AND DILATED PUPIL. RESULT CODE 114 CORRECTED TO RESULT CODE 213. PER UPDATED INFORMATION, THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, LONGER LENGTH, DIFFERENT POWER LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM#(B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED INA MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: RESULT CODE 114 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252703&pc=QCB | 2023826-2019-01960 |  | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+0.5/170 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE PUPIL SIZE WAS LARGE/REACTIVE. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 2 MONTHS. Manufacturer Narrative: CORRECTED DATA: EVENT - "THE PATIENT EXPERIENCES GLARE/HALOES DURING DAY AND NIGHT TIME." SHOULD BE ADDED IN THE INITIAL MDR." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: NA THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252792&pc=QCB | 2023826-2019-01961 |  | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/015 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE PUPIL SIZE WAS LARGE/REACTIVE. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 2 MONTHS. Manufacturer Narrative: CORRECTED DATA: "THE PATIENT EXPERIENCES GLARE/HALOES DURING DAY AND NIGHT TIME." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492810&pc=QCB | 2023826-2019-02418 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -12.00/2.0/091 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT THE EXCHANGED LENS RESOLVED THE PROBLEM BUT "STILL A LITTLE HIGH VAULT". Manufacturer Narrative: EVENT: "GLARE AND HALOS WERE OBSERVED." SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492867&pc=QCB | 2023826-2019-02419 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -12.00/2.0/089 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT THE EXCHANGED LENS RESOLVED THE PROBLEM BUT "STILL A LITTLE HIGH VAULT". ABNORMAL SULCUS TO SULCUS WIDTH WAS REPORTED AS CAUSE OF EVENT. Manufacturer Narrative: CORRECTED DATA: B5- " GLARE AND HALOS WERE OBSERVED." SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498138&pc=QCB | 2023826-2019-02410 | 2019-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARES AND HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO SIZING VARIATION IN OUTCOME. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633747&pc=QCB | 2023826-2020-00160 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.50/4.5/083 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARES/HALOES WERE OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY AFTER EXPLANT (REMOVAL)." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 3191 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881022&pc=QCB | 2023826-2020-00659 | 2019-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916667&pc=QCB | 2023826-2020-00735 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON THE LENS AND LENS WAS RETURNED IN PIECES - UNABLE TO EVALUATE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916671&pc=QCB | 2023826-2020-00736 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+3.5/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105254&pc=QCB | 2023826-2020-01079 | 2019-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.00/1.5/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO GLARE/HALOES AND LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138815&pc=QCB | 2023826-2020-01192 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem; No Apparent Adverse Event | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED AFTER IMPLANTABLE COLLAMER LENS SURGERY. STARBURST AND FLOATERS ARE BEING OBSERVED. REPORTEDLY, HIGH REFRACTIVE POWER (AROUND -16 D) AND PATIENT NOTED STARBURST, FLOATERS AND HALOS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.50/1.5/039 (SPERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT HAS EXPERIENCED STARBURST, FLOATERS, AND GLARE/HALOES FOR THE PAST TWO YEARS. THE PATIENT HAS BEEN REGULARLY VISITING THE SURGEON SINCE IMPLANTATION. THE LENS REMAINS IMPLANTED. H6 - DEVICE CODE 3189 IS NOT APPLICABLE IN INITIAL MDR H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PRECLUDE PERMANENT IMPAIRMENT/DAMAGE IN INITIAL MDR. B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.50/1.5/039 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT HAS EXPERIENCED STARBURST, FLOATERS, AND GLARE/HALOES FOR THE PAST TWO YEARS. THE CLINIC SUGGESTED TO USE PILO (MEDICATION) FOR A MONTH (DURING NIGHT TIME). THE PATIENT WAS FINE FOR 3 DAYS BUT STARTED HAVING THE SAME PROBLEM BEYOND THAT. THE PATIENT HAS BEEN REGULARLY VISITING THE SURGEON SINCE IMPLANTATION. THE LENS REMAINS IMPLANTED. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL AND SUPPLEMENTAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4644 - PILO MEDS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184788&pc=QCB | 2023826-2020-01309 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.00/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184802&pc=QCB | 2023826-2020-01310 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.50/3.0/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491985&pc=QCB | 2023826-2020-01969 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/108 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SMALL ANTERIOR CHAMBER AND GLARE/HALOS. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS ON THE LENS. H6- DEVICE IMPACT CODE 4629- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: "H6- DEVICE HISTORY RECORD- BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK WORD(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." SHOULD HAVE BEEN IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10585633&pc=QCB | 2023826-2020-02222 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Visual Disturbances; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, "PATIENT STILL HAVING SOME GHOSTING ON THE LETTERS ON THE CHART NO MATTER WHERE THE AXIS IS TURNED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -5.00/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THERE WAS LENS ROTATION NOT ASSOCIATED WITH LOW VAULT WITH SOME "GHOSTING OF LETTERS ON THE CHART NO MATTER WHERE THE AXIS IS TURNED", AND ON (B)(6) 2020 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE REPORT INDICATED THAT "NORMAL POST OP COURSE. PATIENT STILL HAS A SUTURE OS AT AXIS 85". CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. G3 - DATE IN INITIAL MEDWATCH REPORT SHOULD BE CORRECTED TO 28-AUG-2020. H6 - HEALTH EFFECT CLINICAL CODE: 2140 SHOULD BE CORRECTED TO 4582. H6 - HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SUTURE. H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823660&pc=QCB | 2023826-2020-02592 | 2020-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: IT WAS LATER CLARIFIED THAT THE PATIENT ALSO EXPERIENCED REFRACTIVE SURPRISE WHICH WAS ALSO RESOLVED WITH THE NOTED LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: H6: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE PREVIOUSLY BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439778&pc=QCB | 2023826-2022-01596 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Unintended Movement; No Apparent Adverse Event | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATES THAT A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -8.0/+2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, THE LENS FLIPPED AND WAS REMOVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SURGEON IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY A 12.6MM VICM5_12.6 -8.00/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 DUE TO LENS BEING IMPLANTED UPSIDE DOWN. THIS RESOLVED THE PROBLEM. THE STATUS OF EYE IS REPORTED AS " BLURRY VISION AT NIGHT, PATIENT SEE HALOES AND GLARE AT NIGHT." CAUSE OF THE EVENT IS REPORTED AS "DEVICE". H6: DEVICE CODE: 1494: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SURGEON IMPLANTED AND FLIPPED A 12.6MM VICM5_12.6 -8.00/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. ON THE SAME DAY BUT DIFFERENT SURGERY, THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. THE STATUS OF EYE IS REPORTED AS " BLURRY VISION AT NIGHT, PATIENT SEE HALOES AND GLARE AT NIGHT." CAUSE OF THE EVENT IS REPORTED AS "DEVICE". CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION- UNK. IMPLANT DATE, EXPLANT DATE-UNK. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14887046&pc=QCB | 2023826-2022-02097 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.00/+4.50/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14889364&pc=QCB | 2023826-2022-02098 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.50/+4.50/81 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8688274&pc=QCB | 2023826-2019-01077 | 2019-04-01 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -10/0.5/75 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER THE EXPLANT, PATIENTS CURRENT CONDITION IS WELL AND HAPPY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8688384&pc=QCB | 2023826-2019-01078 | 2019-04-01 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -8.0/0.5/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER THE EXPLANT, PATIENTS CURRENT CONDITION IS WELL. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10053827&pc=QCB | 2023826-2020-01098 | 2019-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-18.0/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO A PATIENT RELATED ISSUE AND THE PROBLEM RESOLVED. THE REPORTER STATED " THE PATIENT WAS UNABLE TO ADAPT TO THE GLARE AFTER THE OPERATION AND ASKED TO REMOVE THE LENS AND CANCEL THE OPERATION." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- RESULT CODE IN INITIAL MDR SHOULD BE 213. H6- RESULT CODE 114 SHOULD BE IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138800&pc=QCB | 2023826-2020-01190 | 2020-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020 AND (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. IN THE REPORTER'S OPINION, THE CAUSE OF THE EVENT IS ATTRIBUTED TO THE DEVICE. "LENS SIZING IS SMALL, CAUSING LENS TO ROTATE." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B5 - IT WAS REPORTED PATIENT ALSO EXPERIENCED DOUBLE VISION WHICH WAS RESOLVED WITH THE LENS EXCHANGE. THIS SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10369229&pc=QCB | 2023826-2020-01678 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/3.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED HIGH GLARE IMMEDIATELY AFTER REMOVING THE DRESSING ONE WEEK AFTER SURGERY AND IS STILL SUFFERING FROM HIGH GLARE. THE REPORTER STATED "ANNOYING GLARE ESPECIALLY AT NIGHT, ESPECIALLY WHEN DRIVING. PATIENT DECLARES AN IMPROVEMENT DURING THE INSTALLATION OF THE PILOCARPINE BUT UPON STOPPING THE GLARE RESUMES."ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15596965&pc=QCB | MW5112605 | 2022-08-17 04:00 | Injury | STAAR SURGICAL AG | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Compatibility Problem | Visual Impairment |  | 1 | Event Description: THE PATIENT USED ONE DROP IN EACH EYE OF VUITY. SHORTLY AFTER ADMINISTERING THE DROPS, THE REFRACTIVE ERROR HAD A VERY SIGNIFICANT MYOPIC SHIFT. THE PATIENT'S PRESCRIPTION WAS ABOUT -1.00 DIOPTERS PRIOR TO INSTILLING THE MEDICATION. WITHIN 30 MINUTES OF USING THE EYEDROPS, THE PRESCRIPTION CHANGED TO GREATER THAN -6.00 DIOPTERS, RENDERING THE PATIENT ILLEGAL TO DRIVE AND UNABLE TO CLEARLY VIEW ANYTHING BEYOND SIX INCHES FROM HIS NOSE. THE PATIENT HAD IMPLANTABLE COLLAMER LENS SURGERY IN 2012. IT IS PRESUMED THAT THE PUPIL CONSTRICTION CAUSED A SHIFT IN LOCATION OF THE INTRAOCULAR LENS, MAKING A LARGE MYOPIC SHIFT IN THE REFRACTIVE ERROR. FORTUNATELY, THIS WAS REVERSIBLE BY DILATING THE PATIENT'S EYES WITH ONE DROP OF 0.5% TROPICAMIDE AND VISION RETURNED TO ITS PREVIOUS STATE. ALLERGAN. FDA SAFETY REPORT ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11653722&pc=QCB | 2023826-2021-00992 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Visual Impairment | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A CASE REPORT "REPEATED ROTATION OF A TORIC IMPLANTABLE COLLAMER LENS". A (B)(6) MAN UNDERWENT TICL IMPLANTATION (12.1MM) FOR THE CORRECTION OF HIGH MYOPIC ASTIGMATISM IN EYES. HE PRESENTED A SUDDEN DECREASE IN VISUAL ACUITY (VA) OF THE LEFT EYE FOUR MONTHS AFTER TICL IMPLANTATION. A ROTATION OF THE TICL WAS DIAGNOSED. THE TICL WAS RE-SET TO THE ORIGINAL POSITION. TEN MONTHSTICL RELOCATION, THE PATIENT AGAIN PRESENTED A SUDDEN DECREASE IN VA. AGAIN THE TICL HAD ROTATED. THE TICL LENS WAS EXPLANTED AND A NEW TICL LENS WAS IMPLANTED. THE LENS REMAINS STABLE IN THE SIX MONTH FOLLOW-UP PERIOD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628804&pc=QCB | 2023826-2021-03986 | 2021-04-29 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Impairment | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2M VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.50/+6.0/160 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED PERMANENT LOSS OF BCVA. THE LENS WAS REPLACED ON (B)(6) 2021 WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (THE WRONG REFRACTION WAS TYPED INTO THE ICL FORM). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. G4: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628810&pc=QCB | 2023826-2021-03990 | 2021-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Impairment | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+1.0/090, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED ABOUT HIS VISION AND IT WAS FOUND THE WRONG LENS WAS IMPLANTED IN ERROR. THE LENS WAS REPLACED WITH THE CORRECT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: AFTER THE LENS WAS REPLACED WITH THE CORRECT LENS, THE PROBLEM WAS RESOLVED. THERE WERE NO SYMPTOMS AND NO PATIENT COMPLAINT. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. CORRECTED DATA: B3: 13-SEP-2021. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365144&pc=QCB | 2023826-2022-01479 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Impairment; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17200399&pc=QCB | 2023826-2023-02519 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Vitreous Floaters | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED, A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT COMPLAINED OF VITREOUS FLOATERS POST-OP. AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED (PRE-EXISTING CONDITION OF FLOATERS). THE LENS REMAINED IMPLANTED AND THE EYE IS HEALTHY. CLAIM # (B)(4).