## Original_QCB
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13765546&pc=QCB | 2023826-2022-00804 | 2021-10-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event |  |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.0/2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. INTRAOPERATIVELY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10524336&pc=QCB | 2023826-2020-02038 | 2020-08-04 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/1.50/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WITH 'CENTRAL TOUCH TO THE CRYSTALLINE LENS ANTERIOR CAPSULE' WERE REPORTED. THE LENS REMAINS IMPLANTED WITH PLANNED EXCHANGE FOR A LONGER SIZE LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10524379&pc=QCB | 2023826-2020-02037 | 2019-12-14 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.50/2.50/146 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WITH 'CENTRAL TOUCH TO CRYSTALLINE LENS ANTERIOR CAPSULE' WERE REPORTED. THE LENS REMAINS IMPLANTED WITH A PLANNED EXCHANGE FOR A LONGER LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10532967&pc=QCB | 2023826-2020-02043 | 2020-05-31 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/1.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE LENS WAS LATER REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE US.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10529725&pc=QCB | 2023826-2020-02073 | 2020-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS LOW VAULTING WITH LENS ROTATION, AND NARROW ANGLE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO ANOMALY IN THE SULCUS STRUCTURE, SO THE LENS WOULD NOT STABILIZE. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY, REFRACTIVE TREATMENT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548153&pc=QCB | 2023826-2020-02099 | 2020-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.50/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, REFRACTIVE SURPRISE, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS "VAULTING WAS HIGHER THAN EXPECTED. THERE IS NO HEALTH HAZARDS." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: HEALTH EFFECT CLINICAL CODE - 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. SHOULD HAVE BEEN INCLUDED IN H10 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10639547&pc=QCB | 2023826-2020-02308 | 2020-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -14/6/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS DUE TO PATIENT-RELATED FACTOR, THE REPORTER STATED " THE PROBLEM IS WITH THE SULCUS, NARROW SULCUS". Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICO-CENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706892&pc=QCB | 2023826-2020-02408 | 2020-08-29 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.50/+6.0/009 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, LENS ROTATION, REDUCED ANTERIOR CHAMBER ANGLES AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708893&pc=QCB | 2023826-2020-02440 | 2020-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708902&pc=QCB | 2023826-2020-02439 | 2020-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/098 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091992&pc=QCB | 2023826-2020-03222 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.5/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING (THE LENS VAULTED FORWARD), ANGLE CLOSURE AND LENS ROTATION. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S EYE IS QUIET AND THE PATIENT IS HAPPY WITH THE OUTCOME. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.5/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT; ANGLE CLOSURE; ICL ROTATION WITH A HYPEROPIC REFRACTION OUTCOME. THE EXCHANGE RESOLVED THE PROBLEM. H6- HEALTH EFFECT- CLINICAL CODE: 4581- HYPEROPIC REFRACTIVE OUTCOME CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11043450&pc=QCB | 2023826-2020-03083 | 2020-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2,-4.5/+1.0/26 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS PATIENT-RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.5/+1.0/26 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGED RESOLVED THE PROBLEM. ALSO, IN THE REPORTERS OPINION, THE CAUSE OF THIS EVENT IS PATIENT-RELATED FACTOR. H6: IMPACT CODE UPDATED TO 4629. CORRECTIVE DATA: D4: MODEL# WAS UPDATED TO VTICM5_13.2. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980473&pc=QCB | 2023826-2020-02942 | 2020-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE /DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: 4581 (SIGNIFICANT REDUCTION OF ICA). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980497&pc=QCB | 2023826-2020-02925 | 2020-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +3.5/+4.0/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980526&pc=QCB | 2023826-2020-02941 | 2020-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE /DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970021&pc=QCB | 2023826-2020-02891 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852580&pc=QCB | 2023826-2020-02683 | 2020-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+3.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT: CLINICAL CODE: (B)(4); SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852617&pc=QCB | 2023826-2020-02684 | 2020-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+3.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS STABLE BUT THERE STILL WAS A HIGH VAULT - SEE MFR REPORT # 2023826-2020-02685 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND THREADS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10824328&pc=QCB | 2023826-2020-02648 | 2020-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BENT WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116953&pc=QCB | 2023826-2021-00030 | 2020-09-17 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.00/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT DATE: UNK. DATE RECEIVED BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116959&pc=QCB | 2023826-2021-00031 | 2020-09-10 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM513.2, -14.00/4.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT DATE: UNK. DATE RECEIVED BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129384&pc=QCB | 2023826-2021-00082 | 2020-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.50/1.0/080(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11168676&pc=QCB | 2023826-2021-00138 | 2020-10-19 | Injury | STAAR SURGICAL COMPANY | 2021-01-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -3.5/3.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11189749&pc=QCB | 2023826-2021-00148 | 2020-11-16 | Injury | STAAR SURGICAL COMPANY | 2021-01-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.50/+4.0/087 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE VAULT WAS REDUCED AND AC LOOKS DEEP. THE PATIENT IS DOING WELL AND IS HAPPY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11189756&pc=QCB | 2023826-2021-00145 | 2020-11-02 | Injury | STAAR SURGICAL COMPANY | 2021-01-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/067 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT WAS REDUCED AND THE EYE IS QUIET. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11189761&pc=QCB | 2023826-2021-00144 | 2020-11-02 | Injury | STAAR SURGICAL COMPANY | 2021-01-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/104 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT WAS REDUCED AND THE EYE IS QUIET. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11197087&pc=QCB | 2023826-2021-00177 | 2020-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.5/+2.0/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND REPLACED IN A SEPARATE SURGERY WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: UNRELIABLE NOMOGRAM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: 4581 (SIGNIFICANT REDUCTION OF ICA). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11280315&pc=QCB | 2023826-2021-00260 | 2020-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.0/+1.0/177 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ENDOTHELIAL TOUCH AND ANGLE CLOSURE. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY INSIDE A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11290267&pc=QCB | 2023826-2021-00303 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.50/3.5/013 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY WHICH RESOLVED THE PROBLEM. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.50/3.5/013 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY WHICH RESOLVED THE PROBLEM. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11290269&pc=QCB | 2023826-2021-00304 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.50/5.5/018 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY WHICH RESOLVED THE PROBLEM. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_ 12.6, -4.50/5.5/018 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY WHICH RESOLVED THE PROBLEM. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11311339&pc=QCB | 2023826-2021-00379 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.00/3.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AS A REPLACEMENT LENS TO CORRECT EXCESSIVE VAULT AND ELEVATED IOP. IT WAS REPORTED THAT THE LENS RESOLVED THE PROBLEM BUT FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER REPORTED "PIGMENT DISPERSED ON NATURAL LENS AND CORNEAL ENDOTHELIUM." IT WAS REPORTED THAT THE PIGMENT OBSERVED WAS LIKELY FROM PI, ICL EXCHANGE. IT WAS REPORTED THAT THE PIGMENT DISPERSED ON NATURAL LENS AND CORNEAL ENDOTHELIUM HAS RESOLVED, PATIENT IS STABLE AND NO TREATMENT OR INTERVENTION HAS BEEN PERFORMED SINCE EXCHANGE. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSED AND CORNEAL ENDOTHELIUM. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11322755&pc=QCB | 2023826-2021-00421 | 2020-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.5/2.0/088, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343190&pc=QCB | 2023826-2021-00403 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. REPORTEDLY, RESIDUAL ASTIGMATISM WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368811&pc=QCB | 2023826-2021-00467 | 2020-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS NICE AND QUIET. THE CAUSE OF THE EVENT WAS DUE TO WRONG SIZE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368830&pc=QCB | 2023826-2021-00466 | 2020-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS NICE AND QUIET. THE CAUSE OF THE EVENT WAS DUE TO WRONG SIZE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL DRY AND STUCK TO A PIECE OF TAPE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE LENS STUCK TO A PIECE OF TAPE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE RECD BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416124&pc=QCB | 2023826-2021-00502 | 2020-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS "BORDER LINE WTW MEASUREMENT LED TO BIGGER LENS SIZE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11441231&pc=QCB | 2023826-2021-00623 | 2021-01-30 | Injury | STAAR SURGICAL COMPANY | 2021-03-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT GOT REDUCED TO NORMAL AND ANGLES ARE OPEN. PATIENT IS COMFORT AND DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11441236&pc=QCB | 2023826-2021-00622 | 2021-01-30 | Injury | STAAR SURGICAL COMPANY | 2021-03-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT GOT REDUCED TO NORMAL AND ANGLES ARE OPEN. PATIENT IS COMFORT AND DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558793&pc=QCB | 2023826-2021-00719 | 2021-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/003 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "LENS CALCULATION ADVISED A TOO BIG ICL." Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558799&pc=QCB | 2023826-2021-00705 | 2021-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.5/+4.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.5/+4.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES (ICA). THE LENS WAS EXCHANGED FOR A DIFFERENT TORIC MODEL SHORTER IN LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11585607&pc=QCB | 2023826-2021-00809 | 2021-02-06 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+1.0/089 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. HIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11606789&pc=QCB | 2023826-2021-00837 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-04-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/1.5/090 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616791&pc=QCB | 2023826-2021-00843 | 2021-03-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -07.00/+4.0/121 (SPHERE/CYLINDER/AXIS). THE LENS WAS PARTLY INSERTED AND WAS REMOVED, SO IT COULD BE RELOADED. THE SURGEON NOTED AT THAT TIME THAT THERE WAS PIGMENT ON THE LENS. THE REPORTER INDICATED THE LENS WAS LOADED IMPROPERLY. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS TORN UPON DELIVERY INTO THE EYE AND THERE WAS NO PATIENT INJURY. THE PROBLEM WAS RESOLVED WITH THE IMPLANT OF THE BACKUP LENS. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11648974&pc=QCB | 2023826-2021-00948 | 2021-03-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50/+5.0/086 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: A2: DATE OF BIRTH: (B)(6) 2000. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11653731&pc=QCB | 2023826-2021-00949 | 2021-03-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.00/+5.0/087 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: A2: DATE OF BIRTH: (B)(6) 2000. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11653733&pc=QCB | 2023826-2021-00988 | 2021-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+1.5/095 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE ICL WAS TOO LONG FOR THE EYE, WHICH MAY BE RELATED TO THE FACT THE PATIENT WAS (B)(6). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE SURGEON INDICATED THE PATIENT IS DOING FINE NOW. THE LENS WAS EXPLANTED. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11714192&pc=QCB | 2023826-2021-01113 | 2020-12-01 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/1.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS LATER EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11713226&pc=QCB | 2023826-2021-01142 | 2018-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT HAD SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS OVER A PERIOD OF 3 YEARS. THE SURGEON HAS DECIDED TO LEAVE THE LENS IMPLANTED AND THE PATIENT WILL BE MONITORED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: DATE OF EVENT: UNKNOWN. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721499&pc=QCB | 2023826-2021-01159 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2021-04-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -09.00/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11734844&pc=QCB | 2023826-2021-01178 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN THE DELTA JOURNAL OF OPHTHALMOLOGY WAS RECEIVED ON 01-APR-2021, ENTITLED 'ENDOTHELIAL CELL LOSS FOLLOWING TORIC IMPLANTABLE COLLAMER LENS IMPLANTATION FOR CORRECTION OF MYOPIA AND ASTIGMATISM AFTER PENETRATING KERATOPLASTY.' THE ARTICLE STATES THAT VISION TICL WERE IMPLANTED IN 12 EYES OF EIGHT PATIENTS WHO UNDERWENT PKP FOR KERATOCONUS. REPORTEDLY, 6-MONTHS POSTOPERATIVELY ENDOTHELIAL CELL DENSITY (ECD) DECREASED BY -5.09%. AT 1-YEAR POSTOPERATIVELY ECD DECREASED BY -9.61%. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11738879&pc=QCB | 2023826-2021-01193 | 2018-10-30 | Injury | STAAR SURGICAL COMPANY | 2021-04-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -6.50/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PIGMENT DISPERSION. THIS RESOLVED THE PROBLEM. REPORTEDLY, IN ORDER TO PREVENT THE ADVERSE EFFECTS CAUSED BY LONG-TERM EXISTENCE IN THE FUTURE, THE LENS WAS REMOVED FROM POSTERIOR CHAMBER INTRAOCULAR LENS WITH LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11738956&pc=QCB | 2023826-2021-01200 | 2020-09-10 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT. THE REPLACEMENT LENS WAS A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11738883&pc=QCB | 2023826-2021-01190 | 2021-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -7.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED. THE LENS REMAINS IMPLANTED. REPORTEDLY, IRIS ALMOST TOUCHING ENDOTHELIUM IN THE TEMPORAL SECTOR. MEDICATION FOR IOP MANAGEMENT HAS BEEN PRESCRIBED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -7.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT. THE THIS DID NOT RESOLVE THE PROBLEM. REPORTEDLY, IRIS ALMOST TOUCHING ENDOTHELIUM IN THE TEMPORAL SECTOR. MEDICATION FOR IOP MANAGEMENT HAS BEEN PRESCRIBED. PER THE REPORTER, ON 28/APR/2021, THE SPECIALIST PERFORMED LENS REPLACEMENT WITHOUT COMPLICATIONS. PATIENT IS BETTER IN VA AND WITH AN ADEQUATE VAULT. PATIENT UNDER CONTINUOUS OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL AND DRY PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: H10: DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11780602&pc=QCB | 2023826-2021-01369 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL IMPLANTABLE COLLAMER LENS. THE LENS WAS REPORTED AS HAVING AN OVERREFRACTION CALCULATION AFTER IMPLANT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788212&pc=QCB | 2023826-2021-01362 | 2020-10-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+4.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788231&pc=QCB | 2023826-2021-01361 | 2020-10-22 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860720&pc=QCB | 2023826-2021-01581 | 2021-04-09 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/018 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855240&pc=QCB | 2023826-2021-01521 | 2021-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, TRANSIENT LOSS OF BCVA, EXCESSIVE VAULT, IRIS ATROPHY, AND UNREACTIVE(FIXED) PUPIL BEING OBSERVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED STUCK IN A BANDAID, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, WITH RESIDUE ON SURFACE. LENS WAS WRAPPED IN A BANDAID. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894537&pc=QCB | 2023826-2021-01719 | 2021-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/024 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE SURGEON REPORTS EXCESSIVE VAULT AND NARROW IRIS CORNEAL ANGLE. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: UPDATED INFORMATION: ON (B)(6) 2021, THE LENS WAS EXPLANTED. AND THEN REPLACED WITH A SHORTER LENS. AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942670&pc=QCB | 2023826-2021-01764 | 2021-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -15.0/+2.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND NARROW ANGLES. THE PROBLEM IS RESOLVED. Manufacturer Narrative: B5-ADDITIONAL INFO: "DESPITE ALL CALCULATIONS INDICATING 12.6 OR EVEN 13.2 THE LENS WAS CLEARLY OVERSIZED IN THE EYE AND A 12.1 WAS NEEDED. THERE WAS NOTHING IN THE PREOPERATIVE EVALUATION THAT SHOULD HAVE INDICATED THAT RESULT." CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966818&pc=QCB | 2023826-2021-01817 | 2021-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -4.50/6.0/008 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, WTW DIFFERENT FROM STS. PER REPORTER ON (B)(6) 2021, LENS REPLACEMENT PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION : B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -4.50/6.0/008 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4): SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12006531&pc=QCB | 2023826-2021-01910 | 2021-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.5/1.0/75 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009103&pc=QCB | 2023826-2021-01882 | 2021-03-12 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.50/6.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12029494&pc=QCB | 2023826-2021-01915 | 2020-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.5/0.5/004 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. IT WAS REPORTED THAT NO INTERVENTION WAS NECESSARY AND PROBLEM RESOLVED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " QUIET EYE, WE ARE MONITORING AND HAVE DECIDED TO LEAVE THE LENS AS IT IS AND MONITOR". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -07.50/+0.5/004 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT AND SHALLOW ANGLE WAS OBSERVED. REPORTEDLY THE EYE IS QUIET AND UNDER MONITORING AS THE LENS REMAINS IMPLANTED. PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12093373&pc=QCB | 2023826-2021-02188 | 2021-03-15 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+0.5/069 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL AND HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102781&pc=QCB | 2023826-2021-02149 | 2020-12-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/2.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102791&pc=QCB | 2023826-2021-02154 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.50/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222278&pc=QCB | 2023826-2021-02856 | 2021-05-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -6.5/+6.0/148 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 AN ADDITIONAL PI-YAG WAS ADDED, THE SURGEON REPORTING EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "APPARENTLY THE LENS IS LONGER/BIGGER THAN EXPECTED CAUSING A BIGGER VAULT." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233901&pc=QCB | 2023826-2021-02931 | 2021-07-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/1.5/72 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS EXPLANTED DUE TO THE PATIENT COMPLAINED OF EYE DISTENSION AND DEMANDED TO REMOVE THE LENS. IT WAS NOTED "THE EXAMINATION RESULTS WERE NORMAL". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). EYE DISTENSION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12235122&pc=QCB | 2023826-2021-02932 | 2021-07-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.0/1.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS EXPLANTED DUE TO THE PATIENT COMPLAINED OF EYE DISTENSION AND DEMANDED TO REMOVE THE LENS. IT WAS NOTED "THE EXAMINATION RESULTS WERE NORMAL". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12250498&pc=QCB | 2023826-2021-02951 | 2021-05-07 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271333&pc=QCB | 2023826-2021-03014 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+1.5/037 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271661&pc=QCB | 2023826-2021-03015 | 2021-06-11 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322370&pc=QCB | 2023826-2021-03129 | 2021-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LOW VAULT AT 190UM AND ACCOMMODATION ISSUES. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 13.2MM VTICM5_13.2, -12.5/+3.0/117 (SPHERE/CYLINDER/AXIS) ICL WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2021. REPORTEDLY, "THE LENS STAYS IN THE EYE-ACCOMMODATION ISSUES ARE GONE-PATIENT IS HAPPY! CLOSE MONITORING IS PLANNED." H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLO FOR AN UNK OR BLANK ENTRY. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4) (ACCOMMODATION ISSUES). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322374&pc=QCB | 2023826-2021-03119 | 2021-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.5/+3.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS POSSIBLE LENS ROTATION OR MISALIGNMENT AND THE PATIENT IS NOT SATISFIED WITH VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5: PATIENT HAD PRE-EXISTING DRY EYE. A LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2021. THIS DID NOT RESOLVED THE PROBLEM. REPORTEDLY; PATIENT NOT SATISFIED WITH VISION. CLAIM#: (B)(4). Manufacturer Narrative: B5 - PATIENT REPORTED THEY ARE UNABLE TO WEAR GLASSES BECAUSE THEY CAUSE HEADACHES DUE TO PRESSURE ON THE SIDES OF THEIR HEAD. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: (B)(4) (DISSATISFACTION WITH VISION). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322380&pc=QCB | 2023826-2021-03130 | 2021-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS LOW VAULT AT 0UM AND ACCOMMODATION ISSUES. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+4.0/089 (SPHERE/CYLINDER/AXIS) ICL INTO THE PATIENT'S EYE ON (B)(6) 2021. REPORTEDLY, "THE LENS STAYS IN THE EYE-ACCOMODATION ISSUES ARE GONE-PATIENT IS HAPPY! CLOSE MONITORING IS PLANNED." H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLO FOR AN UNK OR BLANK ENTRY. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4) (ACCOMMODATION ISSUES). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12325335&pc=QCB | 2023826-2021-03181 | 2021-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.0/+4.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE SURGEON REPORTS POSSIBLE LENS ROTATION OR MISALIGNMENT AND THE PATIENT IS NOT SATISFIED WITH VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5: PATIENT HAD PRE-EXISTING DRY EYE. A LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2021. THIS DID NOT RESOLVED THE PROBLEM. REPORTEDLY; PATIENT NOT SATISFIED WITH VISION. CLAIM#: (B)(4). Manufacturer Narrative: B5: PATIENT REPORTED THEY ARE UNABLE TO WEAR GLASSES BECAUSE THEY CAUSE HEADACHES DUE TO PRESSURE ON THE SIDES OF THEIR HEAD. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK.. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12411060&pc=QCB | 2023826-2021-03403 | 2021-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (THE LENS WAS IMPLANTED IN THE VERTICAL POSITION) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426112&pc=QCB | 2023826-2021-03456 | 2021-07-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -9.5/+0.5/131 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY THE PUPIL IS FIXED AND DILATED AND PIGMENTED CELLS ARE PRESENT. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE PATIENT CANCELLED SURGERY AS HER PUPIL IS STILL DILATED. THE DOCTOR STATES THE PATIENT'S IOP HAS ALWAYS BEEN NORMAL EVEN AT 2HR POSTOP. THE PATIENT WAS GIVEN PODIAMOX ROUTINELY AND IOP WAS UNREMARKABLE. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443368&pc=QCB | 2023826-2021-03541 | 2021-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -10.5/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS LATER EXPLANTED DUE TO EXCESSIVE VAULT. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS):" HAPPY WITH VISION. SHALLOW AC, VERY HIGH VAULT AND NARROW ANGLES" WAS REPORTED. IT WAS ALSO REPORTED THAT THIS CASE IS PENDING ICL REPLACEMENT FOR SHORTER LENGTH. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443370&pc=QCB | 2023826-2021-03542 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS LATER EXPLANTED DUE TO EXCESSIVE VAULT. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS):" HAPPY WITH VISION. SHALLOW AC, VERY HIGH VAULT AND NARROW ANGLES" WAS REPORTED. IT WAS ALSO REPORTED THAT THIS CASE IS PENDING ICL REPLACEMENT FOR SHORTER LENGTH. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND A OPTIC TORN AND HAPTIC TORN LENS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12454497&pc=QCB | 2023826-2021-03487 | 2021-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+0.5/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: LENS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2021. THE PROBLEM WAS RESOLVED. REPORTEDLY; "PATIENT IS HAPPY." CLAIM# (B)(4) Manufacturer Narrative: D6B: (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, RACE: UNK, ETHNICITY: UNK, EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12454516&pc=QCB | 2023826-2021-03486 | 2021-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/066 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -11.50/+0.5/066 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LONGER LENGTH LENS OF DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT: CLINICAL CODE: (B)(4): REFRACTIVE CHANGE OVERTIME. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479163&pc=QCB | 2023826-2021-03594 | 2021-05-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/5.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). VERY SHALLOW ANTERIOR CHAMBER WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. IMPLANT DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12559075&pc=QCB | 2023826-2021-03832 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -6.5/2.5/90, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT EXCESSIVE VAULT AND ANGLE CLOSURE WITH IOP. FOR STATUS OF THE EYE (SIGNS/ SYMPTOMS): THE REPORTER STATED "2+ CELL, FORMED AC, LENS EDGE VISIBLE SUP AND INF. ANGLE NARROW AT NASAL & TEMP. OPEN BUT SHALLOW, DILATED IRIS." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION " PATIENT IS STABLE AND IS SEEING SURGEON REGULARLY ". CLAIM # (B)(4). Manufacturer Narrative: DATE OF EVENT:UNK. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12573699&pc=QCB | 2023826-2021-03859 | 2021-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+4.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING WITH LENS ROTATION. THE REPORTER INDICATED THE SURGERY WAS PERFORMED WITH SOME TRAUMATISM DUE TO IRIS PROLAPSED INTO BOTH PARACENTESIS DURING SURGERY. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED BY A DIFFERENT SURGEON AND WILL NOT BE RETURNED. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587055&pc=QCB | 2023826-2021-03866 | 2021-09-08 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12637120&pc=QCB | 2023826-2021-03958 | 2021-07-03 | Injury | STAAR SURGICAL COMPANY | 2021-10-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXIS LENS DUE TO THE EXPECTED VISION NOT BEING ACHEIVED. THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR FILE# (B)(4) FOR THE REPLACEMENT LENS. REPORTEDLY, "THE ORIGINAL LENS IS ALWAYS DEFLECTED TO A FIXED POSITON AND THE DOCTOR INSERTS A NEW LENS IN A STABLE POSITION." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12677849&pc=QCB | 2023826-2021-04047 | 2021-09-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT REPORTED POST-OP COLOR DISTORTION, MAINLY IN BLUE TONES. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12677861&pc=QCB | 2023826-2021-04045 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+0.5/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED DYSPHOTOPSIAS AT NIGHT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. IMPLANT DATE: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685424&pc=QCB | 2023826-2021-04076 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12749423&pc=QCB | 2023826-2021-04226 | 2021-07-13 | Injury | STAAR SURGICAL COMPANY | 2021-11-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6,-9.00/1.0/091, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT,SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807721&pc=QCB | 2023826-2021-04270 | 2021-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -12.5/+2.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12833482&pc=QCB | 2023826-2021-04291 | 2021-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+3.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840550&pc=QCB | 2023826-2021-04355 | 2021-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+4.5/103 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE AC IS DEEP, ANGLES ARE OPEN AND PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+4.5/103 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE AC IS DEEP, ANGLES ARE OPEN AND PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840570&pc=QCB | 2023826-2021-04365 | 2021-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS PIGMENT DISPERSION SYNDROME. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: ON 16-NOV-2021 THE LENS WAS EXPLANTED. THE SURGEON STATES AFTER EXPLANT-"NORMAL POSTOP. WAITING FOR PIGMENT TO CLEAR. NO ANTICIPATED ISSUES." THE CAUSE IS REPORTED AS A PATIENT-RELATED FACTOR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840589&pc=QCB | 2023826-2021-04364 | 2021-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS PIGMENT DISPERSION SYNDROME. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: REPORTEDLY ON (B)(6) 2021 THE LENS WAS REPOSITIONED. ON (B)(6) 2021 THE LENS WAS EXPLANTED. THE SURGEON REPORTS PERSISTENT PIGMENT DISPERSION-"PIGMENT DISPERSION FOR 3 MONTHS DESPITE NORMAL ANGLES, 500UM VAULT, AND REPOSITION." AFTER EXPLANT-"WAITING FOR PIGMENT TO CLEAR. NO ANTICIPATED ISSUES." CAUSE IS REPORTED AS A PATIENT-RELATED FACTOR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840698&pc=QCB | 2023826-2021-04358 | 2021-10-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE REPORTER INDICATED THE PATIENT HAD AN EXTENDED PUPIL AFTER 2 WEEKS POST-OP. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +3.00/+3.0/094 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER INDICATED THE PATIENT HAD AN EXTENDED PUPIL AFTER 2 WEEKS POST-OP. AT THE LAST POST-OP VISIT ON (B)(6) 2021, THE PUPIL REACTION WAS NORMAL AND THE PATIENT WAS HAPPY. THE PATIENT WILL BE MONITORED. THE LENS REMAINED IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: D2: COMMON DEVICE NAME - PHAKIC TORIC INTRAOCULAR LENS. PRODUCT CODE: QCB. E1: ADDRESS - (B)(6). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12965261&pc=QCB | 2023826-2021-04649 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/102 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULTING AND REFRACTIVE CHANGE OVER TIME WAS REPORTED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND REPLACED WITH A LONGER LENS ON (B)(6) 2022. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. B4 & G3: CORRECTION TO INITIAL MEDWATCH AWARE DATE: (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12965369&pc=QCB | 2023826-2021-04651 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/074 (SPHERE/ CYLINDER/ AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT COMPLAINED THEIR VISION WAS WORSE. THE VAULT WAS IMPROVED, ANGLE BETTER AND IOP STABLE. THE SURGEON ASKED PATIENT TO ADAPT TO THEIR NEW VISION. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-03865. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972971&pc=QCB | 2023826-2021-04650 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00/+0.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULTING AND REFRACTIVE CHANGE OVER TIME WAS REPORTED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND REPLACED WITH A LONGER LENS ON (B)(6) 2022. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. B4 & G3: CORRECTION TO INITIAL MEDWATCH AWARE DATE: (B)(6) 2021. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973335&pc=QCB | 2023826-2021-04712 | 2021-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND PIGMENT DISPERSION. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "UNKNOWN REASON FOR HIGH VAULT." REPORTEDLY, "ACCORDING TO THE SURGEON THE SUGGESTED SIZE WAS 13.2MM BUT THE VAULT POST OP WAS MORE THAN 1000 MICRON AND WAS SURPRISED FOR THE HIGH VAULT BECAUSE OF THIS." Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE SURGEON STATES THE PREOP PARAMETER SUGGESTED A 13.2MM SIZE BUT THE POSTOP VALUE WAS MORE THAN 1000UM AND HE WAS SURPRISED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973343&pc=QCB | 2023826-2021-04713 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.0/+3.0/018 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS LENS ROTATION AND OVER-REFRACTION OF +0.25/-1.75-144. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023218&pc=QCB | 2023826-2021-04890 | 2021-11-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057227&pc=QCB | 2023826-2021-04910 | 2021-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -03.50/+0.5/014 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS REPLACED WITH A LONGER LENS (THE LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -03.50/+0.5/014 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS REPLACED WITH A LONGER LENS (THE LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082442&pc=QCB | 2023826-2021-04977 | 2021-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -3.0/+4.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS, THE SURGEON REPORTING A VAULT VALUE OF 1288UM, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND PAS. THE VAULT VALUE POST-EXCHANGE REPORTED VALUE IS 746UM; THE PROBLEM IS MARKED AS RESOLVED.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -3.0/+4.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS, THE SURGEON REPORTING A VAULT VALUE OF 1288UM, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND PAS. THE VAULT VALUE POST-EXCHANGE REPORTED VALUE IS 746UM; THE PROBLEM IS MARKED AS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082444&pc=QCB | 2023826-2021-04978 | 2021-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -3.0/+4.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS, THE SURGEON REPORTING EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND PAS. THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082451&pc=QCB | 2023826-2021-04979 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT (OS) EYE ON (B)(6) 2021. REPORTEDLY THE MAIN INCISION NEEDED TO BE SUTURED DUE TO "MANEUVER AND CAUSES WOUND LEAK."Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT (OS) EYE ON (B)(6) 2021. REPORTEDLY THE MAIN INCISION NEEDED TO BE SUTURED DUE TO "MANEUVER AND CAUSES WOUND LEAK." Manufacturer Narrative: H6- MEDICAL DEVICE PROBLEM CODE: REPLACED CODE 3189 WITH 2993. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090715&pc=QCB | 2023826-2021-04996 | 2021-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -9.5/1.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED FOREIGN BODY SENSATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN; BUT NOTED THE DEVICE DID NOT FAIL TO PERFORM AND "PATIENT COMPLAINTS OF FOREIGN BODY SENSATION". Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13102113&pc=QCB | 2023826-2021-05012 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -0.50/+5.5/82 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -0.50/+5.5/82 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13122737&pc=QCB | 2023826-2021-05009 | 2021-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "LENS TOO SMALL SO IT ROTATED 90 DEGREES, PATIENT COMPLAINING OF NOT BEING ABLE TO FOCUS ON OBJECTS UP CLOSE. EXCHANGED FOR LARGER SIZE." Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13122997&pc=QCB | 2023826-2021-04782 | 2021-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.5/0.5/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13123000&pc=QCB | 2023826-2021-04780 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13123001&pc=QCB | 2023826-2021-04781 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/090 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13296116&pc=QCB | 2023826-2022-00242 | 2021-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.5/2.5/178 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. IT WAS REPORTED THE PATIENT EXPERIENCED MACULAR SPLITTING. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. POSTOP COMMENTS NOTED "MACULAR HOLE AND DIOPTER CHANGE OCCURRED AFTER THE FIRST OPERATION". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.5/2.5/178 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. IT WAS REPORTED THE PATIENT EXPERIENCED MACULAR SPLITTING. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. POSTOP COMMENTS NOTED "MACULAR HOLE AND DIOPTER CHANGE OCCURRED AFTER THE FIRST OPERATION". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311248&pc=QCB | 2023826-2022-00222 | 2021-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -17.50/1.5/149 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS PATIENT RELATED FACTOR.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -17.50/1.5/149 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311278&pc=QCB | 2023826-2022-00223 | 2021-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -17.50/1.0/054 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317700&pc=QCB | 2023826-2022-00197 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS THE PATIENT IS NOT SATISFIED WITH THEIR VISION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. EXPIRATION DATE, IMPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13476644&pc=QCB | 2023826-2022-00422 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -14.00/1.0/145 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 202. IT WAS REPORTED THAT THE SIZE IS TOO LARGE CAUSING DANGEROUS VAULT AND INTERMITTENT ANGLE CLOSURE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. WITH NO PATIENT INJURY OR SUTURES. CLAIM#: (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: (B)(4) - ANGLE CLOSURE. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13476664&pc=QCB | 2023826-2022-00421 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -14.00/1.0/026 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED THAT THE SIZE IS TOO LARGE CAUSING DANGEROUS VAULT AND INTERMITTENT ANGLE CLOSURE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. WITH NO PATIENT INJURY OR SUTURES. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13488087&pc=QCB | 2023826-2022-00424 | 2021-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICMO12.6; -8.0/3.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS ALSO INDICATED THE PATIENT'S ANGLE NARROWED SIGNIFICANTLY; AND ANTERIOR CHAMBER WAS 1570UM SHALLOW. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THE PATIENT DECIDED NOT TO IMPLANT ANOTHER LENS DO TO PERSONAL REASONS. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13580848&pc=QCB | 2023826-2022-00569 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.00/+2.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS LATER REPORTED THAT ON (B)(6) 2022; A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. "PATIENT IS HAPPY". CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588732&pc=QCB | 2023826-2022-00577 | 2022-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED TRANSIENT LOSS OF BCVA DUE TO THE SURGEON, TO DILATE THE PUPIL, INTRAOPERATIVELY ADMINISTERED MYDRANE INTRACAMERALLY INTO THE EYE, WHICH MAY HAVE CAUSED VISUAL PROBLEMS. THE PATIENT COMPLAINED OF SHORT-TERM BLINDNESS SYMPTOMS (THERE WAS NO GLAUCOMA OR MACULAR EDEMA NOTICED). THE VISION HAD REHABILITATED SLOWLY BIT BY BIT, TWO HOURS AFTER THE OPERATION ABOUT 20/400. ON THE NEXT DAY, THE VISION WAS BACK TO NORMAL CONDITIONS (UCVA 20/25). THE PATIENT IS HAPPY AND THE LENS REMAINED IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588755&pc=QCB | 2023826-2022-00578 | 2022-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN, AND OPTIC TORN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588761&pc=QCB | 2023826-2022-00489 | 2022-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REMAINED IMPLANTED. THERE WAS AN INCREASE IN MYOPIA AND ASTIGMATISM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: NA. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: B5 - "REPORTEDLY, THERE WAS AN INCREASED CYLINDER CORRECTION WITH LENS ROTATION. THE CAUSE OF THE EVENT WAS DEVICE RELATED AND PROCEDURE RELATED." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: B1 - CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. B5 - "REPORTEDLY, THE TORIC ICL WAS EXPLANTED FROM THE RIGHT EYE. THIS RESOLVED THE PROBLEM. THE SEVERITY OF THIS REPORT IS GRADE 2/MODERATE. " SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO: SERIOUS INJURY IN THE INITAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 2199 SHOULD BE REMOVED FROM THE INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13641144&pc=QCB | 2023826-2022-00654 | 2021-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+2.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: " THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -07.50/+2.5/102 (SPHERE/ CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN." CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE, RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13725932&pc=QCB | 2023826-2022-00744 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SHORTER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. THE REPLACEMENT DID NOT RESOLVE THE PROBLEM. "VAULT STILL SEEMS TOO HIGH." CAUSE REPORTED AS DEVICE. SEE MFR # 2023826-2022-00826. D6B- (B)(6) 2022. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869352&pc=QCB | 2023826-2022-00887 | 2021-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+2.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869607&pc=QCB | 2023826-2022-00684 | 2021-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.00/1.0/056 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, ANGLE PRE-OP: 35.7 - POST-OP: 16 . Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13944046&pc=QCB | 2023826-2022-00980 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -04.50/+3.5/108 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995877&pc=QCB | 2023826-2022-00924 | 2022-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -13.50/3.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: B5: REPORTEDLY, "GOOD JUMP BUT REFRACTIVE ERROR. GOOD LENGTH OF THE IMPLANT BUT NOT HYPERMETROPIA POWER". ADDITIONAL COMMMENTS FROM REPORTER - PATIENT WITH KERATOCONUS OD/OS, STRABISM - RETINA LASER. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. EVENT DATE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14030182&pc=QCB | 2023826-2022-01105 | 2022-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TICL TORIC IMPLANTABLE COLLAMER LENS. WHEN INSERTING THE LENS, THE SURGEON THOUGHT THE LENS MAY HAVE BEEN INSERTING UPSIDE DOWN, BUT WHEN THE LENS UNFOLDED IN THE EYE, THE LENS WAS PROPERLY UNDER THE IRIS IN THE CORRECT ORIENTATION. AT ONE DAY POST-OP, THE EYE HAD A VERY SHALLOW ANTERIOR CHAMBER, INAPPRECIABLE VAULT AND UNRECORDABLE INTRAOCULAR PRESSURE (IOP). THE SURGEON REALIZED THERE WAS A SMALL WOUND LEAK AND HE REPAIRED IT WITH ONE SUTURE. THE ANTERIOR CHAMBER DEEPENED, THERE WAS APPRECIABLE VAULT AND IOP WAS LOW BUT RECORDABLE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/089 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD A LOW VAULT. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS DOING GREAT. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B4 AND G3: 16-MAR-2022. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190873&pc=QCB | 2023826-2022-01292 | 2022-03-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -7.5/+1.0/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) IS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: OCOS. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4). LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14324153&pc=QCB | 2023826-2022-01432 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -12.5/4.0/115 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, WRONG SIZE OF THE LENS, TOO BIG. THE LENS REMAINS IMPLANTED. STATUS OF THE EYE REPORTED AS MIXED ASTIGMATISM. "THE APC IS OPEN IN ALL SEGMENTS, THE PUPILLARY EDGE OF THE IRIS IS RAISED AND IS IN CONTACT WITH THE ANTERIOR SURFACE OF THE IOL IN THE LOWER SEGMENT -0.94 MM FOR 12 H. PART OF THE IOL IS IN CONTACT WITH THE CILIARY PROCESSES (BUTTING AGAINST THE CILIARY PROCESSES, NOT IN THE CILIARY SULCUS). CILIARY SULCUS DIAMETER: VERTICAL SECTION ~ 12.56 - 12.9 MM (SINCE THE CILIARY PROCESSES ARE IN CONTACT WITH THE FIBERS OF THE CILIARY LIGAMENT AND WITH THE PIGMENT LAYER OF THE IRIS, THE VALUE VARIES).CROSS SECTION ~12.8." Manufacturer Narrative: CORRECTED DATA: G2- OTHER: RUSSIA. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLINICAL CODE: (B)(4): APC IS OPEN IN ALL SEGMENT; PUPILLARY EDGE OF THE IRIS IS RAISED AND IN CONTACT WITH THE ANTERIOR SURFACE OF IOL; IOL IS IN CONTACT WITH THE CILIARY PROCESSES. DEVICE CODE (B)(4): OFF-LABEL USE (PATIENT OVER 45 YRS OLD.) CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14350005&pc=QCB | 2023826-2022-01430 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -5.0/5.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS REPORTED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. "LENS TOO BIG." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -5.0/5.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT; REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS REPORTED. LENS REMAINS IMPLANTED. STATUS OF EYE " IOP OK BUT ANGLES SHALLOW." CAUSE REPORTED AS DEVICE. "LENS TOO BIG." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF-5.0/5.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT; REFRACTIVE SURPRISE; SIGNIFICANT REDUCTION OF LRIDOCORNEAL ANGLES AND SHALLOW ANGLES WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED AND PROBLEM WAS RESOLVED. REPORTEDLY "THE EXCHANGE HAS OCCURRED AND PATIENT AND SURGEON HAPPY". CAUSE REPORTED AS DEVICE. "LENS TOO BIG." H6- HEALTH EFFECT-CLINICAL CODE: 4581- SHALLOW ANGLES. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -5.0/5.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT; REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED AND PROBLEM WAS RESOLVED. REPORTEDLY "THE EXCHANGE HAS OCCURRED AND PATIENT AND SURGEON HAPPY". CAUSE REPORTED AS DEVICE. "LENS TOO BIG." CLAIM# (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL WITH RESIDUE AND DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365088&pc=QCB | 2023826-2022-01480 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+2.5/089 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14366518&pc=QCB | 2023826-2022-01484 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: ON 11-APR-2022 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED , ENTITLED, "TORIC INTRAOCULAR LENS FOR ASTIGMATISM CORRECTION FOLLOWING KERATOPLASTY IN PHAKIC AND PSEUDOPHAKIC EYE". THE ARTICLE REPORTS ROTATION AND ENDOTHELIAL CELL LOSS IN EYES IMPLANTED WITH PIGGYBACK TORIC ICL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: AGE OR DATE OF BIRTH-RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14431927&pc=QCB | 2023826-2022-01580 | 2022-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS LOW VAULTING, LENS ROTATION AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR . B5 - "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/ LONGER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. H6 - HEALTH EFFECT - MEDICAL DEVICE CODE: 1401 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/ LONGER LENS DUE TO REFRACTIVE SURPRISE AND LOW VAULT WITH ROTATION. THE PROBLEM IS RESOLVED." IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14432041&pc=QCB | 2023826-2022-01583 | 2022-04-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/066 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (NARROW ANGLES). THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO SIZING ERROR DUE TO ACD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440075&pc=QCB | 2023826-2022-01581 | 2022-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS LOW VAULTING, LENS ROTATION AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR . ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/ LONGER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. H6 - HEALTH EFFECT - MEDICAL DEVICE CODE: 1401 SHOULD BE ADDED TO THE INITAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/ LONGER LENS DUE TO REFRACTIVE SURPRISE AND LOW VAULT WITH ROTATION. THE PROBLEM IS RESOLVED." IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. (B)(4): REFRACTIVE CHANGE OVERTIME. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14455506&pc=QCB | 2023826-2022-01584 | 2022-04-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO SIZING ERROR DUE TO ACD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14504817&pc=QCB | 2023826-2022-01614 | 2022-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS NOT RESOLVED. "VAULT WAS HIGH AND PATIENT WAS ACCEPTING +2 POST OPERATIVELY. HE WAS COMPLETELY UNHAPPY AND WANTED TO REMOVE THE LENS. REMOVED ON (B)(6) 2022 AS PER PATIENT'S DECISION, SINCE HE IS TRAVELING ABROAD IN NEXT WEEK." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14505166&pc=QCB | 2023826-2022-01615 | 2022-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -12.0/5.5/042 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS NOT RESOLVED. "PATIENT REQUESTED TIME TO DECIDE ABOUT EXCHANGE OF SMALLER SIZE LENS.." THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; WITH MOISTURE. VISUAL INSPECTION FOUND NO VISUAL DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14538894&pc=QCB | 2023826-2022-01667 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -08.50/+0.5/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: LENS WAS EXPLANTED AND REPLACED ON (B)(6) 2022 WITH A LONGER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT IS HAPPY." CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14553237&pc=QCB | 2023826-2022-01668 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: LENS WAS EXPLANTED AND REPLACED ON (B)(6) 2022 WITH A LONGER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT IS HAPPY." CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14734164&pc=QCB | 2023826-2022-01852 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -20.00/+3.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2011. THE REPORTER INDICATED THE LENS WILL BE EXPLANTED/EXCHANGED DUE TO PROGRESSIVE MYOPIA AND ASTIGMATISM. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 724687
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14734191&pc=QCB | 2023826-2022-01851 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -19.50/+3.0/130 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2011. THE REPORTER INDICATED THE LENS WILL BE EXPLANTED/EXCHANGED DUE TO PROGRESSIVE MYOPIA AND ASTIGMATISM. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14750598&pc=QCB | 2023826-2022-01924 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -16.0/6.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -16.0/6.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM AND THE REPORTER STATES "CLEAR CORNEA." CAUSE REPORTED AS DEVICE. D6B- (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985511&pc=QCB | 2023826-2022-02119 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985513&pc=QCB | 2023826-2022-02118 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). NO COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985514&pc=QCB | 2023826-2022-02130 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985546&pc=QCB | 2023826-2022-02142 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.0/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985559&pc=QCB | 2023826-2022-02144 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -10.5/3.0/009 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14987248&pc=QCB | 2023826-2022-02070 | 2022-04-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.0/1.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) IS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF EVENT IS REPORTED AS OTHER: ICL CALCULATION SOFTWARE AND/OR WTW FROM DEVICE (ALCON BIOGRAPH) OVER SIZING. Manufacturer Narrative: B5: CORRECTED TO - "THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) IS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF EVENT IS REPORTED AS OTHER: ICL CALCULATION SOFTWARE AND/OR WTW FROM DEVICE (ALCON BIOGRAPH) OVER SIZING." CLAIM # (B)(4). Manufacturer Narrative: B5: "THE REPORTER ORDERED A REPLACEMENT LENS (B)(6) HOWEVER IT WAS NOT USED. THE INITIAL LENS (B)(6) REMAINS IMPLANTED. REPORTEDLY, THE PATIENT IS NOW SATISFIED WITH THE INITIAL LENS AND THE REPLACEMENT LENS IS NOT NEEDED. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15003802&pc=QCB | 2023826-2022-02200 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.00/+1.0/074 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD). THE REPORTER INDICATED AFTER RE-EXAMINATION, THE FREQUENCY CHANGED AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED AFTER RE-EXAMINATION PRIOR TO ICL IMPLANTATION, THE FREQUENCY CHANGED (REFRACTIVE SURPRISE) AND ACD WAS NARROW. A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.00/+1.0/074 (SPHERE/CYLINDER/AXIS), WAS NOT USED OR EMBEDDED IN THE PATIENTS RIGHT EYE (OD) (NOT IMPLANTED). THE SURGEON REQUESTED ANOTHER LENS. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15008811&pc=QCB | 2023826-2022-02204 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15008945&pc=QCB | 2023826-2022-02201 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.00/+1.5/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED AFTER RE-EXAMINATION, THE FREQUENCY CHANGED AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK .SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK .RACE: UNK. DATE OF EVENT: UNK. MODEL #: UNK. SERIAL #: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED AFTER RE-EXAMINATION PRIOR TO ICL IMPLANTATION, THE FREQUENCY CHANGED (REFRACTIVE SURPRISE) AND ACD WAS NARROW. A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.00/+1.5/100 (SPHERE/CYLINDER/AXIS), WAS NOT USED OR EMBEDDED IN THE PATIENTS LEFT EYE (OS) (NOT IMPLANTED). THE SURGEON REQUESTED ANOTHER LENS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15047190&pc=QCB | 2023826-2022-02210 | 2022-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+4.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN AND THE REPORTER STATED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15048209&pc=QCB | 2023826-2022-02303 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.0/1.5/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED HIGH VAULT AND NARROW ANGLES, LENS REMAINS IMPLANTED. THE REPORTER STATED THAT THE PATIENT IS DOING WELL, IOP 17/18 EXCELLENT VISION 20/40. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15080946&pc=QCB | 2023826-2022-02372 | 2022-07-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2; -11.00/2/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND IT WAS NOTED "IRIS TINTED". INTRAOPERATIVELY THE SURGEON REMOVED AND REPLACED THE LENS WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND INDICATED "THE LENS WAS TOO VAULTED AS HE WAS ADJUSTING TO IT AT THAT TIME HE CHOSE TO GO A SMALLER SIZE AND THAT ONE FIT WELL." Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108220&pc=QCB | 2023826-2022-02442 | 2022-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -17.00/+2.50/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON THE SAME DATE BUT DIFFERENT SURGERY THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15126877&pc=QCB | 2023826-2022-02501 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.50/4.0/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). IT WAS REPORTED THAT THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO SLIGHT ANGLE CLOSURE. Manufacturer Narrative: B5- " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/4.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY THE PATIENT EXPERIENCED ANGLE CLOSURE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH THE SAME MODEL/SHORTER LENGTH LENS. REPORTEDLY " DR. (B)(6)HAD TO EXCHANGE ICL DUE TO LENS BEING OVERSIZED. HE ORIGINALLY IMPLANTED 12.6 SIZE AND HAD TO REPLACE WITH 12.1 LENSES DUE TO SLIGHT ANGLE CLOSURE"." CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15126914&pc=QCB | 2023826-2022-02500 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.00/2.5/108 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS REPORTED THAT THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SLIGHT ANGLE CLOSURE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/2.5/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY THE PATIENT EXPERIENCED SLIGHT ANGLE CLOSURE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH THE SAME MODEL/SHORTER LENGTH LENS. REPORTEDLY " DR. (B)(6) HAD TO EXCHANGE ICL DUE TO LENS BEING OVERSIZED. HE ORIGINALLY IMPLANTED 12.6 SIZE AND HAD TO REPLACE WITH 12.1 LENSES DUE TO SLIGHT ANGLE CLOSURE". CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15135400&pc=QCB | 2023826-2022-02453 | 2022-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND NARROW ANGLES WAS OBSERVED, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE CAUSE OF EVENT IS UNKNOWN AND THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/1.0/84 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT; SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND A SHALLOW PERIPHERAL ACD. ON (B)6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN AND THE DEVICE FAILED TO PERFORM AS INTENDED. H6- HEALTH EFFECT- CLINICAL CODE: 4581 SIGNIFICANT REDUCTION OF IRIDOCRNEAL ANGLES AND SHALLOW PERIPHERAL ACD. CLAIM #(B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION ON 03-AUG-2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. H3: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15163864&pc=QCB | 2023826-2022-02586 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.5/+2.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND UNREACTIVE/FIXED PUPIL. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: B3 - CORRECTED TO: DATE OF EVENT: (B)(6) 2022 IN INITIAL MDR. B4 - CORRECTED TO: DATE OF THIS REPORT: (B)(6) 2022 IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15164320&pc=QCB | 2023826-2022-02587 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -16.0/+2.50/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT AND UNREACTIVE/FIXED PUPIL. IT WAS REPORTED THAT THE LENS DID NOT RESOLVE THE PROBLEM, THE REPORTER STATED " PUPIL DID NOT CONSTRICT EVEN AFTER REPLACING WITH SMALLER SIZE.". PER UPDATED INFORMATION THE PATIENT HAS NOT VISIT THE CLININC SINCE MAY. BCVA WAS 20/20 AND VAULT WAS 1CT ON THE LAST VISIT ((B)(6) 2022). IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2022-02586 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182874&pc=QCB | 2023826-2022-02466 | 2022-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.0/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182892&pc=QCB | 2023826-2022-02467 | 2022-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -09.0/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225753&pc=QCB | 2023826-2022-02565 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.5/1.0/40 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT IMPACT CODE (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225754&pc=QCB | 2023826-2022-02564 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.5/3.0/143 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT IMPACT CODE (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15227460&pc=QCB | 2023826-2022-02766 | 2022-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.5/3.0/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION AND COMPLAINED OF DECREASED VISION. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15245958&pc=QCB | 2023826-2022-02734 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER. THE PROBLEM WAS RESOLVED. ANOTHER ICL LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, RACE: UNK, ETHNICITY: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272155&pc=QCB | 2023826-2022-02845 | 2022-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/2.0/107 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE (B)(4). EXPIRATION DATE: 08-31-2024. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272166&pc=QCB | 2023826-2022-02846 | 2022-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). EXPIRATION DATE: 02-28-2025. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15296635&pc=QCB | 2023826-2022-02790 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE PATIENT COMPLAINED OF DOUBLE IMAGES. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS DIOPTER IS -14.50/+2.5/109 (SPHERE/CYLINDER/AXIS). THE METRICS LIKE IOP AND VAULT WERE GOOD, AND VA IS IMPROVING. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299561&pc=QCB | 2023826-2022-02923 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.50/1.00/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED IRIS TRANSILLUMINATION, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: 10/31/2022. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307290&pc=QCB | 2023826-2022-02887 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/3.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE(OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. REPORTEDLY THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307425&pc=QCB | 2023826-2022-02888 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. REPORTEDLY THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307870&pc=QCB | 2023826-2022-02791 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE PATIENT COMPLAINED OF DOUBLE IMAGES. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS DIOPTER IS -14.00/+3.0/088 (SPHERE/CYLINDER/AXIS). THE METRICS LIKE IOP, VAULT WERE GOOD, AND VA IS IMPROVING. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK, EXPIRATION DATE UNK, EXPLANT DATE: UNK, DEVICE MANUFACTURE DATE: UNK (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15362594&pc=QCB | 2023826-2022-02683 | 2022-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS COULD NOT BE PLACED BEHIND THE IRIS IN THE LEFT EYE (OS). ON 13-JUL-2022 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. NO NEW LENS WAS IMPLANTED AND THE CAUSE IS REPORTED AS UNKNOWN. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: D6A: IMPLANT DATE IN INITIAL MDR CORRECTED TO (B)(6) 2022. ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.0/5.5/078 (SPHERE/CYLINDER/AXIS) COULD NOT BE PLACED BEHIND THE IRIS IN THE LEFT EYE (OS). ON (B)(6) 2022 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. A REPLACEMENT LENS OF SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. THE CAUSE WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CODE. 4581- LENS COULD NOT BE PLACED BEHIND IRIS. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15384326&pc=QCB | 2023826-2022-03043 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.00/1.5/064 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS' RIGHT EYE (OD) IN JUNE OF 2021. PER REPORTER PATIENT HAD A HIGH VAULT AT THE FOLLOW UP VISIT AND THE CLINIC DECIDED TO MONITOR PATIENT. LENS REMAINS IMPLANTED AND DOCTOR WANTS TO EXPLANT/EXCHANGE LENS DUE TO THE EXCESSIVE VAULT WHICH CREATE A TOO NARROW ANGLE. IT WAS REPORTED THAT PATIENT IS ON DROPS TO REDUCE EYE PRESSURE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15384329&pc=QCB | 2023826-2022-03044 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.50/2.5/122 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS' LEFT EYE (OS) IN (B)(6) OF 2021. PER REPORTER PATIENT HAD A HIGH VAULT AT THE FOLLOW UP VISIT AND THE CLINIC DECIDED TO MONITOR PATIENT. LENS REMAINS IMPLANTED AND DOCTOR WANTS TO EXPLANT/EXCHANGE LENS DUE TO THE EXCESSIVE VAULT WHICH CREATE A TOO NARROW ANGLE. IT WAS REPORTED THAT PATIENT IS ON DROPS TO REDUCE EYE PRESSURE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15391767&pc=QCB | 2023826-2022-03038 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.5/4.0/74 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND UNREACTIVE (FIXED) PUPIL. IT WAS REPORTED THAT THE PROBLEM DID NOT RESOLVE, FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): PUPIL REMAINED MID-DILATED WAS REPORTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY AND DAMAGED WITH RESIDUE/DEBRIS ON PRODUCT IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIS TORN/BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO DAMAGED STATE OF RETURNED LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423441&pc=QCB | 2023826-2022-03074 | 2022-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS OTHER. DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428728&pc=QCB | 2023826-2022-03096 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.5/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442924&pc=QCB | 2023826-2022-03086 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; AND PIGMENT DISPERSION WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15477263&pc=QCB | 2023826-2022-03280 | 2022-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.0/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THIS SAME DATE A REPLACEMENT LENS OF A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498727&pc=QCB | 2023826-2022-03326 | 2022-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, CORECTOPIA AND PERIPHERAL ANTERIOR SYNECHIAE (PAS) (DEVELOPING IN MULTIPLE AREAS). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK; EXPLANT DATE: NA. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. THE SURGICAL EXPLANT OF THE LENS WENT WELL AND POST-OPERATIVELY THE PATIENT HAD RECOVERED TO A BCVA OF 1.2 WITH THE FOLLOWING SCRIPT: -10.75/-1.00 X15. THE PATIENT WILL BE MONITORED AND THE LENS WAS DISCARDED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506245&pc=QCB | 2023826-2022-03327 | 2022-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON 25-AUG-2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, CORECTOPIA AND PERIPHERAL ANTERIOR SYNECHIAE (PAS) (DEVELOPING IN MULTIPLE AREAS). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, CORECTOPIA AND PERIPHERAL ANTERIOR SYNECHIAE (PAS) (DEVELOPING IN MULTIPLE AREAS). THE LENS WAS EXPLANTED ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. THE SURGICAL EXPLANT OF THE LENS WENT WELL AND POSTOPERATIVELY THE PATIENT HAD RECOVERED TO A BCVA OF 1.2 WITH THE FOLLOWING SCRIPT: -10.75/-1.25 X150. THE PATIENT WILL BE MONITORED AND THE LENS WAS DISCARDED. D6A: ON (B)(6) 2020. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506488&pc=QCB | 2023826-2022-03330 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING, LENS ROTATION AND PERMANENT LOSS OF BCVA. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540750&pc=QCB | 2023826-2022-03399 | 2022-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -10.50/1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15555994&pc=QCB | 2023826-2022-03514 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.50/1.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT. IT WAS REPORTED THAT THE VAULT ISSUE RESOLVED BUT FOR SATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " MILD RESIDUAL CYL." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15559517&pc=QCB | 2023826-2022-03526 | 2022-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.0/063 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - #(B)(4). SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15572321&pc=QCB | 2023826-2022-03409 | 2022-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -06.0/0.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15587255&pc=QCB | 2023826-2022-03410 | 2022-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -06.5/0.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591803&pc=QCB | 2023826-2022-03490 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.50/+0.5/054 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.50/+0.5/054 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15611592&pc=QCB | 2023826-2022-03643 | 2022-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.0/1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; PIGMENT DISPERSION; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620827&pc=QCB | 2023826-2022-03491 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.50/+0.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0350/+05/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME.ON (B)(6) 2023 THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15697250&pc=QCB | 2023826-2022-03856 | 2022-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -10/00/3.0/13 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. VERY SHALLOW CHAMBER; SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES; PIGMENT DISPERSION WAS REPORTED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. "OVER-SIZED LENS". Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: CLINICAL CODE: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15713892&pc=QCB | 2023826-2022-03880 | 2022-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+2.5/089 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, PIGMENT DISPERSION AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT WAS REDUCED TO NORMAL, THE AC IS DEEP AND THE PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15741084&pc=QCB | 2023826-2022-03873 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.0/2.0/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "FOR THE FIRST TIME, THE DOUBLE EYE TICL WAS INSTALLED REVERSELY." Manufacturer Narrative: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITRONAL INFORMATION: TYPE OF INVESTIGATION CODE; 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15741086&pc=QCB | 2023826-2022-03874 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -6.5/3.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "FOR THE FIRST TIME, THE DOUBLE EYE TICL WAS INSTALLED REVERSELY." Manufacturer Narrative: (B)(4). (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: IT WAS LATER REPORTED "DURING THE FIRST OPERATION OF THIS PATIENT THE LEFT AND RIGHT EYES OF THE LENS WERE REVERSED THE LEFT LENS OF THE RIGHT EYE WAS LOADED AND THE RIGHT LENS OF THE LEFT EYE WAS LOADED SO THE REFRACTIVE SURPRISE OCCURRED AFTER THE OPERATION. AFTER THE OPERATION THE DOCTOR FOUND THE CAUSE OF REFRACTIVE SURPRISE WHEN CHECKING THE WHOLE OPERATION PROCESS. LENS OF THE SAME SIZE WERE THEN REORDERED FOR SECONDARY SURGERY." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL/CORRECTED INFORMATION: B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/3.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL/ LENGTH LENS; DIFFERENT AXIS AND THE PROBLEM WAS SOLVED. ADDITIONAL INFORMATION REPORTEDLY WAS " FOR THE FIRST TIME THE DOUBLE EYE TICL WAS INSTALLED REVERSELY". IT WAS LATER REPORTED THAT "DURING THE FIRST OPERATION OF THIS PATIENT THE LEFT AND RIGHT EYES OF THE LENS WERE REVERSED THE LEFT LENS OF THE RIGHT EYE WAS LOADED AND THE RIGHT LENS OF THE LEFT EYE WAS LOADED SO THE REFRACTIVE SURPRISE OCCURRED AFTER THE OPERATION. AFTER THE OPERATION THE DOCTOR FOUND THE CAUSE OF REFRACTIVE SURPRISE WHEN CHECKING THE WHOLE OPERATION PROCESS. LENS OF THE SAME SIZE WERE THEN REORDERED FOR SECONDARY SURGERY." THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: ADDITRONAL INFORMATION: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. DEVICE EVALUATION: H3: TYPE OF INVESTIGATION:10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15765835&pc=QCB | 2023826-2022-04075 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+2.0/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4): SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15767691&pc=QCB | 2023826-2022-04076 | 2022-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15802020&pc=QCB | 2023826-2022-04128 | 2022-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.5/3.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THIS SAME DATE REPLACED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "SUGGESTED SIZE OF ICL BASED ON WTW AND ACD ON YOUR STAAR WEBSITE CALCULATOR FOR ASIAN EYES IS TOO LARGE. SHOULD BE ONE SIZE LESS". Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15810683&pc=QCB | 2023826-2022-04011 | 2022-10-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/1.0/097 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. DOCTOR FOUND A FILAMENTOUS FOREIGN BODY WAS FOUND ON THE LENS DURING SURGERY AND WAS REMOVED WITH FORCEPS AND IMPLANTED IN THE EYE. THE LENS REMAINS IMPLANTED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: FILAMENTOUS FOREIGN BODY ON LENS. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921105&pc=QCB | 2023826-2022-04453 | 2022-11-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -13.5/1.0/060 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE SURGEON REPORTS EXCESSIVE VAULTING AND NARROW ANGLES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION: (B)(4) - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15924103&pc=QCB | 2023826-2022-04457 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -11.50/2.0/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGD FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. FOR STATUS OF THE EYE THE REPORTER STATED: "THE CONJUNCTIVA OF THE RIGHT EYE IS LIGHT HYPEREMIA, THE CORNEAL WOUND IS WELL ALIGNED, THE HEALING IS ACCEPTABLE, THE ANTERIOR CHAMBER DEPTH IS ACCEPTABLE, THE LENS IS PLACED VERTICALLY, THE ARCH HEIGHT IS IDEAL, AND THE ANTERIOR CHAMBER ANGLE WIDTH IS 19.5 DEGREES". AS PER UPDATED INFORMATION THE REPORTER STATED " THE BLEEDING WAS NORMAL BLEEDING FROM THE LIMBAL INCISION DURING THE OPERATION. THE PATIENT HAS NOW HEALED WELL WITHOUT BLEEDING, AND THE PROBLEM HAS BEEN SOLVED." Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15929837&pc=QCB | 2023826-2022-04461 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 -7.00/2.5/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15952729&pc=QCB | 2023826-2022-04358 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO THE PATIENT COULD NOT DIFFERENTIATE BETWEEN COLORS AFTER THE LENS WAS IMPLANTED (COLOR BLIND). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THERE WAS DEBRIS THROUGHOUT THE LENS. DIMENSIONAL AND REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15954065&pc=QCB | 2023826-2022-04355 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15954096&pc=QCB | 2023826-2022-04356 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15954434&pc=QCB | 2023826-2022-04357 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO THE PATIENT COULD NOT DIFFERENTIATE BETWEEN COLORS AFTER THE LENS WAS IMPLANTED (COLOR BLIND). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DAMAGED. THERE WAS RESIDUE AND DEBRIS THROUGHOUT THE LENS. DIMENSIONAL WIDTH AND REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. UNABLE TO PERFORM DIMENSIONAL LENGTH DUE TO HAPTIC DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955851&pc=QCB | 2023826-2022-04521 | 2022-11-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/+0.5/106 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. IT WAS REPORTED THAT THE PATIENT HAS "WRONG REFRACTION" (OVER CORRECTION). LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND NOTED THE DEVICE FAILED TO PERFORM AS INTENDED; "WRONG INITIAL REFRACTION". Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION PROVIDED: " THE PATIENT HAS NOT BEEN AVAILABLE SINCE DECEMBER 2022 AND HAS NOT SPOKEN. I WANT TO CLOSED THE CASE". CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035460&pc=QCB | 2023826-2022-04559 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.0/1.5/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035468&pc=QCB | 2023826-2022-04570 | 2022-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -4.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16036914&pc=QCB | 2023826-2022-04723 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.5/+1.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) AS A REPLACEMENT LENS; ON (B)(6) 2022. THE SURGEON REPORTS ENDOTHELIAL CELLS OF 674 CELLS/MM2. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075815&pc=QCB | 2023826-2022-04651 | 2022-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED/STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. LENS MODEL SIZE RETURNED IS NOT MODEL SIZE INDICATED ON CLAIM. CLAIM# (B)(4). Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075857&pc=QCB | 2023826-2022-04655 | 2022-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/+2.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES SHORTER IN LENGTH. THE REPORTER STATED " LENS TOO BIG. SURGEON PREFERENCE CHANGE." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16087544&pc=QCB | 2023826-2022-04794 | 2022-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -10.50/+1.50/018 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE SURGEON REPORTS AT ONE DAY POSTOPERATIVE; DR. NOTED IRIS PIGMENTATION ON THE LENS IN THE LEFT EYE. PATIENT WAS ALSO CONSIDERED A STEROID RESPONDER AND DOCTOR PUT THE PATIENT ON DIAMOX AND IOP LOWERING DROPS. AT A LATER DATE IT WAS INDICATED THE PATIENT IS DOING WELL WITH GOOD IOP; OFF DIAMOX AND GOOD POSITIONING AND VISION. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. IT WAS NOTED "DOCTOR ROTATED LENS IN SULCUS 180 AND IT CAUSED RELEASE OF IRIS PIGMENTATION. HE ATTRIBUTED THIS TO INADEQUATE DILATION". Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16099316&pc=QCB | 2023826-2023-00063 | 2022-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+3.5/002 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16152680&pc=QCB | 2023826-2023-00125 | 2022-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -17.50-3.0-054 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND A TINY BIT OF PIGMENTATION ON THE ICL WAS OBSERVED. THE LENS REMAINS IMPLANTED. DOCTOR WILL WATCH THE PATIENT FOR NOW SINCE IOP AND VISION IS OPTIMAL. CAUSE REPORTED AS UNKNOWN, AND DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16164430&pc=QCB | 2023826-2023-00143 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -14.50/2.5/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " OS STILL A BIT SORE". PER UPDATED INFORMATION PATIENT IS HAPPY WITH VISION AND SORENESS HAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/2.5/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED EYE SORENESS. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "SHE IS HAPPY WITH HER VISION. SORENESS HAS RESOLVED". CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16165110&pc=QCB | 2023826-2023-00136 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.00/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. "THE LENS WAS REPOSITIONED TO 168 DEGREES BUT ICL BACK TO THE PRE-OP POSITION TWICE." ON (B)(6) 2022 THE LENS WAS EXPLANTED. CAUSE NOT REPORTED. AND IT IS NOTED THAT THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 03-JAN-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLINICAL CODE 4581: LENS DISLOCATION/SUBLUXATION. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16189408&pc=QCB | 2023826-2023-00200 | 2022-09-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS BEFORE THE PROCEDURE; THE SUBJECTIVE REFRACTION ON THE OP WAS -6.0/-2.25/140; THE LENS WAS SUPPOSED TO CORRECT THE CYLINDER IN ITS ENTIRETY; WHILE THE POSTOPERATIVE SUBJECTIVE REFRACTION REMAINED +0.50/-1.50/178; SO A SIGNIFICANT PART OF THE ASTIGMATISM REMAINED; ONLY IN THE UNUSUALLY CHANGED AXIS. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -9.0/-2.50/50 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED ON (B)(6) 2022. REFRACTIVE SURPRISE WAS REPORTED. REPORTEDLY "NO PATHOLOGICAL CHANGES NOTICED DURING LAST EXAMINATION. ICL PLACED IN AXIS 175 (PLANNED: 180)." PROBLEM IS NOT RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16200098&pc=QCB | 2023826-2023-00232 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.0/1.5/78 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THIS SAME DATE THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "DISCREPANCY BETWEEN WTW AND MONOGRAM?". Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16256994&pc=QCB | 2023826-2023-00324 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.00/1.5/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS REMOVED DUE TO NARROW ANGLES. IT WAS NOT REPLACED WITH ANOTHER LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273843&pc=QCB | 2023826-2023-00371 | 2022-12-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.5/3.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2022. THE SURGEON REPORTS THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE WITH LENS DISLOCATION/SUBLUXATION. IT WAS ADDITIONALLY NOTED THERE IS A PLANNED EXCHANGE DUE TO THE PATIENT EXPERIENCING DIPLOPIA. THERE WAS NO PATIENT INJURY. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16351463&pc=QCB | 2023826-2023-00467 | 2022-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -8.00/3.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT/LOSS. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16389094&pc=QCB | 2023826-2023-00650 | 2023-02-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.00/1.0/161 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; AND THE ANGLES ARE NARROW. LENS REMAINS IMPLANTED. THE SURGEON INDICATED THE PATIENT WILL BE MONITORED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "PATIENT BETTER SUITED FOR 12.6 DESPITE FDA GUIDE". PROBLEM IS NOT RESOLVED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; NARROW ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16389105&pc=QCB | 2023826-2023-00649 | 2023-02-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.00/1.0/159 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; AND THE ANGLES ARE NARROW. LENS REMAINS IMPLANTED. THE SURGEON INDICATED THE PATIENT WILL BE MONITORED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "PATIENT BETTER SUITED FOR 12.6 DESPITE FDA GUIDE". PROBLEM IS NOT RESOLVED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; NARROW ANGLES. TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16390507&pc=QCB | 2023826-2023-00594 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -6.00/5.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT OF SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5: LENS DIOPTER SHOULD BE CORRECTED TO STATE: -6.00/0.50/088 (SPHERE/CYLINDER/AXIS). CLAIM# (B)(4). Manufacturer Narrative: CLINICAL SIGNS (B)(4): SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411457&pc=QCB | 2023826-2023-00573 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT PRESENTED WITH A FIXED DILATED PUPIL POST-OP. THE ICL WAS IN GOOD POSITION AND IOP WAS WITHIN NORMAL LIMITS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. HEALTH IMPACT CLINICAL CODE: 4581 UNREACTIVE (FIXED) PUPIL. DEVICE CODE: 1494 THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT POST OP ASTIGMATISM AND A HIGHER THAN EXPECTED VAULT WERE NOTED. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16420139&pc=QCB | 2023826-2023-00611 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -16.00/2.0/086 (SPHERE/CYLINDER AXIS) WAS ATTEMPTED TO BE IMPLANTED ON (B)(6) 2023. "PATIENT PUPIL KEEP CONTRACTING." THE LENS WAS REMOVED AND A BACKUP LENS WAS IMPLANTED. NO PATIENT INJURY. Manufacturer Narrative: CLINICAL CODE 4581: "PATIENT PUPIL KEPT CONTRACTING". INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16507981&pc=QCB | 2023826-2023-00839 | 2023-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2M VTICMO13.2; -4.5/1.0/152 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES IN THE PATIENT. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS REPORTED AS NOT RESOLVED. ADDITIONALLY IT WAS NOTED "PATIENT WEARS RIMMED GLASSES". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 -SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE. MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16507984&pc=QCB | 2023826-2023-00840 | 2023-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -3.0/1.5/028 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES IN THE PATIENT. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS REPORTED AS NOT RESOLVED. ADDITIONALLY IT WAS NOTED "PATIENT WEARS RIMMED GLASSES". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 -SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE. MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16519908&pc=QCB | 2023826-2023-00815 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -0.50/6.0/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(6). HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. STATUS OF EYE REPORTED AS "BINOCULAR UCVA 20/25, ALL FINE, PATIENT IS HAPPY." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16519926&pc=QCB | 2023826-2023-00814 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -0.50/5.5/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER : (B)(6). HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. STATUS OF EYE REPORTED AS "BINOCULAR UCVA 20/25, ALL FINE, PATIENT IS HAPPY." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16532918&pc=QCB | 2023826-2023-00890 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THE PATIENT IS HAVING NEAR VISION PROBLEMS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - NEAR VISION PROBLEMS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16532958&pc=QCB | 2023826-2023-00889 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.00/2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS THE PATIENT IS HAVING NEAR VISION PROBLEMS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - NEAR VISION PROBLEMS. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16550445&pc=QCB | 2023826-2023-00956 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.0/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16603909&pc=QCB | 2023826-2023-00974 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+2.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ADDITIONAL SURGERY WAS PERFORMED WITH THE ADDITION OF NEW PI. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THERE WAS A SMALL CENTRAL ABRASION BUT IT WAS NOT CAUSED BY THE ICL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE SURGEON IMPLANTED A 12.6 VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.00/+2.50/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND YAG WAS PERFORMED. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THE SAME DAY SEPARATE SURGERY A REPLACEMENT LENS OF A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE STATES, "CLEAR, QUIET, SMALL CENTRAL ABRASION". REPORTEDLY " THE CORNEAL ABRASION WAS NOT RELATED TO THE ICL". ADDITIONAL INFORMATION STATES, "M -3.50OU 20/80+ (EXPECTED) WAS 20/70 PRE-OP". THE CAUSE OF THE EVENT IS UNKNOWN AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFO: UNK. DATE OF EVENT: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ADDITION NEW PI. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16625164&pc=QCB | 2023826-2023-01090 | 2023-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER +9.50/+6.0/083 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16633013&pc=QCB | 2023826-2023-01099 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/134 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16650146&pc=QCB | 2023826-2023-01140 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.50/+1.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS IMPLANTED AND REMOVED INTRAOPERATIVELY DUE TO THE SURGEON ACCIDENTALLY PUNCTURING THE ANTERIOR CAPSULE WITH A SCALPEL. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. PROBLEM RESOLVED. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581 CAPSULE TEAR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673868&pc=QCB | 2023826-2023-01035 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.0/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673881&pc=QCB | 2023826-2023-01042 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/2.0/033 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673917&pc=QCB | 2023826-2023-01052 | 2023-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.0/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673936&pc=QCB | 2023826-2023-01054 | 2023-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. HEALTH EFFECT IMPACT CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES. CLAIM # : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16683450&pc=QCB | 2023826-2023-01258 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.00/2.00/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. "DEVICE TOO LARGE" WAS REPORTED FOR CAUSE DETAILS. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16683536&pc=QCB | 2023826-2023-01259 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. "DEVICE TOO LARGE" WAS REPORTED FOR CAUSE DETAILS. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16710057&pc=QCB | 2023826-2023-01326 | 2022-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -4.00/+3.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE SURGEON REPORTS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16710074&pc=QCB | 2023826-2023-01327 | 2022-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.00/+5.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; INFERIOR IRIS; ENDOTHELIAL CORNEAL TOUCH; KERATITIS. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, INFERIOR IRIS, ENDOTHELIAL CORNEAL TOUCH, KERATITIS. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750639&pc=QCB | 2023826-2023-01232 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IN A SECOND SURGERY THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(40MMHG) WITHOUT PUPIL BLOCK AND SINGLE CLOSURE(ASSESSED ON 31/JAN/2023), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFO: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766249&pc=QCB | 2023826-2023-01439 | 2023-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.00/1.5/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. FIXED DILATED PUPIL; UNREACTIVE TO LIGHT WAS OBSERVED. THE REPORTER STATES "FROM DAY 1. PILOCARPINE USED WEEK 1 MINIMAL CONSTRICTION NOTED ON FFUP. PLANNING TO DO INTRACAMERAL DILATION. PRED AND OFTAQUIX WAS USED POST OP" . FOR PATIENTS CURRENT CONDITION THE REPORTER STATED THAT IOP WAS AT 11-13 AT ALL FFUPS; POST OP REFRACTION IS OS PLANO-0.75 X. LENS REMAINS IMPLANTED AND PUPIL IS STILL DILATED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: UNREACTIVE/FIXED PUPIL H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE: 1494: OFF-LABEL USE (PREVIOUS OR PRE-EXISTING OCULAR DISEASE THAT WOULD PRECLUDE POST-OPERATIVE VISUAL ACUITY OF 0.477 LOGMAR (20/60 SNELLEN) OR BETTER. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16769995&pc=QCB | 2023826-2023-01376 | 2023-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773625&pc=QCB | 2023826-2023-01415 | 2022-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/135 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS "OTHER- EXCESSIVE VAULT (SIZING ERROR)". Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16787274&pc=QCB | 2023826-2023-01470 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4.0/+1.0/146 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THAT MONOVISION WAS NOT TOLERATED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: AL THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4.0/+1.0/146 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THAT MONOVISION WAS NOT TOLERATED. LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. D9: RETURN DATE: 19-JUN-2023. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: PROBLEM NOT RESOLVED AFTER LENS WAS EXCHANGED. REPORTEDLY REFRACTIVE SURPRISE WAS OBSERVED ON (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - MONOVISION NOT TOLERATED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811558&pc=QCB | 2023826-2023-01537 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/115 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED. THE LENS WAS REMOVED ON (B)(6) 2023 AND IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850165&pc=QCB | 2023826-2023-01611 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -5.5/1.0/33 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939040&pc=QCB | 2023826-2023-01657 | 2023-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2, -9.0/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2023. PIGMENT DISPERSION AND IRREGULAR PUPIL HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. PATIENT REMAINS UNDER OBSERVATION. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5 - "WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023." - SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: PIGMENT DISPERSION, IRREGULAR PUPIL. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16942439&pc=QCB | 2023826-2023-01855 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50/+3.5/030 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REMOVED ON 30-MAR-2023 DUE TO EXCESSIVE VAULTING AND PIGMENT DISPERSION. THERE WAS VISIBLE IRIS CHAFING AND ANTERIOR LENS SURFACE PIGMENT ADHESIONS. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS IRRITATION-FREE PHAKIC DUOPHAKIA WITH GOOD VAULTING. THERE WAS SOME REFRACTIVE SURPRISE - SEE MFR. REPORT # 2023826-2023-01856 FOR REPLACEMENT LENS. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, IRIS CHAFING. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16960301&pc=QCB | 2023826-2023-01870 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. NO DISCOMFORT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16967951&pc=QCB | 2023826-2023-01914 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/116 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE REPORTED. THE LENS WAS EXPLANTED ON (B)(6)2023 AND IN ANOTHER SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM#(B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989907&pc=QCB | 2023826-2023-02015 | 2023-03-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+2.5/071 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS ALSO NOTED. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16996883&pc=QCB | 2023826-2023-01997 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS OBSERVED, LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTORS, HOWEVER "UNKNOWN" WAS ALSO REPORTED. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER REPORTS "KERATOCONUS". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023.ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGED RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTORS, HOWEVER "UNKNOWN" WAS ALSO REPORTED. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER REPORTS "KERATOCONUS. CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE CODE: 1494: OFF-LABEL USE (SAFETY AND EFFECTIVENESS OF THE LENS HAS NOT BEEN ESTABLISHED IN PATIENTS WITH KERATOCONUS). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17007062&pc=QCB | 2023826-2023-02048 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; 12.50/+3.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; 12.50/+3.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 4629 SHOULD BE OMITTED AND 4627 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17007066&pc=QCB | 2023826-2023-02049 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; 13.50/+3.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; 13.50/+3.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY, THE SURGEON DECIDED NOT TO EXPLANT THE LENS DUE TO DECREASED VAULT AND IOP WAS REDUCED WITHOUT TOPICAL DROPS. PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17013784&pc=QCB | 2023826-2023-01947 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/077 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINED IMPLANTED. THE PATIENT IS HAPPY WITH VERY GOOD VISION AND WILL BE MONITORED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023, WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED AND THE PATIENTS CURRENT CONDITION IS GOOD. CLAIM # (B)(4). Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015322&pc=QCB | 2023826-2023-02068 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039561&pc=QCB | 2023826-2023-01882 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/5.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. PRESSURE CONTROLLED WITH ORAL AND TOPICAL MEDICATIONS. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/5.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. PRESSURE CONTROLLED WITH ORAL AND TOPICAL MEDICATIONS. LENS EXPLANTED ON (B)(6) 2023. IN A SEPARATE SURGERY THAT SAME DATE THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039563&pc=QCB | 2023826-2023-01881 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/3.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. PRESSURE CONTROLLED WITH ORAL AND TOPICAL MEDICATIONS. LENS REMAINS IMPLANTED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/3.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. PRESSURE CONTROLLED WITH ORAL AND TOPICAL MEDICATIONS. THE LENS WAS EXPLANTED ON (B)(6) 2023. IN A SEPARATE SURGERY ON THE SAME DATE THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: D9: RETURN DATE: 19-SEP-2023. H6: 1494: SHOULD BE ADDED FOR CORRECTION- OFF-LABEL ACD<3.00MM AT THE TIME OF IMPLANTATION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084050&pc=QCB | 2023826-2023-02283 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.5/179 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084051&pc=QCB | 2023826-2023-02282 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.0/103 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17091497&pc=QCB | 2023826-2023-02290 | 2023-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/166 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17091777&pc=QCB | 2023826-2023-02288 | 2023-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.5/102 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE REPORTED. THE LENS WAS EXPLANTED ON (B)(6) 2023 AND IN SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17101390&pc=QCB | 2023826-2023-02296 | 2023-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.50/1.0/004 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON INDICATES THE TEXTURE/PATTERN OF THE PERIPHERY OF THE IRIS HAS CHANGED POST ICL SURGERY. IT WAS INDICATED THE LENS WILL NOT BE REMOVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - TEXTURE/PATTERN OF THE PERIPHERY OF THE IRIS CHANGED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- CORRECTED REPORT TYPE FROM PRODUCT PROBLEM TO ADVERSE EVENT. B2-ADD "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE" FOR CORRECTION. H1- CORRECTED TYPE OF REPORT EVENT FROM MALFUNCTION TO SERIOUS INJURY. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112171&pc=QCB | 2023826-2023-02325 | 2023-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112177&pc=QCB | 2023826-2023-02327 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+1.0/071 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112178&pc=QCB | 2023826-2023-02329 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+0.5/112 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112181&pc=QCB | 2023826-2023-02326 | 2023-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/+2.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112202&pc=QCB | 2023826-2023-02328 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/067 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119690&pc=QCB | 2023826-2023-02164 | 2023-05-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6, -17.0/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2023 DUE TO THE LENS TEAR DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT (OS) EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. A NEW SAME SIZE LENS WITH SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE SMALL TEAR NOTICED AFTER IMPLANTATION. NEW LENS IMPLANT WENT SMOOTH AND PATIENT IS DOING GOOD. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17138119&pc=QCB | 2023826-2023-02312 | 2023-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/110 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: A4: UNK A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17138218&pc=QCB | 2023826-2023-02313 | 2023-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: A4: UNK A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210599&pc=QCB | 2023826-2023-02598 | 2023-06-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -5.0/2.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "THE CRYSTAL SIZE IS TOO LARGE". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210601&pc=QCB | 2023826-2023-02599 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10/1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED VERTICALLY, THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: RETURN DATE: 19-SEP-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210782&pc=QCB | 2023826-2023-02536 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/041 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262464&pc=QCB | 2023826-2023-02685 | 2023-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.0/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279001&pc=QCB | 2023826-2023-02614 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE INDIAN J OPHTHALMOL , DATED APRIL-JUNE 2023. TITLED, 'NILATERAL PSEUDOEXFOLIATION DEPOSITS ON AN IMPLANTABLE COLLAMER LENS WITH URETTS-ZAVALIA SYNDROME IN FELLOW EYE.' THE ARTICLE INDICATED THREE GLAUCOMA DRUGS (OD) AND PSEDOEXFOLIATION MATERIAL DEPOSITS ON PULILLARY BORDER AND ON SURFACE OF ICL. PATIENT ON THREE GLAUCOMA DRUGS (OD). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). H6- HEALTH EFFECT IMPACT CODE 4581: URETTS-ZAVALIA SYNDROME, GLAUCOMA DRUGS, PSEDOEXFOLIATION DEPOSITS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4644" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: "4581-PSEUDOEXFOLIATION, PERIPHERAL ANTERIOR SYNECHIAE, PIGMENT DISPERSION" SHOULD BE ADDED. B5- AN ARTICLE WAS PUBLISHED CITING A CASE STUDY OF "UNILATERAL PSEUDOEXFOLIATION DEPOSITS ON AN IMPLANTABLE COLLAMER LENS WITH URETTS-ZAVALIA SYNDROME IN FELLOW EYE". THE PATIENT EXPERIENCED URETTS- ZAVALIA SYNDROME, PSEUDOEXFOLIATION, PERIPHERAL ANTERIOR SYNECHIAE, PIGMENT DISPERSION AND MEDICATION WAS PRESCRIBED. THIS STUDY INCLUDED 'A 33 YEAR OLD WOMAN WHO HAD UNDERGONE IMPLANTABLE COLLAMER LENS (ICL) IMPLANTATION ELSEWHERE IN BOTH EYES 4 YEARS AGO VISITED FOR ROUTINE CHECK UP. SHE HAD A HISTORY OF URETTS ZAVALIA SYNDROME IN THE RIGHT EYE (RE) POST SURGERY AND WAS CURRENTLY ON THREE GLAUCOMA DRUGS HER LEFT EYE (LE) SHOWED PSEUDOEXFOLIATION (PXF) MATERIAL DEPOSITS ON THE PUPILLARY BORDER AND ON THE SURFACE OF ICL, WHEREAS THE RE DID NOT SHOW ANY EVIDENCE OF PXF'. THE ARTICLE STATES THAT 'GONIOSCOPY SHOWED INTERMITTENT PERIPHERAL ANTERIOR SYNECHIAE IN THE RE AND OPEN ANGLES WITH IRREGULAR PIGMENT DEPOSITION IN THE LE. INTRAOCULAR PRESSURE IN THE LE WAS NORMAL WITHOUT MEDICATION'. IT WAS LATER REPORTED 'WE DID NOT INTERVENE' AND 'TO THE BEST OF OUR KNOWLEDGE THE LENSES REMAIN IMPLANTED'. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279005&pc=QCB | 2023826-2023-02613 | 2023-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -3.50/3.5/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. LENS CAUSING EVENT. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17306185&pc=QCB | 2023826-2023-02842 | 2023-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1 -12.0/3.5/016 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. FOR ADDITIONAL INFORMATION THE REPORTER STATED "VERY HIGH VAULT WITH ANGLE CLOSURE. SINCE 12.1MM IS THE LOWEST SIZE, NO EXCHANGE OPTION FOR THIS". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17337301&pc=QCB | 2023826-2023-02930 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/072 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344448&pc=QCB | 2023826-2023-02858 | 2022-09-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM,TMICL12.6, -15.0/2.5/120 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE ASTIGMATISM WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF EVENT REPORTED AS DEVICE, LENS ROTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347859&pc=QCB | 2023826-2023-02992 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.00/3.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD TO BE REMOVED AT THE TIME OF SURGERY DUE TO NOTICEABLY EXCESSIVE VAULT INTRAOPERATIVELY WITH SIGNIFICANT AC SHALLOWING AND IRIS PROLAPSE. THE LENS WAS REMOVED AND SURGERY WAS ABORTED WITHOUT COMPLICATION. FOR STATUS OF THE EYE THE REPORTER STATED " NORMAL, NO SYMPTOMS VISION BACK TO BASELINE". THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO THE DEVICE. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: B1- "PRODUCT PROBLEM" SHOULD BE REMOVED AND "ADVERSE EVENT" SHOULD BE ADDED. B2- "OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS" SHOULD BE ADDED. H1- "MALFUNCTION" SHOULD BE REMOVED AND "SERIOUS INJURY" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "2199" SHOULD BE REMOVED AND "4614" SHOULD BE ADDED.. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/3.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT AC SHALLOWING, AND IRIS PROLAPSE. REPORTEDLY, "ICL HAD TO BE EXPLANTED AT THE TIME OF SURGERY DUE TO NOTICEABLY EXCESSIVE SIZE INTRAOPERATIVELY WITH SIGNIFICANT AC SHALLOWING AND IRIS PROLAPSE. THE LENS WAS EXPLANTED AND SURGERY ABORTED WITHOUT COMPLICATION". THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "NORMAL, NO SYMPTOMS VISION BACK TO BASELINE". THE CAUSE OF THE EVENT IS THE DEVICE AND THE DEVICE FAILED TO PERFORM AS INTENDED. DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION FOUND HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT AC SHALLOWING AND IRIS PROLAPSE TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17366832&pc=QCB | 2023826-2023-03091 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4.50/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023 AS A REPLACEMENT LENS. IT WAS REPORTED THAT THE PATIENT IS "STILL PL-2.00X077" AND IT WAS INDICATED THERE IS RESIDUAL CYLINDER. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - RESIDUAL CYLINDER. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17372945&pc=QCB | 2023826-2023-02890 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. POST-OP, THE MYDRIASIS CONTINUED AFTER GLAUCOMA ATTACK. THE LENS REMAINED IMPLANTED AND THE PATIENT IS UNDER OBSERVATION. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-02889. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B3: UNK. D6B: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - MYDRIASIS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17377645&pc=QCB | 2023826-2023-02962 | 2023-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17385850&pc=QCB | 2023826-2023-03156 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.50/1.5/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON NOTES "POST ROTATION CENTERING ON 16.5. THE PATIENT IS STILL EXPERIENCING SEVERE GLARE AND PERSISTENT TRIPLE IMAGES OS. THE EXISTING LENS WILL BE EXPLANTED SOON". LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS REPOSITIONED. THIS DID NOT SOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SIMILAR LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT CLINICAL CODE: 4581 - TRIPLE IMAGES. H6- TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391257&pc=QCB | 2023826-2023-03032 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -2.5/1.0/040 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391260&pc=QCB | 2023826-2023-03031 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -2.5/1.0/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO UNREACTIVE(FIXED) PUPIL AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: UNREACTIVE (FIXED) PUPIL. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486818&pc=QCB | 2023826-2023-03264 | 2023-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.0/2.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT HAVE BEEN OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486828&pc=QCB | 2023826-2023-03270 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE INT J OPHTHALMOL IN JULY 2023 TITLED, "LONT-TERM OBSERVATION ON SAFETY AND VISUAL QUALITY OF IMPLANTABLE COLLAMER LENS V4C IMPLANTATION FOR MYOPIA CORRECTION: 5-YER FOLLOW-UP.' THE ARTICLE STATES THAT ADVERS NEVENT OCCURRED IN 4 EYES IN 3 PATIENT. FOLLOWING TRAUMA INCIDENT, ONE PATIENT HAD LENS ROTATION IN BOTH EYES(OU). ANOTHER PATIENT HAD IRIS INCARCERATION IN LEFT EYE(OS) AND WAS RELIEVED BY SURGICAL INTERVENTION. THE FINAL PATIENT SUFFERED LENS MOVEMENT IN THE LEFT EYE(OS) FOLLOWING TRAUMA. ALL CASES WERE OPERATED TWICE, POST OPERATIVE CDVA WAS GREATER THAN OR EQUAL TO THE PEROPERATIVE CDVA. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581-ICL DISPLACEMENT" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4625-SURGICAL INTERVENTION" AND "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "2923", "1494- OFF-LABEL USE (ACD<3.0MM)" AND "1494-OFF LABEL USE (AGE<21) " SHOULD BE ADDED. B5- AN ARTICLE WAS PUBLISHED IN INT J OPHTHALMOL CITING A CASE STUDY OF "LONG TERM OBSERVATION ON SAFETY AND VISUAL QUALITY OF IMPLANTABLE COLLAMER LENS V4C IMPLANTATION FOR MYOPIA CORRECTION: A 5 YEAR FOLLOW UP". THE PATIENTS EXPERIENCED IRIS INCARCERATION, ICL DISPLACEMENT, LENS REPOSITIONING, LENS ROTATION AND SURGERY WAS PERFORMED. THE PURPOSE OF THE ARTICLE REPORTED, "ADVERSE EVENTS OCCURRED IN 4 EYES (3.5%) IN 3 PATIENTS AT 5Y AFTER SURGICAL IMPLANTATION. FOLLOWING A TRAUMA INCIDENT, ONE PATIENT HAD TICL ROTATION IN THE RIGHT EYE AND IT WAS ADJUSTED, AND THE SPONTANEOUS TICL ROTATION OF THE OPPOSITE EYE WAS ALSO ADJUSTED. THE SECOND PATIENT (0.9%) HAD IRIS INCARCERATION IN LEFT EYE (0.9%), AND RELIEVED BY SURGICAL INTERVENTION. THE THIRD PATIENT SUFFERED FROM ICL MOVEMENT IN LEFT EYE FOLLOWING TRAUMA. ALL THESE CASES WERE OPERATED TWICE, AND THE POSTOPERATIVE CDVA WAS GREATER THAN OR EQUAL TO THE PREOPERATIVE CDVA". THIS STUDY INCLUDED "58 PATIENTS (114 EYES) WHO WERE FOLLOWED UP FOR AT LEAST 5Y AFTER ICL V4C IMPLANTATION". THE ARTICLE STATES THAT TWO CASES OF TICL ROTATION HAD "ALIGNMENT" PERFORMED, THE IRIS INCARCERATION HAD "SURGERY" PERFORMED AND THE "ICL DISPLACEMENT AFTER TRAUMA" HAD THE "REMOVAL OF ICL". THE AUTHORS CONCLUDE THAT "ICL V4C IMPLANTATION WAS SAFE AND EFFECTIVE FOR THE CORRECTION OF HIGH MYOPIA UP TO 5Y AFTER SURGERY". (B)(4) Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D1 - COPY&PAST BRAND NAME. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. D6B - UNK. G4 - PMA/510(K):THIS PRODUCT IS NOT MARKETED IN THE US. H4 - UNK. H6- HEALTH EFFECT IMPACT CODE 4581: TRAUMA INCIDENT, IRIS INCARCERATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17493561&pc=QCB | 2023826-2023-03435 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO PATIENT SUFFERING FROM PSYCHOLOGICAL ANXIETY, COULD NOT SLEEP, AND SUFFERED FROM SEVERE NEURASTHENIA. REPORTEDLY AFTER THE ICL LENS IMPLANTATION, THE PATIENT HAD A FOREIGN OBJECT IN THE EYE, AND FELT ANXIOUS AND SLEEPLESS! AFTER THE LENS WAS REMOVED THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515502&pc=QCB | 2023826-2023-03412 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE PATIENT WAS NOT HAPPY WITH HER VISION, WHICH CAME OUT TO +2.75 -0.50. THE LENS REMAINED IMPLANTED ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PATIENT NOT HAPPY WITH VISION. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17558286&pc=QCB | 2023826-2023-03477 | 2023-06-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO2.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS ALSO NOTED. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) AND (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17558287&pc=QCB | 2023826-2023-03474 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/033 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17591893&pc=QCB | 2023826-2023-03642 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -17.0/2.50/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "HIGHER SIZE LENS". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619509&pc=QCB | 2023826-2023-03706 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.5/1.5/011 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED TRANSIENT LOSS OF BCVA. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "THE BCVA OF THE LEFT EYE DECREASED AFTER THE IMPLANTATION OF MILD UNDERCORRECTED LENS ACCORDING TO THE RECOMMENDATION OF OCOS IN BOTH EYES". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - TRANSIENT LOSS OF BCVA. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17655763&pc=QCB | 2023826-2023-03798 | 2023-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -4.5/1.0/153 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS SHALLOW CHAMBER, EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER ON THIS SAME DATE REPLACED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER". Manufacturer Narrative: CORRECTED DATA: D6B: EXPLANT DATE: "ON (B)(6) 2023" SHOULD BE REMOVED AND "ON (B)(6) 2023" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2023, THE LENS WAS EXPLANTED AND LATER REPLACED WITH THE SAME MODEL, SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS OTHER. CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SHALLOW CHAMBER. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17657030&pc=QCB | 2023826-2023-03757 | 2023-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669066&pc=QCB | 2023826-2023-03735 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.0/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674681&pc=QCB | 2023826-2023-03765 | 2023-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.5/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704773&pc=QCB | 2023826-2023-03952 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.0/+1.5/4 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON EXPLANTED THE LENS ON (B)(6) 2023, DUE TO THE PATIENT HAVING DISSATISFACTION AND FOREIGN BODY SENSATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: B5 - THE REPORTER INDICATED THE PATIENT ALSO EXPERIENCED PAIN. CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PATIENT DISSATISACTION, FOREIGN SENSATION. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741413&pc=QCB | 2023826-2023-03698 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/021 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SMALLER SIZE AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741535&pc=QCB | 2023826-2023-03699 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/169 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SMALLER SIZE AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17754468&pc=QCB | 2023826-2023-04071 | 2023-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -2.5/1.0/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS REMOVAL AND NO FURTHER TREATMENT REQUESTED BY PATIENT. PATIENT IS HAPPY. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17755067&pc=QCB | 2023826-2023-03983 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2014. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: IT WAS LATER REPORTED THAT ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT POWER WAS IMPLANTED BUT THIS DID NOT RESOLVE THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: B4: DATE OF REPORT: 23-AUG-2023 CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6: HEALTH IMPACT- CLINICAL CODE:(B)(4). H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17761233&pc=QCB | 2023826-2023-04076 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.6, -9.5/3.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS COMPLAINT. IT WAS INDICATED THAT THE PATIENT HAS MILD PHOTOPHOBIA BUT WITH CLEAR VISION AND THE EYE IS IN REALLY GOOD CONDITION. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PHOTOPHOBIA. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770484&pc=QCB | 2023826-2023-04123 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.6, -9.0/2.0/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023 AS A REPLACEMENT LENS COMPLAINT. IT WAS INDICATED THAT THE PATIENT HAS MILD PHOTOPHOBIA BUT WITH CLEAR VISION. "NO SUPPLEMENTATION". LENS REMAINS IMPLANTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.0/2.0/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS COMPLAINT. IT WAS INDICATED THAT THE PATIENT HAS MILD PHOTOPHOBIA BUT WITH CLEAR VISION. "NO SUPPLEMENTATION". LENS REMAINS IMPLANTED. CLAIM #: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PHOTOPHOBIA. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770486&pc=QCB | 2023826-2023-04101 | 2023-03-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.5/+0.5/145 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS LOW VAULTING AND THERE WAS DECREASED NEAR VISION IN BRIGHT LIGHT. LENS REMAINS IMPLANTED WITH A PLANNED LENS EXCHANGED. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER AND NOTED "BRIGHT LIGHT CONSTRICTS THE PUPIL, PRESSING THE ICL TOWARDS THE CRYSTALLINE LENS INHIBITING THE INCREASE IN ANTERIOR CURVATURE WHICH IS NEEDED FOR NEAR VISION. THE ICL FLATTENS THE ANTERIOR LENS CURVATURE CAUSING PRESBYOPIA ONLY OCCURRING IN BRIGHT LIGHT". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - DECREASED VISION IN BRIGHT LIGHTING. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). Manufacturer Narrative: CORRECTED DATA: H4- "01/31/2026" SHOULD BE REMOVED AND "03/03/2023" SHOULD BE ADDED. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822037&pc=QCB | 2023826-2023-04303 | 2023-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -7.50/-1.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED PIGMENT DISPERSION AND DILATED PUPIL. LENS REMAINS IMPLANTED. IT WAS INDICATED THE PATIENT WILL JUST BE MONITORED FOR NOW. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B1: "PRODUCT PROBLEM" SHOULD BE REMOVED AND "ADVERSE EVENT" SHOULD BE ADDED. B2: "OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS" SHOULD BE ADDED. H1: "MALFUNCTION" SHOULD BE REMOVED AND "SERIOUS INJURY" SHOULD BE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION, DILATED PUPIL H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822040&pc=QCB | 2023826-2023-04302 | 2023-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.50/2.50/76 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED PIGMENT DISPERSION AND DILATED PUPIL. LENS REMAINS IMPLANTED. IT WAS INDICATED THE PATIENT WILL JUST BE MONITORED FOR NOW. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT OR PRODUCT PROBLEM: "PRODUCT PROBLEM" SHOULD BE REMOVED AND "ADVERSE EVENT" SHOULD BE ADDED. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: "OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS" SHOULD BE ADDED. H1: TYPE OF REPORTABLE EVENT: "MALFUNCTION" SHOULD BE REMOVED AND "SERIOUS INJURY" SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "4614" SHOULD BE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION, DILATED PUPIL. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822377&pc=QCB | 2023826-2023-04292 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.5/086 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17847262&pc=QCB | 2023826-2023-04323 | 2023-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/+3.0/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858049&pc=QCB | 2023826-2023-04329 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/041 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. CLAIM #(B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) AND (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862023&pc=QCB | 2023826-2023-04067 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.0/2.0/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND REDUCED AC DEPTH; MID-DILATED BUT RESPONSIVE PUPIL WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE; LENS TOO LARGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-"CORNEAL" ANGELS, "REDUCED" AC DEPTH AND MID-DILATED BUT RESPONSIVE PUPIL. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17885043&pc=QCB | 2023826-2023-04318 | 2023-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+3.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17894979&pc=QCB | 2023826-2023-04389 | 2023-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+4.0/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON ((B)(6) 2023 DUE TO LENS ROTATION (LENS DISLOCATION/SUBLUXATION), PIGMENT DISPERSION AND REFRACTIVE CHANGE OVERTIME. THE INCISION WAS ENLARGED AND A SUTURE WAS REQUIRED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS SIGNIFICANT CORNEAL EDEMA AND MARKED INFLAMMATION. SEE MFR. REPORT # 2023826-2023-04425 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, REFRACTIVE CHANGE OVERTIME. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - INCISION ENLARGED, SUTURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17916953&pc=QCB | 2023826-2023-04400 | 2023-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17944001&pc=QCB | 2023826-2023-04577 | 2023-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8/+2.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON NOTES THE PATIENT HAD EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8/+2.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON NOTES THE PATIENT HAD EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17944027&pc=QCB | 2023826-2023-04576 | 2023-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7/+1.5/108 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON NOTES THE PATIENT HAD EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17966721&pc=QCB | 2023826-2023-04640 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/2.0/012 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT'S VISION (UCVA) IS NOT MEETING EXPECTATIONS; AND WILL BE MONITORED. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - UCVA NOT MEETING EXPECTATIONS. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17969208&pc=QCB | 2023826-2023-04665 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18021670&pc=QCB | 2023826-2023-04731 | 2023-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/164 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. AN UNREACTIVE (FIXED) PUPIL IS OBSERVED. REPORTEDLY, "THE PATIENT IS MONITORED CLOSELY." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18021671&pc=QCB | 2023826-2023-04732 | 2023-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -7.0/1.0/158 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. AN UNREACTIVE (FIXED) PUPIL IS OBSERVED. REPORTEDLY, "THE PATIENT IS MONITORED CLOSELY." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067915&pc=QCB | 2023826-2023-04773 | 2023-09-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00/+4.5/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-04772, BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT COMPLAINED OF DOUBLE VISION. AT ONE DAY POST-OP, THE DOUBLE VISION WAS IMPROVING. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: NA. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DOUBLE VISION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18086564&pc=QCB | 2023826-2023-04880 | 2023-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+2.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18093426&pc=QCB | 2023826-2023-05029 | 2023-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/3.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THERE WAS AN UNEXPECTED OUTCOME. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - UNEXPECTED OUTCOME. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18107627&pc=QCB | 2023826-2023-05015 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/011 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18107628&pc=QCB | 2023826-2023-05014 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/177 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023, FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136585&pc=QCB | 2023826-2023-05092 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A CASE REPORT WAS RECEIVED REGARDING "IMPLANTABLE COLLAMER LENS SUBLUXATION IN A PATIENT WITH LENTICULAR COLOBOMA: A CASE REPORT." THE ARTICLE REPORTED A 27-YEAR-OLD MAN WITH A MARFANOID BODY HABITUS WAS SEEKING REFRACTIVE SURGERY FOR CORRECTION OF HIGH ASTIGMATISM IN RIGHT EYE. A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+5.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED. AT THREE WEEKS POST-OP, THE PATIENT PRESENTED WITH DECREASED VISION, THE VAULTING WAS LOW AND THE ICL HAD SUBLUXATED INFERIORLY. AN ULTRASOUND BIO MICROSCOPY CONFIRMED THE PRESENCE OF A CILIARY BODY CYST. THE UNSTABLE ICL WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DECREASED VISION. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18170877&pc=QCB | 2023826-2023-05205 | 2023-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179625&pc=QCB | 2023826-2023-05188 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.00/3.5/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO PATIENT REQUEST. THE PROBLEM WAS RESOLVED AND CAUSE REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-CLINICAL CODE 4581: PATIENT REQUESTED TO REMOVE ICL. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179626&pc=QCB | 2023826-2023-05187 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/3.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO PATIENT REQUEST. THE PROBLEM WAS RESOLVED AND CAUSE REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-CLINICAL CODE 4581: PATIENT REQUESTED TO REMOVE ICL. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188082&pc=QCB | 2023826-2023-05128 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, 11.00/+2.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). POST-OP, THE PATIENT COMPLAINED THAT THE PUPIL WAS NOT CLOSING VERY WELL AND IS WORSE AT NIGHT. THE PATIENT WAS TREATED WITH PILOCARPINE. AFTER OBSERVING A LITTLE RESPONSIVENESS, THE DOCTOR PERFORMED AN OCT AND LENS WAS IS IN A GOOD POSITION. THE PRESSURE IS NORMAL AND PATIENT HAS GOOD QUALITY VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PUPIL NOT CLOSING. H6: HEALTH EFFECT IMPACT CODE: 4644 - PILOCARPINE. . H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258186&pc=QCB | 2023826-2023-05309 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.5/1.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO UNREACTIVE (FIXED) PUPIL. THIS HAS NOT RESOLVED THE PROBLEM. REPORTEDLY, PATIENT PUPIL IS 3-4MM AND CONTINUES TO BE UNDER OBSERVATION. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18340520&pc=QCB | 2023826-2023-05679 | 2023-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/1.0/130 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED; THE SURGEON HAD NOTED "THE PATIENT NEEDS TO CHANGE THE ICL AFTER SURGERY; TAKING ACCOUNT THE IMPACT OF THE SECOND SURGICAL INCISION WHICH WILL GENERATE NEW ASTIGMATISM. COMBINED WITH THE COMPREHENSIVE OPTOMETRY AFTER THE FIRST SURGERY A NEW OPTIMAL DEGREE SHOULD BE SELECTED FOR THE SECOND SURGERY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18341179&pc=QCB | 2023826-2023-05639 | 2023-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+1.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18341181&pc=QCB | 2023826-2023-05638 | 2023-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/109 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18346602&pc=QCB | 2023826-2023-05631 | 2023-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.50/4.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 17/OCT/2023. SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS OBSERVED. ON (B)(6) 2023, AN ANTERIOR CHAMBER FORMATION WITH BSS AND SUTURING OF THE INCISION WAS PERFORMED AS A PROCEDURE FOR A SIGNIFICANTLY NARROW ANGLE CORNER. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023, A SHORTER LENGTH DIFFERENT DIOPTER LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT IMPACT CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS, ANGLE CLOSURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387571&pc=QCB | 2023826-2023-05572 | 2023-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+5.0/103 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON 05-DEC-2023 DUE TO EXCESSIVE VAULTING, IRIDOCORNEAL TOUCH AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. STANDARD POSTOPERATIVE HEALING PROCESS. ALL SCANS AND SYMPTOMS WITHIN NORMAL LIMITS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. THE LENS WAS DISCARDED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387576&pc=QCB | 2023826-2023-05727 | 2023-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00/+1.0/001 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS CALM. THE CAUSE OF THE EVENT WAS DUE TO PROBABLE IRREGULAR EYE; SMALL ANGLE-TO-ANGLE. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387584&pc=QCB | 2023826-2023-05726 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+3.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-03983 BUT THE PROBLEM WAS NOT RESOLVED. THERE WAS STILL A MYOPIC RESIDUAL AFTER THE EXCHANGE. THE LENS REMAINED IMPLANTED AND THE SURGEON WILL OBSERVE THE PATIENT. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - MYOPIC RESIDUAL. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387615&pc=QCB | 2023826-2023-05571 | 2023-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+5.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, IRIDOCORNEAL TOUCH AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. STANDARD POSTOPERATIVE HEALING PROCESS. ALL SCANS AND SYMPTOMS WITHIN NORMAL LIMITS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. THE LENS WAS DISCARDED. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18427342&pc=QCB | 2023826-2024-00008 | 2023-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.0/4.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ANGLE CLOSURE AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, PATIENT RELATED FACTOR. Manufacturer Narrative: A4 - UNK; A5 - UNK; A6 - UNK; H6 - HEALTH EFFECT CLINICAL CODE 4581: ANGLE CLOSURE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18428711&pc=QCB | 2023826-2024-00060 | 2023-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13/2.5/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT HAS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2024 THE SURGEON INDICATES HE WILL EXPLANT AND REPLACE THE LENS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18467571&pc=QCB | 2023826-2024-00150 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/0.5/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE ON (B)(6) 2023 AS A REPLACEMENT LENS. THE SURGEON REPORTS THERE WILL BE ANOTHER ICL EXCHANGE TO TAKE PLACE TO GENERATE A NEW ASTIGMATISM AND "COMBINED WITH THE COMPREHENSIVE OPTOMETRY AFTER THE FIRST SURGERY, A NEW OPTIMAL DEGREE SHOULD BE SELECTED". LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ASTIGMATISM ISSUES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18548443&pc=QCB | 2023826-2024-00294 | 2023-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+3.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18548445&pc=QCB | 2023826-2024-00295 | 2023-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+2.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18565310&pc=QCB | 2023826-2024-00332 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+4.0/005 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED. THE LENS WAS REMOVED ON (B)(6) 2023 AND IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18565318&pc=QCB | 2023826-2024-00326 | 2023-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+5.5/111 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THAT SAME DAY A LENS TWO SIZES SHORTER IN LENGTH WAS IMPLANTED HOWEVER IT IS UNCLEAR IF THIS WAS PERFORMED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515401&pc=QCB | 2023826-2020-02028 | 2020-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Unintended Movement | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.00/2.5/017 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. AFTER INSERTION, THE LENS "FLIPPED IN THE EYE. THE LENS WAS REMOVED AND RE-INSERTED 4 TIMES INTRAOPERATIVELY, EACH TIME FLIPPING INSIDE THE EYE. LENS WAS REMOVED. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE), WITH PATIENT INJURY OF CORNEAL EDEMA DUE TO LENS BEING INJECTED 4 TIMES. PATIENT PRESCRIBED PREDNISOLONE ACETATE AND GATIFLOXACIN DROPS 4 BID. THIS RESOLVED THE PROBLEM. REPORTEDLY, "CORNEAL EDEMA WAS RESOLVED, PATIENT IS IN GOOD CONDITION TO REOPERATE, NO COMPLAINTS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: "A NEW LENS OF SAME LENGTH WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. IT WAS REPORTED THAT CORNEAL EDEMA DISSAPPEARED." SHOULD BE IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "POST-OP STATUS OF THE EYE WAS REPORTED AS MYDRIASIS, CORNEAL EDEMA, PHOTOPHOBIA, AND DESCEMET FOLDS." SHOULD BE IN THE INITIAL MDR. H6- CLINICAL CODE 4581: MYDRIASIS; PHOTOPHOBIA; DESCEMET FOLDS.- SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. MINOR INJURY/ ILLNESS / IMPAIRMENT; LOWERING DROPS AND PILLS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10539565&pc=QCB | 2023826-2020-02150 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.00/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. LENS REMAINS IMPLANTED, THE REPORTER STATED " THE PATIENT IS STILL UNDER OBSERVATION." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.00/2.0/089(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548187&pc=QCB | 2023826-2020-02106 | 2019-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT EXPERIENCED INTENSE PHOTOPHOBIA AND LOCAL ADHESION OF THE ANTERIOR CHAMBER ANGLE WAS OBSERVED POSTOPERATIVELY, SO SURGEON DECIDED TO EXCHANGE THE LENS. AFTER EXCHANGE, THE EYE CONDITION IS GOOD." Manufacturer Narrative: H3 - LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PHOTOPHOBIA, LOCAL ADHESION OF ANTERIOR CHAMBER ANGLE, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678716&pc=QCB | 2023826-2020-02360 | 2020-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH DEBRIS AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678726&pc=QCB | 2023826-2020-02364 | 2020-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014613&pc=QCB | 2023826-2020-03004 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+3.0/137 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS ROTATED AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ANOTHER ICL LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE CAUSE OF THE EVENT WAS PROBABLY DUE TO THE ICL WAS BIGGER THAN THE ANATOMY OF THE EYE. CORRECTED DATA: B3: 20-NOV-2020. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014931&pc=QCB | 2023826-2020-03005 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS ROTATED. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS IMPLANTED ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE CAUSE OF THE EVENT WAS PROBABLY DUE TO THE ICL WAS BIGGER THAN THE ANATOMY OF THE EYE. CORRECTED DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129531&pc=QCB | 2023826-2021-00079 |  | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.5/2.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED THE REPORTER STATED " THERE WAS LOW VAULT, THE LENS WAS DECENTERED." IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS " WEAK ZONULES." IT WAS REPORTED THAT THE PATIENT IS AWAITING LASIK/PRK. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL, ON MEDICAL GAUZE. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND A HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4625 - LASER REFRACTIVE SURGERY WAS ALSO PERFORMED. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ZONULAR LAXITY, LENS TILS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11322454&pc=QCB | 2023826-2021-00418 | 2020-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/1.00/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND PERMANENT LOSS OF BCVA. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS " OTHER, LOW LENS DIAMETER." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MCROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CORRECTED DATA; H6-CLINICAL CODE 4581-PERMANENT LOSS BCVA SHOULD HAVE BEEN ADDED TO INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400369&pc=QCB | 2023826-2021-00529 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON STATES THERE IS A HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ICA AND PUPILLARY ACTION REDUCED. THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE IS REPORTED AS PUPILLARY ACTION WITHOUT ANY PROBLEMS, PHOTOPIC PUPILLARY DIAMETER 2.0MM (MESOPIC DIAMETER 4.0MM). THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B1-UPDATED. B5-UPDATED; PLEASE DISREGARD PREVIOUS B5 STATEMENT. H6-HEALTH CLINICAL CODE UPDATED: 4581 (SIGNIFICANT REDUCTION OF ICA, PUPILLARY ACTION REDUCED). HEALTH IMPACT CODE UPDATED. DEVICE CODE 1494 (PREGNANT AT THE TIME OF IMPLANT). (B)(4) Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11448819&pc=QCB | 2023826-2021-00605 | 2021-02-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, IT WILL POSSIBLY BE EXCHANGED DUE TO THE LENS BEING TOO VAULTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS; -6.0/+1.5/073 (SPHERE/CYLINDER/AXIS); WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2021 WITH A SHORTER LENGTH LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS DEVICE; NOMOGRAM OR PT CORRESPONDENCE TO NOMOGRAM. ICL OVERSIZED CAUSING POOR SEATING AND IRIS LEFT WITH VERY NARROW ANGLES INFERIOR TEMP. QUADRANT. THE REPORTER STATED "ICL NOW LAYS CORRECTLY WITH NO ANGLE COMPROMISE. BEFORE EXCHANGE I DISINSERTED AND INSERTED FOOT PLATES/HEPTICS OF ORIGINAL ICL TO ENSURE NO CURLING OF THE FOOT PLATES- THIS DID NOT SOLVE THE PROBLEM SO THE ICL WAS EXCHANGED SUCCESSFULLY." H6- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIRDO-CORNEAL ANGLES. CLAIM# : (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788209&pc=QCB | 2023826-2021-01360 | 2020-11-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/073 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132577&pc=QCB | 2023826-2021-02385 | 2021-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET, OPEN ANGLES AND PATIENT IS HAPPY WITH OUTCOME. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558411&pc=QCB | 2023826-2021-03827 | 2021-08-11 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.5/1.0/70 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599777&pc=QCB | 2023826-2021-03906 | 2021-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATES A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -10.5/+2.0/087 (SPHERE/ CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. IRIS CHAFING AND LACK OF PUPIL CONSTRICTION TEMPORALLY DUE TO EXCESSIVE VAULT; CAUSED COSMETIC CONCERNS OF ANISOCORIA IS REPORTED. REPORTEDLY LASIK IS BEING CONSIDERED AT A LATER DATE TO CORRECT THE VISION. H6-HEALTH EFFECT CLINICAL CODE 4581 (IRIS CHAFING, LACK OF PUPIL CONSTRICTION). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND THERE WAS RESIDUE THROUGHOUT THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12662187&pc=QCB | 2023826-2021-04056 | 2021-06-26 | Injury | STAAR SURGICAL COMPANY | 2021-10-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/1.5/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666241&pc=QCB | 2023826-2021-04057 | 2021-06-26 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/2.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12708033&pc=QCB | 2023826-2021-04148 | 2021-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -4.5/0.5/050 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50/0.5/050 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT. ON (B)(6) 2021, THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THIS EVENT IS UNKNOWN." H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12741581&pc=QCB | 2023826-2021-04222 | 2021-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-11-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.00/+1.50/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. IT WAS ALSO NOTED THAT "THE RESIDUAL ASTIGMATISM WAS RESOLVED BY BIOPICS". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4): BIOPICS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12741584&pc=QCB | 2023826-2021-04223 | 2021-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-11-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.00/+1.00/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. IT WAS ALSO NOTED THAT "THE RESIDUAL ASTIGMATISM WAS RESOLVED BY BIOPICS". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4): BIOPICS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814434&pc=QCB | 2023826-2021-04312 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -11.00/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - REPORTED 088 AXIS - CORRECT TO 086 AXIS. H6- HEALTH EFFECT IMPACT CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12888619&pc=QCB | 2023826-2021-04565 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.5/1.00/139 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND IRIDO-CORNEAL TOUCH. A REPLACEMENT LENS OF THE SAME MODEL BUT SHORTER LENGTH WAS IMPLANTED ON (B)(6) 2021 AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS "LENS SIZE CHANGE". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12888654&pc=QCB | 2023826-2021-04566 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO_12.6; -12.5/1.00/19 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND IRIDO-CORNEAL TOUCH. A REPLACEMENT LENS OF THE SAME MODEL BUT SHORTER LENGTH WAS IMPLANTED ON (B)(6) 2021 AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS "LENS SIZE CHANGE". Manufacturer Narrative: B5 - CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO_12.6; -12.5/1.00/19 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND IRIDO-CORNEAL TOUCH. A REPLACEMENT LENS OF A DIFFERENT MODEL AND SHORTER LENGTH WAS IMPLANTED ON (B)(6) 2021 AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS "LENS SIZE CHANGE". H4 - CORRECTED TO (B)(6) 2021 IN INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082439&pc=QCB | 2023826-2021-04975 | 2021-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON STATES THERE IS A HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON STATES THERE IS A HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 13.2MM VTICM5_13.2, -9.5/+4.0/085 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE CAUSE IS REPORTED AS DEVICE: "OVERALL LENS DIAMETER TOO BIG." H6-HEALTH CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION ON 18-JAN-2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13173112&pc=QCB | 2023826-2022-00054 | 2021-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/083, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (POTENTIAL RETRO-IRIS CYSTS BUT NOT SURE). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/083, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND PIGMENT DISPERSION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (POTENTIAL RETRO-IRIS CYSTS BUT NOT SURE). D6B: (B)(6)2022. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: N/A. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256336&pc=QCB | 2023826-2022-00136 | 2021-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATES AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK. MODEL #: UNK. IMPLANT DATE: UNK. DEVICE MANUFACTURER DATE: UNK. CLAIM#: (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 13.7MM TMICL13.7, -12.5/+1.0/101 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "THE RECOMMENDATION FROM DFU FOR 13.7 WAS TOO LARGE FOR THE PATIENT." REPORTEDLY THE LENS DID NOT FAIL TO PERFORM AS INTENDED. H6-HEALTH CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13596387&pc=QCB | 2023826-2022-00478 | 2022-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -16.50/3.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THE SAME DAY DIFFERENT SURGERY A SHORTER LENGTH REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -16.50/3.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THE SAME DAY DIFFERENT SURGERY A SHORTER LENGTH REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13670784&pc=QCB | 2023826-2022-00615 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.0/1.0/036 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14115476&pc=QCB | 2023826-2022-00826 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS AN EXCHANGE LENS ON (B)(6) 2022 TO RESOLVE EXCESSIVE VAULT. REPORTER STATES THAT THE VAULT STILL SEEMS TOO HIGH. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: B5- THE PATIENT EXPERIENCED GLAZY, HAZY RINGS IN THE EVENING. PRESCRIBED LUBRICATING EYE DROPS. H6- HEALTH EFFECT- CLINICAL CODE: 4581- GLAZY, HAZY RINGS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14270697&pc=QCB | 2023826-2022-01342 | 2022-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND POWER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. REPORTEDLY EXCESSIVE VAULT WAS OBSERVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1-ADVERSE EVENT B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.0/1.5/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SHORTER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE. D6A- (B)(6) 2022. D6B- (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14270802&pc=QCB | 2023826-2022-01341 | 2022-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND POWER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. REPORTEDLY EXCESSIVE VAULT WAS OBSERVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.00/1.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE REPORTED AS DEVICE. D6A- (B)(6) 2022. D6B- (B)(6) 2022. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14554422&pc=QCB | 2023826-2022-01728 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/0.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS "SIGNIFICANT POSTOPERATIVE VAULTING". LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - PATIENT ALSO EXPERIENCED "RELATIVELY SMALL AC"; NARROW ANGLE; PIGMENT ON ICL - SHOULD HAVE BEEN INCLUDED ON INITIAL MEDWATCH REPORT. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SMALL AC, NARROW ANGLE, PIGMENT ON ICL. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14554431&pc=QCB | 2023826-2022-01729 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -17.50/1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS "SIGNIFICANT POSTOPERATIVE VAULTING". LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED "RELATIVELY SMALL AC"; NARROW ANGLE; PIGMENT ON ICL - SHOULD HAVE BEEN INCLUDED ON INITIAL MEDWATCH REPORT. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SMALL AC, NARROW ANGLE, PIGMENT ON ICL. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639205&pc=QCB | 2023826-2022-01793 | 2022-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/042 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_ 13.7 IMPLANTABLE COLLAMER LENS, -13/2.00/42 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. H6- HEALTH EFFECT- CLINICAL CODE- 4581- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14805947&pc=QCB | 2023826-2022-02007 | 2022-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -15.5/3.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHOTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. REPORTEDLY. "INITIAL LENS REPLACED WITH A SMALLER ONE AND THE VAULT IS IDEAL NOW. PATIENT DOING VERY WELL." Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- " THE PATIENT ALSO EXPERIENCED SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES." H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299314&pc=QCB | 2023826-2022-02903 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT A TORIC 12.6MM SIZE IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S LEFT (OS) EYE. REPORTEDLY, "PT HAS BEEN NOTED TO HAVE A SHALLOW ACD SINCE SHORTLY AFTER SURGERY; I SAW HER TODAY AND NOTED VERY HIGH VAULT WITH SHALLOW AC." Manufacturer Narrative: UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4582 SHALLOW ACD. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506733&pc=QCB | 2023826-2022-03360 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1/109 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENGTH, SPHERICAL LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. REPORTEDLY, "TORIC LENS GOT ROTATED TO 90 DEGREES AND PATIENT WAS UNHAPPY WITH THE OUTCOME. SUSPECTED ZONULAR WEAKNESS. REMOVED TORIC AND IMPLANTED SPHERICAL LENS. PATIENT IS DOING WELL POSTOPERATIVELY WITH 6/9 VISION." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5: LENS MODEL CORRECTED TO VTICM5_13.2 IN INITIAL MDR. D4: CORRECTED TO VTICM5_13.2 IN INITIAL MDR. H4: CORRECTED TO 28-MAR-2022 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- ZONULAR WEAKNESS. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50/1.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND ZONULAR WEAKNESS. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL/ SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED WAS " TORIC LENS GOT ROTATED TO 90 DEGREES AND PATIENT WAS UNHAPPY WITH THE OUTCOME. SUSPECTED ZONULAR WEAKNESS. REMOVED TORIC AND IMPLANTED SPHERICAL LENS. PATIENT IS DOING WELL POSTOPERATIVELY WITH 6/9 VISION." THE CAUSE IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15643337&pc=QCB | 2023826-2022-03627 | 2022-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF -8.0/4.0/021 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO A PATIENT'S EYE. IT WAS THEN REMOVED DUE TO THE LENS BEING, "TOO LARGE." A REPLACEMENT IS PLANNED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: 4581-SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. 2682-PATIENT-RELATED FACTOR REPORTED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15795263&pc=QCB | 2023826-2022-04001 | 2022-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -16.50/3.5/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON AN UNKNOWN DATE. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO SLIGHT CHANGE OF LENS RX. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -16.50/3.5/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15829645&pc=QCB | 2023826-2022-04193 | 2022-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -8.5/2.5/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 ASTIGMATISM WAS ASSESSED. LOW VAULT WAS OBSERVED AND ON 31-AUG-2022 THE LENS WAS EXCHANGED FOR A LONGER LENS OF DIFFERENT POWER AND THE PROBLEM RESOLVED. THE REPORTER STATED, "THE PATIENT IS ANXIOUS FOR SURGERY, SO THE SPHERICAL LENS AND CYLINDRICAL LENS ARE LOWERED BY 50°". IN THE REPORTER'S OPINION THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: HEALTH EFFECT- (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845664&pc=QCB | 2023826-2022-04255 | 2021-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, 10.50/+1.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING (NARROW ANGLE). THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- "NARROW ANGLE." CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15946942&pc=QCB | 2023826-2022-04446 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 03-AUG-2022. ON 23-AUG-2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY STATUS OF THE EYE IS, "DISCOMFORT, GOOD VISION." Manufacturer Narrative: PATIENT INFORMATION:UNK. CLAIM# 727154. Manufacturer Narrative: CORRECTED DATA: B5- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS ADDED FOR CORRECTIONS. D9-RETURN DATE: 20/12/22. H6: 4581 WAS ADDED FOR CORRECTION DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15952648&pc=QCB | 2023826-2022-04350 | 2022-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/080 (SPHERE/CYLINDER/AXIS), FROM THE PATIENTS LEFT EYE (OS) DUE TO THE BIG SIZE OF THE LENS. THE LENS WAS REPLACED WITH A SMALLER SIZE LENS. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SMALLER SIZE LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275251&pc=QCB | 2023826-2023-00259 | 2022-11-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/109 (SPHERE/CYLINDER/AXIS). THE LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES CORRECTED DATA: B1: PRODUCT PROBLEM B5: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/109 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS NORMAL. H1: MALFUNCTION H6: HEALTH EFFECT: IMPACT CODE: 2199 CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16519786&pc=QCB | 2023826-2023-00811 | 2023-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -10.50/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. ON AN UNKNOWN DATE THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS. IT IS UNKNOWN IF THIS RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H6-CLINICAL CODE-4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER - 10.5/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530285&pc=QCB | 2023826-2023-00748 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.0/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, WITH 12.1MM LENS, VAULT GOT REDUCED AND PATIENT IS DOING FINE. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16640712&pc=QCB | 2023826-2023-01014 | 2022-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.0/4.5/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6)2022, THE LENS WAS REMOVED AND IN A SUBSEQUENT SURGERY LATER THAT DAY, A LONGER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIOCORNEAL ANGLES. REPORTEDLY, "SIGNIFICANT LENS VAULT." THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED WAS THE DEVICE. Manufacturer Narrative: CORRECTED DATA: D4- MODEL#: VTICM5_12.6 SHOULD BE CORRECTED TO VTICM5_13.2. H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY & RACE:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678165&pc=QCB | 2023826-2023-01166 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL INTO THE PATIENTS LEFT EYE (OS). THE SURGEON REQUESTED A VECTOR ANALYSIS TO REDUCE THE ASTIGMATISM SO THE PATIENT MIGHT BE SATISFIED WITH A LITTLE MYOPIA. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT AGE OR DATE OF BIRTH : UNKNOWN. SEX : UNKNOWN. WEIGHT : UNKNOWN. ETHNICITY: UNKNOWN. RACE : UNKNOWN. DATE OF EVENT : UNKNOWN. MODEL # : UNKNOWN. IMPLANT DATE : UNKNOWN. EXPLANT DATE : UNKNOWN. DEVICE MANUFACTURE DATE : UNKNOWN CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: G4: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM TICM130V4 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON 02-FEB-2016. THE PATIENT EXPERIENCED A REFRACTIVE CHANGE OVERTIME AND THE LENS REMAINED IMPLANTED. THE UCVA WAS NOT AS SHARP BUT NO OTHER COMPLAINTS OR SYMPTOMS. THE PATIENTS CURRENT REFRACTION IS PLANO -1.23 X 35 20/15. THE CAUSE OF THE EVENT WAS UNKNOWN. G4: PREMARKET INDENTIFICATION: PMA/510(K): UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773857&pc=QCB | 2023826-2023-01422 | 2022-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0/1.0/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. REPORTEDLY, "SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND ANTERIOR SYNECHIA." THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10 DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6 HEALTH EFFECT CLINICAL CODE: 4581 "SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES" AND "ANTERIOR SYNECHIAE." H6 TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - ACD<3.0. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773858&pc=QCB | 2023826-2023-01421 | 2022-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.0/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6)2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS, AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. REPORTEDLY, "SIGNIFICANT ENDOTHELIAL CELL LOSS, AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES." THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- "SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES" AND "SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS". H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811645&pc=QCB | 2023826-2023-01493 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+4.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # :(B)(4). Manufacturer Narrative: H6-HEALTH EFFECT- CLINICAL CODE: 4581- "DYSPHOTOPSIA". B5- THE PATIENT ALSO EXPERIENCED DYSPHOTOPSIA. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16857792&pc=QCB | 2023826-2023-01600 | 2023-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 . THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/90 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND SHALLOW ANTERIOR CHAMBER VOLUME. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL, SIMILAR POWER, SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "BECAUSE THE PATIENT'S ANTERIOR CHAMBER VOLUME IS SMALL, AND THE CRYSTALS ORDERED ACCORDING TO THE SYSTEM CONVERSION ARE SOMEWHAT LARGE, RESULTING IN A SHALLOW POSTOPERATIVE ANTERIOR CHAMBER VOLUME, THE PREOPERATIVE ANTERIOR CHAMBER VOLUME IS 174MM, FOR THE PATIENT'S POSTOPERATIVE SAFETY, CHOOSE TO REPLACE THE SMALLER CRYSTAL". THE CAUSE OF THE EVENT IS UNKNOWN. H6- HEALTH EFFECT- CLINICAL CODE: 4581- SHALLOW ANTERIOR CHAMBER VOLUME. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922359&pc=QCB | 2023826-2023-01778 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -17.50/+3.0/113, INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HIGH VAULTING, WITH NARROW IRIDOCORNEAL ANGLE. THE IRIS IS PUSHED BACK BY THE ICL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK. CLAIM # (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: "4581- NARROW IRIDO-CORNEAL ANGLES" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922573&pc=QCB | 2023826-2023-01779 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -17.50/+3.5/075, INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HIGH VAULTING, WITH NARROW IRIDOCORNEAL ANGLE. THE IRIS IS PUSHED BACK BY THE ICL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "4581- NARROW IRIDO-CORNEAL ANGLES" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924894&pc=QCB | 2023826-2023-01815 | 2022-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.5/170 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: B5: STATEMENT SHOULD INCLUDE SIGNIFICANT REDUCTION OF ICA. H6: HEALTH EFFECT CLINICAL CODE: 4581 (SIGNIFICANT REDUCTION OF ICA). CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066353&pc=QCB | 2023826-2023-02195 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. REPORTEDLY, "ICL IS CENTERED AND POSITIVE." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066783&pc=QCB | 2023826-2023-02194 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. REPORTEDLY, "ICL IS CENTERED AND POSITIVE." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093870&pc=QCB | 2023826-2023-02207 | 2022-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/001 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17195741&pc=QCB | 2023826-2023-02535 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED. THE LENS WAS REMOVED ON (B)(6) 2023 AND A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE REPORTER STATED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE REPORTER STATED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253107&pc=QCB | 2023826-2023-02623 | 2023-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283702&pc=QCB | 2023826-2023-02639 | 2023-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/095 (SPHERE / CYLINDER / AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17375164&pc=QCB | 2023826-2023-02976 | 2023-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.0/6.0/092 (SPHERE / CYLINDER / AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT AND PATIENT INJURY. ADDITIONAL SUTURE WAS PERFORMED. CAUSE OF THE EVENT WAS BOTH USER ERROR AND THE DEVICE. REPORTEDLY, "I ASSUME THIS IN THE BACK-UP (2ND ICL) WHICH ALSO WAS TRANSECTED 50:50 SIMILAR TO 1ST. WITH THIS 2ND ICL TRIPLE CHECKED TO ENSURE NO PLUNGER CAPTURE BEFORE TRYING TO INSERT." ON (B)(6) 2023, THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS, "COMFORTABLE AFTER FAILED INSERTION AND BETTER VISION. BETTER UNAIDED." Manufacturer Narrative: H6: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT FROM INITIAL LOT WAS FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17375205&pc=QCB | 2023826-2023-02975 | 2023-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.0/6.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT AND PATIENT INJURY. IRIS PROLAPSE WAS OBSERVED. ADDITIONAL SUCTURE WAS PERFORMED. CAUSE OF THE EVENT WAS BOTH USER ERROR AND THE DEVICE. REPORTEDLY, "ICL TRANSECTED INTO 2 HALVES. SAME HAPPENED TO BACK UP ICL + OBVIOUSLY DOUBLE & TRIPLE CHECKED NO PLUNGER CAPTURE OF ICL PRIOR TO INJECTION." ON (B)(6) 2023, THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS, "COMFORTABLE DAY! POST OP & HAPPY WITH BINOCULAR IMPROVEMENT." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT FROM INITIAL LOT WAS FOUND. CLAIM#(B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581- IRIS PROLAPSE" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED. B5- THE REPORTER INDICATED THAT THE SURGEON INTRAOPERATIVELY IMPLANTED AND REMOVED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIPPER -5.00/6.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT INJURY. THE PATIENT EXPERIENCED IRIS PROLAPSE. ADDITIONAL SUTURE WAS PERFORMED. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME MODEL AND LENGTH; SIMILAR POWER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS, "COMFORTABLE DAY! POST OP & HAPPY WITH BINOCULAR IMPROVEMENT." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE/USER ERROR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "ICL TRANSECTED INTO 2 HALVES. SAME HAPPENED TO BACK UP ICL + OBVIOUSLY DOUBLE, TRIPLED CHECK. NO PLUNGER CAPTURED ICL PERIOD TO INJECTION". ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17376516&pc=QCB | 2023826-2023-02960 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE PATIENT ALSO EXPERIENCED DECREASED VISION. H6 - CLINICAL CODE 4581, DECREASED VISION. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK . H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410569&pc=QCB | 2023826-2023-03166 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410576&pc=QCB | 2023826-2023-03171 | 2023-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434400&pc=QCB | 2023826-2023-03044 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). HIGH VAULT AND EYEBROW ACHES WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581- EYE BROW ACHE" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023 . THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND 'EYEBROW ACHE'. REPORTEDLY " I HAVE A PATIENT THAT HAS A HIGH VAULT OU AFTER TORIC ICLS 28JUN. SHE IS EXPERIENCING SOME EYEBROW ACHES ONLY OVER OS. " THE LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D4 - UNK. H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434411&pc=QCB | 2023826-2023-03187 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). HIGHT VAULT AND NARROWED ANTERIOR CHAMBERS WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK, A2 - UNK, A3 - UNK, A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY), D4 - UNK, D6A - UNK, D6B - UNK, H4 - UNK, (B)(4). Manufacturer Narrative: H6 HEALTH EFFECT CLINICAL CODE - NARROW ANTERIOR CHAMBERS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434412&pc=QCB | 2023826-2023-03188 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). HIGHT VAULT AND NARROWED ANTERIOR CHAMBERS WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE, NARROW ANTERIOR CHAMBERS. CLAIM#: (B)(4). Manufacturer Narrative: A1 - UNK, A2 - UNK, A3 - UNK, A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY), D4 - UNK, D6A - UNK, D6B - UNK, H4 - UNK. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17456036&pc=QCB | 2023826-2023-03224 | 2023-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -18.0/2.0/107 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED AND REMOVED FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS OBSERVED. LATER, A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "SIZE 13.7 WAS TO LONG. SIZE 13.2 IS O.K." STATUS OF THE EYE WAS STATED, "IT'S OK". Manufacturer Narrative: A4-A6: UNK. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE ADDED. D6A. "(B)(6) 2023" SHOULD BE REMOVED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+2.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH THE SAME MODEL, SIMILAR POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS NOTED TO BE "IT'S OK". ADDITIONAL INFORMATION PROVIDED "SIZE 13.7 WAS WRONG. SIZE 13.2 IT'S OK". CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17456037&pc=QCB | 2023826-2023-03225 | 2023-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -15.0/2.5/076 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED , AND REMOVED FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS OBSERVED. LATER, A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "SIZE 13.7 WAS TO LONG. SIZE 13.2 IS O.K." STATUS OF THE EYE WAS STATED , "IT'S OK." Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE ADDED. D6A. "(B)(6) 2023" SHOULD BE REMOVED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+2.5/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH THE SAME MODEL, SIMILAR POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS NOTED TO BE "IT'S OK". ADDITIONAL INFORMATION PROVIDED - "SIZE 13.7 WAS WRONG. SIZE 13.2 IT'S OK". CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495141&pc=QCB | 2023826-2023-03366 | 2022-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.0/063 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495144&pc=QCB | 2023826-2023-03367 | 2022-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.0/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561396&pc=QCB | 2023826-2023-03496 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/1.0/003 (SPHERE / CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17653872&pc=QCB | 2023826-2023-03439 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -16.00/1.5/167 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. REPORTER STATES. "PATIENT SCHEDULED FOR REPLACEMENT OF ICL." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - LENS WAS IMPLANTED ON (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: CORRECTION DATA: A2- DOB IS (B)(6) 1982. B5-THE REPORTER INDICATED THAT A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -16.00/1.5/167 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED WITH SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE LENS REMAINS IMPLANTED. REPORTER STATES. "PATIENT SCHEDULED FOR REPLACEMENT OF ICL." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6- CLINICLAL CODE 4581- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS EXCHANGED ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17656782&pc=QCB | 2023826-2023-03804 | 2023-08-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17722464&pc=QCB | 2023826-2023-03838 | 2023-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON QUESTIONS HIGH VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.5/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND EXCESSIVE VAULT. A SHORTER LENGTH REPLACEMENT LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE: LENS OVERSIZED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770322&pc=QCB | 2023826-2023-04116 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -10.00/2.0/120 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON AN UNKNOWN DATE. EXCESSIVE VAULT WAS OBSERVED. THE LENS WAS EXPLANTED AND A REPLACEMENT LENS WAS IMPLANTED OF SHORTER LENGTH. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777467&pc=QCB | 2023826-2023-04125 | 2023-09-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848827&pc=QCB | 2023826-2023-04244 | 2023-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "ALL GOOD, PATIENT IS HAPPY." Manufacturer Narrative: H6- MEDICAL IMPACT CODE: "2199" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. H6- MEDICAL IMPACT CODE: "4629" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848836&pc=QCB | 2023826-2023-04243 | 2023-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.0/0.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "ALL GOOD, PATIENT IS HAPPY." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL IMPACT CODE: "2199" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. H6: MEDICAL IMPACT CODE: "4629" SHOULD BE ADDED H6: HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17924961&pc=QCB | 2023826-2023-04523 | 2021-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/0.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B3 - DATE OF EVENT: 21-MAR-2021 SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17962098&pc=QCB | 2023826-2023-04552 | 2023-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD). THE LENS HAD EXCESSIVE VAULTING AND SHALLOW AC. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK A3: UNK A4: UNK A5: UNK A6: UNK B3: UNK D6A: UNK D6B: UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4) Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6)2 023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE AC HAD MORE STABLE DEPTH AFTER EXCHANGE, NORMAL IOP, VAULT AND AC DEPTH MUCH IMPROVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR. THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17978576&pc=QCB | 2023826-2023-04553 | 2023-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE LENS HAD EXCESSIVE VAULTING AND SHALLOW AC. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6A: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE AC HAD MORE STABLE DEPTH AFTER EXCHANGE, NORMAL IOP, VAULT AND AC DEPTH MUCH IMPROVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR. THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17981337&pc=QCB | 2023826-2023-04704 | 2023-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK CLAIM #(B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18022609&pc=QCB | 2023826-2023-04734 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -14.5/4.0/034 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT WAS DISSATISFIED WITH THE RESULTS. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PATIENT DISSATISFIED. CLAIM#: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18022610&pc=QCB | 2023826-2023-04733 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -13.0/3.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT WAS DISSATISFIED WITH THE RESULTS. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PATIENT DISSATISFIED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075206&pc=QCB | 2023826-2023-04921 | 2023-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -5.50/1.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. OVAL SHAPE PUPIL AT TEMPORAL WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: OVAL SHAPE PUPIL AT TEMPORAL. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/+1.0/78 (SPHERE/CYLINDER/AXIS)INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED PUPIL OVALISATION, PIGMENT DISPERSION AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODE, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY "CASE RESOLVED, VA AND VAULTING IS GOOD, IOP STABLE WITH 2 ANTIGLAUCOMA DROPS. DR. WILL MONITOR." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. H6: HEALTH EFFECT- CLINICAL CODE: "4581- PUPIL OVALISATION AND PIGMENT DISPERSION " SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "4629" AND "4644" SHOULD BE ADDED. "2199" SHOULD BE REMOVED. H6: MEDICAL DEVICE PROBLEM CODE: "3189" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175031&pc=QCB | 2023826-2023-05106 | 2023-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -4.0/5.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THIS DID NOT RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE: MISS THE CORRECT SIZING OF THE LENS. REPORTEDLY, REPLACEMENT LENS PLANNED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188487&pc=QCB | 2023826-2023-05286 | 2023-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.0/167 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188488&pc=QCB | 2023826-2023-05285 | 2023-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/155 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18348108&pc=QCB | 2023826-2023-05463 | 2023-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.5/123 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON 28-NOV-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369666&pc=QCB | 2023826-2023-05670 | 2023-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/5.5/047 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18406811&pc=QCB | 2023826-2023-05800 | 2023-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 11-DEC-2023, THE LENS WAS REMOVED DUE TO LOW VAULT WITHOUT ROTATION AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18406816&pc=QCB | 2023826-2023-05799 | 2023-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO LOW VAULT WITHOUT ROTATION AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334474&pc=QCB | 2023826-2021-00366 | 2021-01-20 05:00 | Injury | STAAR SURGICAL COMPANY AG | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/131, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2011. THE PATIENT EXPERIENCED A REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/131, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2011. THE PATIENT EXPERIENCED A REFRACTIVE CHANGE OVERTIME. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL/SIZE SIMILAR POWER AND THIS RESOLVED THE PROBLEM. REPORTER STATES VAULTING MUCH BETTER THEN PRE EXCHANGE. PATIENT IS HAPPY WITH UCVA 20/40. (PRE EXCHANGE 20/30). D6B- (B)(6) 2021. H6-CLINICAL CODE 4581: REFRACTIVE CHANGE OVER TIME. H6-IMPACT CODE 4629; LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8678984&pc=QCB | 2023826-2019-01025 |  | Injury | N/A | 2019-06-07 04:00 | QCB | N/A | Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF ADVANCES IN MEDICINE AND MEDICAL RESEARCH IN MARCH 2019 TITLED, 'TORIC AND PHAKIC IOLS FOR THE TREATMENT OF ASTIGMATISM AND/OR HIGH MYOPIA: OUR EXPERIENCE AT PRINCE HASHEM HOSPITAL ZARQA.' THE ARTICLE STATES THAT 3 PATIENTS WHO UNDERWENT TORIC INTRA-OCULAR LENS IMPLANTATION HAD LENS ROTATION A FEW DAYS AFTER SURGERY THAT RESULTED IN A DECREASE IN THEIR VISUAL ACUITY AND THAT WAS OVERCOME BY REALIGNMENT OF THEIR IOL INTRAOPERATIVELY. ALL PATIENTS HAD SATISFACTORY IMPROVEMENTS IN THEIR VISUAL ACUITIES. THEY WERE FOLLOWED FOR UP TO 3 MONTHS WITH NO CHANGE TO THEIR BCVA OR THEIR LENS STATUS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED TO: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF ADVANCES IN MEDICINE AND MEDICAL RESEARCH IN MARCH 2019 TITLED, 'TORIC AND PHAKIC IOLS FOR THE TREATMENT OF ASTIGMATISM AND OR HIGH MYOPIA: OUR EXPERIENCE AT PRINCE HASHEM HOSPITAL ZARQA.' THE ARTICLE STATES THAT 3 PATIENTS WHO UNDERWENT TORIC INTRA-OCULAR LENS IMPLANTATION HAD LENS ROTATION A FEW DAYS AFTER SURGERY THAT RESULTED IN A DECREASE IN IN THEIR VISUAL ACUITY AND THAT WAS OVERCOME BY REALIGNMENT OF THEIR IOL INTRAOPERATIVELY. IT WAS DETERMINED THAT THE IOLS USED WERE ACRI LISA TORIC (CARL ZEISS MEDITEC) INTRAOCULAR LENSES AND ARE THEREFORE NOT MANUFACTURED BY STAAR. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372525&pc=QCB | 2023826-2019-00315 | 2019-01-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0/+2.5/146 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS DEVICE CODE 1494: PREVIOUS CODE NOT APPLICABLE CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8498014&pc=QCB | 2023826-2019-00583 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-04-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS THE VAULT IS A "BIT HIGH" AND THAT THE PATIENT COMPLAINS OF BLURRED VISION IN SLIGHTLY DIM LIGHTING "LIKE A LAYER OF OIL COVERED ON YOUR GLASSES." LENS REMAINS IMPLANTED. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8500103&pc=QCB | 2023826-2019-00558 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-04-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON'S "FIRST TORIC PATIENT HAS AN OUTCOME THAT IS -2.00 +3.25 AT 128," (REFRACTIVE SURPRISE). REPORTER STATED "THE LENS POWER IS -10.50 +3.50 AT 92." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511547&pc=QCB | 2023826-2019-00600 | 2019-03-04 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE REPORTER INDICATED THERE WAS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511615&pc=QCB | 2023826-2019-00601 | 2019-03-04 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+4.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE REPORTER INDICATED THERE WAS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519002&pc=QCB | 2023826-2019-00613 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS AN UNEXPECTED RESIDUAL REFRACTION OF -1.25 CYL POSTOPERATIVELY. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: IMPLANT DATE: PREVIOUSLY MARKED UNK IN ERROR. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8524100&pc=QCB | 2023826-2019-00619 | 2018-09-05 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM,VTICM5_13.2, -11.0/5.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018 AND REFRACTIVE SURPRISE HAS BEEN OBSERVED. THE REPORTER STATED THAT THE LENS REMIANS IMPLANTED WITH POOR VA (VISUAL ACUITY). PATIENT BCVA 1.0 WITH OVER CORRECTION +0.75-1.25X63. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8571515&pc=QCB | 2023826-2019-00698 | 2019-03-18 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +2.0/+2.0/174 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611967&pc=QCB | 2023826-2019-00675 | 2019-04-03 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. THE DOCTOR STATES "THE PATIENT IS HAPPY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672597&pc=QCB | 2023826-2019-01033 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-06-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +4.00/+3.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD. THE REPORTER INDICATED THE PATIENT HAS RESIDUAL ASTIGMATISM AFTER THE ICL IMPLANTATION. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672629&pc=QCB | 2023826-2019-01034 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-06-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +6.00/+3.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED THE PATIENT HAS RESIDUAL ASTIGMATISM AFTER THE ICL IMPLANTATION. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692892&pc=QCB | 2023826-2019-01070 | 2017-10-19 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-12 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.0/+0.5/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8708611&pc=QCB | 2023826-2019-01090 | 2019-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1MM, -1.00/1.0/032 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE REPORTER STATED " THE PATIENT HAD A SMALL REFRACTIVE ERROR ON THE RIGHT EYE PRE-OP, BUT IT HAS BEEN WORSE AFTER ICL IMPLANT ". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED THAT THE PATIENT HAS NOT BEEN TO THE CLININC SINCE MAR 20TH, BCVA WAS REPORTED AS 1.2 , PATIENT WANTS A LENS EXCHANGE. PATIENT WAS TRAINED WITH A COVER AS A CHILD, RIGHT EYE DOMINANT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8708784&pc=QCB | 2023826-2019-01091 | 2019-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1MM, -2.00/2.0/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER STATED THAT THE PATIENT HAS NOT BEEN TO THE CLININC SINCE MAR 20TH, BCVA WAS REPORTED AS 1.2 , PATIENT WANTS A LENS EXCHANGE. PATIENT WAS TRAINED WITH A COVER AS A CHILD, RIGHT EYE DOMINANT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8773867&pc=QCB | 2023826-2019-01187 | 2019-03-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.5/1.0/144 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2019 AND REFRACTIVE SURPRISE WAS OBSERVED. PATIENT IS UNHAPPY WITH NEAR VISION, AWAITING ICL EXCHANGE FOR HYPEROPIC SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED FOR THIS CASE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANTED DATE IS (B)(6) 2019; THE EXACT DAY IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE FOR ACD <3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795444&pc=QCB | 2023826-2019-01242 | 2019-05-20 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8798875&pc=QCB | 2023826-2019-01267 | 2019-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -0.50/6.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8798976&pc=QCB | 2023826-2019-01266 | 2019-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -0.50/6.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813862&pc=QCB | 2023826-2019-01290 | 2019-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/020 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813865&pc=QCB | 2023826-2019-01289 | 2019-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/139 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852208&pc=QCB | 2023826-2019-01323 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-08-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/141 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852214&pc=QCB | 2023826-2019-01324 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-08-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/079 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871962&pc=QCB | 2023826-2019-01395 | 2019-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894318&pc=QCB | 2023826-2019-01433 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.0/111 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE. LOSS OF BCVA IS ALSO REPORTED, HOWEVER THE PRE-OP BCVA IS LISTED AS "?" AND POST-OP IS "20/25 GLASSES." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ONLY YEAR OF BIRTH REPORTED; EXACT DATE NOT KNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966498&pc=QCB | 2023826-2019-01648 | 2019-07-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -0.5/+6.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND THAT UCVA HAS WORSENED FROM THE 1ST WEEK OF SURGERY. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD CODE 4110: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9088369&pc=QCB | 2023826-2019-01670 | 2019-07-03 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -9.50/3.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9100732&pc=QCB | 2023826-2019-01733 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +1.5/+6.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER INDICATED THIS LENS REPLACED ANOTHER LENS (12.1MM) BUT THE PATIENT STILL HAD SOME ASTIGMATISM LEFT AND WOULD BE HAVING A LASER PROCEDURE TO RESOLVE THE ASTIGMATISM. SEE MFR. REPORT # 2023826-2019-01688 FOR 12.1MM LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136940&pc=QCB | 2023826-2019-01809 | 2019-08-30 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-30 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -3.5/+6.0/118 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED AND NO SURGICAL INTERVENTION WILL BE PERFORMED AS "PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9148352&pc=QCB | 2023826-2019-01853 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-10-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT WITH POST LASIK ECTASIA WHO HE CROSS LINKED AND IMPLANTED 13.7MM, VTICMO13.7, -2.00/4.0/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE RIGHT EYE (OD) HAD AN EXCELLENT OUTCOME, BUT OVER THE PAST 2 YEARS HIS CORNEAL SHAPE HAS CHANGED SLIGHTLY AND HAS BECOME MORE MYOPIC. PATIENT HAS NO CATARACT AND HIS CORNEAL ENDOTHELIAL CELL COUNT IS PERFECT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454032&pc=QCB | 2023826-2019-02344 | 2019-09-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -8.50/0.5/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. IT WAS REPORTED THAT THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1- "PRODUCT PROBLEM" SHOULD BE MARKED IN THE INITIAL MDR. B2- "REQUIRED INTERVENTION" IS NOT APPLICABLE IN THE INITIAL MDR. H6- WORK ORDER SEARCH SHOULD SAY: "NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9469760&pc=QCB | 2023826-2019-02383 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-12-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/065 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476270&pc=QCB | 2023826-2019-02391 | 2019-11-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.5/+3.5/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - POSTOPERATIVE REPORT INDICATED PATIENT'S IOP IS STABLE AND REPORTED BLURRED VISION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476271&pc=QCB | 2023826-2019-02390 | 2019-11-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.5/+4.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - POSTOPERATIVE REPORT STATED PATIENT'S IOP IS STABLE AND REPORTS BLURRED VISION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476283&pc=QCB | 2023826-2019-02389 | 2017-01-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS "HYPEROPIC SHIFT 2D." Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES THE "PATIENT HAS REFRACTIVE CHANGES 'SIGNIFICANT' OVER TIME AFTER 3 YEARS OF ICL IMPLANTATION." ALSO STATED, "PATIENT IS DOING FINE NOW WITH ACCOMMODATION AND BRAIN ADAPTATION." REFRACTIVE SURPRISE IS ALSO REPORTED. THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9491791&pc=QCB | 2023826-2019-02421 | 2020-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/3.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS IMPLANTED ON (B)(6) 2019. REFRACTIVE SURPRISE WAS ASSESSED ON (B)(6) 2020. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. THE LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516576&pc=QCB | 2023826-2019-02450 | 2019-12-02 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.00/6.0/156 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9536347&pc=QCB | 2023826-2019-02508 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-12-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.50/+6.0/025 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9588359&pc=QCB | 2023826-2020-00120 | 2019-11-04 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/5.00/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 REFRACTIVE SURPRISE WAS NOTED. LENS REMAINS IMPLANTED, PATIENTS CURRENT CONDITION IS GOOD, BCVA 20/20. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9614033&pc=QCB | 2023826-2020-00138 | 2019-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.5/+6.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REPORTED AS HAVING ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. AT POST-OP VISIT ON (B)(6) 2020, THE REFRACTION WAS STABLE. THE PATIENT WAS PRESCRIBED SPECTACLES TO CORRECT RESIDUAL ASTIGMATISM WHILE DRIVING. THE PATIENT IS SATISFIED WITH THE FINAL OUTCOME. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9614063&pc=QCB | 2023826-2020-00139 | 2019-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -1.5/+6.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. AT POST-OP VISIT ON (B)(6) 2020, THE REFRACTION WAS STABLE. THE PATIENT WAS PRESCRIBED SPECTACLES TO CORRECT RESIDUAL ASTIGMATISM WHILE DRIVING. THE PATIENT IS SATISFIED WITH THE FINAL OUTCOME. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689160&pc=QCB | 2023826-2020-00282 | 2018-02-09 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.00/2.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT "ENDED UP WITH SMALL MINUS ON ONE EYE AND SMALL PLUS ON OTHER EYE. WE WILL CHANGE RIGHT LENS AND HOPE THAT THE BALANCE OF THE EYES WILL RELIEVE SYMPTOMS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706694&pc=QCB | 2023826-2020-00306 | 2020-01-13 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -5.00/2.0/084 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. "VERY LOW VAULT AND REFRACTIVE SURPRISE" WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706704&pc=QCB | 2023826-2020-00305 | 2020-01-13 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -5.50/2.5/087 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. "VERY LOW VAULT AND REFRACTIVE SURPRISE" WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721915&pc=QCB | 2023826-2020-00308 | 2020-01-22 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-18 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+4.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721913&pc=QCB | 2023826-2020-00309 | 2020-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Appropriate Term/Code Not Available | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.5/+4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION:: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE . B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6; -11.5/4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME MODEL/SIZE LENS AND THE PROBLEM WAS RESOLVED. D7- EXPLANT DATE:(B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- IN PREVIOUS MDR RESULT CODE 114 IS CORRECTED TO INVESTIGATION FINDINGS 213. H6- IN PREVIOUS MDR INVESTIGATION CONCLUSION CODE 4310 NEEDED TO BE REPORTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9759568&pc=QCB | 2023826-2020-00325 | 2019-07-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.00/+6.0/180 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019. PRE-EXISTING: PREVIOUS CORNEAL OR REFRACTIVE SURGERY; KERATOCONUS; POST PK SURGERY. IMPLANT DATE: (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.00/+6.0/180 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6)2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME AND TRANSIENT LOSS OF BCVA. THE LENS REMAINS IMPLANTED. AT THE TIME OF IMPLANTATION WITH HISTORY OF KERATOCONUS AND POST PK SURGERY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9850132&pc=QCB | 2023826-2020-00571 | 2020-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6-DEVICE CODE 2993 IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME MODEL/SIZE AND DIFFERENT POWER LENS. ON (B)(6) 2020 THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. D6A: (B)(6) 2019. D6B: (B)(6) 2020. H1- SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6)2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME MODEL/SIZE LENS AND DIFFERENT POWER LENS. THE EXCHANGED RESOLVED THE PROBLEM. H6-IMPACT CODE 4629 IS NOT APPLICABLE; THE EXCHANGE LENS WAS NOT REPOSITONED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960153&pc=QCB | 2023826-2020-00853 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-04-14 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.5/+3.5/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND POSSIBLE LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960163&pc=QCB | 2023826-2020-00854 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-04-14 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.0/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND POSSIBLE LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025646&pc=QCB | 2023826-2020-01007 | 2020-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-05 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/+2.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025656&pc=QCB | 2023826-2020-01006 | 2020-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-05 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10130964&pc=QCB | 2023826-2020-01217 | 2020-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -1700/3.50/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED HIS OR IS (B)(6). HE HAS NOT RECEIVED ANY TREATMENT AND INTERVENTION AT THIS TIME. THE REFRACTIVE SURPRISE WAS FIRST NOTED ON (B)(6) 2020, 10 DAYS POST-OP. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - CORRECT LENS (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO -17.00/+3.50/89. B5 - DATE OF IMPLANT SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10131012&pc=QCB | 2023826-2020-01218 | 2020-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.00/3.00/77 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED HIS OR IS (B)(6). HE HAS NOT RECEIVED ANY TREATMENT AND INTERVENTION AT THIS TIME. THE REFRACTIVE SURPRISE WAS FIRST NOTED ON (B)(6) 2020, 10 DAYS POST-OP. CLAIM # (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10162715&pc=QCB | 2023826-2020-01294 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision | NA | 1 | Event Description: THE REPORTER REPORTED REFRACTIVE CHANGE OVERTIME IN THE LEFT EYE (OS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.00/0.50/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE, ASSESSED ON (B)(6) 2020 WAS REPORTED. LENS REMAINS IMPLANTED. THE REPORTER STATED " WE PLANNED A TARGET REFRACTION THE LEFT EYE FOR PROXIMITY, THE RESULT WAS EMMETROPIA. WE ASK FOR A LENS EXCHANGE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. D2: CORRECTED FROM PHAKIC INTRAOCULAR LENS,MTA TO PHAKIC TORIC INTRAOCULAR LENS,QCB. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- ADVERSE EVENT SHOULD BE CHECKED IN PREVIOUS MDR. B2- REQUIRED SURGICAL INTERVENTION SHOULD BE CHECKED IN PREVIOUS MDR. B5- "LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH LENS ON (B)(6) 2020, AND THE PROBLEM WAS RESOLVED." SHOULD BE IN THE PREVIOUS MDR. D6B- (B)(6) 2020, SHOULD BE IN THE PREVIOUS MDR. H1- "SERIOUS INJURY" SHOULD BE CHECKED IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE NAME: UNK, NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222177&pc=QCB | 2023826-2020-01420 | 2020-06-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-01 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -14.0/+5.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222170&pc=QCB | 2023826-2020-01418 | 2020-06-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.5/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO PREVIOUS MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254180&pc=QCB | 2023826-2020-01444 | 2020-03-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -12.5/+4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B1- CORRECTED TO PRODUCT PROBLEM-MALFUNCTION. H1- CORRECTED TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332955&pc=QCB | 2023826-2020-01568 |  | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -12.50/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION : B5- THE LENS WAS ROTATED ON (B)(6) 2020 8 DEGREES. IT IS REPORTED THAT THE PATIENT IS STILL NOT SATISFIED WITH RESULT. H1- SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: UNK, UNK, UNK DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10356745&pc=QCB | 2023826-2020-01608 | 2020-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00/+4.0/065 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED A 20 DEGREE MISALIGNMENT HAD LED TO THE SURPRISE REFRACTION. DOCTOR DECIDED TO RE-ALIGN LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403862&pc=QCB | 2023826-2020-01730 | 2020-07-15 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.0/+4.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483346&pc=QCB | 2023826-2020-01947 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: REPORTER INDICATED SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO PATIENTS EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.00 DIOPTER INTO THE PATIENT RIGHT EYE ON AN UNKNOWN DATE. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON AN UNKNOWN DATE FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON. D4- MODEL#: VICMO 13.2. SERIAL #: (B)(6). EXPIRATION DATE: 30-APR-2023 D6A- UNK. D6B-UNK. H1- SERIOUS INJURY. H6-IMPACT CODE 4629: LENS WAS EXCHANGED. H6-METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.00 DIOPTER INTO THE PATIENT RIGHT EYE ON (B)(6) 2020. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON D6A: ON (B)(6) 2020, D6B: ON (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY. VISUAL INSPECTION FOUND HAPTIC TORN WITH THREADS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. H6 -TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483351&pc=QCB | 2023826-2020-01948 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: REPORTER INDICATED SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION B1- ADVERSE EVENT B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.50 DIOPTER INTO THE PATIENT LEFT EYE ON AN UNKNOWN DATE. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON AN UNKNOWN DATE FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON. D4- MODEL#: VICMO 13.2. SERIAL #: (B)(6),. EXPIRATION DATE: 30-APR-2023 H1- SERIOUS INJURY H6-IMPACT CODE 4629: LENS WAS EXCHANGED. H6-METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.50 DIOPTER INTO THE PATIENT LEFT EYE ON (B)(6) 2020. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON D6A:10-AUG-2020. D6B: 31-AUG-2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL WIDTH VERIFICATION INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. H6 -TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS (B)(4): DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10508720&pc=QCB | 2023826-2020-02061 | 2020-07-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/2.00/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE RECEIVED BY MFR: THIS PRODUCT IS NOT MARKETED IN THE U.S.. HEALTH EFFECT IMPACT : LENS REMAINS IMPLANTED, NO SECONDARY INTERVENTION HAS BEEN PERFORMED. UNABLE TO ENTER FIELD. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10508814&pc=QCB | 2023826-2020-02063 | 2020-07-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/2.00/61 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10508931&pc=QCB | 2023826-2020-02065 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -12.50/4.5/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REPOSITIONED DUE TO REFRACTIVE SURPRISE. IT WAS REPORTED THAT THE PROBLEM RESOLVED, PATIENT IS HAPPY. USER ERROR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10524412&pc=QCB | 2023826-2020-02039 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.5/2.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE CHANGE OVER TIME WAS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER/AXIS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551357&pc=QCB | 2023826-2020-02122 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/+3.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS LENS ROTATION AND REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. THE SURGEON STATES THE LENS WAS REPOSITIONED 3 TIMES AND THE LENS ROTATED AT THE SAME ANGLE AND DIRECTIN AFTERWARD. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- LENS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL/LENGTH LENS. THIS RESOLVED THE PROBLEM. - SHOUDL BE IN THE PREVIOUS MDR. D6B- (B)(6) 2020 - SHOULD BE IN PREVIOUS MDR. H6- HEALTH IMPACT CODE 4629 - SHOULD BE IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10556954&pc=QCB | 2023826-2020-02154 | 2020-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/1.50/73 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED, NO FURTHER INTERVENTION HAS BEEN PERFORMED, VISION REMAINS THE SAME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. REPORTEDLY, THE "PATIENT VISION IS 6/6 AND IS HAPPY". NO INTERVENTION WAS DONE. H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10560526&pc=QCB | 2023826-2020-02112 | 2020-08-26 | Injury | STAAR SURGICAL COMPANY | 2020-09-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 12.6MM TMICL12.6, -10.50/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. BLURRED VISION WAS OBSERVED AND LENS WAS REPOSITIONED ON (B)(6) 2020, THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE, SPHERICAL LENS. THIS RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10600979&pc=QCB | 2023826-2020-02232 | 2020-06-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/1.5/021 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, SURGEON "NOT YET DECIDED WHETHER TO TRY ROTATION, REPLACEMENT, OR LASER CORRECTION." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: LENS ROTATION ALSO REPORTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10617676&pc=QCB | 2023826-2020-02269 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN TORIC IMPLANTABLE COLLAMER LENS (TICL) INTO THE PATIENT'S EYE. THE SURGEON REPORTS A RESIDUAL CYLINDER OF 2D. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635460&pc=QCB | 2023826-2020-02247 | 2020-09-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -17.50/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER REPORTER ON (B)(6) 2020, SURGEON "HAS GIVEN PATIENT CONTACT LENSES TO CORRECT THE TINY RESIDUAL. PATIENT CONTINUES TO BE UNDER OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H10 "(B)(4)" CORRECTED TO "CLAIM#: (B)(4)" IN INITAL MDR. Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT 13.2MM VTICMO_13.2 -17.50/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER REPORTER ON (B)(6) 2020, THE SURGEON "HAS GIVEN PATIENT CONTACT LENSES TO CORRECT THE TINY RESIDUAL." PATIENT CONTINUES TO BE UNDER OBSERVATION. THE PATIENT IS UNHAPPY WITH VISION BUT REPORTEDLY THE PATIENT'S VA IS GOOD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED." IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635463&pc=QCB | 2023826-2020-02246 | 2020-09-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -17.50/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER REPORTER ON (B)(6) 2020, SURGEON "HAS GIVEN PATIENT CONTACT LENSES TO CORRECT THE TINY RESIDUAL. "PATIENT CONTINUES TO BE UNDER OBSERVATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT 13.2MM VTICMO_13.2 -17.50/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER REPORTER ON (B)(6) 2020, THE SURGEON "HAS GIVEN PATIENT CONTACT LENSES TO CORRECT THE TINY RESIDUAL." PATIENT CONTINUES TO BE UNDER OBSERVATION. THE PATIENT IS UNHAPPY WITH VISION BUT REPORTEDLY THE PATIENT'S VA IS GOOD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED." IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H10 - "CLAIM# (B)(4)" CORRECTED TO "CLAIM#: (B)(4)" IN INITAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659329&pc=QCB | 2023826-2020-02317 | 2020-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+3.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6: METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION IS UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659337&pc=QCB | 2023826-2020-02313 | 2020-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Patient-Device Incompatibility | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -11.5/+3.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659339&pc=QCB | 2023826-2020-02310 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/+2.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS RESIDUAL ASTIGMATISM, AND BLURRY VISION. THE LENS WAS REPOSITIONED, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.5/+2.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTED RESIDUAL ASTIGMATISM AND BLURRY VISION. THE LENS WAS REPOSITIONED ON (B)(6) 2020, AND THIS RESOLVED THE PROBLEM. LENS REMAINS IMPLANTED. SECONDARY SURGICAL INTERVENTION OCCURRED ON (B)(6) 2023, FOR LASIK ENHANCEMENT FOR RESIDUAL ASTIGMATISM AFFECTING VISUAL OUTCOME. REPORTEDLY THIS WAS NOT RELATED TO THE LENS OR PROCEDURE. THIS RESOLVED THE PROBLEM. H6: 4625-LASIK ENHANCEMENT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678723&pc=QCB | 2023826-2020-02357 | 2020-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -7.00/2.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. BLURRED VISION AND POSSIBLE REFRACTIVE SURPRISE WAS OBSERVED. REPORTEDLY, "A WHITE FIBROUS MEMBRANE ADHERENT TO POSTERIOR ICL." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 OF -7.0/2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS A WHITE FIBROUS MEMBRANE ADHERENT TO POSTERIOR ICL. THE PATIENT EXPERIENCED BLURRED VISION AND DUREZOL DROPS WERE PRESCRIBED. THE SURGEON TRIED TO REMOVE MEMBRANE WITH LOW POWER LASER ENERGY ANTERIOR TO ICL WITH NO SUCCESS. THE LENS WAS EXPLANTED ON (B)(6) 2020. D6A- (B)(6) 2020. H3- DEVICE EVALUATION: LENS RETURNED IN A SPECIMEN CUP WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. H6- HEALTH EFFECT: 4624- LOW POWER LASER ENERGY ANTERIOR TO ICL. CORRECTED DATA: G4- IN PREVIOUS MDR SHOULD BE CORRECTED TO 14-OCT-2020. H6-TYPE OF INVESTIGATION CODE 4110: UPDATED ERIAL NUMBER LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D4: REPORTED AS (B)(6). UPDATED / CORRECTED TO (B)(6). D9: REPORTED AS NO. UPDATED / CORRECTED TO YES, 04/NOV/2020. H4: REPORTED AS 11/OCT/2018. UPDATED / CORRECTED TO 29/OCT/2018. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687742&pc=QCB | 2023826-2020-02382 | 2020-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.5/+6.0/175 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXPLANTED WITH A PLANNED REPLACEMENT IMPLANTATION IN 2 MONTHS. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTION INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749959&pc=QCB | 2023826-2020-02527 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/+5.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT , ETHNICITY, RACE-UNK. DATE OF EVENT-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749979&pc=QCB | 2023826-2020-02528 | 2020-09-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.5/+4.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 -8.5/+4.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY, THE LENS WAS REPOSITIONED. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098630&pc=QCB | 2023826-2020-03245 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+3.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE HISTPRY RECORD: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK WORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BENN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 TYPE OF INVESTIGATION- "4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT." SHOULD BE IN THE PREVIOUS MDR. H6- TYPE OF INVESTIGATION: 3331 SHOULD BE IN THE PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065615&pc=QCB | 2023826-2020-03110 | 2020-11-26 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.50/6.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024114&pc=QCB | 2023826-2020-03041 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -6.5/+2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS RESIDUAL ERROR AND LENS ROTATION. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - "-6.5/+2.5/095" FOR IMPLANTED LENS SPHERE/CYLINDER/AXIS SHOULD BE CORRECTED TO "-8.50/2.5/095" IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024133&pc=QCB | 2023826-2020-03040 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -6.0/+2.5/065 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS RESIDUAL ERROR AND LENS ROTATION. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: BIRTH YEAR: (B)(6). DATE OF EVENT- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024129&pc=QCB | 2023826-2020-03049 | 2020-06-29 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.5/+3.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS POSTOP UCVA OF 20/50. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SPHERE, AND AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR, THE WRONG POWER LENS WAS CHOSEN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010445&pc=QCB | 2023826-2020-02979 | 2020-11-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -5.50/0.5/030 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED, THE REPORTER STATED "THE DOCTOR DECIDED TO WAIT 3 MONTHS FOR THE PATIENT TO GET USED TO THE CONDITION. THEN A DECISION IS MADE ABOUT FURTHER MEASURES. THE PATIENT IS DOING WELL SO FAR." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -5.50/0.5/030 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED, THE REPORTER STATED "THE DOCTOR DECIDED TO WAIT 3 MONTHS FOR THE PATIENT TO GET USED TO THE CONDITION. THEN A DECISION IS MADE ABOUT FURTHER MEASURES. THE PATIENT IS DOING WELL SO FAR." PER LATEST UPDATED MEDICAL OR SURGICAL INTERVENTION WAS NOT NECESSARY. IT WAS REPORTED THAT THE PROBLEM RESOLVED. FOR PATIENT CONDITION (LAST VISIT) BCVA: 20/22, VAULT 740 AND IOP 12MMHG WAS REPORTED. THE REPORTER STATED " A BETTER RESULT IS NOT POSSIBLE DUE TO THE ORIGINAL CONDITIONS." CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012259&pc=QCB | 2023826-2020-03019 | 2020-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -7.50/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2020, "PATIENT IN ACCIDENT. EYE SURGERY IS NOT POSSIBLE FOR THE NEXT TWO MONTHS." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -7.50/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2020, "PATIENT IN ACCIDENT. EYE SURGERY IS NOT POSSIBLE FOR THE NEXT TWO MONTHS. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2021 WHICH RESOLVED THE PROBLEM. REPORTEDLY, THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012186&pc=QCB | 2023826-2020-03008 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/2.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984844&pc=QCB | 2023826-2020-02879 | 2018-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/1.0/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LOW VAULT WAS OBSERVED AND REFRACTIVE CHANGE OVERTIME WAS ASSESSED ON (B)(6) 2018 AND (B)(6) 2020. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/1.0/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LOW VAULT WAS OBSERVED AND REFRACTIVE CHANGE OVERTIME WAS ASSESSED ON (B)(6) 2018 AND (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6- HEALTH EFFECT IMPACT CODE REPORTED AS 2199 - UPDATE TO 4629. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980485&pc=QCB | 2023826-2020-02934 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/020 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME LENGTH AND MODEL, DIFFERENT SPHERE AND AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITIH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980505&pc=QCB | 2023826-2020-02920 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -1.0/+2.5/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WASEXPLANTED AND REPLACED WITH A LONGER LENGTH LENS DUE TO BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: B5-CLAIM IS FOR LEFT (OS) AND NOT RIGHT (OD) EYE AS PREVIOUSLY REPORTED. ADDITIONAL SYMPTOMS: LENS ROTATION, VISION LOSS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INPSECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: H1- IN PREVIOUS MDR SHOULD BE SERIOUS INJURY. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980516&pc=QCB | 2023826-2020-02910 | 2020-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.0/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS THE LENS WAS IMPLANTED UPSIDE DOWN, LEADING TO REFRACTIVE SURPRISE. REPORTEDLY, THE LENS WAS EXCHANGED. Manufacturer Narrative: B5 - "REPORTEDLY OD EYE WAS DONE." REPLACES "REPORTEDLY, THE LENS WAS EXCHANGED." H10 - 4629 CORRECTED TO SURGICAL CODE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980478&pc=QCB | 2023826-2020-02911 | 2020-10-26 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS THE LENS WAS IMPLANTED UPSIDE DOWN, LEADING TO REFRACTIVE SURPRISE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10942061&pc=QCB | 2023826-2020-02845 | 2020-11-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.0/+5.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT REFRACTIVE TREATMENT WAS PERFORMED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - "REFRACTIVE TREATMENT PERFORMED" SHOULD BE REMOVED AS IT IS NOT APPLICABLE. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED. H6 - HEALTH EFFECT IMPACT CODE: CORRECTED TO 2199, AS 4625 IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10942069&pc=QCB | 2023826-2020-02846 | 2020-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -14.0/+2.00/119. (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS TILTED WAS LISTED AS CAUSE OF EVENT, THE REPORTER STATED " OBSERVED DURING SURGERY AND THIS TILT CAN ALSO EXPLAIN THE REFRACTIVE SURPRISE. TILTED CERTAINLY BY BIG VAULTING." IT WAS REPORTED THAT THE PATIENT WEARS GLASSES AND THAT NO ADDITIONAL INTERVENTION IS BEING PERFORMED FOR THE MOMENT. "PUPIL CLOSED AND NOT VERY REACTIVE" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS), NOT PRODUCT RELATED "ALWAYS NOT VERY REACTIVE". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ETHNICITY & RACE: UNK. DATE REC'D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10942215&pc=QCB | 2023826-2020-02847 | 2020-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -15.50/+3.0/060. (SPHERE/CYLINDER/AXIS),IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS TILTED WAS LISTED AS CAUSE OF EVENT, THE REPORTER STATED " OBSERVED DURING SURGERY AND THIS TILT CAN ALSO EXPLAIN THE REFRACTIVE SURPRISE. TILTED CERTAINLY BY BIG VAULTING." IT WAS REPORTED THAT THE PATIENT WEARS GLASSES AND THAT NO ADDITIONAL INTERVENTION IS BEING PERFORMED FOR THE MOMENT. "PUPIL CLOSED AND NOT VERY REACTIVE" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS), NOT PRODUCT RELATED "ALWAYS NOT VERY REACTIVE". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ETHNICITY & RACE: UNK. DATE REC'D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937922&pc=QCB | 2023826-2020-02787 | 2020-09-30 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -5.50/2.5/161 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE, GLARE/HALOES, AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR (USER TRANSPOSE ERROR, PLUS TO MINUS CYLINDER RESULTED IN THE INCORRECT DATA ENTRY.). REPORTEDLY, PATIENT SATISFIED WITH VISION. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. IMPLANT DATE - UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930495&pc=QCB | 2023826-2020-02836 | 2020-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, SIMILAR SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10899949&pc=QCB | 2023826-2020-02740 | 2020-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED POOR VISUAL ACUITY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. CORRECTED DATA: B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10899956&pc=QCB | 2023826-2020-02741 | 2020-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -09.50/+4.0/171 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED POOR VISUAL ACUITY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. CORRECTED DATA: B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905489&pc=QCB | 2023826-2020-02738 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT REFRACTIVE SURPRISE HAS BEEN OBSERVED IN PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.00/0.5/122 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10: "EXCESSIVE VAULT" SHOULD BE CORRECTED TO "LOW VAULT" IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNKNOWN. SEX: UNKNOWN. WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE OF EVENT: UNKNOWN. COMMON DEVICE NAME-PRODUCT CODE: UNKNOWN. MODEL #: UNKNOWN. IMPLANT DATE: UNKNOWN. PMA/510K-THIS PRODUCT IS NOT MARKETED IN THE US, DEVICE MANUFACTURE DATE: UNKNOWN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905491&pc=QCB | 2023826-2020-02726 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/2.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905564&pc=QCB | 2023826-2020-02725 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/1.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: D6A - CORRECTED TO 07-MAY-2020 IN INITIAL MDR. G4 - COMBINATION PRODUCTION IS N/A IN INITIAL MDR. CLAIM#: (B)(4) . Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851802&pc=QCB | 2023826-2020-02723 | 2020-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.50/4.0/001 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS EXCHANGED FOR SAME LENGTH LENS ON (B)(6) 2020 RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, "SURGEON FORGOT TO ENTER THE NEGATIVE SIGN BEFORE ASTIGMATISM DIOPTER WHEN ENTERING DATA." REPORTEDLY, "THE CORNEA IS CLEAR, THE ANTERIOR CHAMBER IS DEEP, AND THE LENS IS CENTERED." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852506&pc=QCB | 2023826-2020-02673 | 2020-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856344&pc=QCB | 2023826-2020-02700 | 2020-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6)2020, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS OUT OF SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851799&pc=QCB | 2023826-2020-02664 | 2020-10-13 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE OVERTIME WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2020, SURGEON "HAS DECIDED TO MONITOR AND ON HOLD FOR THE LENS EXCHANGE AT THE MOMENT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851807&pc=QCB | 2023826-2020-02724 | 2020-10-21 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7 -16.00/3.5/034 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, "CONSIDERING ROTATION IF NEXT EXAM REVEALS ROTATION OFF AXIS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823677&pc=QCB | 2023826-2020-02598 | 2020-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/3.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- BS THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 OF -7.00/3.0/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE DID NOT SOLVE THE PROBLEM. D6B. (B)(6) 2021. H6-IMPACT CODE 4629; LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. SEE MFR# 2023826-2021-01504 (FOR EXCHANGE LENS). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10824095&pc=QCB | 2023826-2020-02661 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.5/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS OUT OF SPECIFICATIONS. CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10795929&pc=QCB | 2023826-2020-02572 | 2020-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+6.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, LOSS OF BCVA, BLURRED VISION AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN (WRONG SIZE). Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED. CORRECTED DATA: D10: INJECTOR MODEL MSI-PF. CLAIM # (B)(4). Manufacturer Narrative: B5: A LONGER LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE PATIENT HAS NO PROBLEMS WITH THE NEW LENS AND IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779886&pc=QCB | 2023826-2020-02544 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A TORIC ICL INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS LENS ROTATION AND RESIDUAL ASTIGMATISM. THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, LENS REMAINS IMPLANTED.Event Description: B5-THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2020 WHICH DID NOT RESOLVE THE PROBLEM. THE CAUSE IS REPORTED AS DEVICE, "IN LEFT EYE RESIDUAL ASTIGMATISM +1.0/-1.75/40 AFTER ICL IMPLANTATION." Manufacturer Narrative: B5-PLEASE DISREGARD PREVIOUS B5 STATEMENT. H6-DEVICE CODE 2923 REPORTED IN ERROR. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2020 WHICH DID NOT RESOLVE THE PROBLEM. THE CAUSE IS REPORTED AS DEVICE, "IN LEFT EYE RESIDUAL ASTIGMATISM +1.0/-1.75/40 AFTER ICL IMPLANTATION." THE ICL WAS ROTATED 13 DEGREES WITHOUT COMPLICATIONS ON (B)(6) 2021. REPORTEDLY, "POSTO RESULT IS UCVA LEFT EYE 0,6 AND 1,5 (+1,0-2,0X125) WHICH HAS BEEN STABLE SINCE REOP." THE DOCTOR IS PLANNING AN ICL EXCHANGE AFTER THE PATIENT'S FOLLOW UP VISIT ON (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: B5 - PER EMAIL ON 03/SEP/2021, "THE LENS STAYED ALIGNED ALL THE TIME. THE SIZING IS NOT THE PROBLEM. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2021." CLAIM#: (B)(4). Manufacturer Narrative: B5- "THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "AFTER 2 ATTEMPTS TO REPOSITION THE LENS, THE CLINIC ORDERED A NEW REPLACEMENT LENS, WITH THE SAME SIZE BUT WITH A 0.5D IN DIFFERENCE. THE RESULT OF THIS WAS THAT THE PATIENT IS SEEING 1.2 AND ARE REALLY SATISFIED. SO END RESULT WAS GREAT. BUT THIS REPLACEMENT LENS "STRENGTHEN" THEIR CLAIM IN THAT THE FIRST LENS WAS LABELED WITH THE INCORRECT INFORMATION SINCE IT DID NOT ROTATE OR ANYTHING LIKE THAT." SHOULD BE ADDED TO THE INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON THE LENS IN A VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11122217&pc=QCB | 2023826-2021-00072 | 2020-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129308&pc=QCB | 2023826-2021-00078 | 2020-11-19 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -3.0/1.5/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT ,REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED. ON (B)(6) 2020 LENS REPOSITIONING WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159195&pc=QCB | 2023826-2021-00128 | 2020-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.5/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR: THE WRONG CALCULATED LENS WAS IMPLANTED. Manufacturer Narrative: B1-UPDATED. B5-UPDATED INFORMATION: ON (B)(6) 2021, THE LENS WAS ECHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/CYLINDER/AXIS LENS AND PATIENT IS HAPPY. H6-HEALTH IMPACT CODE UPDATED . CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159208&pc=QCB | 2023826-2021-00121 | 2020-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+4.0/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+4.0/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS; DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159253&pc=QCB | 2023826-2021-00123 | 2020-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+6.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS RESPOSITIONED DUE TO REFRACTIVE SURPRISE WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED; REFERENCE MFR REPORT# 2023826-2021-00124 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CORRECTED DATA: B5 - "ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED." SHOULD BE CORRECTED TO "ON (B)(6) 2020 THE LENS WAS EXPLANTED AND LATER THE SAME DATE WAS REPLACED". H6 - HEALTH EFFECT IMPACT CODE: 4627. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159266&pc=QCB | 2023826-2021-00124 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+6.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G6/H10 - PREVIOUS SUPPLEMENTAL #1 IS TO BE DISREGARDED. CONTENTS OF THAT REPORT ARE NOT ASSOCIATED WITH THIS CLAIM. H1 - CHANGED TO MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION (B)(4)-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11161084&pc=QCB | 2023826-2021-00130 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.50/1.0/090(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN DEBRIS ON THE LENS. ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS DIOPTER -5.50/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS LATER EXCHANGED ON (B)(6) 2021 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11161342&pc=QCB | 2023826-2021-00131 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.00/1.0/090(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN DEBRIS ON THE LENS. ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS DIOPTER -5.0/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS LATER EXCHANGED ON (B)(6) 2021 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11168375&pc=QCB | 2023826-2021-00142 | 2017-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+3.0/152 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL DEVICE PROBLEM CODE 1494 THIS CODE WAS USED IN ERROR. THE PATIENTS AGE WAS WITHIN LABELED PARAMETERS. CLAIM (B)(4). Manufacturer Narrative: B5: A POSSIBLE CAUSE OF THE EVENT WOULD BE ICL ROTATION CAUSED THE REFRACTIVE CHANGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11169229&pc=QCB | 2023826-2021-00070 | 2020-03-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S EYE. ACCOMMODATIVE SPASMS AND FOGGY VISION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM, TICML12.1, 13.00/3.5/105, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2020. ACCOMMODATIVE SPASMS, REFRACTIVE CHANGE OVERTIME, AND FOGGY VISION WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "AT (B)(6) 2020 VISIT, PATIENT VOICED COMPLAINT OF BLURRY VISION ONLY AFTER WORK WHICH WOULD PERSIST FOR MINUTES TO HOURS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11196998&pc=QCB | 2023826-2021-00180 | 2020-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11197000&pc=QCB | 2023826-2021-00179 | 2020-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+2.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231160&pc=QCB | 2023826-2021-00208 | 2020-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -05.50/+4.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VAIL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238059&pc=QCB | 2023826-2021-00174 | 2020-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -13.00/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -13.00/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238071&pc=QCB | 2023826-2021-00166 | 2020-03-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.00/2.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.00/2.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2021, "PATIENT WILL GET LASIK ENHANCEMENT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC'D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11239144&pc=QCB | 2023826-2021-00215 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -13.5/+3.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES THAT THE PATIENT ISN'T COMING IN FOR A FOLLOW-UP. THE PATIENT IS HAPPY WITH THEIR VISION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11239149&pc=QCB | 2023826-2021-00214 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.5/+3.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: ACCORDING TO THE REPORTER, THE PATIENT ISN'T COMING IN FOR A FOLLOW-UP. THE PATIENT IS HAPPY WITH THEIR VISION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273779&pc=QCB | 2023826-2021-00228 | 2020-11-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.50/1.5/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, ROTATION SUSPECT. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11277279&pc=QCB | 2023826-2021-00298 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/4.0/73, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. " WOUND STABLE, ICL ON AXIS, EYE COMFORTABLE PER PT. AC CLEAR" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/4.0/73, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE CHANGE OVERTIME. THIS EXCHANGE RESOLVED THE PROBLEM. " WOUND STABLE, ICL ON AXIS, EYE COMFORTABLE PER PT., AC CLEAR" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). CAUSE OF EVENT IS UNKNOWN. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11280319&pc=QCB | 2023826-2021-00252 | 2021-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHTT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING, WITH LENS ROTATION AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DO TO THE LENS ROTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION NOT REQUIRED. B5 STATEMENT CORRECTED TO: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS; -14.00/+1.0/107 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTER HAVING LOW VAULTING WITH LENS ROTATION AND BLURRED VISION. ON (B)(6) 2020, THE LENS WAS REPOSITIONED; AND THE ISSUE WAS NOT RESOLVED. ON (B)(6) 2021; THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS STATED TO BE UNKNOWN: ICL 80 DEGREE ROTATION FROM HORIZONTAL TO VERTICAL. CLAIM: (B)(4). Manufacturer Narrative: D9: LENS RETURNED (B)(6) 2021. B5: LENS EXCHANGED ON (B)(6) 2021 FOR A LONGER LENTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299007&pc=QCB | 2023826-2021-00349 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR: THE SURGEON SWAPPED THE LENSES FOR THE LEFT AND RIGHT EYE AND IMPLANTED THE INCORRECT LENS. THE LENS WAS CHANGED FOR THE OD EYE, WHICH WAS ASTIGMATIC AND HAD POOR QUALITY OF VISION. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11310525&pc=QCB | 2023826-2021-00253 | 2020-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/176 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT COMPLAINED OF BLURRED VISION, THE PATIENT IS AMBLYOPIA AND FELT UNSATISFACTION ABOUT FAR. THE PROBLEM WAS RESOLVED. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11342672&pc=QCB | 2023826-2021-00436 | 2020-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -7.0/1.0/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THIS EXPLANT RESOLVED THE PROBLEM, THE REPORTER STATED " THE PATIENT HAS RECOVERED HIS PREOPERATIVE VISION AND NO DECISION HAS BEEN MADE TO OPERATE AGAIN." CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION AND FUNCTIONAL VERIFICATION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11352786&pc=QCB | 2023826-2021-00493 | 2019-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT, AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11358972&pc=QCB | 2023826-2021-00498 | 2021-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.00/5.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366596&pc=QCB | 2023826-2021-00486 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11.0/+1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTSBLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366580&pc=QCB | 2023826-2021-00482 | 2021-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372200&pc=QCB | 2023826-2021-00441 | 2020-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.00/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. "USED WRONG REFRACTION FOR THE CALCULATION." .REPORTEDLY, EXCHANGE IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400355&pc=QCB | 2023826-2021-00526 | 2020-12-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11.5/+1.0/018 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400354&pc=QCB | 2023826-2021-00527 | 2020-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/115 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11423056&pc=QCB | 2023826-2021-00573 | 2020-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/2.0/031 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. "NORMAL" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM,VTICM5_13.2, -14.00/2.0/031 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN LENS VIAL. VISUAL INSPECTION FOUND THE HAPTICS BENT/DEFORMED AND FOREIGN MATERIAL (FIBERS) ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453117&pc=QCB | 2023826-2021-00601 | 2020-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -13.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "NO REPOSITION: THE TICL ROTATES UNARTIFICIALLY COUNTERCLOCKWISE 92°AFTER THE SURGERY. THE PATIENT COMPLAINED OF BLURRED VISION AND THE AUTOMATIC REFRACTOR RESULTS SHOWS A CHANGE IN THE ASTIGMATISM AXIS.(THE POSITION ARE STABLE DURING FOLLOW-UP.) Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-RESULT CODE 213 SHOULD HAVE BEEN REPORTED IN INITIAL MDR. RESULT CODE 114 SHOULD HAVE BEEN REPORTED IN SUPPLEMENTAL MDR #1. CONCLUSION CODE 4315 SHOULD HAVE BEEN REPORTED IN INITIAL MDR CLAIM# (B)(4) Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US.(B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11467987&pc=QCB | 2023826-2021-00632 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AS A REPLACEMENT LENS DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. "BLURRINESS" WAS REPORTER FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: PER UPDATED INFORMATION THE PUPIL IS REACTING, IOP 24MMHG. BLURINESS HAS NOT RESOLVED. PATIENT DOES HAVE A LEFT OVER RX OF : -0.25 -0.75 X 023. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND PROBLEM RESOLVED. FINAL POST-OP ON (B)(6) 2021 REPORTS: UCVA 20/20 AND IOP 18MMHG AND FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTED STATED " EXCELLENT OUTCOME. PUPIL ROUND AND REACTIVE WITH NORMAL IOP." CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507463&pc=QCB | 2023826-2021-00677 | 2021-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS OUT OF SPECIFICATIONS. ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507471&pc=QCB | 2023826-2021-00675 | 2021-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: INITIALLY THE PATIENT'S OBJECTIVE REFRACTION VALUE WAS DIFFERENT FROM THE SUBJECTIVE REFRACTION VALUE. THE DOCTOR PERFORMED MANY TESTS, BUT THE SUBJECTIVE REFRACTION VALUE SHOWED THE SAME VALUE. THEREFORE, THE DOCTOR DECIDED THE LENS POWER WITH AN EMPHASIS ON THE SUBJECTIVE REFRACTION VALUE. HOWEVER, THE CYLINDRICAL POWER WAS INSUFFICIENT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507474&pc=QCB | 2023826-2021-00662 | 2021-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+6.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5- THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+6.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. ON (B)(6) 2021 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A SAME MODEL AND SIZE BUT DIFFERENT POWER REPLACEMENT LENS WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. D6B-(B)(6) 2021. H1- SERIOUS INJURY. H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/ DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507487&pc=QCB | 2023826-2021-00676 | 2021-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT SPHERE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515292&pc=QCB | 2023826-2021-00586 | 2020-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/1.0/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED (B)(6) 2021, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC SPLIT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/1.0/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED (B)(6) 2021. THE LENS WAS REPLACED WITH A SAME LENGTH, HIGHT POWER (SPHERE/CYLINDER/AXIS) LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "PATIENT HEALED OFF TARGET." REPORTEDLY, "HEALING WELL, APPEARS UNTARGET AT INTIAL POST OP." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558783&pc=QCB | 2023826-2021-00723 | 2021-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE 2 MONTHS POSTOPERATIVELY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1-UPDATED. B5-UPDATED: ON (B)(6) 2021 THE LENS WS REPOSITIONED. THERE IS RESIDUAL REFRACTION AND THE LENS "APPEARS TO BE SITTING OFF-AXIS A BIT." H6-HEALTH IMPACT CODE UPDATED. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+0.5/173 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED ON (B)(6) 2021. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE-UNK. DATE OF EVENT-UNK. DATE REC¿D BY MFR-UNK. DATE REC¿D BY MFR-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE -UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563257&pc=QCB | 2023826-2021-00763 | 2020-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -04.50/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH, LENS WITH A -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) POWER ON (B)(6) 2020 TO RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 -04.50/1.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SUPRISE REPORTED. THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) POWER ON (B)(6) 2020 DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563279&pc=QCB | 2023826-2021-00753 | 2021-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, -8.00/2.0/054 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER (10.00/2.0/050(SPHERE/CYLINDER/AXIS)) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, "STILL LETTLE EDEMA OF THE CORNEA, IMPROVING.". PER THE REPORTER ON (B)(6) 2021, "ACTUALLY THE PATIENT CAN SEE 20/20 WITHOUT CORRECTION AND VAULT IS CORRECT. HE PREVIOUSLY HAD A LITTLE BIT OF A HIPEROPIA AND A VERY UNSATISFACTORY CLOSE VISION. MUCH BETTER NOW." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND WITH SPOTS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTION: DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11606821&pc=QCB | 2023826-2021-00834 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: IT WAS REPORTED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REFRACTIVE SURPRISE WAS NOTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOTS. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620609&pc=QCB | 2023826-2021-00880 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -3.0/+4.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SUPRISE, RESIDUAL ASTIGMATISM, AND BLURRED VISION WAS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS DEVICE: PREOPERATIVE AND POSTOPERATIVE COURSE WAS SMOOTH WITHOUT ANY COMPLICATION. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: "HIGHER TOTAL ASTIGMATISM MORE DUE TO LENTAL ASTIGMATISM; VOD=0.05CC+5.50X90=0.1." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11632409&pc=QCB | 2023826-2021-00940 | 2021-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN 2012. ZERO VAULTING AND REFRACTIVE CHANGE OVERTIME WAS REPORTED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THA A 12.6MM, VTICM12.6, -5.5/6.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. LOW VAULT WAS OBSERVED, PATIENT IS FINE AND THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654115&pc=QCB | 2023826-2021-00960 | 2021-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -13.5/+1.5/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WAS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11676163&pc=QCB | 2023826-2021-01056 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/4.0/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS IMPLANTED AS A REPLACEMENT LENS TO CORRECT REFRACTIVE SURPRISE AND LOW VAULT WITH ROTATION BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2021 THE LENS WAS REPOSITIONED, IT WAS REPORTED THAT AFTER REPOSITIONING THE LENS UCVA WAS 20/25. THE REPORTER STATED " THE CURRENT VISUAL OUTCOME IS GOOD AND THE REFRACTIVE SURPRISE PROBLEM IS RESOLVED WITH REPOSITIONING AS OF NOW. PROBLEM WAS RESOLVED AND HE WILL CONTINUE TO MONITOR". Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11676122&pc=QCB | 2023826-2021-01055 | 2020-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.00/4.0/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE AND LOW VAULT WITH ROTATION WAS OBSERVED. LENS REPOSITIOING WAS PERFORMED BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME DIOPTER, LONGER LENGTH LENS BUT IT WAS REPORTED THAT THE PROBLEM DID NOT RESOLVE. THE REPORTER STATED " THE SURGEON DECIDED THAT THE VISUAL ACUITY WAS NOT GOOD DUE TO THE ROTATION AT THAT TIME. HE PERFORMED REPOSITIONING SURGERY, BUT THE OUTCOME WAS NOT BETTER. SO, HE EXCHANGED THE LENS WITH SAME DIOPTER. THE DOCTOR MEASURED THE DATA AGAIN BEFORE THE EXCHANGE SURGERY, AND VAULT WAS LOWER THAN BEFORE. SO, HE DECIDED TO REPLACE WITH A LARGER LENGTH LENS. AFTER EXCHANGING THE LENS, THE VAULT WAS 0.94CT ((B)(6) 2021)." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION AND FUNCTIONAL VERFICATION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697123&pc=QCB | 2023826-2021-01081 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -8.5/+5.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B1-UPDATED. B5-UPDATED INFO: THE REPORTER STATES THE "DOCTOR SOLVED THIS CASE ROTATING THE LENS." H6-HEALTH IMPACT CODE UPDATED. INVESTIGATION CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11704363&pc=QCB | 2023826-2021-01102 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-04-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, 2.5/6.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2021. RESIDUAL EYE DEFECT, ASTIGMATISM WAS REPORTED. LENS REMAINS IN THE EYE. CAUSE OF EVENT WAS REPORTED AS UNKNOWN " VISUAL ACUITY BEFORE SURGERY LOG 0.0 AFTER SURGERY 0.6- RESIDUAL EYE DEFECT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11714600&pc=QCB | 2023826-2021-01131 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.50/2.0/079 9SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721900&pc=QCB | 2023826-2021-01017 | 2019-05-25 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/3.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. REPORTEDLY, BLURRED VISION AND CHANGE IN THE ASTIGMATISM AXIS WAS OBSERVED. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS AND , THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS POSITION IS STABLE DURING FOLLOW-UP. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721910&pc=QCB | 2023826-2021-01026 | 2018-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -18.00/2.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS ROTATED CLOCKWISE 90 DEGREES AFTER SURGERY. BLURRED VISION AND CHANGE IN THE ASTIGMATISM AXIS WERE OBSERVED. ON (B)(6) 2018 THE LENS WAS REPLACED FOR A SAME LENGTH LENS POSITIONED VERTICALLY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721911&pc=QCB | 2023826-2021-01027 | 2018-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -18.00/3.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS ROTATED CLOCKWISE 90 DEGREES AFTER SURGERY. BLURRED VISION AND CHANGE IN THE ASTIGMATISM AXIS WERE OBSERVED. ON (B)(6) 2018 THE LENS WAS REPLACED FOR A SAME LENGTH LENS POSITIONED VERTICALLY; THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721923&pc=QCB | 2023826-2021-01045 | 2021-03-23 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/0.5/148 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721928&pc=QCB | 2023826-2021-01046 | 2021-03-23 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/0.5/176 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721885&pc=QCB | 2023826-2021-01004 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.50/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA; H3-DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721890&pc=QCB | 2023826-2021-01008 | 2021-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/6.0/123 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REPORTEDLY, "PATIENT IS NOT HAPPY WITH VISION." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 13.2 VTICMO 13.2 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2018. REPORTEDLY THE PATIENT EXPERIENCED "SURPRISE REFRACTION OR REFRACTIVE CHANGE OVER TIME AFTER ALMOST 2 YEARS." PATIENT NOT HAPPY WITH HIS VISION. THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721908&pc=QCB | 2023826-2021-01024 | 2018-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2017. LEN OPACITY (NS) WAS OBSERVED ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSEVERED. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US . CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721919&pc=QCB | 2023826-2021-01033 | 2020-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.00/1.5/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER LENS REPLACEMENT, PATIENT HAD 1.0 VISION AND FELT VERY SATISFIED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11729939&pc=QCB | 2023826-2021-01238 | 2020-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/1.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: XIAOJUN821229. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # ((B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11734843&pc=QCB | 2023826-2021-01180 | 2020-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+5.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS REPOSITIONED DUE TO REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: LOW VAULT WAS ALSO OBSERVED. THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11734852&pc=QCB | 2023826-2021-01179 | 2018-01-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS RESIDUAL ASTIGMATISM AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5- ADDITIONAL INFO: A 12.5MM TICM12.5, -7.5/+2.0/164 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2012. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 12.5MM TICM 12.5 OF -7.5/2.0/164 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2012. TICL ROTATED 1 DEGREE COUNTERCLOCKWISE AFTER IMPLANTATION. REFRACTIVE SURPRISE WAS NOTED. THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11750600&pc=QCB | 2023826-2021-01297 | 2019-02-21 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPOSITIONED . THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS (DIFFERENT AXIS) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11750604&pc=QCB | 2023826-2021-01298 | 2019-01-14 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/115 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11751162&pc=QCB | 2023826-2021-01282 | 2018-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.0/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778982&pc=QCB | 2023826-2021-01323 | 2021-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6 -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WERE OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PRESCRIBED DISTANCE GLASSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PATIENT WAS PRESCRIBED DISTANCE GLASSES. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6 - THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION ON THE OPTOCRAFT FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING; THE SPHERE WAS OUT OF TOLERANCE BY 0.19D. H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #719586. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779013&pc=QCB | 2023826-2021-01322 | 2021-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6 -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, "PRESCRIBED DISTANCE GLASSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PATIENT WAS PRESCRIBED DISTANCE GLASSES. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6 - LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION ON THE OPTOCRAFT FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VAULES MEASURED AT THE TIME OF MANUFACTURING; THE SPHERE WAS OUT OF TOLERANCE BY 0.08D. H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: H3 - RESULT CODE: 114 IN SUPPLEMENTAL MDR #3 CLAIM # (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779017&pc=QCB | 2023826-2021-01329 | 2021-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OD -8.50/2.5/097 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE AND DIFFERENT POWER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. D6B- (B)(6) 2021. H6- HEALTH IMPACT CODE 4629- LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779026&pc=QCB | 2023826-2021-01330 | 2021-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OD -8.0/2.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE AND DIFFERENT POWER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. D6B- (B)(6) 2021. H6- HEALTH IMPACT CODE 4629- LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11787301&pc=QCB | 2023826-2021-01387 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/0.5/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. SECONDARY INTERVENTION/ ADDITIONAL TREATMENT HAS NOT BEEN PERFORMED OR DECIDED UPON. FOR STATUS OF THE EYE " PATIENT CANNOT READ ANYMORE " WAS REPORTED. FOR ADDITIONAL INFORMATION THE REPORTER REPORTED " REFRACTIVE SURPRISE. DOCTOR WANTS TO DO A LASER TO SOLVE THE REFRACTIVE SURPRISE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11781988&pc=QCB | 2023826-2021-01384 | 2021-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.0/4.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTD INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. "SIGNIFICANT REFRACTIVE SURPRISE AFTER RE ICL IMPLANTATION" WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITRIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/4.0/155 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE (INCREASED ASTIGMATISM). THE LENS WAS EXPLANTED AND A REPLACEMENT LENS WAS IMPLANTED. REPORTEDLY "THE ORIGINAL LENS WAS EXPLANTED AND REPLACED WITH A NON-TORIC - VA WAS AROUND (B)(6) UNAIDED AND VAULT OK". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788224&pc=QCB | 2023826-2021-01349 | 2021-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -4.0/+1.0/168 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) -2020. ON (B)(6) -2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788229&pc=QCB | 2023826-2021-01348 | 2021-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4) : LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820908&pc=QCB | 2023826-2021-01406 | 2018-07-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.50/1.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820910&pc=QCB | 2023826-2021-01407 | 2018-07-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.50/2.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823183&pc=QCB | 2023826-2021-01480 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS -13.0/+1.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS RESIDUAL ASTIGMATISM AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823341&pc=QCB | 2023826-2021-01455 | 2020-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+0.5/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6)2020, THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823528&pc=QCB | 2023826-2021-01440 | 2021-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.00/3.00/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WA REPORTED TO BE UNKNOWN, THE REPORTER STATED "IN DR. OPINION MUST HAVE A MISTAKE ABOUT ASTIGMATISM. THE MANIFEST REFRACTION AND THE ORDERED IOL WAS -3.00 D OF ASTIGMATISM AND THAT IS ALSO THE ASTIGMATISM THE PATIENT HAVE IN HIS GLASSES. HOWEVER TOPOGRAPHIC ASTIGMATISM WAS NEAR 4.00D. ALSO IN THE CONTRALATERAL EYE ( RIGHT ) BOTH MANIFEST AND TOPOGRAPHIC ASTIGMATISM WERE 4.25D AND WITH 4.50D ORDERED IN THE ICL LENSES THE PATIENT IS FINE. DR. (B)(6) BELIEVES THAT A REPLACEMENT OF THE IOL FOR A NEW ONE WITH 4.00 D OR 4.25 D OF ASTIGMATISM WILL BE OK." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021 AND LATER REPLACED WITH A LONGER LENGTH LENS OF DIFFERENT POWER AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823603&pc=QCB | 2023826-2021-01441 | 2020-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 12.6 MM, TMICL12.6, -9.00/3.00/067(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION (TORIC ICL CONTINUED TO ROTATE) AND BLURRED VISION. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CONTAINER, IN A STERILE POUCH. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: B5 - IT WAS REPORTED THAT A 12.6MM TMICL12.6 -9.00/3.00/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LENS SIZE TOO SMALL; BLURRED VISION AND LENS ROTATION WERE REPORTED. ON (B)(6) 2020 LENS REPOSITIONING WAS PERFORMED BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11829659&pc=QCB | 2023826-2021-01538 | 2021-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AS A REPLACEMENT LENS DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. PER UPDATED INFORMATION PATIENT HAS LEFT OVER RX OF -0.25 -0.50X026. PUPIL IS REACTING IOP 23MMHG. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER LATEST UPDATE, A MEDICAL/SECONDARY SURGICAL INTERVENTION WAS PERFORMED AND THE PROBLEM RESOLVED; HOWEVER NO DETAILS WERE PROVIDED. FINAL POST-OP ON (B)(6) 2021 REPORTS: UCVA 20/25 AND IOP 15MMHG AND FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTED STATED " EXCELLENT OUTCOME. PUPIL ROUND AND REACTIVE WITH NORMAL IOP. PATIENT SATISFIED. PRN." CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11839592&pc=QCB | 2023826-2021-01558 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6MM, -4.50/+4/71 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11839593&pc=QCB | 2023826-2021-01557 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6MM, -4.50/+4.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855234&pc=QCB | 2023826-2021-01516 | 2020-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.50/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL OBSERVATIONS FOUND HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: DATE RECD BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11886266&pc=QCB | 2023826-2021-01614 | 2021-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WILL BE EXPLANTED ON (B)(6) 2021. DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT IS HAPPY. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891207&pc=QCB | 2023826-2021-01699 | 2021-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -10.0/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891209&pc=QCB | 2023826-2021-01698 | 2021-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -9.5/1.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11899111&pc=QCB | 2023826-2021-01682 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -15.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE LENS HAPTIC, BENT DEFORMED AND STRETCHED OUT. DIMENSIONAL INSPECTION WAS PERFORMED, AND THE DIAMETER VERIFICATION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION PERFORMED FOUND THE LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H6- WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11904673&pc=QCB | 2023826-2021-01689 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICMO13.2, -7.0/2.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED DRY, INSIDE A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC STRETCHED OUT, HAPTIC DEFORMED, AND HAPTIC STRETCHED OUT. DIMENSIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE OUT OF SPECIFICATION (ONLY LENS SPHERE, CYLINDER AND AXIS WITHIN SPECIFICATION).. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: A2: DOB (B)(6) 1992. CLAIM (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11904725&pc=QCB | 2023826-2021-01690 | 2020-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.0/1.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BENT, DEFORMED AND STRETCHED OUT. DIMENSIONAL WIDTH AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. UNABLE TO PERFORM LENGTH VERIFICATION, DUE TO HAPTIC FOOTPLATES DAMAGE. CLAIM (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938949&pc=QCB | 2023826-2021-01648 | 2021-05-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.50/5.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938956&pc=QCB | 2023826-2021-01649 | 2021-05-10 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.50/4.0/009 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939030&pc=QCB | 2023826-2021-01673 | 2021-05-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.00/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938963&pc=QCB | 2023826-2021-01658 | 2020-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION NOT REQUIRED . B3: DATE CHANGED TO 11/10/2020. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938964&pc=QCB | 2023826-2021-01657 | 2021-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.00/0.05/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH LIQUID ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942673&pc=QCB | 2023826-2021-01780 | 2020-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942675&pc=QCB | 2023826-2021-01778 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942687&pc=QCB | 2023826-2021-01772 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATION. H6-METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT; UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966802&pc=QCB | 2023826-2021-01813 | 2021-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.50/1.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020 . REFRACTIVE SURPRISE WAS OBSERVED. REFRACTIVE TREATMENT BY CORNEAL T-CUT WAS PERFORMED. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, AN EXCHANGE IS PLANNED AND WILL UPDATE WHEN EXCHANGE IS DONE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: ADVERSE EVENT. H1: SERIOUS INJURY. CLAIM : (B)(4). Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK, DATE OF EVENT- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986307&pc=QCB | 2023826-2021-01875 | 2021-05-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+2.0/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS BLURRED NIGHT VISION DUE TO LARGE PUPIL SIZE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005350&pc=QCB | 2023826-2021-01939 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-06-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+4.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT WAS REPORTED AS COMPLAINING OF BLURRED VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005489&pc=QCB | 2023826-2021-01940 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-06-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+4.5/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT WAS REPORTED AS COMPLAINING OF BLURRED VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009118&pc=QCB | 2023826-2021-01896 | 2021-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/.05/058 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, THERE IS STILL A RESIDUAL REFRACTION, AN EXCHANGE IS PLANNED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/.05/058 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, THERE IS STILL A RESIDUAL REFRACTION, AN EXCHANGE IS PLANNED." ON (B)(6) 2021 THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. REPORTEDLY, ALL FINE. CLINIC TOLD THERE IS GOOD VAULTING. PATIENT IS HAPPY!. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021861&pc=QCB | 2023826-2021-02001 | 2021-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/176 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS ON THE LENS. DIMENSION INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12029743&pc=QCB | 2023826-2021-01916 | 2021-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.0/1.0/16 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM, VTICM5_ 12.6, -5.0/1.0/16 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT POWER DUE TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME. THIS EXCHANGED RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060456&pc=QCB | 2023826-2021-01953 | 2021-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.00/2.0/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BROKEN, BENT, DEFORMED, AND STRETCHED OUT. CLAIM#: (B)(4). Manufacturer Narrative: H3: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060470&pc=QCB | 2023826-2021-01954 | 2021-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: H3: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060480&pc=QCB | 2023826-2021-01966 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060483&pc=QCB | 2023826-2021-01967 | 2020-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.00/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE, SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: H3: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064274&pc=QCB | 2023826-2021-02062 | 2021-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+5.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING, WITH LENS ROTATION. THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED, THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENS. AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND LENS IS IN POSITION. PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. H3: DEVICE EVALUATION IS NOT REQUIRED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072254&pc=QCB | 2023826-2021-02088 | 2020-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072262&pc=QCB | 2023826-2021-02075 | 2021-05-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.5/+3.0/049 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE SURGEON STATES, "THE LENS IS WELL ALIGNED IN PLACE AND THE EYE IS QUITE AND STABLE BUT THE SURPRISE REFRACTION IS UNEXPLAINABLE TO US." Manufacturer Narrative: B5- SURGEON PLANS TO REMOVE THE LENS. RECEIVED UPDATE ON 24-MAR-2022 "PATIENT JUST SHOW UP TODAY SAME POST OP REFRACTION 3 D ASTIGMATISM. SURGEON LOOKING FOR ADVISE, HE IS PLANNING TO REMOVE THE LENS AND ACCESS AGAIN". CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072269&pc=QCB | 2023826-2021-02097 | 2021-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. PER EMAIL, "ALL IS WELL WITH THE PATIENT." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072274&pc=QCB | 2023826-2021-02072 | 2021-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.5/+1.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B4: IN FOLLOW-UP 1 ENTERED IN ERROR. SHOULD BE BLANK AND DATE ENTERED IN D6B. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -9.5/+1.0/170 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS REPORTED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH/DIFFERENT DIOPTER LENS. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - CORRECTION TO INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102507&pc=QCB | 2023826-2021-02132 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVERTIME WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND POOR UNAIDED VISION WERE OBSERVED. LENS REMAINS IMPLANTED AND AN EXCHANGE IS PLANNED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2019. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: 1 SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102535&pc=QCB | 2023826-2021-02131 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.50/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVERTIME WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE, "CHANGED POSITION AND NOW OFF AXIS AND POOR UNAIDED VISION. " ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.50/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, REFRACTIVE CHANGE OVERTIME AND POOR UNAIDED VISION WERE OBSERVED. LENS REMAINS IMPLANTED AND A LENS EXCHANGE IS PLANNED. TH REPORTER MARKED THE CAUSE OF THE EVENT AS UNKNOWN AND STATED "BASICALLY VISION AND TICL POSITION CHANGED AFTER 6/12 AND NOW OFF AXIS AND POOR UNAIDED VISION." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED CQ: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.50/+4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2019. LOW VAULT WITH ROTATION, REFRACTIVE CHANGE OVERTIME AND POOR UNAIDED VISION WERE OBSERVED. LENS WAS EXPLANTED IN (B)(6) 2021 AND REPLACED WITH A LONGER LENGTH LENS WITH SIMILAR POWER (VTICMO 13.2, -15.50/+4.0/095) ON (B)(6) 2021. THE PROBLEM WAS RESOLVED AND "VAULT HAS IMPROVED AND PATIENT IS HAPPY". CAUSE OF EVENTS WAS REPORTED AS UNKNOWN. D6B- (B)(6) 2021. H6 DEVICE CODE: 1583 ADDED; LOW VAULT WITH LENS ROTATION REPORTED. H6 DEVICE CODE: 1494: OFF LABEL USE- AGE<21 YEARS OLD. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2019. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102798&pc=QCB | 2023826-2021-02145 | 2021-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: HE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.50/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D4: WRONG SERIAL NUMBER PROVIDED IN INITIAL MDR B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -2.50/+6.0/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). REFRACTIVE SURPRISE AND BLURRY VISION REPORTED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 WHICH DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. REPORTEDLY MISTAKE WAS MADE WITH MIXING UP THE SERIAL NUMBERS. THEY ORDERED THE SAME LENS FOR THE SAME PATIENT UNDER 2 DIFFERENT DOCTOR NAMES. AN ANALYSIS WAS MADE BASED ON THE WRONG LENS. THE INFORMATION CORRECT LENS IMPLANTED HAS A DIFFERENT IOD AND NOW PROBABLY A NEW ANALYSIS IS MADE WITH THE CORRECT IOD. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12112717&pc=QCB | 2023826-2021-02226 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/1.0/90, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AFTER IMPLANT WAS OBSERVED. " THE PATIENT HAS GOT AN HYPEROPIA AND STILL SMOOTH LOSS OF UCVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: B5- LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE PATEINT IS HAPPY WITH THE RESULTS AND WEARS GLASSES IF NECESSARY. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12112726&pc=QCB | 2023826-2021-02225 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/1.0/86, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AFTER IMPLANT WAS OBSERVED. " THE PATIENT HAS GOT AN HYPEROPIA AND STILL SMOOTH LOSS OF UCVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: B5- LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE PATIENT IS HAPPY WITH THE RESULTS AND WEARS GLASSES IF NECESSARY. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119274&pc=QCB | 2023826-2021-02282 | 2020-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: THE LENS HAS ROTATED OUT OF POSITION. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE ON LENS. CLAIM# : (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. COMPLAINT# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132564&pc=QCB | 2023826-2021-02384 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). AT A POST-OP VISIT, THERE WAS AN INCREASE IN REFRACTIVE ASTIGMATISM AND PATIENT WAS NOT HAPPY WITH THEIR UNCORRECTED DISTANCE VISUAL ACUITY. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140273&pc=QCB | 2023826-2021-02413 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 TH LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED " ON THE FIRST DAY AFTER OPERATION, THE VISUAL ACUITY WAS 0.6 AND THE CORNEAL EDEMA WAS MILD." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM,VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED " ON THE FIRST DAY AFTER OPERATION, THE VISUAL ACUITY WAS 0.6 AND THE CORNEAL EDEMA WAS MILD." NO SPECIAL TREATMENT WAS NEEDED FOR THE EDEMA. THE EDEMA DISAPPEARED THE NEXT DAY, AND THERE WAS NO SIGNIFICANT CHANGE BETWEEN CORNEAL ENDOTHELIUM AND PREOPERATIVE EXAMINATION. CLAIM#: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147536&pc=QCB | 2023826-2021-02268 | 2021-05-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, "PENTACAM HAS NOT BEEN PERFORMED YET." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147853&pc=QCB | 2023826-2021-02481 | 2021-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.00/4.00/092, IMPLANTABLE COLLAMER LENS WAS IMLPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS NOTED, LENS REMAINS IMPLANTED. . IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: ADVERSE EVENT; B2: REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE; B5: THE REPORTER INDICATED, THAT A 12.6MM, VTICM5 12.6, -11.00/4.00/092, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME SIZE SPHERICAL LENS, DUE TO REFRACTIVE SURPRISE. RESIDUAL ASTIGMATISM WAS CORRECTED BY EXCIMER LASER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS DID NOT MEET THE ORIGINAL VALUES MEASURED AT TIME OF MANUFACTURING. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147855&pc=QCB | 2023826-2021-02497 | 2021-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.0/2.0/100, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H10: IT WAS LATER CLARIFIED THAT THIS CLAIM WAS REPORTED IN ERROR WITH INCORRECT INFORMATION AND SEE MFR#: 2023826-2021-04168 WHICH HAS THE CORRECT DETAILS (CLAIM#: (B)(4). DISREGARD INITIAL MDR SUBMISSION. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION WERE UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148279&pc=QCB | 2023826-2021-02476 | 2020-10-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -13.5/+6.0/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND POOR NEAR VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148267&pc=QCB | 2023826-2021-02470 | 2021-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+6.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS LENS ROTATION AND REFRACTIVE SURPRISE. REPORTEDLY AN ADDITIONAL SUTURE WAS ADDED AND ON (B)(6) 2021 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: REPORTEDLY THE LENS WAS REPOSITIONED AGAIN AND THE PATIENT IS "FINE FOR NOW." CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148274&pc=QCB | 2023826-2021-02471 | 2021-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/001 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155600&pc=QCB | 2023826-2021-02446 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN EYE AND VISION WAS RECEIVED ON 14-JUN-2021, ENTITLED 'SEVEN-YEAR FOLLOW-UP OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS WITH CENTRAL PORT DESIGN.' THE ARTICLE STATES THAT 84 EYES OF 52 PATIENTS UNDERWENT IMPLANTATION OF THE V4C LENS TO CORRECT MYOPIA. TWO EYES OF THE SAME PATIENT REQUIRED LASER TOUCH UP TO CORRECT RESIDUAL REFRACTIVE ERROR 2 YEARS AFTER ICL IMPLANT. 4 EYES HAD THE ICL ROTATED 90 DEGREES TO A VERTICAL ORIENTATION BECAUSE OF HIGH VAULT AT THE 1-DAY POSTOP VISIT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT IMPACT CODE: 4628 SHOULD BE REMOVED AS IT'S N/A. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE-UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12160678&pc=QCB | 2023826-2021-02556 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2; -10.0/4.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH AND MODEL LENS BUT DIFFERENT SPHERE/AXIS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND DEPOSIT ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167087&pc=QCB | 2023826-2021-02510 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. LENS ROTATION WAS OBSERVED. MAKING PATIENT'S VISION WORSE AND BLURRY. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167123&pc=QCB | 2023826-2021-02530 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, LENS TOO SHORT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167124&pc=QCB | 2023826-2021-02531 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM0 12.6, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -15.00/4.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, LENS TOO SHORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12208007&pc=QCB | 2023826-2021-02760 | 2021-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -14.5/2.0/162, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT AXIS LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN, THE REPORTER STATED " THE SURGEON INPUT THE WRONG ASTIGMATISM AXIS IN OCOS SYSTEM". Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12211083&pc=QCB | 2023826-2021-02758 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.0/89, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN /STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214331&pc=QCB | 2023826-2021-02694 | 2021-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND FIBER ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214338&pc=QCB | 2023826-2021-02696 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.50/3.50/175 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: ON (B)(6) 2021, THE LENS WAS EXPLANTED AND AT A UNKNOWN DATE, PHACO-EMULSIFICATION AND IOL IMPLANTATON WAS PERFORMED. THE PROBLEM RESOLVED. REPORTEDLY, "PHACO WITH MIOL. PATIENT IS VERY HAPPY!" CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214343&pc=QCB | 2023826-2021-02695 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/3.00/034 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - ON (B)(6) 2021, THE LENS WAS EXPLANTED. CLAIM #: (B)(4). Manufacturer Narrative: B5: ON (B)(6) 2021, THE LENS WAS EXPLANTED AND AT A UNKNOWN DATE, PHACO-EMULSIFICATION AND IOL IMPLANTATON WAS PERFORMED. THE PROBLEM RESOLVED. REPORTEDLY, "PHACO WITH MIOL. PATIENT IS VERY HAPPY!" CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222277&pc=QCB | 2023826-2021-02815 | 2021-06-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/117 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222317&pc=QCB | 2023826-2021-02805 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS -14.5/+2.0/064 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS A SMALL REFRACTIVE ERROR AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12255959&pc=QCB | 2023826-2021-02976 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/015 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED BLURRED VISION. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR BUT THE DEVICE DID FAIL TO PERFORM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257216&pc=QCB | 2023826-2021-02971 | 2021-06-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -17.5/+1.0/015 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257224&pc=QCB | 2023826-2021-02964 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+5.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT-UNK IMPLANT DATE-UNK PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257227&pc=QCB | 2023826-2021-02970 | 2021-06-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -16.5/+0.5/178 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12263420&pc=QCB | 2023826-2021-02555 | 2021-07-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.00/+2.50/73 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT "THE RESULT IS OK FOR THE PATIENT AND WISH NO FURTHER SURGICAL INTERVENTION". PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278288&pc=QCB | 2023826-2021-03053 | 2021-07-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -4.0/+4.0/140 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. MANIFEST REFRACTION IS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-ADDITIONAL AND CORRECTED INFORMATION: THE LENS WAS IMPLANTED ON (B)(6) 2014. REPORTEDLY, "THE RESULT WAS OK AND THE PATIENT WAS HAPPY" AFTER THE INITIAL LENS WAS IMPLANTED. REFRACTIVE CHANGE OVER TIME IS REPORTED DUE TO MYOPIA GROWTH. (B)(4). Manufacturer Narrative: SEX WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281699&pc=QCB | 2023826-2021-03066 | 2021-06-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.MM, VTICH13.2, -5.00/5.5/004 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281714&pc=QCB | 2023826-2021-03067 | 2021-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.MM, VTICH13.2, -5.00/5.5/004 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12321646&pc=QCB | 2023826-2021-03148 | 2021-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.50/1.0/171, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT ID: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12326006&pc=QCB | 2023826-2021-03182 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2MM; -1.0/+6.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE; LENS ROTATION WHICH THE SURGEON PERFORMED REPOSITIONING ON (B)(6) 2021 WHICH DID IMPROVE VISION AND REFRACTIVE BUT AGAIN THE VISION HAS BLURRED AND LENS HAD ROTATED AGAIN. THE PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378554&pc=QCB | 2023826-2021-03370 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -10.5/+0.5/025 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SUPRISE AND EXCESSIVE VAULT ARE REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -10.5/+0.5/025 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE AND EXCESSIVE VAULT ARE REPORTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378674&pc=QCB | 2023826-2021-03369 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -9.0/+0.5/172 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SUPRISE AND EXCESSIVE VAULT ARE REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -9.0/+0.5/172 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE AND EXCESSIVE VAULT ARE REPORTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426115&pc=QCB | 2023826-2021-03452 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS "REDIALED" DUE TO LENS ROTATION AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS ROTATED AGAIN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443369&pc=QCB | 2023826-2021-03537 | 2021-08-03 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021.ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVERTIME. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. REPORT SOURCE: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457493&pc=QCB | 2023826-2021-03557 | 2021-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.0/+6.0/044 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. AT A LATER DATE A SAME SIZE DIFFERENT SPHERE/AXIS LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON REPORTS, "WE IMPLANTED THE TICL AND WE GOT A REFRACTIVE SURPRISE AND LOW VA OF 0.1...SO WE REMOVED THE LENS AND DID SUBJECTIVE REFRACTION, THERE WAS A CHANGE IN REFRACTION, SO I REQUESTED A TRANSPOSED BCVA TEST, WE ACHIEVED THE 20/40 WHICH IS THE TARGET." Manufacturer Narrative: B5-THE REPORTER NOW REPORTS PERMANENT LOSS OF BCVA AS THE REASON FOR THE LENS CHANGE. ALSO REPORTED, "ICL WAS IMPLANTED AND HAD A REFRACTIVE ERROR. I REQUESTED A NEW MANIFEST REFRACTION AND BCVA WITH A TRANSPOSED REFRACTION AND CONCLUDED THAT THE OLD LENS REFRACTION WAS NOT ACHIEVING THE TARGET SO WE HAD TO REDO THE REFRACTION AGAIN." CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES, "AFTER THE PATIENT IMPLANTED THE FIRST ICL HE HAD REDUCTION IN THE BCVA. AFTER EXCHANGING IT THE NEW LENS RETRIEVED THE BCVA." AFTER EXCHANGE THE BCVA IS 0.5. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION CODE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457494&pc=QCB | 2023826-2021-03558 | 2021-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/+2.0/059 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER DIFFERENT MODEL LENS DUE TO LENS ROTATION AND LOSS OF BCVA. THE PROBLEM WAS RESOLVED. THE REPORTER STATES, "THERE WERE BIG DIFFERENCES IN WTW MEASUREMENT BY OPD, IOL MASTER AND CALIPER." Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. THE LENS RETURNED IS NOT THE LENGTH STATED IN THE CLAIM. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479353&pc=QCB | 2023826-2021-03598 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.5/2/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AS REPLACEMENT LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. THE LENS ROTATED AND THE PATIENTS VISION AT EARLY POSTOPERATIVE EXAMINATION WAS ONLY 20/40. THE REPORTER STATED "THE DOCTOR DECIDED TO CONTINUE THE OBSERVATION. THE PATIENT WAS LAST REVIEWED ON (B)(6) 2021 AND HIS VISUAL ACUITY REACHED 20/20. THE SURGEON CONSIDERED THAT THE REASON FOR 20/20 VISION WAS DUE TO THE EFFECT OF THE CORNEAL INCISION AND THE POSITION OF THE LENS BUT DID NOT DILATE THE PUPIL." THE CORNEAL INCISION WAS PERFORMED WITHIN THE SAME SURGERY AS IMPLANTATION OF THE LENS. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479078&pc=QCB | 2023826-2021-03590 | 2021-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM512.1, -8.00/2.50/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479327&pc=QCB | 2023826-2021-03657 | 2021-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.5/2.0/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS LATER REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS BUT THE PROBLEM DID NOT RESOLVE. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12489291&pc=QCB | 2023826-2021-03574 | 2021-03-20 | Injury | STAAR SURGICAL COMPANY | 2021-09-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, 07.00/1.0/133 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, BLURRED VISION AND DECREASED VISUAL ACUITY. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12492856&pc=QCB | 2023826-2021-03651 | 2021-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND OPTIC AND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL AND FUNCTIONAL VERIFICATION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S.. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12517049&pc=QCB | 2023826-2021-03704 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: A REFRACTIVE SURPRISE OF 3D OF RESIDUAL CYLINDER WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. CORRECTED TO SERIOUS INJURY IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: AGE OR DOB, SEX, WEIGHT, ETHNICITY, RACE: UNK. EVENT DATE: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524398&pc=QCB | 2023826-2021-03718 | 2021-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-11.5/1.0/68, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND REFRACTIVE SURPRISE. ON (B)(6) 2021 A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT DIOPTER WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12537688&pc=QCB | 2023826-2021-03697 | 2021-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED. THE PLAN IS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPLACED WITH ANOTHER ICL LENS (SPHERICAL) AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6 - LENS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE : NA. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12550693&pc=QCB | 2023826-2021-03794 | 2021-07-22 | Injury | STAAR SURGICAL COMPANY | 2021-09-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.50/5.0/02 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2021 DUE TO A MINOR ROTATION. IT WAS REPORTED THAT THE REPOSITIONING DID NOT RESOLVE THE PROBLEM. THE LENS WAS REMOVED AND ON (B)(6) 2021 A REPLACEMENT LENS OF DIFFERENT POWER WAS IMPLANTED, PATIENT IS IN EXCELLENT CONDITION AND SATISFIED. CAUSE OF EVENT WAS REPORTED TO BE USER ERROR, LENS POWER WAS NOT CALCULATED CORRECTLY. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT WAS FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558537&pc=QCB | 2023826-2021-03808 | 2020-10-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR, "MAY BE CAUSE OF THE OCULAR ANATOMY." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599770&pc=QCB | 2023826-2021-03908 | 2021-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.5/+2.0/020 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES, "PATIENT IS UNHAPPY POST OPERATIVELY AS HE ACCEPTS +1/-1.75@145. WITH THIS CORRECTION PATIENT IMPROVES TO 20/20." Manufacturer Narrative: B5-UPDATED INFORMATION: REPORTEDLY THE LENS WAS REPOSITIONED. AS OF 1-DAY POSTOP THE PATIENT IS "DOING WELL.". CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12643335&pc=QCB | 2023826-2021-04012 | 2021-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.5/2.0/112, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. IT WAS REPORTED THAT IT WAS UNKNOWN WHEN THE DAMAGE OCCUR AND THAT THE DAMAGE WAS NOTED WHEN WITHDRAWING THE LENS. LOSS OF BCVA WAS ALSO NOTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-PER UPDATED INFORMATION THE REPORTER STATED " THE DOCTOR NOTICED THAT THE LENS WAS BROKEN, WHEN THE NEXT DAY, WHEN CHECKING THE EYE. THE PATIENT'S EYE IS FINE. HE BELIEVES THAT THE FAILURE MUST HAVE BEEN WHEN INSERTING THE LENS INTO THE EYE WAS BROKEN IN THE INJECTION." CLAIM # (B)(4). Manufacturer Narrative: G3- CORRECTED TO (B)(6) 2021, IN INITIAL MDR. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12643795&pc=QCB | 2023826-2021-04018 | 2021-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.00/+3.0/083, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 03 SEP 2021. REFRACTIVE SURPRISE WAS NOTED, LENS REMAINS IMPLANTED. " PATIENT IS ACCEPTING +1/-2.5@12 AND IMPROVING TO 6/5" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 - "REPORTEDLY PER EMAIL ON 06/06/2022, 'SURGEON RE-ROTATED THE LENS ON (B)(6) 2022 TO POSITION. THE REFRACTION WAS REPORTED AS +1.5/-2.5/30. THE PATIENT IS HAPPY WITH 6/6 VISION.'" SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4628 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B5 - "REPORTEDLY, 'LAST MONTH THIS PATIENT VISITED THE SURGEON ON (B)(6) 2022 AND THAT TIME THE REFRACTION WAS +1.5/-2.5/30. SURGEON RE-ROTATED THE LENS ON (B)(6) 2022 TO POSITION AND THE PATIENT IS HAPPY WITH 6/6 VISION.'" SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12678194&pc=QCB | 2023826-2021-04106 | 2021-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/3.0/122, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS NOTED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 - "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/ MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 2923 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: .UNK PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12724193&pc=QCB | 2023826-2021-04156 |  | Injury | STAAR SURGICAL COMPANY | 2021-10-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -7.5/+2.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION AND REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED AND THE LENS ROTATED AGAIN. REPORTEDLY, "THE PATIENT COMPLAINED OF REDUCTION OF VISION. AFTER THE FIRST SURGERY THE IOL WAS AT 4 DEGREES. THE TARGET WAS 178 DEGREES. REFRACTION WAS +0.50/-1.00X160." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725216&pc=QCB | 2023826-2021-04179 | 2021-09-06 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/2.50/018, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PER UPDATED INFORMATION AN EXCHANGED WAS CARRIED OUT ON (B)(6) 2021. THE REPORTER STATED THAT THE SURGEON HAS CONFIRMED THAT THE PATIENT IS NOW DOING WELL. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725051&pc=QCB | 2023826-2021-04168 | 2021-03-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/1.0/102, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE PRODUCT WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE RETURNED LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725068&pc=QCB | 2023826-2021-04167 | 2021-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.5/172, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS USER ERROR, THE REPORTER STATED "THE DOCTOR DID NOT ENTER WHEN CALCULATING THE CYLINDER." Manufacturer Narrative: B3- CORRECTED TO (B)(6) 2021 IN INITIAL MDR. D4 MODEL #: CORRECTED TO VTICMO12.6 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE PRODUCT WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED LENS FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING; H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12749569&pc=QCB | 2023826-2021-04228 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.5/1.5/101, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE LENS CYLINDER TO BE OUT OF SPECIFICATION. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755752&pc=QCB | 2023826-2021-04209 | 2021-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.5/+0.5/116 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: H6: INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY THE SYMPTOMS HAVE RESOLVED AND NO INTERVENTION IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755753&pc=QCB | 2023826-2021-04208 | 2021-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.5/+0.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES, "THE REFRACTION WAS REMINDED UNCORRECTED (OR WORSE) AT 1D POST-OPE AND THE PATIENT GOT VERY ANGRY ABOUT THE RESULT, BUT AT 1W, THE REFRACTION BECAME NORMAL AS EXPECTED. WE DON'T HAVE ANY ESTIMATIONS OF REASON WHY THE TRANSIENT REFRACTIVE SURPRISE WAS HAPPENED." Manufacturer Narrative: B5 ADDITIONAL INFORMATION: REPORTEDLY THE PATIENT'S SYMPTOMS HAVE RESOLVED AND NO INTERVENTION IS PLANNED. (B)(4). Manufacturer Narrative: H6: INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763784&pc=QCB | 2023826-2021-04231 | 2021-10-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.5/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807798&pc=QCB | 2023826-2021-04263 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/174 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12808809&pc=QCB | 2023826-2021-04333 | 2021-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/+0.5/138, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION ON 27-DEC-2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND WIDTH TO BE WITHIN SPECIFICATION; HOWEVER, FUNCTIONAL INSPECTION FOUND THE LENS AXIS OUT OF SPECIFICATION. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12849444&pc=QCB | 2023826-2021-04442 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+1.50/97, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872781&pc=QCB | 2023826-2021-04487 | 2021-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.00/+2.5/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING, LENS ROTATION AND BLURRED VISION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872828&pc=QCB | 2023826-2021-04488 | 2021-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50/+4.0/079 (SPHERE/ CYLINDER/ AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING, LENS ROTATION AND BLURRED VISION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872841&pc=QCB | 2023826-2021-04562 | 2021-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; 10.0/1.5/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE REPLACED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER/AXIS DUE TO THE PATIENT EXPERIENCING REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; 10.0/1.5/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE REPLACED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER/AXIS DUE TO THE PATIENT EXPERIENCING REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12879840&pc=QCB | 2023826-2021-04526 | 2021-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -16.5/5.0/100, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN/BROKEN . DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12911285&pc=QCB | 2023826-2021-04623 | 2021-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/3.5/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH SIMILAR POWER DIFFERENT CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973012&pc=QCB | 2023826-2021-04653 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.50/+5.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973346&pc=QCB | 2023826-2021-04707 | 2021-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/000 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISIBLE INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973347&pc=QCB | 2023826-2021-04711 | 2020-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-REPORTEDLY THE SURGEON REPOSITIONED THE LENS AND AS OF (B)(6) 2021 THE UCVA IS 20/20. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973353&pc=QCB | 2023826-2021-04697 | 2020-03-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6)2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023235&pc=QCB | 2023826-2021-04887 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7, -12.0/+0.5/119 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13040123&pc=QCB | 2023826-2021-04906 | 2021-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.5/3.0/95 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND IT WAS INDICATED THAT THE 13.2MM LENS CAUSED A OVERCORRECTION; WHICH MADE IT DIFFICULT TO READ UP CLOSE FOR THE PATIENT. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM AND NOTED "MYOPIA UNDER CORRECTION". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MDR # 2023826-2021-04907 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082112&pc=QCB | 2023826-2021-04987 | 2021-12-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -5.5/5.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATE RECEIVED ON 09-MARCH-2022 "NO REPLACEMENT WILL BE PLANNED. VAULTING WILL BE MONITORED. RESIDUAL REFRACTION WILL BE TREATED BY LASER TOUCH UP." THE PROBLEM WAS MARKED AS RESOLVED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082432&pc=QCB | 2023826-2021-04962 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/046 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. REPORTEDLY THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/046 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. REPORTEDLY THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AGE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AGE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082443&pc=QCB | 2023826-2021-04968 | 2021-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.5/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS REPOSITIONED DUE TO REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT AXIS LENS AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "CILIARY BODY CYSTS ROTATING ICL TO 165 DEGREES AND ROTATION ICL ON (B)(6) 2021 DIDN'T RESOLVED THE PROBLEM. ON 1ST DAY AFTER SURGERY ICL WAS AGAIN ROTATED TO 165 DEGREES." Manufacturer Narrative: (B)(6). PATIENT EIGHT, ETHNICITY AND RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090705&pc=QCB | 2023826-2021-04997 | 2021-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/3.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13115042&pc=QCB | 2023826-2021-05018 | 2020-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.5/1.0/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME MODEL BUT DIFFERENT LENGTH/POWER LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13115047&pc=QCB | 2023826-2021-05021 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.50/+0.50/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN, BENT, DEFORMED, AND STRETCHED OUT. A BROKEN LENS HAPTIC PIECE WAS FOUND ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13122738&pc=QCB | 2023826-2021-05007 | 2021-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0/+2.0/005 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6-INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13144019&pc=QCB | 2023826-2022-00016 | 2021-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/6.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/6.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS WAS REPOSITIONED ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. LENS REMAINS IMPLANTED. PATIENT IS HAPPY. CAUSE OF THE EVENT WAS REPORTED AS UKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- PER CQ RECEIVED ON 02-FEB-2022: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/6.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS WAS REPOSITIONED ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. LENS REMAINS IMPLANTED. PATIENT IS HAPPY. CAUSE OF THE EVENT WAS REPORTED AS UKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13160463&pc=QCB | 2023826-2022-00077 | 2021-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7; -14.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME WERE ASSESSED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182567&pc=QCB | 2023826-2022-00026 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11.5/+3.5/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. BLURRED VISION IN THE IMMEDIATE POST-OPERATIVE CONTROL IS REPORTED. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5- ADDITIONAL INFORMATION: PER THE REPORTER BLURRED VISION IMMEDIATELY POST-OP IS AN EXPECTED OCCURRENCE. THE PATIENT IS "HAPPY AND ADAPTED." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13233320&pc=QCB | 2023826-2022-00154 | 2021-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.5/1.0/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13233423&pc=QCB | 2023826-2022-00153 | 2021-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.00/4.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE . VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256317&pc=QCB | 2023826-2022-00140 | 2021-11-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.0/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE AND BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256345&pc=QCB | 2023826-2022-00139 | 2022-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS AXIAL MISALIGNMENT. IT WAS "CORRECTED" AND THE LENS ROTATED AGAIN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -4.5/+5.0/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 AND (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO AXIAL MISALIGNMENT AND ROTATION BUT PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED WITH LONGER LENS AND PROBLEM WAS RESOLVED. "LENS SIZE IS NOT A PROBLEM. A LITTLE ASTIGMATISM REMAINS. CONSIDERING WHETHER TO RECORRECT THE AXIS OR GET LASIK." THE CAUSE OF THE EVENT WAS NOT REPORTED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256368&pc=QCB | 2023826-2022-00152 | 2021-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -7.0/+4.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A DIFFERENT SPHERE/CYLINDER LENS. THE PROBLEM WAS RESOLVED. CLAIM# (B)(4) Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286530&pc=QCB | 2023826-2022-00179 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE (VA WAS 20/400) AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: H3) DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. CLAIM # (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286581&pc=QCB | 2023826-2022-00194 | 2021-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -18.0/4.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS REPORTED THE LENS HAD ROTATED 30 DEGREES COUNTER CLOCKWISE AND THE PATIENT COMPLAINED OF BLURRY VISION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS AND DIFFERENT AXIS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13296036&pc=QCB | 2023826-2022-00241 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.00/+1.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED BLURRED VISION AND WAS DISSATISFIED WITH THE QUALITY OF VISION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. ON (B)(6) 2021 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH BUT DIFFERENT POWER/AXIS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. PATIENT IS MORE SATISFIED AND EYE IS QUITE. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND NOTED "IN MY OPINION MAYBE INTRAOCULAR REACTION LED TO THIS BAD QUALITY OF VISION OR LENS DEPOSITS". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: D4: CORRECTED TO VTICM5_13.2. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13389436&pc=QCB | 2023826-2022-00356 | 2020-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.5/2 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.5/2 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13389437&pc=QCB | 2023826-2022-00355 | 2020-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. THE LENS RETURNED IS NOT THE SIZE STATED IN THE CLAIM. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459226&pc=QCB | 2023826-2022-00393 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/+5.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5 - "THE SURGEON REPOSITIONED THE LENS TWO DAYS LATER." SHOULD BE ADDED TO THE INITIAL MDR. H1- CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4628 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: D4- MODEL# CORRECTED TO VTICMO12.6 IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13488089&pc=QCB | 2023826-2022-00426 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -4.00/-3.00/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS REPORTED A LENS EXCHANGED IS PLANNED TO RESOLVE THE ISSUE. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13542049&pc=QCB | 2023826-2022-00439 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO 12.1, -17.50/1.0/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE SPHERICAL LENS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588892&pc=QCB | 2023826-2022-00599 | 2017-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND NOTED "OVER CORRECTION +1.75/0.25/45-0.7. IN A SEPARATE SURGERY THAT SAME DAY, THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. POSTOP STATUS INDICATED "ECD 2283/MM2". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED THAT THE LENS WAS NOT EXCHANGED ON THE SAME DAY OF (B)(6) 2017, BUT EXPLANTED ON 05-JUL-2017 AND LATER ON THIS SAME DATE REPLACED WITH A SAME LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13608356&pc=QCB | 2023826-2022-00618 | 2021-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.5/2.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAS REFRACTIVE SURPRISE AND THERE ARE "NO ISSUES. VAULT WAS NOT RECORDED ON MOST RECENT VISION; BUT AT A PREVIOUS VISIT IT WAS 3+ ICL THICKNESS; VECTOR ANALYSIS WAS DONE BUT NO DECISION ON FURTHER TREATMENT". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS INDICATED THAT THE PATIENT HAD LENS ROTATION. THIS SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13608364&pc=QCB | 2023826-2022-00619 | 2021-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.5/2.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAS REFRACTIVE SURPRISE AND THERE ARE "NO ISSUES. VAULT WAS NOT RECORDED ON MOST RECENT VISION; BUT AT A PREVIOUS VISIT IT WAS 3+ ICL THICKNESS; VECTOR ANALYSIS WAS DONE BUT NO DECISION ON FURTHER TREATMENT". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#:(B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS INDICATED THAT THE PATIENT HAD LENS ROTATION. THIS SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13712176&pc=QCB | 2023826-2022-00760 | 2021-11-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/+4.0/146 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13725581&pc=QCB | 2023826-2022-00772 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2; 5.50/3.0/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION AND LOSS OF VISUAL ACUITY WHICH THE SURGEON REPOSITIONED THE LENS ON (B)(6) 2022; BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS AND THIS RESOLVED THE PROBLEM. IT WAS ALSO INDICATED THAT THE RESIDUAL CYLINDER WILL BE TREATED BY LASER TOUCH UP AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831254&pc=QCB | 2023826-2022-00882 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.00/1.5/92 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO CORRECT REFRACTIVE SURPRISE BUT IT DID NOT RESOLVE THE PROBLEM, UNAIDED VA RE 6/9+1 WAS REPORTED. LENS ROTATION WAS OBSERVED AT 20 DAYS POST-OP. IT WAS REPORTED THAT LENS REPOSITIONING WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED, UCVA 20/20-2 WAS REPORTED AT PATIENTS LAS VISIT ON (B)(6) 2022. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2022-00435 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13872480&pc=QCB | 2023826-2022-00903 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL/SIZE BUT DIFFERENT POWER -10.5/2.0/90 (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND HAPATIC BROKEN. UNABLE TO IDENTIFY LENS MODEL OR PROCEED WITH DIMENSIONAL AND REFRACTIVE MEASUREMENTS DUE TO LENS RETURNED SPLIT IN TWO. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13878813&pc=QCB | 2023826-2022-00943 | 2022-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.50/+3.50/77 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2022. THE PATIENT EXPERIENCED BLURRED VISION. THEN ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS NOTED "THERE IS NO PLAN TO IMPLANT OTHER LENS". THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911552&pc=QCB | 2023826-2022-00964 | 2022-02-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 -7.0/+0.5/051 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911603&pc=QCB | 2023826-2022-00963 | 2022-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7, -6.0/+1.5/004 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - "ON (B)(6) 2022 THE LENS WAS REPLACED WITH A DIFFERENT MODEL/SAME LENGTH LENS. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: B5 - "ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS NO LONER REMAINS IMPLANTED. PATIENT IS HAPPY. THE PROBLEM IS RESOVED." CLAIM # (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911619&pc=QCB | 2023826-2022-00965 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -15.5/+2.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13932582&pc=QCB | 2023826-2022-00948 | 2021-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -16.0/5.0/012 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS REMOVED DUE TO REFRACTIVE SURPRISE. REPORTER STATES " IN THE FIRST POSTOPERATIVE DAY FOLLOWING PHAKIC LENS EXPLANTATION CORNEA AND CRYSTALLINE WERE CLEAR. IT IS EXPECTED THAT THE CORNEA DEFLATES AFTER THE EXTRACTION OF THE LENS TO TAKE MEASUREMENTS AGAIN AND REQUEST THE NEW REPLACEMENT LENS." CAUSE WAS REPORTED AS THE PATIENT RELATED FACTOR AND USER ERROR. "DUE TO PREVIOUS KERATOCONUS IT WAS POSSIBLE THAT THE PRE-OPERATIVE MANIFEST REFRACTION WAS NOT ACCURATE." Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13997876&pc=QCB | 2023826-2022-00836 | 2021-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -10.00/2.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME SIZE DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. SEE MFR. REP#2023826-2022-00838 FOR EXCHANGE LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998048&pc=QCB | 2023826-2022-00838 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.00/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS. THIS DID NOT RESOLVE REFRACTIVE SURPRISE AND IS UNDER FURTHER OBSERVATION. LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT-RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14006177&pc=QCB | 2023826-2022-01047 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.50/2.5/060 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 LENS REPOSITIONING WAS PERFORMED DUE TO REFRACTIVE SURPRISE AND ROTATION. IT WAS REPORTED THAT AT 1-DAY POST-REPOSITIONING THE PATIENT WAS FINE AND FELT BETTER. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061348&pc=QCB | 2023826-2022-01129 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND WITH RESIDUE/DEBRIS. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL ANF FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061353&pc=QCB | 2023826-2022-01135 | 2022-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 -14.5/+2.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SUPRISE, BLURRED VISION, AND LENS ROTATION ARE REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS REPOSITIONED. THE "PATIENT'S UCVA IS 20/20. OS MR IS PLANO/-0.75X180 FOR THE FIRST REVIEW." CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED EMAIL, "PATIENT CAME BACK ON (B)(6) 2022 AND REPORTED BLURRED VISION IN THE LEFT EYE. AFTER DILATION, DR LIM HAS NOTICED THAT THE LENS HAS YET AGAIN ROTATED OUT OF POSITION." CLAIM# (B)(4). Manufacturer Narrative: B5 - ADDITIONAL INFORMATION: THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED LATER ON THE SAME DAY WITH A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. "ICL ON AXIS AT 170". CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138859&pc=QCB | 2023826-2022-01232 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -12.0/6.0/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A NEW MODEL LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14145816&pc=QCB | 2023826-2022-01204 | 2022-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS -12.5/+4.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED, THE SURGEON REPORTING LENS ROTATION, REFRACTIVE SURPRISE, AND DROP OF UCVA. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "1ST ROTATION OF THE IOL (20DEGREES) AND RE-ROTATION AFTER REPOSITIONING, NO CORRECTION OF THE ASTIGMATISM; POSTOP REFRACTION=-4X40DEGREES." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14164434&pc=QCB | 2023826-2022-01272 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: RESIDUAL ASTIGMATISM 2 MONTHS POST-OP ON A PATIENT WITH TORIC IMPLANTABLE COLLAMER LENS WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER-RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. DEVICE MANUFACTURING DATE: NO SERIAL NUMBERS REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14164530&pc=QCB | 2023826-2022-01271 | 2022-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6) 2022 LASIK ENHANCEMENT WAS PERFORMED. ON (B)(6) 2022 THE LENS ROTATED AGAIN, IT WAS REPORTED THAT LENS REPOSITIONING WAS PERFORMED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME LENGTH DIFFERENT MODEL AND POWER LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION PROBLEM HAS BEEN RESOLVED . CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190875&pc=QCB | 2023826-2022-01295 | 2021-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -9.5/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B3- DATE OF EVENT- (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B3-DATE OF EVENT: 03-DEC-2021. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICTY:UNKNOWN, RACE:UNKNOWN. DATE OF EVENT:UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190880&pc=QCB | 2023826-2022-01291 | 2021-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) IN (B)(6) 2021. THE SURGEON REPORTS BLURRED VISION AND LENS ROTATION. REPORTEDLY THE LENS WAS REPOSITIONED AND IT ROTATED AGAIN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT IDENTIFIER: UNKNOWN, WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN. DATE OF EVENT: UNKNOWN. DEVICE MANUFACTURE DATE: UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPORTER STATED THAT A 13.7MM, VTICMO13.7, -14.5/+3.0/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, BLURRED VISION AND INDUCED ASTIGMATISM. THE REPORTER STATED THAT THE PROBLEM RESOLVED BUT AFTER THREE MONTHS THE LENS ROTATED AGAIN. ON (B)(6) 2022 LASIK SURGERY WAS PERFORMED TO CORRECT THE ASTIGMATISM. IT WAS REPORTED THAT THE PROBLEM RESOLVED, EYE IS STABLE POST LASIK ENHANCEMENT. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14197862&pc=QCB | 2023826-2022-01309 | 2022-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6; -12.0/4.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - UPDATED DEVICE EVALUATION: LENS WAS RETURNED IN AN EYE DROP CONTRAINER IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A EYE DROP CONTAINER IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14274820&pc=QCB | 2023826-2022-01393 | 2022-03-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/+1.0/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. DEVICE AND PATIENT-RELATED FACTOR WAS REPORTED AS CAUSE OF EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14274990&pc=QCB | 2023826-2022-01392 | 2022-03-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.5/+1.5/157 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. DEVICE AND PATIENT-RELATED FACTOR WAS REPORTED AS CAUSE OF EVENT. THE REPORTER STATED "REGARDING THE REFRACTIVE ERROR, HER RIGHT EYE REFRACTIONS WERE A BIT UNRELIABLE PRE-OP AS MY OPTOMETRIST NOTED HER RIGHT EYE WAS ACCOMMODATING VERY STRONGLY." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14288001&pc=QCB | 2023826-2022-01302 | 2022-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+1.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE LENS ROTATED. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED ON (B)(6) 2022 WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14311898&pc=QCB | 2023826-2022-01242 | 2022-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 -17.5/2.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A NEW MODEL LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3:DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISABLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342556&pc=QCB | 2023826-2022-01475 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -3.5/+1.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT NOW HAS LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS DID NOT RESOLVE THE ISSUE AS THE SURGEON NOTED "UCVA IS VERY GOOD; VAULTING STILL VERY LOW. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND DO A CLOSE MONITORING". CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342603&pc=QCB | 2023826-2022-01474 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -3.50/+1.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT NOW HAS LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS DID NOT RESOLVE THE ISSUE AS THE SURGEON NOTED "UCVA IS VERY GOOD; VAULTING STILL VERY LOW. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND DO A CLOSE MONITORING". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439265&pc=QCB | 2023826-2022-01369 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -6.50/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440972&pc=QCB | 2023826-2022-01603 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: BS THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 31-DEC-2021. ON 26-MAR-2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14466440&pc=QCB | 2023826-2022-01379 | 2022-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -8.50/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -8.50/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LENS OF DIFFERENT POWER. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -8.50/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LENS OF DIFFERENT POWER. CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VIISBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14493479&pc=QCB | 2023826-2022-01685 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -2.00/4.5/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14493491&pc=QCB | 2023826-2022-01686 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -1.50/4.00/68 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS OTHER. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14523289&pc=QCB | 2023826-2022-01699 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -15.0/2.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT AND BLURRED VISION. IT WAS REPORTED THAT BLURRY VISION WAS STILL PRESENT AFTER IMPLANTATION OF THIS REPLACEMENT LENS. THE REPORTER STATED, "WILL CONTINUE TO MONITOR". SEE MFR # 2023826-2022-01316 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14525830&pc=QCB | 2023826-2022-01545 | 2022-04-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH13.2 +1.5/6.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE AND LENS ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14525956&pc=QCB | 2023826-2022-01546 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH13.2 +3.0/6.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE AND LENS ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14570417&pc=QCB | 2023826-2022-01761 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/1.5/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND ON (B)(6) 2022; THE SURGEON EXCHANGED THE LENS WITH A SAME LENGTH/MODEL BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. IT WAS NOTED "DUE TO UNDERCORRECTED 100 DEGREE MYOPIA; HIS VISION HAS BEEN POOR AFTER THE OPERATION; AND HE HAS NOT IMPROVED AFTER VISUAL FUNCTION TRAINING. HE REPLACED THE LENS ON (B)(6) 2022 AND HIS VISION REACHED 1.0". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 -DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH"FIBER" ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITIONAL INFORMATION: H6- TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14605631&pc=QCB | 2023826-2022-01800 | 2021-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -1.00/+2.50/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "APPEARANCE IS THAT THE IMPLANTED ICL DID NOT CORRESPOND TO THE PATIENT'S RX". Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, AND DRY. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, THE HAPTIC BROKEN,BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATION. WIDTH AND FUNCTIONAL INSPECTION COULD NOT BE PERFORMED DUE TO HAPTIC/OPTIC DAMAGE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14650054&pc=QCB | 2023826-2022-01838 | 2021-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -3.0/1.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14650083&pc=QCB | 2023826-2022-01837 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.50/+5.0/127 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14706977&pc=QCB | 2023826-2022-01750 | 2004-03-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.5MM TICM12.5 -13.0/1.2/064 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2004. REFRACTIVE CHANGE OVERTIME WAS REPORTED. THE LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: AGE OR DATE OF BIRTH, WEIGHT-RACE:UNK. DATE OF EVENT: UNK CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14707086&pc=QCB | 2023826-2022-01888 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -5.50/+5.50/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2022, THE SURGEON REPOSITIONED THE LENS. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: . CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14707407&pc=QCB | 2023826-2022-01863 | 2021-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.0/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WITH ROTATION WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS THE DEVICE. FOR CAUSE DETAILS THE REPORTER STATED " THE TORIC LENS IS CONTINUOUSLY ROTATING WITHIN 5 MONTH FOLLOW-UP INTERVAL, OVER 330 DEGREES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862313&pc=QCB | 2023826-2022-02092 | 2022-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.5/2.5/170 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862676&pc=QCB | 2023826-2022-01908 | 2022-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 -5.5/1.0/066 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DAY IN A SEPARATE SURGERY A BACKUP LENS OF DIFFERENT AXIS WAS IMPLANTED. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: PER ADDITIONAL INFORMATION, "UNDESIRABLE REFRACTIVE OUTCOME" AFTER THE LENS WAS EXCHANGED. REFRACTIVE SURPRISE WAS ALSO REPORTED. PROBLEM WAS NOT RESOLVED. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14885725&pc=QCB | 2023826-2022-02095 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/54 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. ON (B)(6) 2022 THE LENS WAS REPLACED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1: SHOULD BE CORRECTED TO: ADVERSE EVENT. B2: SHOULD BE CORRECTED TO: REQUIRED INTERVENTION PREVENT PERMANENT IMPAIRMENT/DAMAGE. H1: SHOULD BE CORRECTED TO: SERIOUS INJURY. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE LENS WAS REMOVED AND INTRAOPERATIVELY REPLACED ON (B)(6) 2022. D6B: ON (B)(6) 2022. CORRECTED DATA: H6: 4627 SHOULD BE OMITTED AND REPLACED WITH 4629 FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE:23-MAR-2023 H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN DEVICE SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15010429&pc=QCB | 2023826-2022-02233 | 2020-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.5/0.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED, HAPTIC BENT AND STRETCHED OUT TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15017540&pc=QCB | 2023826-2022-02026 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM TICM12.0 -15.0/2.0/122 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. REFRACTIVE CHANGE OVERTIME AND PROGRESSION OF MYOPIA WAS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15022376&pc=QCB | 2023826-2022-02270 | 2021-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7 VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT POWER. THE EXCHANGE LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA; H10 IN PREVIOUS MDR CLAIM NUMBER (B)(4), SHOULD BE CORRECTED TO CLAIM NUMBER (B)(4). DEVICE EVALUATION: H6-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISYURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15097257&pc=QCB | 2023826-2022-02279 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/3.5/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE INITIAL LENS WAS EXPLANTED AND A REPLACEMENT LENS OF SAME MODEL/SIZE AND DIFFERENT POWER WAS IMPLANTED DUE TO REFRACTIVE SURPRISE. THE CAUSE IS REPORTED AS DEVICE "MISMARKED LENS RESULTED IN REFRACTIVE SURPRISE." Manufacturer Narrative: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15102877&pc=QCB | 2023826-2022-02436 | 2021-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -17.0/1.5/119 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(6). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15103623&pc=QCB | 2023826-2022-02448 | 2022-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -10.5/1.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND DIFFERENT POWER LENS DUE TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATIION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIMN# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118601&pc=QCB | 2023826-2022-02443 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2M VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.0/+1.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 . THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CASUSE OF THE EVENT WAS USER ERROR. FOR CAUSE DETAILS THE REPORTER STATED " SUMM UP OF MANY FACTORES: ERROR IN REFRACTION (CYCLOPLEGIC REFRACTION WAS NOT NOTICED), INTRAOPERATIVELY INDUCED CYLINDER AND CYLINDRIC LENSPOWER GAIND A RESIDUAL REFRACTIVE ERROR." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2022. THE LENS WAS REPLACED LATER IN ANOTHER SURGERY ON (B)(6) 2022 WITH A SAME LENGTH DIFFERENT DIOPTER LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT IS HAPPY BUT DISTANT VISION HAS YET TO STABILIZE." CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118602&pc=QCB | 2023826-2022-02445 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6M VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT CYLINDER AND AXIS LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN (FIBERS)ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118603&pc=QCB | 2023826-2022-02444 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6M VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/077 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT CYLINDER AND AXIS LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED AND THE HAPTIC BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: H3: FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153360&pc=QCB | 2023826-2022-02581 | 2022-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS OF LONGER LENGTH DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. BLOOD STAINS WERE OBSERVED ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15180105&pc=QCB | 2023826-2022-02630 | 2021-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/98 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY A DIFFERENT SURGERY A REPLACEMENT LENS OF THE SAME MODEL; SIZE LENS BUT DIFFERENT POWER LENS WAS IMPLANTED DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15180107&pc=QCB | 2023826-2022-02629 | 2021-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/88 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY A DIFFERENT SURGERY A REPLACEMENT LENS OF THE SAME MODEL; SIZE LENS BUT DIFFERENT POWER LENS WAS IMPLANTED DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#(B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 -DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15204034&pc=QCB | 2023826-2022-02701 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/4.00/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND UNDESIRABLE REFRACTIVE OUTCOME WAS REPORTER, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5: LENS WAS REPOSITIONED ON (B)(6) 2022 AND PROBLEM WAS RESOLVED. STATUS OF THE EYE IS "QUIET". DEVICE DID NOT FAIL TO PERFORM. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15204622&pc=QCB | 2023826-2022-02627 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -8.00/2.50/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022; AS A REPLACEMENT LENS TO A PREVIOUS LOW VAULT WITH ROTATION ISSUE. THE PATIENT NOW HAS REFRACTIVE SURPRISE; AND REFRACTIVE CHANGE OVER TIME. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: "THE LENS WAS EXPLANTED ON AN UNKNOWN DATE" SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 12-APR-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15211435&pc=QCB | 2023826-2022-02702 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED LENS ROTATION AND LOSS OF CORRECTION SHORTLY AFTER ICL SURGERY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. DEVICE MANUFACTURING DATE: NO SERIAL NUMBERS REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15211436&pc=QCB | 2023826-2022-02703 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED LENS ROTATION AND LOSS OF CORRECTION SHORTLY AFTER ICL SURGERY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNKNOWN, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNKNOWN. DEVICE MANUFACTURING DATE: NO SERIAL NUMBERS REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15217420&pc=QCB | 2023826-2022-02695 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.5/+2.0/134 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR DIFFERENT POWER LENS DUE TO THE DOCTOR SELECTING THE WRONG LENS POWER. THE EXCHANGE RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS USER ERROR AND THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED . Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15227610&pc=QCB | 2023826-2022-02783 | 2022-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.5/3.86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND OPTIC TORN ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.50/3.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022." CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15276388&pc=QCB | 2023826-2022-02826 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: D4. MODEL # VTICMO12.6, EXPIRATION DATE:(B)(6) 2023. Manufacturer Narrative: CORRECTED DATA: H4: DEVICE MANUFACTURE DATE: 02-SEP-2020. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299051&pc=QCB | 2023826-2022-02919 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -7.5/1.5/168 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 21-OCT-2022 . H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN BROKEN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN DEVICE SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: EXPIRATION DATE: 12-31-2024. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299255&pc=QCB | 2023826-2022-02530 | 2022-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH LENS ROTATION AND BLURRED VISION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299617&pc=QCB | 2023826-2022-02925 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/161 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT . A REPLACEMENT LENS OF SAME LENGTH, DIFFERENT MODEL AND POWER WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. THE REPORTER STATED " TOO MUCH VAULT AND ASTIGMATISM CORRECTION IN THE LENS. WITH THE SECOND LENS IMPLANTED WITHOUT ASTIGMATISM CORRECTION THE FINAL REFRACTION WAS CLOSE TO EMETROPY. VAULT ALSO WAS BETTER WITH THE SECOND IMPLANT." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 02/28/2025. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299656&pc=QCB | 2023826-2022-02927 | 2022-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/076 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 12/31/2024. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325099&pc=QCB | 2023826-2022-02977 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.0/1.5/065 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 03-31-2025. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325103&pc=QCB | 2023826-2022-02976 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.0/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 01-31-2025. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15342016&pc=QCB | 2023826-2022-02969 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -10.50/0.5/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. ON (B)(6) 2022, REFRACTIVE CHANGE OVER TIME WITH LOW VAULT WAS RECORDED. A LENS EXHANGE IS PLANNED. LENS REMAINS IMPLANTED. Manufacturer Narrative: B1 - CORRECTED TO PRODUCT PROBLEM (E.G. DEFECTS/MALFUNCTIONS) IN INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE REMOVED FROM INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B1 CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/ LONGER LENS DUE TO REFRACTIVE CHANGE OVERTIME AND LOW VAULT. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 HEALTH EFFECT IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15342195&pc=QCB | 2023826-2022-02970 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2022, REFRACTIVE CHANGE OVER TIME WITH LOW VAULT WAS RECORDED. A LENS EXHANGE IS PLANNED. LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO PRODUCT PROBLEM (E.G. DEFECTS/MALFUNCTIONS) IN INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE REMOVED FROM INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR . B5 "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/ LONGER LENS DUE TO REFRACTIVE CHANGE OVERTIME AND LOW VAULT. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 HEALTH EFFECT IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15385938&pc=QCB | 2023826-2022-03048 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.00/5.0/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 02-28-2025. (B)(4). Manufacturer Narrative: B1: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5: "ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE." SHOULD BE ADDED TO THE INITIAL MDR. D6B: EXPLANT DATE: ON (B)(6) 2022, SHOULD BE ADDED TO INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6: HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15407358&pc=QCB | 2023826-2022-03069 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -1.5/1.5/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED; ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE ON THE LENS SPECIFICALLY FIBRE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423219&pc=QCB | 2023826-2022-03072 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/4.0/58 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED AND BLURRED VISION.. THE VAULT WAS GOOD BUT THE LENS ROTATED OFF AXIS AND AFFECTED THE REFRACTION. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF LONGER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS OTHER- LENS ROTATED DUE TO SMALL LENS. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED TO - "THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF 11.0/4.0/58 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED AND BLURRED VISION.. THE VAULT WAS GOOD BUT THE LENS ROTATED OFF AXIS AND AFFECTED THE REFRACTION. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022, THE LENS WAS REMOVED AND ON SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF LONGER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS OTHER- LENS ROTATED DUE TO SMALL LENS. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. D4: CORRECTED TO MODEL#: VTICM5 12.6 IN THE INITIAL MDR. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423220&pc=QCB | 2023826-2022-03071 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.5/4.5/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED AND BLURRED VISION. THE VAULT WAS GOOD BUT THE LENS ROTATED OFF AXIS AND AFFECTED THE REFRACTION. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF LONGER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS OTHER- LENS ROTATED DUE TO SMALL LENS. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423292&pc=QCB | 2023826-2022-03053 | 2022-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS REMOVED AND IN A SUBSEQUENT SURGERY LATER THAT DAY, A LONGER LENGTH LENS WAS IMPLANTED DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. REPORTEDLY, "AFTER THE FIRST OPERATON, THE VAULT OF THE LENS WAS LOW , AND THE SIZE OF THE REPLACEMENT LENS WAS CHANGED FROM SMALL (12.1MM) TO LARGE (12.6MM). CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423498&pc=QCB | 2023826-2022-03054 | 2022-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS REMOVED AND IN A SUBSEQUENT SURGERY LATER THAT DAY, A LONGER LENGTH LENS WAS IMPLANTED DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. REPORTEDLY, "AFTER THE FIRST OPERATON, THE VAULT OF THE LENS WAS LOW , AND THE SIZE OF THE REPLACEMENT LENS WAS CHANGED FROM SMALL (12.1MM) TO LARGE (12.6MM). CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN AND FIBER ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15441785&pc=QCB | 2023826-2022-03151 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.5/2.0/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A REPLACEMENT LENS OF DIFFERENT POWER WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF LABEL USE (ACD<3.0MM) AND (AGE<21). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15459789&pc=QCB | 2023826-2022-03239 | 2022-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.50/+3.0/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B1: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5: "ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS DUE TO REFRACTIVE SURPRISE." SHOULD BE ADDED TO THE INITIAL MDR. D6B: EXPLANT DATE: ON (B)(6) 2022, SHOULD BE ADDED TO INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO INITIAL MDR. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498967&pc=QCB | 2023826-2022-03371 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -13.0/1.5/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498988&pc=QCB | 2023826-2022-03369 | 2022-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -2.50/+2.5/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15499015&pc=QCB | 2023826-2022-03370 | 2022-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -4.00/+4.0/174 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15508658&pc=QCB | 2023826-2022-03344 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.5/3.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE LENS HAPTIC BENT AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15524739&pc=QCB | 2023826-2022-03366 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -15.0/5.0/59 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. REPORTEDLY, "THE DOCTOR PROBABLY DID NOT POSITION LENS CORRECTLY." CAUSE OF THE EVENT WAS USER ERROR. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS EXPLANTED ON 12 DEC 2022 AND PROBLEM WAS RESOLVED. PATIENT IS HAPPY. CLAIM#(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15534242&pc=QCB | 2023826-2022-03428 | 2022-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED, BUT WAS NOT RESOLVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "TWO WEEKS AFTER THE OPERATION, THE PATIENT FOUND THAT THE LENS OF HIS RIGHT EYE WAS ROTATED, AND THE VISUAL ACUITY DECREASED TO 0.6. AFTER CONSULTATION BY THE OPERATING DOCTOR, HE DECIDED TO PERFORM THE ADJUSTMENT OPERATION. HOWEVER, IT WAS FOUND THAT THE LENS ROTATED GAIN 3 MONTHS AFTER THE ALIGNMENT OPERATION, AND THE VISUAL ACUITY DECREASED TO 0.6.0. AFTER CONSULTATION AND COMMUNICATION WITH THE PATIENT, THE DOCTOR DECIDED TO CUSTOMIZE THE LENS WITH SPECIAL AXIAL POSITION. HOWEVER, DUE TO THE LONG TIME OF CRYSTAL CUSTOMIZATION AND THE EPIDEMIC SITUATION, THE PATIENT UNDERWENT THE RIGHT EYE TICL LENS REPLACMENT ON AUGEST 19." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS, CLAIM# (B)(4) 4B LOW VAULT, CLAIM#(B)(4)4A EXCESSIVE VAULT. CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED, BUT WAS NOT RESOLVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "TWO WEEKS AFTER THE OPERATION, THE PATIENT FOUND THAT THE LENS OF HIS RIGHT EYE WAS ROTATED, AND THE VISUAL ACUITY DECREASED TO 0.6. AFTER CONSULTATION BY THE OPERATING DOCTOR, HE DECIDED TO PERFORM THE ADJUSTMENT OPERATION. HOWEVER, IT WAS FOUND THAT THE LENS ROTATED GAIN 3 MONTHS AFTER THE ALIGNMENT OPERATION, AND THE VISUAL ACUITY DECREASED TO 0.6.0. AFTER CONSULTATION AND COMMUNICATION WITH THE PATIENT, THE DOCTOR DECIDED TO CUSTOMIZE THE LENS WITH SPECIAL AXIAL POSITION. HOWEVER, DUE TO THE LONG TIME OF CRYSTAL CUSTOMIZATION AND THE EPIDEMIC SITUATION, THE PATIENT UNDERWENT THE RIGHT EYE TICL LENS REPLACEMENT ON (B)(6)." CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15534408&pc=QCB | 2023826-2022-03427 | 2022-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITION AND WAS NOT RESOLVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL, SAME SIZE LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "ON (B)(6), THE PATIENT IMPLANTED TICL CRYSTAL IN HIS LEFT EYE. AFTER THE OPERATION, IT WAS FOUND THAT THE CRYSTAL WAS ROTATED 36 DEGREES ANTICLOCKWISE. AFTER THE FACE-TO-FACE CONSULTATION OF THE OPERATIVE DOCTOR, THE PATIENT DECIDED TO PERFORM TICL CRYSTAL ADJUSTMENT. AFTER FACE-TO-FACE REVIEW, IT WAS FOUND THAT THE CRYSTAL WAS ROTATED AGAIN. AFTER THE FACE-TO-FACE CONSULTATION, THE OPERATING DOCTOR COMMUNICATED WITH THE PATIENT AND FINALLY DECIDED TO REPLACE THE ICL CRYSTAL. ASTIGMATISM WAS CORRECTED BY LASER SURGERY." Manufacturer Narrative: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15557486&pc=QCB | 2023826-2022-03452 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON OBSERVED REFRACTIVE CHANGE OVER TIME IN THE PATIENT'S LEFT EYE (OS). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK CLAIM# (B)(4). Manufacturer Narrative: B1: PRODUCT PROBLEM. B2: ENTRY SHOULD BE LEFT BLANK . CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS WITH A -3.0/05/082 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT ON (B)(6) 2022. ON (B)(6) 2022, REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "MONOVISION ON THE LEFT EYE WAS PLANNED. THE RESULT IS NOT ACHIEVED WITH THE VALUE -9.25-9.50/88. WE ASK FOR EXCHANGE. TARGET 1.25." H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND WITH ASSOCIATED LOTS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS WITH A -3.0/0.5/082 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. RETURN DATE- 12-APR-2023 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL/RESIDUE/DEBRIS ON LENS SURFACE, SPECIFICALLY FIBERS. NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15609430&pc=QCB | 2023826-2022-03642 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -18.0/6.0/80 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "THE FIRST TIME, AN INCORRECT LENS WAS IMPLANTED, THE REASON WAS THAT THE CALCULATION WERE NOT DONE CORRECTLY." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15629399&pc=QCB | 2023826-2022-03705 | 2022-09-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.5/+3.0/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H1: CORRECTED TO - MALFUNCTION IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15650328&pc=QCB | 2023826-2022-03704 | 2022-09-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.00/+2.5/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D4 - CORRECTED TO : MODEL # VTICM5 13.2 IN THE INTIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15651860&pc=QCB | 2023826-2022-03636 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.50/1.0/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS REPLACED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: PRODUCT EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H2: DEVICE EVALUATION SHOULD HAVE BEEN SELECTED IN PREVIOUS SUPPLEMENTAL. H3: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND RESIDUE/DEBRIS ON THE LENS - SHOULD HAVE ADDITIONALLY BEEN INCLUDED IN PREVIOUSLY SUPPLEMENTAL. ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15651919&pc=QCB | 2023826-2022-03633 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: IN THE ARTICAL, "IMPLANTABLE COLLAMER LENS SIZING BASED ON MEASUREMENT OF THE SULCUS-TO-SULCUS DISTANCE IN ULTRASOUND BIOMICROSCOPY VIDEO CLIPS AND ZZ ICL FORMULA," A STUDY TO EVALUATE A NEW METHOD OF IMPLANTABLE COLLAMER LENS (ICL) SIZING BASED ON ULTRASOUND BIOMICROSCOPY (UBM) VIDEO CLIPS. THE STUDY INCLUDED CONSECUTIVE PATIENTS WITH MYOPIA AND MYOPIC ASTIGMATISM SCHEDULED FOR V4C TORIC ICL (TICL) IMPLANTATION FROM OCT 2020 TO NOV-2020. THE ANALYSIS INCLUDED 168 EYES IN 84 PATIENTS. POSTOPERATIVE VAULT SIE WAS COMPARABLE TO THAT PREDICTED BY THE ZZ ICL FORMULA. VAULT PREDICTION ERROR (PE) BY THE ZZ ICL FORMULA WAS WITHIN 100, 300, AND 500 UM IN 40.48%, 88.10%, AND 100% OF EYES. SPHERICAL EQUIVALENT (SE) AND ABSOLUTE CYLINDRICAL REFRACTIVE ERROR WERE .36+/- 0.48 AND 0.40 +/- 0.31 AT 3 MONTHS POSTOPERATIVELY. THE SE PE, ABSOLUTE CYLINDERICAL PE, AND PERCENTAGES OF EYES WITH AN ABSOULTE CYLINDERICAL PE WITHIN +/- 0.50 D AND +/- 1.00 D WERE LOWER FOR THE ZZ ICL FORMULA THAN FOR THE STAAR SOFTWARE (P < 0.01). THE STUDY CONCLUDED THAT COMBINING MEASUREMENTS OBTAINED IN UBM VIDEO CLIPS WITH THE ZZ ICL FORMULA PROVIDES AN EFFECTIVE METHOD OF SIZING TICLS AND PREDICTING VAULT HEIGHT AND RESIDUAL REFRACTIVE ERROR. Manufacturer Narrative: CORRECTED DATA: B1/B2 - CHANGED TO PRODUCT PROBLEM/MALFUNCTION, REQUIRED INTERVENTION TO PREVENT PERMANENT DAMAGE/IMPAIRMENT SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15697248&pc=QCB | 2023826-2022-03860 | 2022-02-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 2.6 IMPLANTABLE COLLAMER LENS OF -12.50/2.5/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. A MEDICAL CONSULT IS BEING PROVIDED TO GIVE GUIDANCE ON HOW TO CORRECT THE PATIENT'S VISION. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15727569&pc=QCB | 2023826-2022-03976 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2; -6.50/2.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. . Manufacturer Narrative: B5 - THE LENS WAS REPLACED ON (B)(6) 2022 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15728310&pc=QCB | 2023826-2022-03975 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.50/2.0/103 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE LENS WAS REPLACED ON (B)(6) 2022 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15782580&pc=QCB | 2023826-2022-03959 | 2022-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -1.00/4.0/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. CAUSE IS REPORTED AS "USER ERROR." REPORTEDLY, "AWAITING ADVICE FROM MEDICAL AFFERS". Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE: UNK. PATIENT RELATED FACTOR (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- "ADVERSE EVENT" SHOULD BE CORRECTED TO "PRODUCT PROBLEM". H6- MEDICAL DEVICE PROBLEM CODE- REMOVE CODE: 2682. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.00/+4.00/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY "AWAITING ADVICE FROM MEDICAL AFFAIRS" AND "AN 10º ERROR WAS MADE WHEN PUTTING THE SUBJECTIVE REFRACTION IN OCOS." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS USER ERROR AND ITS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15842667&pc=QCB | 2023826-2022-04287 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -7.00/1.50/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 13-JAN-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15843892&pc=QCB | 2023826-2022-04289 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15850084&pc=QCB | 2023826-2022-04265 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: D4- UNIQUE IDENTIFIER (UDI) #: (B)(4). ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE" AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15850228&pc=QCB | 2023826-2022-04267 | 2022-10-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -7.50/0.50/80 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTEDLY PROBLEM RESOLVED. THE PATIENT DOES NOT WANT A TREATMENT AND THE PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15854790&pc=QCB | 2023826-2022-04293 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00/+3.00/003 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/3.00/003 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE REPLACEMENT LENS HAD "LOSS OF STERILIZATION", AND WAS NOT IMPLANTED. H6- MEDICAL DEVICE CODE: "3189" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. D6B. "12/13/2022" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15874950&pc=QCB | 2023826-2022-04304 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.5/1.5/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5 - "ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH/ MODEL LENS DUE TO REFRACTIVE CHANGE OVERTIME. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2023 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 2923 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15875543&pc=QCB | 2023826-2022-04305 | 2021-01-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.5/1.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021 . THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - CORRECTED TO: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15891930&pc=QCB | 2023826-2022-04367 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.50/0.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF 4.50/0.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL SAME LENGTH LENS. THE PROBLEM RESOLVED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15912687&pc=QCB | 2023826-2022-04282 | 2022-11-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.00/1.0/069 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975198&pc=QCB | 2023826-2022-04416 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/5.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975200&pc=QCB | 2023826-2022-04414 | 2022-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -14.00/4.5/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5: CORRECTED TO : "THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -14.00/4.5/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022. THE LENS REMAINS IMPLANTED." SHOULD BE ADDED TO THE INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6: HEALTH EFFECT - IMPACT CODE: 4628 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: INVESTIGATION TYPE (B)(4). NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15990827&pc=QCB | 2023826-2022-04591 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -18.00/2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM SO THE LENS WAS EXPLANTED. A SPHERICAL LENS OF LONGER LENGTH WAS LATER IMPLANTED, AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT LASIK SURGERY WAS NECESSARY DUE TO EXISTING RESIDUAL ASTIGMATISM. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15991920&pc=QCB | 2023826-2022-04417 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/4.5/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000596&pc=QCB | 2023826-2022-04599 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -10.5/3.0/038 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. CAUSE OF EVENT WAS UNKNOWN. SEE MFR # 2023826-2022-02093 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16019046&pc=QCB | 2023826-2022-04708 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/+1.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON(B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "USER ERROR" AND NOTED "THERE IS A SUBJECTIVE OVERCORRECTION IN LENS POWER". Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SIMILAR LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16024296&pc=QCB | 2023826-2022-04693 | 2020-11-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED WITH THE REPLACEMENT LENS DUE TO A LENS TEAR/BREAK DURING INJECTION DELIVERY INTO THE EYE AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT/ ETHINICITY/ RACE:UNK WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: BASED ON CORRECTED ADDITIONAL INFORMATION FROM THE REPORTER, REFRACTIVE SURPRISE WAS NOT OBSERVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: 26-NOV-2022 SHOULD BE CORRECTED TO 26-NOV-2020. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030437&pc=QCB | 2023826-2022-04435 | 2020-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -7.5/1.5/173 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2020 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF SHORTER LENGTH AND DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030860&pc=QCB | 2023826-2022-04694 | 2021-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0/3.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTER AS, OTHER. Manufacturer Narrative: B5: THE CAUSE OF THE EVENT IS UNCERTAIN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030861&pc=QCB | 2023826-2022-04695 | 2021-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/2.5/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS OTHER. Manufacturer Narrative: B5: THE CAUSE OF THE EVENT IS UNCERTAIN. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16032380&pc=QCB | 2023826-2022-04646 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS A COMPLICATED REFRACTION AND REQUESTED GUIDANCE FROM A CONSULTATION DOCTOR FOR MANAGING THE RESIDUAL REFRACTIVE ERROR. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16034661&pc=QCB | 2023826-2022-04644 | 2021-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/72 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: D9: RETURN DATE: 13-JAN-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057495&pc=QCB | 2023826-2022-04733 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/1.0/70 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16064724&pc=QCB | 2023826-2022-04691 | 2021-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075117&pc=QCB | 2023826-2022-04773 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/1.0/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LENS ROTATION AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075678&pc=QCB | 2023826-2022-04747 | 2022-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS WITH A -6.0/1.0/089 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED AS A REPLACEMENT INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2022. REPORTEDLY, "LENS TOO BIG. SURGEON PREFERENCE CHANGE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. "BLURRY VISION" AND "VAULT SMALL" WAS ALSO REPORTED. Manufacturer Narrative: B3 - DATE OF EVENT: 20-MAY-2022 SHOULD BE ADDED TO THE INITIAL MDR. B5 - "THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO "SULCUS TOO SMALL FOR 12.1." AND BLURRY VISION. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE: (B)(6) 2022 SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO: SERIOUS INJURY IN THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075681&pc=QCB | 2023826-2022-04746 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS WITH A -6.0/1.0/107 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. REPORTEDLY, "LENS TOO BIG. SURGEON PREFERENCE CHANGE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. "BLURRY VISION" AND "VAULT SMALL" WAS ALSO REPORTED. Manufacturer Narrative: DEVICE EVALUATED BY MFR: OFF-LABEL: PRE-EXISTING PROGRESSIVE MYOPIA. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16077729&pc=QCB | 2023826-2022-04792 | 2022-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -6.5/1.5/11 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A LENS OF DIFFERENT POWER WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-FEB-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16099355&pc=QCB | 2023826-2023-00083 | 2022-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13/1.5/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "ERROR IN ASTIGMATISM DATA". Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16099943&pc=QCB | 2023826-2023-00078 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTES A POSSIBLE SECONDARY SURGICAL INTERVENTION DUE TO LENS ROTATION AND BLURRY VISION AND REQUESTED ASSISTANCE FROM A CONSULTATION DOCTOR. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANT DATE: (B)(6) 2022. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16099944&pc=QCB | 2023826-2023-00077 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTES A POSSIBLE SECONDARY SURGICAL INTERVENTION DUE TO LENS ROTATION AND BLURRY VISION AND REQUESTED ASSISTANCE FROM A CONSULTATION DOCTOR. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(6). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16112609&pc=QCB | 2023826-2023-00079 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -0.50/6.00/134 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. THE REPORTER STATED THAT THE EYE IS QUIET. IT WAS REPORTED THAT VISION IS 20/60 UNDILATED WITH A VAULT OF 660 AND 20/30 DILATED WITH A VAULT OF 442. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16113739&pc=QCB | 2023826-2023-00081 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/167 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16117774&pc=QCB | 2023826-2023-00054 | 2022-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.0/130 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "SURGEON USED WRONG DATAS FOR CALCULATION IN OCOS, DATA FOR OD USED FOR OS. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD HAVE BEEN PRODUCT PROBLEM IN THE ORIGINAL MDR SUBMISSION. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD HAVE BEEN OMITTED FROM THE ORIGINAL MDR SUBMISSION. B5: REFRACTIVE CHANGE OVER TIME SHOULD BE CORRECTED TO STATE REFRACTIVE SURPRISE. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.0/130 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT WAS USER ERROR. ON (B)(6) 2022 THE LENS WAS REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL AND DIOPTER LENS. REPORTEDLY, "SURGEON USED WRONG DATA FOR CALCULATION IN OCOS, DATA FOR OD USED FOR OS. D6B: EXPLANT DATE: ON (B)(6) 2023. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16117832&pc=QCB | 2023826-2023-00053 | 2022-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.5/78 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "SURGEON USED WRONG DATAS FOR CALCULATION IN OCOS, DATA FOR OS USED FOR OD. Manufacturer Narrative: B1: ADVERSE EVENT SHOULD HAVE BEEN PRODUCT PROBLEM IN THE ORIGINAL MDR SUBMISSION. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD HAVE BEEN OMITTED FROM THE ORIGINAL MDR SUBMISSION. B5: REFRACTIVE CHANGE OVER TIME SHOULD BE CORRECTED TO STATE REFRACTIVE SURPRISE. H6: OMIT 1494 FOR CORRECTION. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.5/78 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPLACED WITH A SAME MODEL AND LENGTH BUT DIFFERENT DIOPTER LENS. ON (B)(6) 2022, THE LENS WAS REPLACED WITH A SAME MODEL AND LENGTH BUT DIFFERENT DIOPTER LENS. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "SURGEON USED WRONG DATA FOR CALCULATION IN OCOS, DATA FOR OS USED FOR OD. D6B: EXPLANT DATE: ON (B)(6) 2023. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16141698&pc=QCB | 2023826-2023-00120 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICM513.2, -7.50/3.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITONAL INFORMATION: B5- ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE" AND "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16142821&pc=QCB | 2023826-2023-00123 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -3.50/1.0/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY SIZE BUT A DIFFERENT MODEL AND POWER LENS WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: D4 SERIAL # - CORRECTED TO: (B)(6) IN THE INITIAL MDR. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16142822&pc=QCB | 2023826-2023-00124 | 2022-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.0/1.50/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED. ON A LATER DATE A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9:29-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER SHOULD BE OMITTED FOR CORRECTION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153694&pc=QCB | 2023826-2023-00107 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -3.0/2.0/024 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT WAS USER UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -3.0/2.0/024 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPOSITIONED; THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16154820&pc=QCB | 2023826-2023-00130 | 2022-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.50/2.00/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9:29-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16189407&pc=QCB | 2023826-2023-00173 | 2022-10-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.50/+3.50/97 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED THAT CALCULATION IS BEING DONE IN PLAN TO REPOSITION THE LENS. LENS REMAINS IMPLANTED. THE PROBLEM IS REPORTED AS NOT RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16189410&pc=QCB | 2023826-2023-00199 | 2022-10-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/3.50/64 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE LEFT RIGHT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED THAT CALCULATION IS BEING DONE IN PLAN TO REPOSITION THE LENS. LENS REMAINS IMPLANTED. THE PROBLEM IS REPORTED AS NOT RESOLVED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16199991&pc=QCB | 2023826-2023-00202 | 2022-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -4.50/1.5/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT - IMPACT CODE 4628. CLAIM (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16199992&pc=QCB | 2023826-2023-00203 | 2022-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM13.7 IMPLANTABLE COLLAMER LENS OF -13.50/4.00/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. PER UPDATED INFORMATION THE LENS WAS ROTATED ON (B)(6) 2022 AND THE PATIENT NOW HAS 1.2 VA. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED AFTER THE LENS ROTATION. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16200368&pc=QCB | 2023826-2023-00109 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A TORIC IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "HER OS EYE IS NOT AS CLEAR AS THE OD EYE." ALSO, "IT WAS POSSIBLE THAT IT ROTATED." Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16210693&pc=QCB | 2023826-2023-00244 | 2022-11-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -1/+1/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THE PATIENT HAS REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED "UDVA IS NOT SUFFICIENT DUE TO RESIDUAL MYOPIA AND ASTIGMATISM. WE SHOULD REPLACE THE ICL". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16214291&pc=QCB | 2023826-2023-00204 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -1.50/2.0/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON(B)(6)2022. REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "2993" SHOULD BE ADDED. D6B. "(B)(6) 2023" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_ 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.50/2.0/174(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY "ALL FINE!" AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16219689&pc=QCB | 2023826-2023-00251 | 2022-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.00/2.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED.THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS OTHER. DEVICE DID NOT FAULT O PERFORM AS INTENDED. "THERE WAS A MISUNDERSTANDING ABOUT ALIGNMENT." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16248972&pc=QCB | 2023826-2023-00337 | 2022-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE DID NOT RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FIBRE ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16249317&pc=QCB | 2023826-2023-00338 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.5/3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS REFRACTIVE SURPRISE BUT IT DID NOT RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2023-00337 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/3.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE; BUT THAT DID NOT RESOLVE THE PROBLEM. THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16297884&pc=QCB | 2023826-2023-00435 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -2.50/5.5/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION AND BAD VISION. ON (B)(6) 2023 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: EXPLANT DATE: UNKNOWN. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16300819&pc=QCB | 2023826-2023-00440 | 2022-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.00/1.50/133 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301296&pc=QCB | 2023826-2023-00442 | 2021-10-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16328105&pc=QCB | 2023826-2023-00380 | 2022-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -15.0/3.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, IMPLANTED VERTICALLY DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16343807&pc=QCB | 2023826-2023-00390 | 2023-01-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.00/4.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16343985&pc=QCB | 2023826-2023-00391 | 2023-01-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM 12.6 IMPLANTABLE COLLAMER LENS OF -10.50/1.50/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. LOW VAULT WITH REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM VTICM 12.6 IMPLANTABLE COLLAMER LENS OF -10.50/1.50/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. LOW VAULT WAS OBSERVED.THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED." IN THE INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16344319&pc=QCB | 2023826-2023-00402 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -4.5/2.0/83 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: UNK. OFF-LABEL - AGE<21. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16359096&pc=QCB | 2023826-2023-00487 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MMVTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/87(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD)ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MMVTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/87(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD)ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SIMILAR MAKE AND MODEL BUT DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE-12-APR-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376165&pc=QCB | 2023826-2023-00563 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/1.0/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DIFFERENT POWER; THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16385392&pc=QCB | 2023826-2023-00561 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT POWER DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: ADVERSE EVENT PROBLEM - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16390126&pc=QCB | 2023826-2023-00591 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN/BROKEN TO THE LENS. REHYDRATION OF A LENS HAS CONFIRMED INADEQUACY OF PERFORMING DIM/REF INSPECTION DUE TO SIGNIFICANT DAMAGE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16390268&pc=QCB | 2023826-2023-00592 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: THE DHR REVIEW INDICATES THAT THE PRODUCT HAS BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS AND THAT THERE IS NO INDICATION THAT THE MANUFACTURING AND/OR PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16406313&pc=QCB | 2023826-2023-00599 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020 AS AND EXCHANGE LENS TO RESOLVE LOW VAULT. TRANSIENT LOSS OF BCVA WAS OBSERVED. LENS REMAINS IMPLANTED. "A LASIK TOUCH UP IS PLANNED. SEE MDR # (B)(4) FOR INITIAL LENS. Manufacturer Narrative: B5: CORRECTED TO "THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS AN EXCHANGE LENS TO RESOLVE LOW VAULT. TRANSIENT LOSS OF BCVA WAS OBSERVED. LENS REMAINS IMPLANTED. "A LASIK TOUCH UP IS PLANNED." IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFOMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/1.5/153 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS. TRANSIENT LOSS OF BCVA WAS OBSERVED. ON (B)(6) 2023 LASIK WAS PERFORMED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "PATIENT IS HAPPY". THE REPORTER MARKED THE CAUSE OF THE EVENT AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411082&pc=QCB | 2023826-2023-00569 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.5/160 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6)2022. POST-OP THE PATIENTS VA WAS BLURRY AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00568. THE PATIENT HAD INTACS IMPLANTED AND WAS USING ARTIFICIAL TEARS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411155&pc=QCB | 2023826-2023-00567 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. POST-OP THE PATIENTS VA WAS BLURRY. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00566. THE PATIENT HAD INTACS IMPLANTED AND WAS USING ARTIFICIAL TEARS. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16417982&pc=QCB | 2023826-2023-00608 | 2023-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/064 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT POWER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC DEFORMED TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16425177&pc=QCB | 2023826-2023-00620 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH 13.2 IMPLANTABLE COLLAMER LENS OF -3.50/2.5/034 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. VERY HIGH RESIDUAL ASTIGMATISM WAS PRESENT. A BAROMETRY DEVICE WAS USED WHICH SHOWED THAT THERE WAS NO INTERNAL CYLINDER POWER. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433883&pc=QCB | 2023826-2023-00693 | 2022-05-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50/6.0/008 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433993&pc=QCB | 2023826-2023-00694 | 2022-05-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.50/5.0/165 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16471700&pc=QCB | 2023826-2023-00724 | 2021-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.0/4.0/58 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WITH ROTATION AND REFRACTIVE SUPRISE WERE OBSERVED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: PATIENT IS FINE WITH THE RESULT. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16471706&pc=QCB | 2023826-2023-00723 | 2021-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -5.5/3.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WERE OBSERVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI)# (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO: SERIOUS INJURY IN THE INITIAL MDR. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.5/3.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021.LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WERE OBSERVED. THE LENS WAS REPOSITIONED BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THE PROBLEM WAS NOT RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 11-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: SERIAL NUMBER SHOULD BE CORRECTED TO (B)(6). EXPIRATION DATE: SHOULD BE CORRECTED TO 31-MAR-2022. UDI: SHOULD BE OMITTED FOR CORRECTION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16492175&pc=QCB | 2023826-2023-00779 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -11.5/2.5/170 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023, AND IN A SUBSEQUENT SURGERY LATER THAT DAY, WAS REPLACED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "NO PROBLEM" IS STATED FOR STATUS OF THE EYE. Manufacturer Narrative: CORRECTED DATE: EXPLANT DATE: (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAM#: (B)(4). Manufacturer Narrative: PATIENT INFO:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16492446&pc=QCB | 2023826-2023-00780 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.50/2.5/178 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "POSTOPERATIVE DAY 1 IOP IS HIGH AND I/A IS PERFORMED. CURRENTLY, DOCTOR IS TAKING A WAIT-AND-SEE APPROACH." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATE: EXPLANT DATE: (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE- 27-MAR-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16492703&pc=QCB | 2023826-2023-00781 | 2022-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -16.5/2.0/179 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND REPLACED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY STATUS OF THE EYE IS, "STABLE." Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/ DEFORMED WITH FIBRE ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16507974&pc=QCB | 2023826-2023-00835 | 2023-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.50/1.50/29 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS A REFRACTIVE ERROR DUE TO WRONG LENS CALCULATION; AXIS 117 INSTEAD OF 177 WAS USED. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS USER ERROR. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16508352&pc=QCB | 2023826-2023-00788 | 2022-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.5/3/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT DISSATISFACTION/STORE AND BLURRY SYMPTOM. PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16519835&pc=QCB | 2023826-2023-00813 | 2023-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -8.50/2.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT MODEL AND POWER ON (B)(6) 2023 AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION-SHOULD BE CORRECTED TO STATE "FOUND THE OPTIC AND THE HAPTIC TORN WITH RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS". CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16521994&pc=QCB | 2023826-2023-00789 | 2022-09-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WITHOUT ROTATION AND ON (B)(6) 2022, REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. THE LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: PATIENT INFORMATION: UNK. EVENT PROBLEM AND EVALUATION CODES: OFF-LABEL AGE<21 CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16521995&pc=QCB | 2023826-2023-00790 | 2022-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Improper or Incorrect Procedure or Method | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS OF +4/+6/60 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT'S WITH PRE-EXISTING KERATOCONUS ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO REFRACTIVE SURPRISE. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "INCORRECT INPUT INTO OCOS. CYLINDER SHOULD HAVE BEEN -5.5 BUT WAS INSTEAD PUT IN AS 5.5 WHICH DEFAULTED TO +5.5 FOR CALCULATION." THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED, "NO ICL IN EYE," AND "WILL ALLOW THE EYE TO SETTLE POST EXPLANTATION, WILL REMEASURE BIOMETRY AND REFRACTION AFTER EYE IS STABLE FOR FUTURE IMPLANT." Manufacturer Narrative: OFF-LABEL - PRE-EXISTING KERATOCONUS - USER ERROR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16521996&pc=QCB | 2023826-2023-00791 | 2022-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS OF +1/1.5/32 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WITH LENS ROTATION AND REFRACTIVE SURPRISE WERE OBSERVED. ON (B)(6) 2022, THE LENS WAS REPOSITIONED AND IT DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2022, THE LENS WAS REMOVED. CAUSE OF THE EVENT WAS PATIENT RELATED. REPORTEDLY, "PATIENT HAD EXCESSIVE ODEMA AND ICL ROTATION POST-OP." THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED, "NO ICL IN EYE," AND "WILL ALLOW THE EYE TO SETTLE POST EXPLANTATION, WILL REMEASURE BIOMETRY AND REFRACTION AFTER EYE IS STABLE FOR FUTURE IMPLANT." Manufacturer Narrative: PATIENT RELATED FACTOR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16538767&pc=QCB | 2023826-2023-00792 | 2023-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: OFF-LABEL - ACD < 3.0 AGE > 45. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 03/23/2023. H3:DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: 728275 2023 00792 SUP2. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16549892&pc=QCB | 2023826-2023-00905 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/4.5/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE WERE OBSERVED. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "GOOD HOWEVER ICL IN @ 158 POST OPERATIVELY AND WAS SUPPOSED TO BE AT 167. I WANT TO MOVE LENS BUT WHICH ONE WOULD BE BEST?" AND "ICL HAS NOT BEEN EXPLANTED. DR (B)(6) FIRST WANTS TO SEE, IF BY ROTATING THE ICL HE CAN IMPROVE THE REFRACTIVE OUTCOME." Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CORRECTED DATA: G2: SOUTH AFRICA. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16567677&pc=QCB | 2023826-2023-00940 | 2023-02-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND THE PATIENT COMPLAINED OF BLURRY VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.50/+3.50/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, BLURRED VISION, AND RESIDUAL REFRACTIVE ERROR. REPORTEDLY " PATIENT HAS A HIGH VAULT OU BUT THERE HAVE BEEN NO ISSUES WITH IOP. HIS PATIENT ALSO HAS RESIDUAL REFRACTIVE ERROR OD GREATER THAN OS. HOWEVER, THE PATIENT WAS AT UPPER EXTREME OF MYOPIA OD VS OS. HE IS NOT COMPLAINING BUT WANTED TO RECEIVE SOME ADVICE ON FOLLOWING AND OPTIONS FOR THE PATIENT." THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: " BLURRY VISION". THE CAUSE OF THE EVENT IS UNKNOWN. CAUSE DETAILS- "HIGH VAULT". CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16584143&pc=QCB | 2023826-2023-00987 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.5/+0.5/48 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED LENS ROTATION AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 LENS REPLACED ON (B)(6) 2023, WITH A DIFFERENT POWER LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ALL GOOD. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16584242&pc=QCB | 2023826-2023-00986 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.0/+0.5/68 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - LENS REPLACED ON (B)(6) 2023 WITH A DIFFERENT POWER LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ALL GOOD. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16606958&pc=QCB | 2023826-2023-01003 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -6.5/3.5/016 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO DOUBLE VISION. Manufacturer Narrative: CORRECTED DATA: D6A: SHOULD BE CORRECTED TO 07-NOV-2022. D9: RETURN DATE: 15-APR-2023 G2: SHOULD BE OMITTED FOR CORRECTION. ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A ZIPLOCK BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#(B)(4). Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. ADVERSE EVENT PROBLEM WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16610082&pc=QCB | 2023826-2023-00980 | 2023-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED, LENS REMAINS IMPLANTED. UNKNOWN WAS REPORTED AS CAUSE OF EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS OF DIFFERENT POWER ON (B)(6) 2023 AND THE PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16610212&pc=QCB | 2023826-2023-00981 | 2023-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED, LENS REMAINS IMPLANTED. UNKNOWN WAS REPORTED AS CAUSE OF EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS OF DIFFERENT POWER ON (B)(6) 2023 AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16610530&pc=QCB | 2023826-2023-01082 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS A REFRACTIVE SURPRISE AND REQUESTED ASSISTANCE FROM A CONSULTATION DOCTOR. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16633210&pc=QCB | 2023826-2023-01010 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/3.0/061 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "SO FAR PATIENT NOT COME FOR FOLLOW UP AND MORE INVESTIGATION." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16633211&pc=QCB | 2023826-2023-01009 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/4.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "SO FAR PATIENT NOT COME FOR FOLLOW UP AND MORE INVESTIGATION." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16640306&pc=QCB | 2023826-2023-01122 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.0/+3.0/002 (SPHERE,CYLINDER,AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. IT WAS ADDITIONALLY NOTED "PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16649748&pc=QCB | 2023826-2023-01078 | 2023-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/013 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16668634&pc=QCB | 2023826-2023-01193 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A TICM12.0; -13.5/+3.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE SURGEON NOTES THE PATIENT HAS REFRACTIVE CHANGE OVER TIME WHICH WAS ASSESSED ON (B)(6) 2023. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVEN WAS REPORTED AS UNKNOWN. THE PROBLEM IS NOT RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) . CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16670694&pc=QCB | 2023826-2023-01198 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -7.00/1.0/012 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE. A PLANNED LENS REMOVAL IS REPORTED DUE TO POSTOPERATIVE REFRACTION (INTENDED PLANO), LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678777&pc=QCB | 2023826-2023-01181 | 2023-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.5/099 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE CHANGE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "REPOSITIONING OF THE LENS BY PROF. PARASTA WAS NOT SUCCESSFUL. PATIENT IS WEARING GLASSES AGAIN. COLE MONITORING. IF THE PATIENT WANT A LENS EXCHANGE THE SURGEON WOULD LIKE TO EXPLANT THE CURRENT LENS. WAIT FOR 3 MONTHS AND THEN REMEASURE THE EYE AND RECALCULATE THE LENS FOR EXCHANGE. WE WILL KEEP YOU UPDATED". THE CAUSE OF THE EVENT AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-8.5/+1.5/099 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678779&pc=QCB | 2023826-2023-01184 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -6.0/1.0/76 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 08-FEB-2023. ON 17-FEB-2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT INFORMATION:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS EXPLANTED ON (B)(4) 2023 DUE TO REFRACTIVE SURPRISE. PROBLEM WAS RESOLVED. H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678780&pc=QCB | 2023826-2023-01180 | 2023-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE CHANGE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "REPOSITIONING OF THE LENS BY PROF. PARASTA WAS NOT SUCCESSFUL. PATIENT IS WEARING GLASSES AGAIN. COLE MONITORING. IF THE PATIENT WANT A LENS EXCHANGE THE SURGEON WOULD LIKE TO EXPLANT THE CURRENT LENS. WAIT FOR 3 MONTHS AND THEN REMEASURE THE EYE AND RECALCULATE THE LENS FOR EXCHANGE. WE WILL KEEP YOU UPDATED". THE CAUSE OF THE EVENT AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678782&pc=QCB | 2023826-2023-01182 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION, WHICH DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2022, THE LENS WAS REPLACED FOR A SAME SIZE LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SUPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "AFTER IMPLANTATION OF TICL LENS IN THE RIGHT EYE, THE PATIENT'S VISION WAS DECREASED AND THE LENS POSITION WAS CHECKED FOR DEVIATION. AFTER THE SURGEON'S INTERVIEW, HE DECIDED TO PERFORM THE REPOSITIONING OPERATION ON (B)(6), AFTER THE REPOSITIONING OPERATION, THE PATIENT'S EYE CONDITION WAS RECHECKED AGAIN. THE INSPECTION RESULTS SHOWED THAT THE VISION WAS STILL NOT IMPROVED, THE LENS POSITION WAS NOT ROTATED, AND THE ARCH HEIGHT WAS NORMAL. AFTER CONSULTATION WITH THE SURGEON AND EXPERTS, THE CRYSTAL WAS INITIALLY SUSPECTED. AFTER COMMUNICATION WITH THE PATIENT, THE NEW TICL LENS WAS IMPLANTED AGAIN ON SEPTEMBER 30. THE VISION WAS STILL NOT IMPROVED ONE DAY AFTER SURGERY." Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-APR-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). Manufacturer Narrative: PATIENT DEMOGRAPHICS:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678785&pc=QCB | 2023826-2023-01177 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/0.5/140 (SPHERE,CYLINDER,AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SPHERIC MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678786&pc=QCB | 2023826-2023-01183 | 2022-09-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE EXCHANGE WAS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "AFTER IMPLANTATION OF TICL LENS IN THE RIGHT EYE, THE PATIENT'S VISION WAS DECREASED AND THE LENS POSITION WAS CHECKED FOR DEVIATION. AFTER THE SURGEON'S INTERVIEW, HE DECIDED TO PERFORM THE REPOSITIONING OPERATION ON AUGUST 5. AFTER THE REPOSITIONING OPERATION, THE PATIENT'S EYE CONDITION WAS RECHECKED AGAIN. THE INSPECTION RESULTS SHOWED THAT THE VISION WAS STILL NOT IMPROVED, THE LENS POSITION WAS NOT ROTATED, AND THE ARCH HEIGHT WAS NORMAL. AFTER CONSULTATION WITH THE SURGEON AND EXPERTS, THE CRYSTAL WAS INITIALLY SUSPECTED. AFTER COMMUNICATION WITH THE PATIENT, THE NEW TICL LENS WAS IMPLANTED AGAIN ON SEPTEMBER 30. THE VISION WAS STILL NOT IMPROVED ONE DAY AFTER SURGERY." Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728332&pc=QCB | 2023826-2023-01362 | 2021-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/2.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE SURGEON REPORTS LENS ROTATION; REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773435&pc=QCB | 2023826-2023-01464 | 2022-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -10.5/+0.50/78 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791810&pc=QCB | 2023826-2023-01446 | 2023-03-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -6.00/1.5/159 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. THE CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - REPORTEDLY, MEDICAL OR SURGICAL INTERVENTION WERE NOT NECESSARY. PER THE COMPLAINT QUESTIONNAIRE THE ISSUE WAS REPORTED AS "RESOLVED". THE STATUS OF THE EYE IS "ALL FINE". NO INJURY REPORTED. ADDITIONAL INFORMATION PROVIDED STATES "THE CLINIC PLANNED AN OUTCOME OF -0.75DPT BUT THE FINAL OUTCOME WAS ABOUT 0.00DPT. IT'S OK FOR THE PATIENT AND CURRENTLY WISH NO FURTHER TREATMENT". CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791940&pc=QCB | 2023826-2023-01447 | 2023-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -4.00/1.0/29 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. THE CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2023. H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. OVERALL LENGTH (DIAMETER) VERIFICATION AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WIDTH VERIFICATION COULD NOT BE PERFORMED DUE TO HAPTIC DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792761&pc=QCB | 2023826-2023-01428 | 2022-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: UNK WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE/DEBRIS ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811280&pc=QCB | 2023826-2023-01517 | 2022-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -2.5/1.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811756&pc=QCB | 2023826-2023-01558 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.00/+4.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+4.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022.THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE!" AND "PATIENT IS HAPPY". THE THE CAUSE OF THE EVENT IS UNKNOWN. H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "2993" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16828371&pc=QCB | 2023826-2023-01531 | 2022-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.5/1.5/167 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16839441&pc=QCB | 2023826-2023-01568 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.50/+3.0/068 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULTING; BLURRY VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "ORIGINAL CALCULATION & RECOMMENDED 12.6 SIZING TOO SMALL; = +1- 90 DEGREES ROTATION CAUSING BLURRY VISION. Manufacturer Narrative: CORRECTED DATA: HEALTH EFFECT - IMPACT CODE: "4628" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.50/+3.00/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND BLURRED VISION. ON (B)(6) 2023 THE LENS WAS REPOSITIONED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS 'OTHER'. CAUSE DETAILS: "ORIGINAL CALCULATION AND RECOMMENDED (12.6) SIZING TOO SMALL= +/- 90 DEGREES. ROTATION CAUSING BLURRY VISION". DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION WILL NOT BE PERFORMED DUE TO SIGNIFICANT HAPTIC DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16839442&pc=QCB | 2023826-2023-01569 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.00/+4.0/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULTING; BLURRY VISION. IT WAS NOTED THAT AFTER IMPLANT; THE PATIENT WENT TO THE EXAM ROOM TO BE EVALUATED AND THE SURGEON NOTICED THE LEFT ICL HAD ROTATED. THE SURGEON HAD REPOSITIONED THE ICL TO IT'S ORIGINAL POSITION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "ORIGINAL CALCULATION 12.6 SIZING TOO SMALL; = +1- 90 DEGREES ROTATION CAUSING BLURRY VISION". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850168&pc=QCB | 2023826-2023-01598 | 2023-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. ON THIS SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED WITH NO PATIENT INJURY. IT WAS ADDITIONALLY NOTED "AFTER SURGERY THE PATIENT COMPLAINED OF VISUAL PROBLEMS IN THE OPTICAL ZONE AND STRONGLY REQUESTED TO REMOVE THE LENS. AFTER REMOVAL A TRACE WAS FOUND NEAR THE OPTICAL ZONE MARKER LINE". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16864790&pc=QCB | 2023826-2023-01663 | 2023-03-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -14.00/4.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS REFRACTIVE SURPRISE AND BLURRY VISION. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16888245&pc=QCB | 2023826-2023-01638 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF A -15.0/3.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2023, DUE TO THE INITIAL LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. BLURRED VISON AND MILD CORNEAL EDEMA WAS OBERSRVED. THE STATUS OF THE EYE REPORTED, "MILD CORNEA EDEMA." Manufacturer Narrative: CORRECTED DATA: B3: DATE OF EVENT: SHOULD BE CORRECTED TO 08-04-2023. B5: OD SHOULD BE CORRECTED TO OS. D6B: EXPLANT DATE SHOULD BE OMITTED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16890526&pc=QCB | 2023826-2023-01746 | 2023-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM013.2, -15.50/1.5/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO CORRECT LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM, " LENS HAS ROTATED/ INDUCED ASTIGMATISM" WAS REPORTED. LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTOR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16892246&pc=QCB | 2023826-2023-01747 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.5/1.0/068 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894692&pc=QCB | 2023826-2023-01556 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT REFRACTIVE SURPRISE WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICM12.6 -3.0/2.0/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED ON (B)(6) 2023 INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS WAS REPOSITIONED ON (B)(6) 2023 AND THIS RESOLVED THE PROBLEM. STATUS OF THE EYE IS "EXPECTANT". REPORTEDLY, "IN FUTURE VISITS WE WILL SEE THE EVOLUTION." CLAIM# (B)(4). Manufacturer Narrative: H4 - CORRECTED TO: 09-AUG-2022 IN THE INTIAL MDR. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT SHOULD BE ADDED TO THE INITIAL MDR. (B)(4) Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894694&pc=QCB | 2023826-2023-01555 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.0/1.5/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - REPORTEDLY "THE PATIENT IS ALREADY OPERATED ON AND SEEING 20/20. IT IS CONFIRMED THAT THE INTIAL LENSES WERE SWAPPED." CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894700&pc=QCB | 2023826-2023-01554 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.0/1.5/041 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - REPORTEDLY "THE PATIENT IS ALREADY OPERATED ON AND SEEING 20/20. IT IS CONFIRMED THAT THE INTIAL LENSES WERE SWAPPED." CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16897246&pc=QCB | 2023826-2023-01750 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 -9.5/1.0/019 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND DIFFERENT POWER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTICS BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16908166&pc=QCB | 2023826-2023-01776 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.0/048 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPORTED AS LOW VAULTING AND THE PATIENT EXPERIENCED BLURRED VISION AND A REFRACTIVE SURPRISE. THE LENS VAULT WAS LOW AND SAT TOO HIGH IN THE SULCUS CAUSING A "SMUDGE" IN VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023 FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16920640&pc=QCB | 2023826-2023-01802 | 2023-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.50/0.5/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16920684&pc=QCB | 2023826-2023-01801 | 2023-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.50/0.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922176&pc=QCB | 2023826-2023-01781 | 2023-04-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+4.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD A 1 + CELL AND THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-01687. Manufacturer Narrative: B5: THE PATIENT WAS SEEN ON (B)(6) 2023 AT POST-OP VISIT. THE PATIENT IS HAPPY WITH IMPROVED VISION, VA UNCORRECTED 20/30-2, BCVA 20/30. VISUAL ACUITY IS AT POTENTIAL DUE TO AMBLYOPIA. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16934647&pc=QCB | 2023826-2023-01876 | 2023-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/+1.0/142 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16962318&pc=QCB | 2023826-2023-01924 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -5.5/1.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991212&pc=QCB | 2023826-2023-01771 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -14.5/2.5/055 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "2993" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+2.5/55 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2022 .THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REPLACED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16994612&pc=QCB | 2023826-2023-01993 | 2020-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998681&pc=QCB | 2023826-2023-01998 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.00/3.0/178 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH,DIFFERENT MODEL AND POWER LENS DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT ASTIGMATISM WAS TREATED WITH RELAXING INISION. PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED, IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17007679&pc=QCB | 2023826-2023-02045 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.5/1.5/066 (SPHERE/CYLINDER/AXIS), INTO SON'S LEFT EYE (OS) ON (B)(6) 2023 FOR CORRECTION OF -12 DIOPTERS OF MYOPIA AND -2 DIOPTERS OF ASTIGMATISM. THE REPORTER STATES " AFTER THE SECOND POST-OPERATIVE VISIT MY SON COMPLAINS OF BLURRY VISION AND THE SURGEON REALIZES THAT THE 2 DIOPTERS OF ASTIGMATISM HAVE NOT BEEN CORRECTED ". LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - REFRACTIVE SURPRISE WAS REPORTED. CAUSE OF EVENT IS UNKNOWN WITH LENS ROTATION NOTED. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039566&pc=QCB | 2023826-2023-01885 | 2022-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated; Optical Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.00/4.5/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- MEDICAL DEVICE PROBLEM CODE: "3001" SHOULD BE REMOVED AND "1401" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE<21) . CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039571&pc=QCB | 2023826-2023-01888 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.5/3.0/022 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - A6: UNK. B3: UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D6A - UNK. D6B - UNK. H4 - UNK. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.5/3.0/022 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE (B)(6) 2023. REFRACTION WAS OBSERVED. THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND LENGTH LENS WITH A DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039736&pc=QCB | 2023826-2023-02153 | 2023-04-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.00/+1.0/174 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17054422&pc=QCB | 2023826-2023-02173 | 2023-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -10.0/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH, MODEL, AND POWER WAS IMPLANTED BUT THIS DID NOT RESOLVE THE PROBLEM. IN WAS ADDITIONALLY NOTED AFTER THE REPLACEMENT LENS WAS IMPLANTED THE PATIENT HAS CORNEAL EDEMA, AND THE POST OP REFRACTION WAS +0.75/-1.75/165. SURGEON ADVISED THE NEXT REVIEW AFTER 6 WEEKS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MFR# 2023826-2023-02174 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "4627" SHOULD BE REMOVED AND "4629" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION: "PATIENT IS HAVING RESIDUAL OF +1.75/3@175. CHECKED POST OP ON (B)(6) WITH CORNEAL EDEMA, AND THE POST OP REFRACTION WAS +0.75/-1.75/165. SURGEON ADVISED THE NEXT REVIEW AFTER 6 WEEKS." IT WAS ALSO REPORTED THAT "SUGGEST REPOSITIONING OF THIS SAME LENS TO CORRECT THE RESIDUAL ERROR". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17061058&pc=QCB | 2023826-2023-02039 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -5.0/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SAME LENGTH DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17062290&pc=QCB | 2023826-2023-02038 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.5/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066749&pc=QCB | 2023826-2023-02221 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL USE (ACD<3.0MM). (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076192&pc=QCB | 2023826-2023-02226 | 2023-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -7.00/1.0/070 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT WAS: DEVICE, A PATIENT RELATED FACTOR, AND UNKNOWN. THE DEVICE DID NOT FAIL AS INTENDED. REPORTEDLY, "NO PROBLEM WITH THE PROCEDURE AND TEST DATA, BUT NO IMPROVEMENT IN VISION." THE PROBLEM HAS NOT RESOLVED. STATUS OF THE EYE SAID, "THERE IS NOTHING WRONG WITH THE EYE, BUT THE VISION IS NOT GOOD," AND "DOCTOR SUSPECTS A PARALYSIS OF THE CILIARY MUSCLE OR A PROBLEM WITH THE LENS ITSELF." Manufacturer Narrative: B5 - LENS WAS EXPLANTED. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID WITHIN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BROKEN LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS REPLACED WITH A DIFFERENT POWER LENS ON (B)(6) 2023. STATUS OF THE EYE IS "GOOD". CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, THE DHR REVIEW INDICATES THAT THE PRODUCT HAS NOT BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS AND THAT THERE IS INDICATION THAT THE MANUFACTURING AND/OR PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. ROOT CAUSE WAS NOT DETERMINED. PROBABLE CAUSE IS LIKELY ASSOCIATED TO THE LENS LABEL REWORK PROCESS. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084116&pc=QCB | 2023826-2023-02230 | 2022-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -.05/6/140 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON 18-MAY-2023, REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT WAS REPORTED AS BOTH THE DEVICE AND A PATIENT RELATED FACTOR. REPORTEDLY, "LENS WAS NOT AVAILABLE. HAD TO WAIT FOR STAAR TO MANUFACTURE (NOT AN OFF SHELF LENS). OUTCOME IS OFF TARGET." STATUS OF THE EYE IS STATED, "UNHAPPY PATIENT BECAUSE OF THE RESULTANT REFRACTIVE SURPRISE ANISOMETROPIA MAKES IT DIFFICULT TO CORRECT WITH GLASSES CL INTOLERANT." THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5: DIOPTER SHOULD BE CORRECTED TO: -.50/6/140. CLAIM#(B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17089408&pc=QCB | 2023826-2023-02232 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/4.5/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS, IMPLANTED VERTICALLY DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, "BOTH LENS ARE PLACED VERTICALLY." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093105&pc=QCB | 2023826-2023-02236 | 2023-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/035 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO BLURRED VISION. REPORTEDLY, "THE PATIENT WAS NOT SATISFIED WITH THE POSTOPERATIVE VISUAL QUALITY, AND THE CLINICIAN FOUND THAT THE PATIENT LACKED CORRECTION FOR 0.5D THROUGH MYDRIASIS AND OPTOMETRY, AND FULLY COMMUNICATED THAT THE PATIENT WAS EXTREMELY REQUIRED TO REPLACE THE LENS." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS OTHER. THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. STATUS OF THE EYE IS REPORTED AS, "OK.". Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - ACD <3.0. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17102366&pc=QCB | 2023826-2023-02186 | 2023-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023 FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS. THE PROBLEM WAS RESOLVED AND THE EYE WAS QUIET. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 729514
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17105578&pc=QCB | 2023826-2023-02268 | 2023-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -09.50/4.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 24-FEB-2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17105579&pc=QCB | 2023826-2023-02269 | 2023-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.50/3.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119691&pc=QCB | 2023826-2023-02163 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM0 12.6, -9.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. LENS WAS RETURNED IN ONE HALF. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL/ FUNCTIONAL INSPECTION DUE TO LENS DAMAGE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140261&pc=QCB | 2023826-2023-02405 | 2019-08-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -09.00/+20/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. THE PATIENT NOW HAS REFRACTIVE SURPRISE. LEN REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17159843&pc=QCB | 2023826-2023-02439 | 2023-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.0/+2.5/76 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.0/+2.5/76 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS INSERTED UPSIDE DOWN. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SIMILAR LENGTH AND MODEL LENS. THIS DID NOT RESOLVE THE PROBLEM. LASER TOUCH UP SCHEDULED. D6B: EXPLANT DATE: (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17185665&pc=QCB | 2023826-2023-02524 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/2.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE WHICH WAS NOTED ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4) Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210598&pc=QCB | 2023826-2023-02570 | 2023-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13.0/3.0/7 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, WITH FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216367&pc=QCB | 2023826-2023-02428 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.0/1.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. REFRACTIVE SURPRISE WAS OBSERVED. LASER TOUCH UP WAS PERFORMED ON (B)(6)2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT IS FINE AND ALL IS FINE. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/78 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LASER VISION CORRECTION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "CORRECTED WITH TOUCH UP LVC". THE STATUS OF THE EYE AS "PATIENT IS HAPPY & ALL IS FINE!". ADDITIONAL INFORMATION PROVIDED: "AFTER THE SURGERY THERE WAS RESIDUAL REFRACTION OF -0.50/-0.25/35 UCVA 20/30" THE CAUSE OF THE EVENT IS UNKNOWN, AND ITS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT IMPACT CODE 4625: REFRACTIVE TREATMENT(LASER TOUCH UP). H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17217896&pc=QCB | 2023826-2023-02504 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -5.0/1.0/82 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND REPLACED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY "DARE TO CORRECT LOW MONOVISION." Manufacturer Narrative: H6: OFF-LABEL AGE<45. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17222268&pc=QCB | 2023826-2023-02601 | 2021-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.5/3/009 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REFRACTIVE TREATMENT WAS PERFORMED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "ALL IS GOOD PATIENT IS HAPPY," AND "THE RESIDUAL RX: +1.25/-1.25/189 (UCVA 20/40) WAS CORRECTED BY LVC TOUCH UP ON (B)(6) 2022." Manufacturer Narrative: A4-A6:UNK. H3: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253208&pc=QCB | 2023826-2023-02693 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT LENGTH AND POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17261568&pc=QCB | 2023826-2023-02739 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THERE IS EXCESSIVE VAULTING AND HIGH REMAINING RESIDUAL ASTIGMATISM WITH MENTION OF POSSIBLE OF ROTATING THE LENS OR EXCHANGING THE LENS TO A SMALLER SIZE TO REDUCE THE HIGH VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE (RESIDUAL ASTIGMATISM). REPORTEDLY "SHE HAS A VERY HIGH VAULT AND HIGH REMAINING RESIDUAL ASTIGMATISM". THE LENS REMAINS IMPLANTED. THE SERIAL NUMBER WAS NOT PROVIDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17261853&pc=QCB | 2023826-2023-02737 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -10.0/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS WAS EXPLANTED ON (B)(6) 2022. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17273273&pc=QCB | 2023826-2023-02722 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.5/+4.5/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE SURGEON PERFORMED PRK SURGERY ON THE PATIENT. LENS REMAINS IMPLANTED. THE PROBLEM WAS REPORTED AS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17278996&pc=QCB | 2023826-2023-02610 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM0_12.6, -10.0/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DUE TO SIGNIFICANT LENS HAPTIC DAMAGE, DIMENSIONAL INSPECTION NOT POSSIBLE TO PERFORM. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283187&pc=QCB | 2023826-2023-02724 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.5/2.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283882&pc=QCB | 2023826-2023-02676 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -5.00/2.0/092(SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "DATA STATUS IS FINE. THE PATIENT IS UNDER OBSERVATION BECAUSE OF LACK OF VISUAL ACUITY". Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283884&pc=QCB | 2023826-2023-02675 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.5/1.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS LATER IMPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "DATA STATUS IS FINE. THE PATIENT IS UNDER OBSERVATION BECAUSE OF LACK OF VISUAL ACUITY." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17297503&pc=QCB | 2023826-2023-02785 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.0/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A DIFFERENT POWER, SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17297504&pc=QCB | 2023826-2023-02786 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.0/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A DIFFERENT POWER, SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS SPECIFICALLY FIBRE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS SPECIFICALLY FIBRE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: A4 - UNK; A5 - UNK; A6 - UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313215&pc=QCB | 2023826-2023-02878 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -3.50/2.5/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE A REPLACEMENT LENS OF A LONGER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322358&pc=QCB | 2023826-2023-02846 | 2023-06-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -14.00/2.0/163 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND RESIDUAL MANIFEST REFRACTION. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. " ERROR IN MRX VS POOR ICL SELECTION" WAS REPORTED FOR CAUSE DETAILS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PER UPDATED INFORMATION THE PATIENT IS CURRENTLY STABLE REFRACTIVELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322375&pc=QCB | 2023826-2023-02847 | 2023-06-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00/1.0/032 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON \(B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND RESIDUAL MANIFEST REFRACTION. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. " ERROR IN MRX VS POOR ICL SELECTION" WAS REPORTED FOR CAUSE DETAILS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PATIENT IS CURRENTLY STABLE REFRACTIVELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17330143&pc=QCB | 2023826-2023-02929 | 2023-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: RETURNED DEVICE EVALUATION, LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS AND FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17330144&pc=QCB | 2023826-2023-02927 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, RETURNED WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17338891&pc=QCB | 2023826-2023-02980 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.50/6.0/138 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN 2016. THE SURGEON REPORTS THE ICL AXIS IS AT 5 DEGREES BUT REFRACTIVE IS +2.50-6.25@157 AND INDICATED THE PATIENT HAS ASTIGMATISM. POSSIBLE LENS ROTATION PLANNED. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (KERATOCONUS). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17380973&pc=QCB | 2023826-2023-03138 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -2.00/3.5/162 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.0/+3.5/162 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "THE DOCTOR WANTS TO REMOVE THE ICL AND IMPLANT ANOTHER ONE". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17383221&pc=QCB | 2023826-2023-03139 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.50/3.5/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS REPORTED THAT ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17385329&pc=QCB | 2023826-2023-03154 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.50/+5.0/58 (SPHERE / CYLINDER / AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391261&pc=QCB | 2023826-2023-03039 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - (B)(6) 2023; A3 - M; B3 - 10-MAY-2023. B5 - REPORTED AS: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. - UPDATE TO: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.0/2.0/128 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING D4 - VTICM5_13.2, T1347711, 31-JAN-2026; D6A - (B)(6) 2023; D10 - INJECTOR MODEL#MSI-PF,LOT#UNK, CARTRIDGE MODEL#SFC-45,LOT#UNK, FOAM TIP PLUNGER MODEL#FTP,LOT#UNK; H4 - 07-FEB-2023; H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: A2 - UNK A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. D6B - UNK. H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17417150&pc=QCB | 2023826-2023-03133 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.0/3.0/002 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND LATER REPLACED A SAME LENGTH SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. REPORTEDLY, "THE PATIENT WENT TO THE HOSPITAL FOR PREOPERATIVE EXAMINATION OF ICL SURGERY ON (B)(6) 2023. THE EXAMINATION FOUND THAT THERE WAS A SIGNIFICANT DIFFERENCE BETWEEN COMPUTER OPTOMETRY AND COMPREHENSIVE OPTOMETRY, AND REPEATED COMPREHENSIVE OPTOMETRY RESULTS WERE CONSISTENT. ACCORDING TO THE COMPREHENSIVE OPTOMETRY RESULTS BEFORE SURGERY, THE VISUAL ACUITY COULD REACH NORMAL CORRECTED VISION. POSTOPERATIVE VISUAL ACUITY: 1.0 IN THE RIGHT EYE AND 0.6 IN THE LEFT EYE. COMPUTER OPTOMETRY SHOWED THAT BOTH EYES WERE UNDER CORRECTION FOR ASTIGMATISM. THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED A DECREASE IN BINOCULAR VISION, AND REPEATED COMPUTER OPTOMETRY RESULTED IN UNDERCORRECTED MYOPIA AND ASTIGMATISM IN BOTH EYES. THE SUFFICIENT DIOPTER WAS CORRECTED TO 1.0. AFTER COMMUNICATING WITH THE PATIENT, THE PATIENT UNDERWENT LENS REPLACEMENT SURGERY. BILATERAL TICL LENS REPLACEMENT SURGERY WAS PERFORMED ON (B)(6) 2023. POSTOPERATIVE VISION IMPROVED, WITH SLIGHT CORNEAL EDEMA. THE PATIENT WAS INSTRUCTED TO TAKE MEDICATION ON TIME AND UNDERGO REGULAR FOLLOW UP." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CORRECTED DATA: 4627-LENS EXPLANT ADDED FOR CORRECTION. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17417153&pc=QCB | 2023826-2023-03132 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.0/2.5/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND LATER REPLACED A SAME LENGTH SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. REPORTEDLY, "THE PATIENT WENT TO THE HOSPITAL FOR PREOPERATIVE EXAMINATION OF ICL SURGERY ON (B)(6) 2023. THE EXAMINATION FOUND THAT THERE WAS A SIGNIFICANT DIFFERENCE BETWEEN COMPUTER OPTOMETRY AND COMPREHENSIVE OPTOMETRY, AND REPEATED COMPREHENSIVE OPTOMETRY RESULTS WERE CONSISTENT. ACCORDING TO THE COMPREHENSIVE OPTOMETRY RESULTS BEFORE SURGERY, THE VISUAL ACUITY COULD REACH NORMAL CORRECTED VISION. POSTOPERATIVE VISUAL ACUITY: 1.0 IN THE RIGHT EYE AND 0.6 IN THE LEFT EYE. COMPUTER OPTOMETRY SHOWED THAT BOTH EYES WERE UNDER CORRECTION FOR ASTIGMATISM. THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED A DECREASE IN BINOCULAR VISION, AND REPEATED COMPUTER OPTOMETRY RESULTED IN UNDERCORRECTED MYOPIA AND ASTIGMATISM IN BOTH EYES. THE SUFFICIENT DIOPTER WAS CORRECTED TO 1.0. AFTER COMMUNICATING WITH THE PATIENT, THE PATIENT UNDERWENT LENS REPLACEMENT SURGERY. BILATERAL TICL LENS REPLACEMENT SURGERY WAS PERFORMED ON (B)(6) 2023. POSTOPERATIVE VISION IMPROVED, WITH SLIGHT CORNEAL EDEMA. THE PATIENT WAS INSTRUCTED TO TAKE MEDICATION ON TIME AND UNDERGO REGULAR FOLLOW UP." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING CORRECTED DATA: 4627 EXPLANT ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: 1 SIMILAR COMPLAINT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM# (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17440469&pc=QCB | 2023826-2023-03207 | 2023-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.00/1.00/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SUPRISE HAS BEEN OBSERVED. REPORTEDLY, "PATIENT UNDERWENT SURGERY ON (B)(6) 2023, THE NEXT DAY SHE HAD VA 0.4. IN THE REVISION OF (B)(6) 2023, SHE STILL HAD VA 0.4 WITH RESIDUAL REFRACTION OF +1.00-0.75X165 VA 1.0. AT THE LAST EXAMINATION ON (B)(6) 2023, SHE HAS A +1.25-1.00X165 AV 1.0. LENS IS AT 163." THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446725&pc=QCB | 2023826-2023-03291 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -5.50/1.0/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE WHICH WAS NOTED ON (B)(6) 2023. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023, THE LENS WAS EXPLANTED AND REPLACED WITH A SIMILAR LENS. PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17447261&pc=QCB | 2023826-2023-03213 | 2023-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.0/0.5/010 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, "BOTH TIMES IN HORIZONTAL POSITION." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 "THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE" SHOULD BE CORRECTED TO "THE LENS WAS EXCHANGED FOR A SAME SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE". CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17476696&pc=QCB | 2023826-2023-03334 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.; -2.50/+6.0/068 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477169&pc=QCB | 2023826-2023-03333 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -0.50/+2.0/143 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17483721&pc=QCB | 2023826-2023-03348 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484473&pc=QCB | 2023826-2023-03355 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -14.00/2.5/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. PATIENTS VISION AFTER SURGERY WAS 20/50 +1 WITH VAULT AROUND 1000. AFTER 2 WEEKS WITH NO VISUAL IMPROVEMENT, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.00/2.5/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. REPORTEDLY "NO VISUAL IMPROVEMENT. (VA 20/50). HIGHER VAULTING WAS ALSO NOTED". ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENS LENGTH AND A DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486826&pc=QCB | 2023826-2023-03273 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE(OD). REFRACTIVE CHANGE OVERTIME HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2 - UNK, A3 - UNK, A4 - UNK , A5 - UNK, A6 - UNK, B3 - UNK, D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY), D6A - UNK, H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486827&pc=QCB | 2023826-2023-03266 | 2023-07-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -12.50/4.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) ADDED TO INITIAL MDR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486829&pc=QCB | 2023826-2023-03274 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE(OS). REFRACTIVE CHANGE OVERTIME HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D6A - UNK. H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17503977&pc=QCB | 2023826-2023-03359 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE REPORTER REQUESTED ASSISTANCE FROM A CONSULTATION DOCTOR DUE TO ASTIGMATISM. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE REPORTER REQUESTED ASSISTANCE FROM A CONSULTATION DOCTOR DUE TO LENS ROTATION. THE REPORTER STATED "ROTATED 8 DEGREE CLOCKWISE AFTER HORIZONTAL IMPLANTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM #(B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17505812&pc=QCB | 2023826-2023-03324 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.0/0.5/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL ACD<3.0. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17542184&pc=QCB | 2023826-2023-03515 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -2.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION (PT. RIGHT EYE FIELD IS BLURRED) AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45 YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17545344&pc=QCB | 2023826-2023-03436 | 2023-06-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.50/6.00/047 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: A4, A5, A6-UNK. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17546728&pc=QCB | 2023826-2023-03519 | 2023-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/157 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IT WAS REPORTED THAT LENS REPOSITIONING WAS PERFORMED DUE TO REFRACTIVE SURPRISE. PER SURGEON ROTATION WENT WELL, NO PROBLEMS ENCOUNTERED, PATIENT NOTICED INSTANTANEOUS IMPROVEMENT POST OP. PER UPDATED INFORMATION TORIC MARKINGS ARE @ 98 RATHER THAN 103, THE SURGEON STATES " 5 DEGREE OUT IS MINIMAL. IT SHOULDN'T BE SIGNIFICANT. I'D WAIT AND REVIEW MONTHLY." THE CAUSE OF THIS EVENT WAS DUE TO USER ERROR (TYPING ERROR), THE AXIS ENTERED WAS INCORRECT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17546992&pc=QCB | 2023826-2023-03520 | 2023-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS OF 6.5/3.0/146 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - REPORETEDLY, "THE SURGEON EXPLANT THE INITIAL LENS ON (B)(6) 2023; DURING INJECTION/IMPLANTATION OF REPLACEMENT LENS THERE WAS A LENS BREAK. SO HE DIRECTLY EXPLANT THE REPLACEMENT LENS AND IMPLANT AGAIN THE INITIAL LENS." REFRACTIVE SURPRISE IS OBSERVED. EXCHANGE IS PLANNED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549362&pc=QCB | 2023826-2023-03447 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -4.0/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# 730787. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549364&pc=QCB | 2023826-2023-03446 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/5.0/060 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D6A - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: ADDITIONAL AND CORRECTIONS: ADDITIONAL INFORMATION: B1 - REPORTED AS PRODUCT PROBLEM - UPDATE TO ADVERSE EVENT. B3 - DATE OF EVENT (B)(6) 2023. B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/5.0/060 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, EYE IS QUIET AND LENS IN POSITION AT DESIRED AXIS. PATIENT AND DOCTOR ARE HAPPY. D6A - (B)(6) 2023. H6 - HEALTH EFFECT - IMPACTT CODE REPORTED AS 2199 - UPDATE TO 4627. CORRECTION: A1 - REPORTED AS (B)(6). CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT - IMPACT CODE REPORTED AS 4627 IN SUPP1 - UPDATE TO 4628. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549367&pc=QCB | 2023826-2023-03452 | 2023-05-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.0/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549368&pc=QCB | 2023826-2023-03451 | 2023-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -11.50/3.5/048 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549381&pc=QCB | 2023826-2023-03453 | 2023-05-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.0/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17552477&pc=QCB | 2023826-2023-03524 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: IT WAS REPORTED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -5.0/1.0/003 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS MARKED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: B5: LENS EXPLANTED ON (B)(6) 2023 AND REPLACED THE SAME DAY WITH SAME LENGTH DIFFERENT DIOPTER LENS. PROBLEM RESOLVED. STATUS OF EYE "ALL FINE! PATIENT IS HAPPY! SUBJ REFRACTION: -1.50DPT." H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17568859&pc=QCB | 2023826-2023-03512 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A REPLACEMENT 13.2MM VTICMO13.2 WITH A -15.0/1.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY, "PT NOTES LESS ABBERRATIONS, SORE, STILL BLURE BUT BETTER." Manufacturer Narrative: A4-A6:UNK. H6: OFF LABEL: ACD<3.0. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581263&pc=QCB | 2023826-2023-03595 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.50/0.50/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17595281&pc=QCB | 2023826-2023-03531 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED ON (B)(6) 2023 TO RESOLVE THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. LENS MODEL SIZE RETURNED IS NOT LENS MODEL SIZE STATED IN CLAIM. DIMENSIONAL INSPECTION FOUND THE OVERALL WIDTH WITHIN SPECIFICATIONS BUT OVERALL LENGTH FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17599751&pc=QCB | 2023826-2023-03624 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; No Apparent Adverse Event | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). A SMALL REFRACTIVE ERROR WAS OBSERVED. REPORTEDLY, "THE TORIC ICL REMAINS IN PLACE AND THE PATIENT DOES NOT WANT TO DO ANYTHING AS SHE IS SO MUCH BETTER THAN SHE WAS BEFORE SURGERY." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5: A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -10.0/2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17627778&pc=QCB | 2023826-2023-03720 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18.0/1.5/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A DIFFERENT POWER/BRAND ICL AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BENT/STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17627779&pc=QCB | 2023826-2023-03721 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18.0/1.0/058 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A DIFFERENT POWER/BRAND ICL AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17628351&pc=QCB | 2023826-2023-03605 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/169 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17637145&pc=QCB | 2023826-2023-03733 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6; +2.5/+6.0/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023 AS A REPLACEMENT LENS. IT WAS INDICATED THE PATIENT STILL HAS LOW RESIDUAL ASTIGMATISM BUT THE PATIENT DOES NOT WANT A LASER TOUCH UP, THERE WILL BE A CHECK UP IN 3 MONTHS. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494, OFF-LABEL USE, ACD < 3.0MM CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17637154&pc=QCB | 2023826-2023-03753 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.0/6.0/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE SURGEON REPOSITIONED THE LENS AND THE PATIENT IS PLANO (NO VISUAL PROBLEMS). THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17644938&pc=QCB | 2023826-2023-03693 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/3.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: A4, A5, A6-UNK. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+3.0/94 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SLIGHTLY WEAKER POWER; SAME MODEL AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17646185&pc=QCB | 2023826-2023-03794 | 2022-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.0/1.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH/DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17655764&pc=QCB | 2023826-2023-03800 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/1.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED AND NOTED " DUE TO PSYCHOLOGICAL REASONS, THE PATIENT COMPLAINT OF SUBNORMAL VISION AND SORE EYES MULTIPLE TIMES WITHIN A MONTH AFTER SURGERY. SHE ALSO REFUSED CONSERVATIVE TREATMENT AND REQUESTED THE REMOVAL OF THE ICL FROM BOTH EYES". Manufacturer Narrative: (B)(4) Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H4 - 30-JUN-2025 CORRECTED TO 22-JUL-2022. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17655766&pc=QCB | 2023826-2023-03802 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/0.5/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED AND NOTED "DUE TO PSYCHOLOGICAL REASONS, THE PATIENT COMPLAINT OF SUBNORMAL VISION AND SORE EYES MULTIPLE TIMES WITHIN A MONTH AFTER SURGERY. SHE ALSO REFUSED CONSERVATIVE TREATMENT AND REQUESTED THE REMOVAL OF THE ICL FROM BOTH EYES". Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC SPLIT AND HAPTIC BROKEN. ADDITIONAL COMMENTS STATE, "UNABLE TO PERFORM DIMENSIONAL AND FUNCTIONAL VERIFICATION DUE TO SIGNIFICANT LENS DAMAGE". CLAIM #(B)(4). Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669071&pc=QCB | 2023826-2023-03737 | 2023-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/5.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS ROTATED DUE TO DUE TO LARGE AMOUNT OF ASTIGMATISM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, AFTER ROTATION THE ASTIGMATISM WAS HIGHLY IMPROVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17671454&pc=QCB | 2023826-2023-03839 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.50/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695113&pc=QCB | 2023826-2023-03946 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.0/+2.5/75 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6). 2023 AS A REPLACEMENT LENS. IT WAS INDICATED THAT THE PATIENT STILL HAS THE SAME CHECK UP DATA TO A PRIOR UPSIDE-DOWN LENS IMPLANT AND THE PATIENT CURRENTLY HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IT WAS REPORTED THE SURGEON HAS PLANS TO PERFORM A LASER TOUCH UP ON THE PATIENT. THE PROBLEM IS NOT RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17696270&pc=QCB | 2023826-2023-03773 | 2023-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: PRODUCT WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5: LENS ROTATION AND REPOSITIONING SHOULD BE INCLUDED. CLAIM #: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716004&pc=QCB | 2023826-2023-03955 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.50/3.5/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS VAULT AND 20/50 +1 VISION. THE REPORTER STATES THAT THIS REPLACEMENT RESOLVED THE PROBLEM BUT THERE IS SOME RESIDUAL REFRACTIVE ERROR. SURGEON WILL MONITOR FOR STABILITY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716019&pc=QCB | 2023826-2023-03956 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -13.50/3.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS VAULT AND 20/50 +1 VISION. THE REPORTER STATES THAT THIS REPLACEMENT RESOLVED THE PROBLEM BUT THERE IS SOME RESIDUAL REFRACTIVE ERROR. SURGEON WILL MONITOR FOR STABILITY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716134&pc=QCB | 2023826-2023-03696 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT AN EVO IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. REFRACTIVE SURPRISE WAS REPORTED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716167&pc=QCB | 2023826-2023-03957 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.00/1.0/17 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17722426&pc=QCB | 2023826-2023-03740 | 2022-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO BLURRED VISION AND LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17725920&pc=QCB | 2023826-2023-04007 | 2023-07-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.5/+0.5/158 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME WHICH WAS ASSESSED ON (B)(6) 2023. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733557&pc=QCB | 2023826-2023-04035 | 2023-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE SURGEON INDICATED THAT A 13.2MM IMPLANTABLE COLLAMER LENS OF -7.5/1.0/066 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON STATES THAT THE LENS ROTATED 15 MONTHS AFTER SURGERY, BLURRED VISION WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT HAD REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THIS WAS ADDITIONALLY REPORTED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - IT WAS LATER ADDITIONALLY REPORTED THAT THE SURGEON PERFORMED LENS REPOSITIONING ON (B)(6) 2023 AND "THE PATIENT HAD A POSTOP ON (B)(6) 2023 AND IS NOW SEEING 20/20 AGAIN. IF ANYTHING ELSE CHANGES AFTER HIS POSTOP ON (B)(6) 2023, DR. (B)(6) WILL HAVE US REACH OUT." CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733896&pc=QCB | 2023826-2023-04023 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/+1.5/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS REPORTED THE PATIENT HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH THE SAME MODEL AND LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6A: "ON (B)(6) 2023" SHOULD BE ADDED. D6B: "ON (B)(6) 2023" SHOULD BE ADDED. D3: "08/18/2023" SHOULD BE REMOVED AND "07/26/2023" SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733897&pc=QCB | 2023826-2023-04024 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/+2.0/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. IT WAS REPORTED THE PATIENT HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17751507&pc=QCB | 2023826-2023-04000 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL, SAME LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "LRI WAS PERFORMED FOR ASTIGMATISM." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN WITH FIBER ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17769660&pc=QCB | 2023826-2023-04103 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- SHOULD BE CORRECTED TO STATE: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IN A SEPARATE SURGERY THAT SAME DATE REPLACED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS UNKNOWN. H6: 4627-SHOULD BE ADDED AND 4629 SHOULD BE OMITTED FOR CORRECTION. H6: 1494-OFF-LABEL : IT SHOULD BE NOTED THAT THE PATIENT'S AGE FELL OUTSIDE THE INDICATED AGE RANGE AT THE TIME OF IMPLANTATION. ADDITIONAL INFORMATION: D9: 15-SEP-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DAMAGED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND SPLIT AND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE UNABLE TO BE PERFORMED DUE TO THE DAMAGED STATE OF THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17778161&pc=QCB | 2023826-2023-04110 | 2023-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -13.0/4.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17799577&pc=QCB | 2023826-2023-04172 | 2023-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.0/+6.0/23 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED THERE WAS DEVICE CAUSALITY AND "REFRACTIVE IN OCOS WAS -0.25/-5.0X108 BUT SHOULD HAVE BEEN -0.25/-5.0X168 (60 DEGREES DIFFERENT FROM WHAT IT SHOULD HAVE BEEN)". Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT ON 11-SEP-2023 THE LENS WAS REPOSTIONED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814792&pc=QCB | 2023826-2023-04276 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 31-JUL-2023. THE REPORTER REPORTS EXCESSIVE VAULT AND STATES THAT THERE IS STILL RESIDUAL ASTIGMATISM, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/85 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814804&pc=QCB | 2023826-2023-04277 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE REPORTER REPORTS EXCESSIVE VAULT AND STATES THAT THERE IS STILL RESIDUAL ASTIGMATISM, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/92 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814901&pc=QCB | 2023826-2023-04230 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. CAUSE OF THE EVENT IS UNKNOWN. THE DEVICE DID NOT PERFORM AS INTENDED. REPORTEDLY,"OS ICL HAS ROTATED AT 8 WEEKS POST-OP. CAUSE UNKNOWN." STATUS OF EYE IS STATED, "DECREASED VISION, BURNING SENSATION," AND "PATIENT HAD ICL INSERTED ON (B)(6) 2023. SURGERY WAS ROUTINE, AND ICL WAS INSERTED INTO THE SULCUS AT THE TARGET AXIS OF 175 DEGREES (5 DEGREE ROTATION). VISION WAS 6/6 (20/20) POST OP, INCLUDING UPON EXAMINATION ON (B)(6) 2023. ON (B)(6) 2023 THE PATIENT PRESENTED TO THE CLINIC WITH BLURRED VISION WHICH HAD STARTED THAT MORNING AND WAS NOT PRESENT BEFORE GOING TO BE THE NIGHT BEFORE. NO PAIN WAS NOTED BUT PATIENT MENTIONED THE EYE FELT A BURNING SENSATION 4 DAYS EARLIER, PRIOR TO VISION CHANGE. NO TRAUMA WAS NOTED BY THE PATIENT. UPON EXAMINATION THE ICL HAD ROTATED TO AN AXIS OF 70 DEGREES SINCE THE LAST POST-OP VISIT." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822039&pc=QCB | 2023826-2023-04304 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/+2.0/014 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822446&pc=QCB | 2023826-2023-04281 | 2022-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -3.50/1.0/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PER REPORTER THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED "PATIENT DIDNT REACH EXPECTED VA" WAS REPORTED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840454&pc=QCB | 2023826-2023-04286 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -15.50/+3.50/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2014 LENS REPOSITIONING WAS PERFORMED DUE TO MISALIGNMENT AFTER TRAUMA, "POSSIBLE SMALL LENS" REPORTED. PER REPORTER PATIENT WAS RECENTLY SEEN WITH DECREASED VA (NOTICED 1 MONTH AGO), EXAMINTATION SHOWED A DISLOCATED IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840487&pc=QCB | 2023826-2023-04287 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -15.50/+2.50/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. ON (B)(6) 2018 LENS REPOSITIONING WAS PERFORMED DUE TO MISALIGNMENT AFTER TRAUMA, "POSSIBLE SMALL LENS" REPORTED. PER REPORTER PATIENT WAS RECENTLY SEEN WITH DECREASED VA (NOTICED 1 MONTH AGO), EXAMINATION SHOWED A DISLOCATED IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17846529&pc=QCB | 2023826-2023-04241 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS USER ERROR. Manufacturer Narrative: A4-A6:UNK. H6: USER ERROR . CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 4627-EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND RESIDUE ON THE LENS, SPECIFICALLY FIBRE. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848603&pc=QCB | 2023826-2023-04371 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.5/4.0/86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17857611&pc=QCB | 2023826-2023-04373 | 2023-09-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -7.50/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE CHANGE OVER TIME WAS OBSERVED. THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17857612&pc=QCB | 2023826-2023-04374 | 2023-09-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -8.50/1.5/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17857763&pc=QCB | 2023826-2023-04405 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -2.00/1.00/078 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS THE PATIENT WAS UNHAPPY AND THERE WERE SOME VAGUE VISUAL COMPLAINTS. THE SURGEON INDICATED THAT THERE ARE GLARE/HALOES AND BLURRED VISION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE, AND A PATIENT RELATED FACTOR AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862021&pc=QCB | 2023826-2023-04064 | 2022-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -12.0/3.5/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS WAS REPOSITIONED ON (B)(6) 2023, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE(ASSESSED ON (B)(6) 2022). THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE-14-NOV-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#731573. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17874041&pc=QCB | 2023826-2023-04416 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/3.5/175 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION. THE EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATED THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882808&pc=QCB | 2023826-2023-04422 | 2023-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17883015&pc=QCB | 2023826-2023-04448 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON AN UNKNOWN DATE. THE REPORTER INDICATES THE FOLLOWING STATEMENT "ICL POWER IN OS IS MAXIMAL AND CANT BE INCREASED TO BUILD IN CURRENT RESIDUAL REFRACTIVE ERROR/SURPRISE (MYOPIA). LENS REMAIN IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17890610&pc=QCB | 2023826-2023-04449 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -13.0/1.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING, GLARE/HALOES, BLURRED VISION, ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND A LONGER LENGTH LENS WAS IMPLANTED ON THIS SAME DATE AND RESOLVED THE PROBLEM. AN ADDITIONAL SUTURE WAS ALSO LATER USED ON (B)(6) 2023. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "ICL TOO SMALL". Manufacturer Narrative: B5 - NO REPOSITIONING WAS DONE FOR RIGHT EYE (OD). TRACE CATARACT WAS NOTED AS WELL. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17892475&pc=QCB | 2023826-2023-04454 | 2023-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -13.00/1.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. EXCESSIVE VAULT WAS REPORTED AND BLURRED VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936453&pc=QCB | 2023826-2023-04364 | 2023-09-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.3 VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.00/4.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE AND ROTATION WAS OBSERVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6 VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.00/4.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. CLAIM #(B)(4). Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936654&pc=QCB | 2023826-2023-04366 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -16.50/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936660&pc=QCB | 2023826-2023-04367 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -17.50/2.50/091(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936727&pc=QCB | 2023826-2023-04544 | 2023-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/012 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL OBSERVATION FOUND HAPTIC TORN. THE RETURNED LENS DID NOT MEET THE ORIGINAL VALUE MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954380&pc=QCB | 2023826-2023-04626 | 2023-09-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -13.50/4.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17966720&pc=QCB | 2023826-2023-04648 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -5.0/1.0/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT), ACD <3.00MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17970392&pc=QCB | 2023826-2023-04668 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.5/2.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH, SPHERICAL LENS WITH DIFFERENT POWER AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (AGE < 21). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971214&pc=QCB | 2023826-2023-04634 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A TORIC IMPLANTABLE COLLAMER LENS PATIENT HAS A 2D CLY REMAINING. IT WAS REPORTED THAT THE PATIENT IS ABOUT ONE MONTH POST-OP. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17981364&pc=QCB | 2023826-2023-04658 | 2023-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -6.00/4.0/024 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-DEVICE CODE 1494: OFF-LABEL USE: PRE-EXISTING KERATOCONUS; PREVIOUS CORNEAL OR REFRACTIVE SURGERY. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17981402&pc=QCB | 2023826-2023-04659 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -6.00/4.0/024 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-DEVICE CODE 1494: OFF-LABEL USE: PRE-EXISTING KERATOCONUS; PREVIOUS CORNEAL OR REFRACTIVE SURGERY. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986850&pc=QCB | 2023826-2023-04595 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH DIFFERENT MODEL AND LONGER LENGTH LENS DUE TO ROTATION ASSOCIATED WITH LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6: PATIENT CODE: 2137 - REFRACTIVE SURPRISE CORRECTED TO 2923 LENS ROTATION MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) ADDED TO INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17998684&pc=QCB | 2023826-2023-04774 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.50/2.5/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999785&pc=QCB | 2023826-2023-04730 | 2023-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN WITH THE DEVICE NOT PERFORMING AS INTENDED. Manufacturer Narrative: WEIGHT-RACE:UNK WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18008643&pc=QCB | 2023826-2023-04811 | 2023-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE CHANGE OVERTIME. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017561&pc=QCB | 2023826-2023-04789 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN EVO VISIAN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON AN UNKNOWN DATE. REPORTER STATED "SEEING THINGS A LITTLE FUZZY." THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18030398&pc=QCB | 2023826-2023-04843 | 2023-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/2.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036112&pc=QCB | 2023826-2023-04812 | 2023-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/+3.5/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, LIGHT ANISOCORIA/BLURRY VISION. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18052034&pc=QCB | 2023826-2023-04847 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS, IMPLANTED VERTICALLY DUE TO REFRACTIVE SUPRISE AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18058847&pc=QCB | 2023826-2023-04795 | 2023-09-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -5.50/4.5/047 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED.THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18066237&pc=QCB | 2023826-2023-04856 | 2023-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.5/2.5/007 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE FIRST TICL IS IN A VERTICAL POSITION." Manufacturer Narrative: H6: OFF-LABEL - ACD<3.0. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075201&pc=QCB | 2023826-2023-04920 | 2023-09-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.0/1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE, BLURRED VISION, AND EXCESSIVE VAULT WAS OBSERVED. LENS REMIANS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-12.0/+1.0/80 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075202&pc=QCB | 2023826-2023-04919 | 2023-09-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.0/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE, BLURRED VISION, AND EXCESSIVE VAULT WAS OBSERVED. LENS REMIANS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075217&pc=QCB | 2023826-2023-04927 | 2023-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.5/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2023. A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18077635&pc=QCB | 2023826-2023-04797 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE(OD) ON (B)(6) 2019. REFRACTIVE CHANGE OVER TIME WAS REPORTED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18077667&pc=QCB | 2023826-2023-04798 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)V2019. REFRACTIVE CHANGE OVER TIME WAS REPORTED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18080760&pc=QCB | 2023826-2023-04994 | 2023-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.5/009 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18101041&pc=QCB | 2023826-2023-05056 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -8.00/5.5/020 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. THE REPORTER STATES " ASTIGMATISM AXIS WAS WRONG". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18125522&pc=QCB | 2023826-2023-05097 | 2023-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4.50/+1.0/125 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT HAD LOW VAULTING; REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED THAT "CONSULTING WITH MA WHAT TO DO WITH THE VAULT AND REFRACTIVE SURPRISE". LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129704&pc=QCB | 2023826-2023-05001 | 2023-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -13.0/2.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129705&pc=QCB | 2023826-2023-05004 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6. -3/0.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE/CHANGE OVERTIME WAS ASSESSED ON (B)(6) 2020 AND (B)(6) 2023. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6. -3/0.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE/CHANGE OVERTIME WAS ASSESSED ON (B)(6) 2020 AND (B)(6) 2023. LENS WAS EXCHANGED FOR A SAME LENGTH DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION ADDED TO FORM. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+0.5/96 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129709&pc=QCB | 2023826-2023-05000 | 2023-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -13.0/1.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136692&pc=QCB | 2023826-2023-05059 | 2023-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.50/4.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS REPORTED. PER UPDATED INFORMATION LENS ROTATION WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18151978&pc=QCB | 2023826-2023-05087 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.0/3.0/132 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE IS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18166188&pc=MTA | 2023826-2023-05169 | 2023-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THERE IS REFRACTIVE SURPRISE IN A PATIENT WHO HAD SURGERY IN LATE MARCH/EARLY (B)(6) 2023. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.5/2.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON PERFORMED LASIK SURGERY TO REDUCE REMAINING CYLINDER POST ICL AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18166188&pc=QCB | 2023826-2023-05169 | 2023-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THERE IS REFRACTIVE SURPRISE IN A PATIENT WHO HAD SURGERY IN LATE MARCH/EARLY (B)(6) 2023. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.5/2.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON PERFORMED LASIK SURGERY TO REDUCE REMAINING CYLINDER POST ICL AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18170503&pc=QCB | 2023826-2023-05148 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.00/+4.0/141 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE, LENS CONTINUES TO ROTATE. LENS REMAINS IMPLANTED. FOR CAUSE DETAILS THE REPORTER STATES THAT THE LENS MAY BE A SMALL SIZE FOR THE PATIENTS EYE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18170905&pc=QCB | 2023826-2023-05149 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.50/+2.5/045 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE, LENS CONTINUES TO ROTATE. LENS REMAINS IMPLANTED. FOR CAUSE DETAILS THE REPORTER STATES THAT THE LENS MAY BE A SMALL SIZE FOR THE PATIENTS EYE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18173823&pc=QCB | 2023826-2023-05218 | 2023-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.5/+5.5/73 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6 2023. THE SURGEON REPORTS THERE IS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING; REFRACTIVE SURPRISE AND ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND NOTED "CALCULATION ERROR CAUSED BY USE OF THE WRONG AXIS; AFTER REPOSITION INTO THE INTENDED AXIS; THE LENS ROTATES AGAIN WITHOUT LOW VAULT". Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18173829&pc=QCB | 2023826-2023-05217 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179664&pc=QCB | 2023826-2023-05220 | 2023-10-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -2.00/1.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188684&pc=QCB | 2023826-2023-05254 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4.50/-1.50/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER AND SPHERICAL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188687&pc=QCB | 2023826-2023-05255 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -2.00/-0.50/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER AND SPHERICAL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND THE DEVICE. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18230975&pc=QCB | 2023826-2023-05364 | 2023-10-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.60/4.5/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 05-JUL-2023. REFRACTIVE SURPRISE IS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258190&pc=QCB | 2023826-2023-05310 | 2023-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.0/2.5/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT DIOPTER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE(ASSESSED ON (B)(6) 2023). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18268549&pc=QCB | 2023826-2023-05386 | 2023-11-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.5/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2023 AS AN EXCHANGE LENS. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MDR # 2023826-2023-01948 FOR INITIAL LENS. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271595&pc=QCB | 2023826-2023-05503 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.00/0.5/021 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18279460&pc=QCB | 2023826-2023-05477 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.5/2.5/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A LONGER LEGNTH LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18298063&pc=QCB | 2023826-2023-05433 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MMVTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -6.00/1.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18298064&pc=QCB | 2023826-2023-05434 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MMVTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -6.00/1.0/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18314824&pc=QCB | 2023826-2023-05633 | 2023-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/1.5/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER AND AXIS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18314840&pc=QCB | 2023826-2023-05632 | 2023-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -5.5/2.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON 26-OCT-2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18332220&pc=QCB | 2023826-2023-05584 | 2023-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.0/3.5/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS A PATIENT-RELATED-FACTOR WITH THE DEVICE NOT PERFORMING AS INTENDED. STATUS OF THE EYE IS REPORTED AS, "OK." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. PATIENT-RELATED-FACTOR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18340519&pc=QCB | 2023826-2023-05680 | 2023-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -6.0/1.0/050 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18352639&pc=QCB | 2023826-2023-05699 | 2023-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 13.7MM, VTICM513.7, -8.5/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18352702&pc=QCB | 2023826-2023-05700 | 2023-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 13.7MM, VTICM513.7, -9.00/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369677&pc=QCB | 2023826-2023-05707 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: WE WERE MADE AWARE OF AN ARTICLE THAT WAS PUBLISHED TITLED, 'CLINICAL MANAGEMENT OF A NONOPTIMAL VAULT AND UNEXPECTED POSTOPERATIVE REFRACTION IN A PATIENT IMPLANTED WITH A PHAKIC INTRAOCULAR LENS'. THE ARTICLE REPORTS "BILATERAL LOW VAULT AND UNILATERAL VA REDUCTION (20/40) BECAUSE OF POSTOPERATIVE OBLIQUE ASTIGMATISM". AN EXCHANGE FOR A LONGER LENGTH LENS OF DIFFERENT POWER WAS PERFORMED FOR THE LEFT EYE, EXCHANGE INCREASED THE VAULT AND IMPROVED VA. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369992&pc=QCB | 2023826-2023-05623 | 2023-11-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -10.00/3.5/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18378510&pc=QCB | 2023826-2023-05717 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.5/3.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION AND REFRACTIVE SURPRISE BUT IT DID NOT RESOLVE THE PROBLEM, LENS ROTATED AGAIN. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387275&pc=QCB | 2023826-2023-05783 | 2023-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.0/1.0/56 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE REPLACED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18432640&pc=QCB | 2023826-2024-00106 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/3.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LENS ROTATION AND REFRACTIVE SURPRISE. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "BLUR VISION FOR DISTANCE AND NEAR". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18438447&pc=QCB | 2023826-2024-00120 | 2023-10-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -6.0/+6.0/19 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON NOTES THERE IS LENS ROTATION AND THE PATIENTS VISION IS BLURRY. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18438448&pc=QCB | 2023826-2024-00121 | 2023-09-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -11.00/+1.5/118 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT HAS LOW VAULTING, REFRACTIVE SURPRISE AND BLURRED VISION. LENS REMAINS IMPLANTED. THE SURGEON ADDITIONALLY NOTES THE MAIN CONCERN IS THE LOW VAULTING. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18452220&pc=QCB | 2023826-2024-00108 | 2023-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -14.50/2.0/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON 12-NOV-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE.THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. IT WAS REPORTED THAT THE FIRST LENS WAS IN VERTICAL POSITION. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18454578&pc=QCB | 2023826-2024-00109 | 2023-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -13.50/1.5/007 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 12-NOV-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. IT WAS REPORTED THAT BOTH LENSES WAS IN VERTICAL POSITION. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18458880&pc=QCB | 2023826-2024-00111 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Blurred Vision | P030016 | 1 | Event Description: LENS ROTATION POST TORIC ICL IMPLANTATION RESULTING IN AN RX OF -0.25+1.50X142 WAS REPORTED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473804&pc=QCB | 2023826-2024-00135 | 2021-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/4.0/082 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2015. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18499009&pc=QCB | 2023826-2024-00211 | 2023-12-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICM513.2, -5.5/0.5/042 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18505077&pc=QCB | 2023826-2024-00118 | 2023-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.00/2.00/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. IT WAS REPORTED THAT THE PROBLEM RESOLVED. PATIENT-RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THE EVENT. HOWEVER, IT WAS REPORTED THAT THE LENS DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18521107&pc=QCB | 2023826-2024-00229 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM, VTICM512.6 -1.00/5.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525784&pc=QCB | 2023826-2024-00174 | 2023-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -6.0/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 05/DEC/2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014078&pc=QCB | 2023826-2020-02992 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.50/+6.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE LENS WAS DISCARDED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11270315&pc=QCB | 2023826-2021-00283 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/110 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS TEMPORAL IRIS ATROPHY-MYDRIASIS ASPECT AND IRIS PIGMENT DISPERSION. THE PATIENT COMPLAINS ABOUT THE QUALITY OF VISION. Manufacturer Narrative: DATE OF EVENT-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788215&pc=QCB | 2023826-2021-01350 | 2021-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+6.0/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DATE A SURGICAL PI WAS ADDED, THE SURGEON REPORTING EXCESSIVE VAULT, REFRACTIVE SURPRISE, SIGNIFICANT REDUCTION OF ICA, AND BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED IS DEVICE: "THE ICL PHAKIC LENS IS LONGER THAN THE WTW MEASURE IN HER EYE, THE EXCESSIVE VAULT IS CAUSING SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES MERITING A PERIPHERAL IRIDECTOMY. ALSO THERE WAS A REFRACTIVE SURPRISE." THE SURGEON HAS DECIDED TO LEAVE THE LENS IMPLANTED, KEEPING THE PATIENT UNDER SURVEILLANCE. Manufacturer Narrative: B5-ADDITIONAL INFO: THE DOCTOR REPORTS THE PATIENT IS IMPROVING REFRACTION AND LEAVES THE PATIENT WITH A SPHERE OF +0.50 WHICH IS BETTER THAN THE PUBLICATION RESULT. IRIDOTOMIES WERE PERFORMED AND THE PATIENT IS STABLE. THE DOCTOR CONSIDERS THE CASE CLOSED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147111&pc=QCB | 2023826-2021-02410 | 2021-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/161 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVERTIME AND TRANSIENT LOSS OF BCVA. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE, RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS FIBER ON LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4) Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. REFRACTIVE CHANGE OVERTIME. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12182327&pc=QCB | 2023826-2021-02729 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 13.2MM, VTICM5_13.2, -11.0/+6.0/148, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH KERATOCONUS. IT WAS REPORTED THAT EXCESSIVE VAULT, REFRACTIVE SURPRISE, NARROWING OF ANGLES DUE TO VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: ON (B)(6) 2021;. LENS WAS REMOVED AND A SHORTER LENGTH LENS WAS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12874528&pc=QCB | 2023826-2021-04516 | 2021-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -15.0/2.0/88, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS. ON PATIENTS POST-OP VISIT ON (B)(6) 2021 FOR STATUS OF THE EYE THE REPORTER REPORTED "BLURRY VISION". ON (B)(6) 2021 VISIT "DECOMPRESSION RETINOPATHY AFTER ICL EXCHANGE" WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPORTER STATED "THE INITIAL ICL WAS TOO LARGE. AFTER REPLACEMENT EMERGENTLY, THE RETINOPATHY OCCURRED." STEROIDS AND NSAIDS WERE PRESCRIBED, IT WAS REPORTED THAT APPROXIMATELY 90% OF THE BLURRY VISION AND RETINOPATHY HAS RESOLVED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12879841&pc=QCB | 2023826-2021-04537 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -6.50/1.5/92, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, BLURRED VISION AND UNREACTIVE PUPIL. THIS EXCHANGE RESOLVED THE PROBLEM. NO PATIENT INJURY WAS REPORTED. "OTHER" WAS REPORTED TO BE THE CAUSE OF THIS EVENT, FOR CAUSE DETAILS THE REPORTER STATED " ICL TOO LARGE." Manufacturer Narrative: B5 - CORRECTED TO :"THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -6.50/1.5/90, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, BLURRED VISION AND UNREACTIVE PUPIL. THIS EXCHANGE RESOLVED THE PROBLEM. NO PATIENT INJURY WAS REPORTED. "OTHER" WAS REPORTED TO BE THE CAUSE OF THIS EVENT, FOR CAUSE DETAILS THE REPORTER STATED " ICL TOO LARGE." CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286578&pc=QCB | 2023826-2022-00175 | 2021-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -6.50/+2.50/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; BLURRY VISION AND THE LENS WAS EXCHANGED ON (B)(6) 2022 BUT THIS DID NOT RESOLVE THE PROBLEM. IT WAS LATER CLARIFIED THAT THE PATIENT STILL HAS FUZZY VISION; PUPILS ARE DILATED AND PRK SURGERY WAS PERFORMED. PATIENT STILL HAS BLURRY VISION. CAUSE OF THE EVENT WAS REPORTED AS "OTHER- UNEXPECTED SHORT CILLARY SULCUS SPACE". SEE MDR# 2023826-2022-00177 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311391&pc=QCB | 2023826-2022-00177 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6; -6.00/+2.50/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. IT WAS REPORTED THE PATIENT FELT IMPROVEMENT BUT VISION IS STILL FUZZY; THE PUPIL WAS STILL DILATED AFTER SURGERY. AT A LATER POSTOP VISIT THE PATIENT HAD -1.00/+1.00/107 VISION;THE SURGEON HAD PERFORMED PRK SURGERY. PATIENT HAS REPORTED VISION IS STILL BLURRY. LENS REMAINS IMPLANTED. SEE MDR# 2023826-2022-00175 FOR INITIAL LENS IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459220&pc=QCB | 2023826-2022-00387 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/113 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE AND DIZZINESS. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES, "AS OVERCORRECTION OF ASTIGMATISM, THE PATIENT FELT DIZZY AND COULDN'T ADAPT TO IT, SO A NON-TORIC ICL WAS REPLACED." Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869343&pc=QCB | 2023826-2022-00886 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+6.0/159 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT FELT HIS VISION WAS DISTORTED AND BLURRED. HIS FINAL ASTIGMATISM WAS EXTREMELY HIGH AT -6.50D. THE SURGEON PRESCRIBED PILOCARPINE TO SEE IF THIS WOULD IMPROVE HIS VISION. THE LENS REMAINED IMPLANTED. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02729. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061336&pc=QCB | 2023826-2022-01131 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -12.5/+1.0/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS BLURRED VISION, IRIS SPHINCTER LOSS AT 5 O'CLOCK, HEAVILY PIGMENTS TM, KIV EXPLANT, AND IRREGULAR PUPIL. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: B5- PER EMAIL ON 23-JUN-2022 THE REPORTER STATED THAT "THERE IS NO ADVERSE EVENT FOR OS". CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14863456&pc=QCB | 2023826-2022-02037 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS OF -9.0/4.0/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, LENS DISLOCATION/SUBLUXATION, AND REFRACTIVE CHANGE OVERTIME WERE OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DAY IN A SEPARATE SURGERY A NEW MODEL/SAME LENGTH LENS OF DIFFERENT POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: 0000001876. UNK. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15386494&pc=QCB | 2023826-2022-03049 | 2022-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/+0.5/004 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "THE SURGEON SUSPECTED THAT THE LENS WAS WITH THE WRONG LABEL". Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN, AND THE RETURNED LENS IS NOT THE SIZE STATED IN THE CLAIM. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465942&pc=QCB | 2023826-2022-03251 | 2022-08-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/+1.0/103 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN 2019. IT WAS REPORTED THE PATIENT HAS REFRACTIVE SURPRISE AND "PATIENT HAS BEEN PUNCHED IN THE EYE". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801908&pc=QCB | 2023826-2022-04084 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED O (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, BLURRED VISION AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THE INTRAOCULAR PRESSURE (IOP) WAS ELEVATED. INITIATED COMBIGAN BID AND PRED FORTE BID. THE ELEVATED IOP WAS RESOLVED AND WAS MEASURED WITHIN NORMAL RANGE ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS DUE TO LENS WAS TOO SMALL. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15813337&pc=QCB | 2023826-2022-03932 | 2022-09-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -8.5/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION IN IRIDO-CORNEAL ANGELS, AND BLURRED VISION.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTION DATA D4: REPORTED EXPIRATION DATE WAS 31-MAY-2022 - CORRECT TO 31-MAY-2025. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856726&pc=QCB | 2023826-2022-04213 | 2022-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: 1494- OFF LABEL USE: PRE-EXISTING CATARACTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856830&pc=QCB | 2023826-2022-04214 | 2022-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: 1494- OFF LABEL USE: PRE-EXISTING CATARACTS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15882930&pc=QCB | 2023826-2022-04339 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/1.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; GLARE/HALOES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "MOST PROBABLY THE POSTERIOR SULCUS IS NARROWER THAN AVERAGE STATISTICALLY EXPECTED"; THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15882947&pc=QCB | 2023826-2022-04338 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/1.5/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; GLARE/HALOES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "MOST PROBABLY THE POSTERIOR SULCUS IS NARROWER THEN AVERAGE STATISTICALLY EXPECTED"; THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15898502&pc=QCB | 2023826-2022-04377 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND COMPLAINED OF DOUBLE VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16309073&pc=QCB | 2023826-2023-00443 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THERE WAS A HYPEROPIC OUTCOME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VTICM5_13.2; -6.00/+1.00/157 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTES THERE IS A HYPEROPIC OUTCOME/REFRACTIVE SURPRISE; IT WAS ADDITIONALLY INDICATED LATER THAT LENS HAD ALSO ROTATED. LENS REMAINS IMPLANTED WITH A PLANNED LENS EXCHANGE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45 YRS OF AGE AT DATE OF IMPLANT AND HISTORY OF REFRACTIVE SURGERY). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16358874&pc=QCB | 2023826-2023-00497 | 2022-12-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON NOTES THERE IS A HYPEROPIC OUTCOME; IT WAS ADDITIONALLY INDICATED LATER THAT LENS HAD ALSO ROTATED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION: H3: HEALTH EFFECT CLINICAL CODE: 4581 - HYPEROPIC SHOULD BE CORRECTION TO H6 -HEALTH EFFECT CLINICAL CODE: 4581 - HYPEROPIC. ADDITONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VTICM5_13.2; -6.00/+1.00/157 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTES THERE IS A HYPEROPIC OUTCOME/REFRACTIVE SURPRISE; IT WAS ADDITIONALLY INDICATED LATER THAT LENS HAD ALSO ROTATED. LENS REMAINS IMPLANTED WITH A PLANNED LENS EXCHANGE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: 13.2MM VTICM5_13.2; -6.00/+1.00/157 SHOULD BE UPDATED TO 13.2MM VTICM5_13.2; -6.00/+1.00/162. CLAIM#: (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - HYPEROPIC. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17083434&pc=QCB | 2023826-2023-02276 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.5/3.5/77 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT EXPERIENCED EXCESSIVE VAULTING, HALO/GLARES AND IRIS TRANSILLUMINATION. IT WAS INDICATED THAT IT IS PLANNED TO MONITOR THE PATIENT OVER 6 MONTHS. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - IRIS TRANSILLUMINATION. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295356&pc=QCB | 2023826-2023-02833 | 2023-07-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; Optical Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -12.0/+1.0/135 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON NOTES THERE IS LENS GLISTENING'S ON THE ICL, THE LENS WAS OBSERVED WITH GLISTENING'S/MILD OPACIFICATION MATERIAL ON THE LENS. PATIENT COMPLAINTS OF SLIGHT GLARE/DISTURBANCE DURING DIM LIGHTS CONDITION. LENS REMAINS IMPLANTED. THE SURGEON REPORTS THE PATIENT WILL BE MONITORED WITH NO PLANNED INTERVENTION AT THIS TIME. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIPPER -12.0/+1.0/135 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED GLARES/HALOES, BLURRED VISION, AND 'LENS GLISTENINGS'. REPORTEDLY "BLURRING OF VISION DURING DIM LIGHT CONDITION AND DR YONG FINDS THERE IS A LAYER OF GLISTENING/MILD OPACIFICATION ON THE LENS SURFACE." THE LENS REMAINS IMPLANTED. REPORTEDLY "THERE IS NO PLANNED SURGICAL INTERVENTION AT THIS TIME". THE CAUSE OF THE EVENT IS UNKNOWN. CAUSE DETAILS PROVIDED: "THE ICL LENS IS OBSERVED WITH GLISTENINGS/MILD OPACIFICATION WHERE PT COMPLAINTS SLIGHT MORE GLARE AND DISTURBANCE DURING DIM LIGHTS CONDITION." HEALTH EFFECT CLINICAL CODE: "4581 LENS GLISTENING" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18080354&pc=QCB | 2023826-2023-04850 | 2023-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem; Unintended Movement | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.5/122 (SPHERE/CYLINDER/AXIS) LENS TORE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA, ON (B)(6) 2023. LENS DISLOCATION/SUBLUXATION AND DOUBLE VISION ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "HAPTICS CUT." THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: A4-A6: UNK. H6: OFF-LABEL: PRE-EXISTING MYOPIA. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136617&pc=QCB | 2023826-2023-05053 | 2023-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/100, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: "4581-DOUBLE VISION" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-10.0/+1.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REPORTEDLY THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE PATIENT EXPERIENCED BLURRED VISION AND DOUBLE VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LENGTH, AND POWER LENS AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18437834&pc=QCB | 2023826-2024-00029 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/101(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND A SMALLER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS OTHER. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH - ONE SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND: 731987. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18446384&pc=QCB | 2023826-2024-00050 | 2023-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -5.5/1.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPORTEDLY IMPLANTED UPSIDE DOWN; SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES; BLURRED VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS AND THE PROBLEM WAS RESOLVED. CAUSE IS USER ERROR. . Manufacturer Narrative: H6-CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082445&pc=QCB | 2023826-2021-04976 | 2021-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON STATES THERE IS A HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON STATES THERE IS A HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT INFORMATION - UNK. DATE OF THE EVENT -UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE - UNK. (B)(4). Manufacturer Narrative: PATIENT INFORMATION - UNK. DATE OF THE EVENT -UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4). Manufacturer Narrative: B5- ADDITIONAL INFORMATION: THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+3.5/083 (SPHERE/CYLINDER/AXIS) LENS INTO THE PATIENT'S EYE ON (B)(6)2021. THE SURGEON REPORTS EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE CAUSE IS REPORTED AS DEVICE: "OVERALL LENS DIAMETER TOO BIG." H6-CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER AND THE PROBLEM RESOLVED. FOR STATUS OF THE EYE (SIGN/SYMPTOMS) "MILD IRIS ATROPHY SUBINCISIONAL" WAS REPORTED. HOWEVER, PER UPDATED INFORMATION THE REPORTER STATED "IT'S JUST PAR FOR THE COURSE" CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627814&pc=QCB | 2023826-2022-01824 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -9.50/1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO HIGH VAULT. Manufacturer Narrative: ADDITO NAL INFORMATION: B5 - IT WAS ADDITIONALLY LATER REPORTED THAT THE PATIENT ALSO EXPERIENCED BLURRED VISION AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "SIZING". THE PROBLEM WAS RESOLVED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627817&pc=QCB | 2023826-2022-01825 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -9.00/1.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO HIGH VAULT. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - IT WAS ADDITIONALLY LATER REPORTED THAT THE PATIENT ALSO EXPERIENCED BLURRED VISION AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16251274&pc=QCB | 2023826-2023-00242 | 2022-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -4.5/1.0/018 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). RESIDUAL MYOPIA WAS OBSERVED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50/1.0/018 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND "RESIDUAL MYOPIA". REPORTEDLY " RESIDUAL MYOPIA AND HE WOULD LIKE TO DO AN ICL EXCHANGE". ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS "WELL PLACED EVO ICL; EYE IS QUIET". ADDITIONAL INFORMATION PROVIDED WAS "PT'S COMPLAINTS RESOLVED WITH ICL EXCHANGE". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS "OTHER" AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY "OVER ACCOMMODATION DURING PRE-OP MR". CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4581- "RESIDUAL MYOPIA." CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411020&pc=QCB | 2023826-2023-00568 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/179 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. POST-OP, THE PATIENT'S VA WAS BLURRY, SEE MFR. REPORT # 2023826-2023-00569 FOR REPLACEMENT LENS. THE PATIENT HAD INTACS IMPLANTED AND WAS USING ARTIFICIAL TEARS. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY & RACE UNKNOWN. EVENT DATE UNKNOWN. ADVERSE EVENT PROBLEM WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS ROTATED (REPOSITIONED) DESPITE NORMAL VAULT. PATIENT'S VA WAS MUCH BLURRIER AT INITIAL POST-OP, WHICH IS WHY THE LENS WAS REMOVED AND REPLACED WITH LONGER LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.00/2.5/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT, REPOSITIONING, BLURRED VISION AND INTACS INTRACORNEAL RINGS. ON (B)(6) 2022 THE LENS WAS REPOSITIONED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE; "CLEAR VA AND HAD ADJUSTED TO HAZE/SHADOW USING ARTIFICIAL TEARS PRN." ADDITIONAL INFORMATION PROVIDED WAS "CLEAR VA AT NIGHT, DILATED EXAM SHOWS INTACS INTACT". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN- " ROTATED VISIAN ICL DESPITE NORMAL VAULT. EXPLANT AND INSERTION OF LONGER ICL TO HELP PREVENT ROTATION. VA MUCH BLURRIER DURING POST OP". H6- HEALTH EFFECT- CLINICAL CODE: 4581-INTACS INTRACORNEAL RINGS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411273&pc=QCB | 2023826-2023-00566 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. POST-OP, THE PATIENT'S VA WAS BLURRY, SEE MFR. REPORT # 2023826-2023-00567 FOR REPLACEMENT LENS. THE PATIENT HAD INTACS IMPLANTED AND WAS USING ARTIFICIAL TEARS. Manufacturer Narrative: WEIGHT: UNKNOWN; ETHNICITY: UNKNOWN; RACE: UNKNOWN; DATE OF EVENT: UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS ROTATED (REPOSITIONED) DESPITE NORMAL VAULT. PATIENT'S VA WAS MUCH BLURRIER AT INITIAL POST-OP, WHICH IS WHY THE LENS WAS REMOVED AND REPLACED WITH LONGER LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/3.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT, LENS REPOSITIONING, BLURRED VISION AND INTACS (INTRACORNEAL RINGS). ON (B)(6) 2022, THE LENS WAS REPOSITIONED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE; "CLEAR VA AND HAD ADJUSTED TO INTACS USING ARTIFICIAL TEARS PRN." ADDITIONAL INFORMATION PROVIDED WAS "CLEAR VA AT NIGHT, DILATED EXAM SHOWS INTACS INTACT". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN- " ROTATED VISIAN ICL DESPITE NORMAL VAULT. EXPLANT AND INSERTION OF LONGER ICL TO HELP PREVENT ROTATION. VA MUCH BLURRIER DURING POST OP". H6- HEALTH EFFECT- CLINICAL CODE: 4581-INTACS INTRACORNEAL RINGS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093103&pc=QCB | 2023826-2023-02235 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.5/3.5/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, "1 WEEK POST OP KERATOCONUS PATIENT'S OUTCOME". Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "4581- KERATOCONUS" AND "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/3.5/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD).THE PATIENT EXPERIENCED KERATOCONUS AND RESIDUAL REFRACTIVE ERROR. REPORTEDLY "1 WEEK POST OP KERATOCONUS PATIENT'S OUTCOME". THE LENS REMAINS IMPLANTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093104&pc=QCB | 2023826-2023-02234 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -14.50/3.5/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "1 WEEK POST OP KERATOCONUS PATIENT'S OUTCOME". Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: "4581- KERATOCONUS" AND "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/3.5/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED KERATOCONUS AND RESIDUAL REFRACTIVE ERROR. REPORTEDLY "1 WEEK POST OP KERATOCONUS PATIENT'S OUTCOME". THE LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18008540&pc=QCB | 2023826-2023-04810 | 2023-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND SEEMS TO HAVE HIGH VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL "KERATOCONUS" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.00/2.0/165 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. REPORTEDLY "POST-OPERATIVELY "HIGH VAULTS & REFRACTIVE SURPRISE" AND "2 PATIENTS I RECENTLY DID TORIC ICLS AND UNFORTUNATELY THEY BOTH SEEM TO HAVE HIGH VAULT. THEY ARE BOTH KERATOCONIC WITH HIGH PRESCRIPTIONS SO THEY HAD HIGH POWER LENSES, AND ALTHOUGH THEIR VAULT IS >1000MICRONS THEY HAVE GOOD IOPS AND OPEN ANGLES."THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.00/2.0/165 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND SEEMS TO HAVE HIGH VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540747&pc=QCB | 2023826-2022-03403 | 2022-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.0/1.5/024 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO OVERCORRECTION AND DISCOMFORT IN LOOKING CLOSE. THE EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500060&pc=QCB | 2023826-2023-00803 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.50/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 TO ADDRESS EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. IT WAS REPORTED THAT THE VAULT ISSUE RESOLVED BUT REPORTER STATES THAT THE PATIENT IS VERY UNCOMFORTABLE AND DISAPPOINTED WITH HER REFRACTION RESULTS, LENS REMAINS IMPLANTED. ADDITIONAL COMMENTS FROM REPORTER- "GOOD JUMP BUT REFRACTIVE ERROR. GOOD LENGTH OF THE IMPLANT BUT NOT HYPERMETROPIA POWER. PATIENT WITH KERATOCONUS OD/OS, STRABISM - RETINA LASER." Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND DISCOMFORT. REPORTEDLY THE PATIENT IS "VERY UNCOMFORTABLE AND DISAPPOINTED WITH REFRACTION RESULTS". ADDITIONAL COMMENTS FROM REPORTER- "GOOD JUMP BUT REFRACTIVE ERROR. GOOD LENGTH OF THE IMPLANT BUT NOT HYPERMETROPIA POWER. PATIENT WITH KERATOCONUS OD/OS, STRABISM - RETINA LASER." THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9541911&pc=QCB | 2023826-2020-00022 | 2019-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/6.00/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. IT WAS REPORTED THAT THE PATIENT COMPLAINED OF DISCOMFORT RIGHT AFTER SURGERY. THE REPORTER STATED " THE PATIENT CONDITION IS GOOD AND HAVE SAME CONDITION BEFORE THE SURGERY, BCVA 20/25." Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATENT CODE 2330 SHOULD BE ADDED TO THE INITIAL MDR. H6- CONCLUSION CODE 4310 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9541936&pc=QCB | 2023826-2020-00023 | 2019-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/6.00/78 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. IT WAS REPORTED THAT THE PATIENT COMPLAINED OF DISCOMFORT RIGHT AFTER SURGERY. THE REPORTER STATED " THE PATIENT CONDITION IS GOOD AND HAVE SAME CONDITION BEFORE THE SURGERY, BCVA 20/25." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 2330 SHOULD BE ADDED TO THE INITIAL MDR. H6- CONCLUSION CODE 4310 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882069&pc=QCB | 2023826-2023-04269 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -4.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2023. BLURRED VISION; SEVERE DISTENDING SENSATION;PAIN; DIZZINESS WAS REPORTED AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS REPOSITIONED THEN ON SAME DAY REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: H6-CLINICAL CODE 4581- SEVERE DISTENDING SENSATION; DIZZINESS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882376&pc=QCB | 2023826-2023-04270 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -4.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE(OS) ON 07-JUL-2023. BLURRED VISION; SEVERE DISTENDING SENSATION;PAIN; DIZZINESS WAS REPORTED AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS REPOSITIONED THEN ON SAME DAY REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLINICAL CODE 4581- SEVERE DISTENDING SENSATION; DIZZINESS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515456&pc=QCB | 2023826-2020-02030 | 2020-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Halo | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, 3.00/4.0/120 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH AND SIMILAR POWER LENS WHICH RESOLVED THE ISSUE. CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED LENS ROTATION. THIS SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563252&pc=QCB | 2023826-2021-00692 | 2019-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.50/5.5/122 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME WERE OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS REMOVED ON (B)(6) 2021. CAUSE OF THE EVENT IS REPORTED AS DEVICE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -5.50/5.5/122 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PATIENT HAD PRE-EXISTING KERATOCONUS AND HISTORY OF PREVIOUS NON-REFRACTIVE CORNEAL SURGERY PRIOR TO IMPLANT. REFRACTIVE SURPRISE; REFRACTIVE CHANGE OVERTIME; BLURRED VISION (HAZE); HALOES WERE OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14962949&pc=QCB | 2023826-2022-02195 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Halo | P030016 | 1 | Event Description: AN ARTICLE PUBLISHED TITLED "A NATIONWIDE MULTICENTER STUDY ON 1-YEAR OUTCOMES OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS IMPLANTATION FOR LOW MYOPIA"; WAS TO ASSES THE OUTCOME OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS IMPLANTATION WITH A CENTER HOLE (EVO-ICL; STAAR SURGICAL) FOR PATIENT'S WITH LOW MYOPIA. THE STUDY COVERS 172 EYES OF 111 PATIENTS. IT MENTIONS FOUR EYES REQUIRED REROTATION OF TORIC ICLS; FOUR EYES COMPLAINED OF SYMPTOMATIC GLARE OR HALO DURING THE EARLY POSTOPERATIVE PERIOD BUT THESE SYMPTOMS RECOVERED OVER TIME. LENSES REMAIN IMPLANTED. SEE MFR# 2023826-2022-02196 FOR AN ADDITIONAL EYE RELATED TO THIS ARTICLE. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567718&pc=QCB | 2023826-2022-03573 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 -16.0/2.0/124 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. PATIENT EXPERIENCED HALOS AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15847565&pc=QCB | 2023826-2022-04169 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A SMALLER SIZE LENS. THE PROBLEM WAS RESOLVED. THE POST-OP RESULTS WERE VAULT IS 75%, UCVA 20/100, BCVA 20/80, NO EDEMA AND AC IS CLEAR. AT THE (B)(6) 2022 POST-OP VISIT, THE PATIENT WAS UCVA 20/70 -2 @ DISTANCE AND NEAR WAS CF. HIS REFRACTION WAS +1.25 -0.50 X 077, BCVA 20/70 AND THE VAULT WAS AT 100%. PATIENT WILL RETURN FOR A YAG PC ON (B)(6) 2023 DUE TO PIGMENT DUSTING ACROSS THE ICL. THE PATIENT IS DOING WELL. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376164&pc=QCB | 2023826-2023-00550 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS NOTED "REMOVE AND NO LONGER IMPLANT". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376167&pc=QCB | 2023826-2023-00551 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS NOTED "REMOVE AND NO LONGER IMPLANT". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792366&pc=QCB | 2023826-2023-01443 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Halo; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO HIS LEFT EYE (OS) IN (B)(6) 2022. THE PATIENT IS DISSATISFIED, COMPLAINS OF HALOES, PAIN AND FEELS NO OPTIMAL VISUAL ACUITY. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002296&pc=QCB | 2023826-2020-00994 | 2017-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -10.5/1.5/49 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT; B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES); B5- BS THE REPORTER INDICATED THAT A VTICM5_12.6, -10.5/1.5/49 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2020, THE LENS WAS EXCHANGED DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- (B)(6) 2020. H1- SERIOUS . CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION- LENS RETURNED IN A MICRO- CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002391&pc=QCB | 2023826-2020-00995 | 2017-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -11.5/1.0/18 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-BS THE REPORTER INDICATED THAT A VTICM5_12.6, -11.5/1.0/18 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS EXCHANGED DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION- LENS RETURNED IN A MICRO- CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708891&pc=QCB | 2023826-2020-02442 | 2019-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/+2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LENS ROTATION AND BLURRED VISION. REPORTEDLY, THE PATIENT MAY HAVE "OVERLY RUBBED HIS EYES" AFTER SCUBA DIVING. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/+2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PATIENT MAY HAVE "OVERLY RUBBED HIS EYES" AFTER SCUBA DIVING. THE EXCHANGED RESOLVED THE PROBLEM. D6B- (B)(6) 2020. H6-IMPACT CODE 4629; LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED MICROCENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708899&pc=QCB | 2023826-2020-02441 | 2019-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LENS ROTATION AND BLURRED VISION. REPORTEDLY, THE PATIENT MAY HAVE "OVERLY RUBBED HIS EYES" AFTER SCUBA DIVING. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 OF -12/2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. D6B- (B)(6) 2020. H6-IMPACT CODE 4629; LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED MICROCENTRIFUGE VIAL; RETURN DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT WITH DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11033257&pc=QCB | 2023826-2020-03079 | 2020-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -13.00/+3.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B(6) 2020 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED, THE REPORTER STATED "VAULT IS 0.75CT, TO OBSERVE IN 3 MONTHS". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - REFRACTIVE SURPRISE WAS REPORTED. APPLYING FOR EXCHANGE, REPLACEMENT SURGERY. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852537&pc=QCB | 2023826-2020-02679 | 2020-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED DUE TO BLURRED VISION. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11158612&pc=QCB | 2023826-2021-00094 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/162 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020 DUE TO BEING TOO SHORT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED AND EXCHANGED DUE TO LOW VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: A2: DATE OF BIRTH - (B)(6) 1979 B3: DATE OF EVENT -(B)(6) 2020 CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231262&pc=QCB | 2023826-2021-00192 |  | Injury | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A VISIAN TORIC IMPLANTABLE COLLAMER LENS, AND THE LENS ROTATED. THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE LENS WAS REPOSITIONED BACK INTO POSITION BUT AGAIN ROTATED BACK TO THE WRONG POSITION. THE PATIENT COMPLAINED OF BLURRED VISION. THE PATIENT CHOSE TO HAVE PHOTOREFRACTIVE KERATECTOMY (PRK) SURGERY AND THE LENS REMAINED IMPLANTED. THE PATIENT IS HAPPY WITH THE VISUAL OUTCOME AFTER PRK. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PRK SURGERY. CLAIM # (B)(4). Manufacturer Narrative: EXPIRATION DATE UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238061&pc=QCB | 2023826-2021-00175 | 2020-12-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -12.00/2.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -12.00/2.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651612&pc=QCB | 2023826-2021-00914 | 2021-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.50/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: G3: DATE ON FOLLOW-UP #1 CHANGED FROM (B)(6) 2021 TO (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: B5 - INCLUDE BLURRED VISION WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651657&pc=QCB | 2023826-2021-00915 | 2021-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -10.00/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5 - INCLUDE BLURRED VISION WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654117&pc=QCB | 2023826-2021-00986 | 2020-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A TMICL12.6 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2020. LOW VAULT, REFRACTIVE SURPRISE, AND LENS ROTATION WERE OBSERVED. REPORTEDLY, LENS REPOSITION IS PLANNED TO TAKE PLACE ON (B)(6) 2021. THE LENS REMAINS IMPLANTED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED, THAT THE SURGEON IMPLANTED A TMICL12.6 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2020. LOW VAULT, REFRACTIVE SURPRISE, AND LENS ROTATION WERE OBSERVED. ON (B)(6) 2021, THE LENS WAS EXCHANGED. REPORTEDLY, THE SURGERY WENT WELL. AND THE PATIENT IS HAPPY. THE REPORTER STATES, THE PATIENT'S RIGHT EYE IS 20/20. CLAIM#: (B)(4). Manufacturer Narrative: B5 - LENS REPOSITIONING WAS ALSO PERFORMED ON (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION DID NOT FIND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683204&pc=QCB | 2023826-2021-01077 | 2020-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 13.7MM, TMICL13.7, -14.0/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. FOR CAUSE OF EVENT OTHER: THE LENS WAS TOO LARGE FOR THE PATIENT. WAS REPORTED. Manufacturer Narrative: B5- IT WAS REPORTED THAT A 13. 7MM, TMICL 13. 7, -14.0/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND BLURRED VISION. A LENS OF SHORTER LENGTH WAS LATER IMPLANTED, AND THE PROBLEM RESOLVED. FOR CAUSE OF EVENT OTHER: THE LENS WAS TOO LARGE FOR THE PATIENT. WAS REPORTED. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11750617&pc=QCB | 2023826-2021-01295 | 2017-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B3: DATE OF EVENT IS CORRECTED TO (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED, THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. "THE PATIENT COMPLAINED OF BLURRED VISION. AND THE AUTOMATIC REFRACTOR RESULTS SHOWS -1.00DC*10 -> 1.0. THE PATIENT WHO STRONGLY DEMANDED, FOR TICL REPLACEMENT SURGERY". ON (B)(6) 2017, THE LENS WAS REMOVED. AND ON THE SAME DAY, BUT DIFFERENT SURGERY, A REPLACEMENT LENS OF THE SAME MODEL/SIZE AND DIFFERENT POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K : THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11781966&pc=QCB | 2023826-2021-01397 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: A TICL ROTATION CASE WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.00/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT REFRACTIVE TREATMENT (PRK) WAS PERFORMED AND THE PROBLEM RESOLVED. CAUSE FOR THIS EVENT WAS PATIENT RELATED FACTOR AND IT WAS ALSO REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: - RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE (2021 REPORTS): UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891166&pc=QCB | 2023826-2021-01675 | 2021-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -13/2/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (0S) ON (B)(6) 2021. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION (UNKNOWN WHETHER VAULT WAS LOW), IT WAS REPORTED THAT REPOSITIONING RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGN/SYMPTOMS): "RESOLVED" WAS REPORTED. PER STUDY INVESTIGATOR, THE CAUSE OF THIS EVENT WAS RELATED TO THE DEVICE. FOR ADDITIONAL INFORMATION THE REPORTER STATED "PER MEDICAL MONITOR OF NOTE, LENS POWER CALCULATION RECOMMENDED BY STAAR WAS TMICL 13.7, -13.00/2.0/080. HOWEVER, THE DOCTOR CHOSE TO GO WITH THE TMICL 13.2." Manufacturer Narrative: B5- PER UPDATED INFORMATION "AS OF (B)(6) 2021, LENS ROTATED AGAIN AND SUBJECT HAD INCREASED CYLINDER TO BASELINE (2.75 VS. BASELINE 2.25). SLE: ROTATED ICL @ 087 (83 DEGREES OF ROTATION). CLAIM # (B)(4). Manufacturer Narrative: B5 - PER UPDATED EMAIL, PROBLEM IS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12104944&pc=QCB | 2023826-2021-02220 | 2021-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/3.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "BLURRED VISION, RESIDUAL REFRACTION, AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED." SHOULD BE ADDED TO INITIAL MDR. H6- MEDICAL DEVICE PROBLEM CODE: 1401 - MISFOCUSING ADDED TO INITIAL MDR. H6- CLINICAL CODE: 2137 - BLURRED VISION ADDED TO INITIAL MDR CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171281&pc=QCB | 2023826-2021-02615 | 2021-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A TORIC IMPLANTABLE COLLAMER LENS ROTATED AT ONE WEEK POST-OP. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.00/3.0/094, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THIS EXCHANGED RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED AND THAT PATIENT RELATED-FACTOR WAS THE CAUSE OF THIS EVENT. THE REPORTER STATED " WTW LESS THAN ACTUAL SULCUS SIZE". WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171287&pc=QCB | 2023826-2021-02614 | 2021-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A TORIC IMPLANTABLE COLLAMER LENS ROTATED AT ONE WEEK POST-OP. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.00/2.0/086, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THIS EXCHANGED RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED AND THAT PATIENT RELATED-FACTOR WAS THE CAUSE OF THIS EVENT. THE REPORTER STATED " WTW LESS THAN ACTUAL SULCUS SIZE". H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478395&pc=QCB | 2023826-2021-03564 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON IS PLANNING AN EXPLANT AND EXCHANGE OF THE LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/035 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND PATIENT HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: AGE OR DOB: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EVENT DATE: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973345&pc=QCB | 2023826-2021-04718 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +3.5/+3.0/124 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +3.5/+3.0/124 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: B5-PREVIOUSLY STATED THE LENS ROTATED. CORRECTED INFORMATION: ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973359&pc=QCB | 2023826-2021-04717 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +4.5/+1.5/038 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-PREVIOUSLY STATED THE LENS ROTATED. CORRECTED INFORMATION: ON (B)(6) 2021 THE LENS WAS REPOSTIONED DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. THE LENS WAS INITIALLY IMPLANTED ON (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13515929&pc=QCB | 2023826-2022-00417 | 2021-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+5.5/111 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+5.5/111 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+5.50/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS OF THE SAME LENGTH AND DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON IT. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13546061&pc=QCB | 2023826-2022-00435 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.00/1.5/99 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IN THE REPORTERS OPINION "PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.00/1.5/99 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. IN THE REPORTER'S OPINION "PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. H6- MEDICAL DEVICE PROBLEM CODE: 2923 TO 1401- PREVIOUS CODE NOT APPLICABLE, THERE WAS NO LENS ROTATION PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13657019&pc=QCB | 2023826-2022-00667 | 2021-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.0/1.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION WHICH WAS NOTED ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS WHICH WAS ALSO IMPLANTED VERTICALLY WHICH HAD RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: CORRECTION: B5 - PATIENT ALSO EXPERIENCED BLURRED VISION - SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664194&pc=QCB | 2023826-2022-00659 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: REPORTEDLY, AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE AND THE LENS ROTATED 32 DEGREES 5 DAYS POSTOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 12.6MM VTICMO12.6 -14.5/+2.5/066 (SPHERE/CYLINDER/AXIS) ICL WAS IMPLANTED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE SURGEON REPORTS ADDITIONAL SYMPTOMS OF LOW VAULT, UNSTABLE REFRACTIVE CHANGE OVER TIME, BLURRED VISION, AND "REFRACTIVE ERROR KEEPS CHANGING, ICL IS PROPELLERING." THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY EXPIRATION DATE, IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14657149&pc=QCB | 2023826-2022-01841 | 2021-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1; -10/4.0/123 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THERE WAS LENS ROTATION AND ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT HAD BLURRED VISION. AC IRRIGATION WAS ALSO PERFORMED. -THIS SHOULD HAVE BEEN INCLUDED IN PREVIOUS MFR. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15022388&pc=QCB | 2023826-2022-02269 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -6.50/2.0/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- "PATIENT ALSO EXPERIENCED REFRACTIVE CHANGE OVER TIME." CLAIM: (B)(4). Manufacturer Narrative: CORRECTED DATA: H4- "03/NOV/2013 SHOULD BE CORRECTED TO 11/MAR/2013." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: "THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/2.0/89 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULT WITHOUT ROTATION AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H6- HEALTH EFFECT- IMPACT CODE: 2199 SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. INVESTIGATION TYPE- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071012&pc=QCB | 2023826-2022-02325 | 2021-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Insufficient Information | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND IT WAS REPORTED THAT THERE IS A PLANNED ICL EXCHANGE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.0/4.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. THE REPORTER ALSO STATED THAT THE LENS DID NOT FAIL TO PERFORM AS INTENDED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15103186&pc=QCB | 2023826-2022-02446 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -13.0/3.5/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -13.0/3.5/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A VERTICALLY IMPLANTED; SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. H6- DEVICE CODE: 1494- OFF LABEL USE (AGE<21) . CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6- TYPE OF INVESTIGATION CODE; 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118180&pc=QCB | 2023826-2022-02283 | 2021-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/171 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH THE SAME LENGTH BUT DIFFERENT AXIS LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: D6A (B)(6) 2017. D6B- UNK. B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/171 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON AN UNKNOWN DATE THE LENS WAS REPLACED WITH THE SAME LENGTH BUT DIFFERENT AXIS LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/171 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH THE SAME LENGTH BUT DIFFERENT AXIS LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136758&pc=QCB | 2023826-2022-02345 | 2021-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -3.5/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233578&pc=QCB | 2023826-2022-02726 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+4.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT (UNDESIRABLE REFRACTIVE OUTCOME). THE LENS WAS REPOSITIONED ON (B)(6) 2022 AND THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINED IMPLANTED. THE AC IS DEEP AND QUIET. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233848&pc=QCB | 2023826-2022-02738 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+4.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 -11.0/+4.50/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS REPLACED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: D9: RETURN DATE: 16-JAN-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15246687&pc=QCB | 2023826-2022-02795 | 2022-06-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -11.0/4.0/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN THE YEAR 2014. REFRACTIVE SURPRISE; REFRACTIVE CHANGE OVER TIME WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15246688&pc=QCB | 2023826-2022-02796 | 2022-06-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -8.0/3.50/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN THE YEAR 2014. REFRACTIVE SURPRISE; REFRACTIVE CHANGE OVER TIME WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: 1991. IMPLANT DATE: 2014. (B)(4). NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15285627&pc=QCB | 2023826-2022-02521 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -1.0/2.5/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED DATA: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15294643&pc=QCB | 2023826-2022-02798 | 2022-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -4.0/1.5/63 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299118&pc=QCB | 2023826-2022-02801 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/6.0/025 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. STATUS OF THE EYE IS NORMAL AC/ NORMAL PUPIL SIZE/ V. HIGH ASTIGMATISM. " IT IS A 5Y POST PKP CASE A PRE-OP CORRECTED VISION REACHES 20/20, FOLLOW UP DONE OVER 2 MONTHS WITH STABLE REF. OF 0.0/-8.0/AX 105." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL/CORRECTED DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/6.0/015 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND PRK WAS PERFORMED. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "NORMAL AC/ NORMAL PUPIL SIZE/ V. HIGH ASTIGMATISM". ADDITIONAL INFORMATION REPORTED WAS "IT IS 5Y POST PKP CASE WITH A PRE OP CORRECTED VISION REACHED 20/20. FOLLOW UP DONE OVER 2 MONTHS WITH A STABLE REF OF 0.0/-18.0/AX 105 ALTHOUGH THE LEFT EYE DONE WITH AN EXCELLENT RESULT." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15480682&pc=QCB | 2023826-2022-03282 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: REPORTEDLY, "UNSTERILE." Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS WITH A, -14.0/2.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15555029&pc=QCB | 2023826-2022-03516 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS DID NOT RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "THE PATIENT IS 45 YEARS OLD, PRESBYOPIA -1.50D, PROFESSIONAL ACCOUNTANT, USUALLY SEE MORE CLOSE, THE MAIN COMPLAINT OF NEAR VISION IS STILL POOR VISION TO SOLVE THE NEAR VISION, SO 2 OPERATIONS." SEE MFR# 2023826-2022-03312 FOR EXCHANGE LENS. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "THE PATIENT IS 45 YEARS OLD, PRESBYOPIA -1.50D, PROFESSIONAL ACCOUNTANT, USUALLY SEE MORE CLOSE, THE MAIN COMPLAINT OF NEAR VISION IS STILL POOR VISION TO SOLVE THE NEAR VISION, SO 2 OPERATIONS." MFR# 2023826-2022-03312 FOR EXCHANGE LENS IS NOT A COMPLAINT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15566370&pc=QCB | 2023826-2022-03332 | 2022-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+2.0/155 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE REPORTER INDICATED THE LENS HAD ROTATED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: UDI# - (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+2.0/155 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED AND REPLACED FOR A NEW LENS ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM # (B)(4). Manufacturer Narrative: B5 UPDATE - THE LENS WAS EXCHANGED FOR A FOR A NEW LENS OF DIFFERENT POWER ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCING A REFRACTIVE SURPRISE AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK , RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624407&pc=QCB | 2023826-2022-03547 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -08.50/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCAHNGED WITH THE SAME MODEL; STRONGER POWER; SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15795261&pc=QCB | 2023826-2022-04002 | 2022-09-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.50/2.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON AN UNKNOWN DATE. ON 01-SEPT-2022 THE LENS WAS EXCHANGED FOR A SAME MODEL/ SIZE LENS BUT DIFFERENT POWER OF -10.50/3.0/061 (SPHERE/CYLINDER/AXIS) DUE TO REFRACTIVE SURPRISE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.50/2.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS IMPLANTED; REMOVED AND REPLACED INTRAOPERATIVELY FOR A SAME MODEL/ SIZE LENS BUT DIFFERENT POWER OF -10.50/3.0/061 (SPHERE/CYLINDER/AXIS) DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. " A WRONG LENS WAS ORDERED RESULTING IN A REFRACTIVE SURPRISE." H6: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15849968&pc=QCB | 2023826-2022-04260 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.5/5.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS). THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: D6A- IMPLANT DATE IS (B)(6) 2022. CORRECTED DATA: H10 IN PREVIOUS MDR SHOULD CONTAIN THE FOLLOWING WORK ORDER SEARCH RESULTS. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+5.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. REPORTEDLY "HE DID NOT HAVE CYCLOREFRACTION. THE PREOP SPECTACLE FOR THE LEFT EYE WAS -8.00/-6.50 X 80 (BACK VERTEX DISTANCE 13MM). THE BEST CORRECTED LEFT VISION WAS 6/7.5 HIS REFRACTION WAS CHECKED AND THE CYL CONFIRMED AS STABLE. YES THERE IS A SIGNIFICANT RESIDUAL OCULAR ASTIGMATISM (ROA) BUT WE DECIDED TO TREAT ON REFRACTION. THERE WAS NO CAUSE FOUND TO ACCOUNT FOR THE DISPARITY. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4- MODEL#: "VTICM5 12.6" SHOULD BE CORRECTED TO "VTICMO 12.6" CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035423&pc=QCB | 2023826-2022-04548 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO : " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, "THE PATIENT'S POSTOPERATIVE UCVA WAS NOT GOOD AND COULDN'T REACH THE PREOPERATIVE CDVA. THERE WAS HIGH VAULT AND REFRACTIVE RESIDUAL, SO THE REFRACTIVE POWER WAS ADJUSTED AT THE SAME TIME WHEN THE LENS WAS REPLACED'. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN." IN THE INITIAL MDR. H6 - HEALTH EFFECT - CLINICAL CODE: 2137 SHOULD BE ADDED TO THE INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16065467&pc=QCB | 2023826-2022-04741 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: D9-RETURN DATE: 12-JAN-2023 H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND HAPTIC DEFORMED AND STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075819&pc=QCB | 2023826-2022-04772 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON NOTES A POSSIBLE SECONDARY INTERVENTION DUE TO RESIDUAL CYL. AND REQUESTED ASSISTANCE FROM A CONSULTATION DOCTOR. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H6: CLINICAL CODE: 2137 SHOULD BE REMOVED FROM THE INITIAL MDR. H6: CLINICAL CODE: 4582 SHOULD BE ADDED TO THE INITIAL MDR. H6: MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE REMOVED FROM THE INITIAL MDR. H6: MEDICAL DEVICE PROBLEM CODE: 2923 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16383938&pc=QCB | 2023826-2023-00488 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 12-JUL-2022. LOW VAULT WAS OBSERVED. ON 15-JAN-2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT AND REFRACTIVE SUPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - HEALTH EFFECT CLINICAL CODE: 2137 SHOULD BE ADDED TO THE INITIAL MDR. H6- MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16397843&pc=QCB | 2023826-2023-00564 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+1.0/149 (SPHERE/CYLINDER/AXIS), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED THAT VISION WAS NOT GOOD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND SIGNIFICANT BLURRING OF VA. POST-OP, THE BCVA WAS MUCH IMPROVED. THE CAUSE OF THE EVENT WAS THE DEVICE AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A HAPTIC TORN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CORRECTED DATA: B3: DATE OF EVENT: (B)(6) 2023. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16742891&pc=QCB | 2023826-2023-01252 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REMOVED ON(B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: UNK. UNK. UNK. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE:"2993" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16809012&pc=QCB | 2023826-2023-01477 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -10.5/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON 02-MAR-2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO LENS ROTATION. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS EXCHANGED DUE TO REFRACTIVE CHANGE OVERTIME. REPORTEDLY, "LOW VAULT RESULTED IN ROTATION." LENS WAS REMOVED AND REPLACED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS "DEVICE" AND "OTHER." THE REPORTER NOTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. STATUS OF THE EYE IS REPORTED AS "DEEP AND QUIET." CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H3: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A VIAL. RETURNED PRODUCT EVALUATION FOUND NO VISIBLE DAMAGE TO THE LENS. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16871559&pc=QCB | 2023826-2023-01595 | 2023-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.50/+2.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BENT. DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16945974&pc=QCB | 2023826-2023-01839 | 2023-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. Manufacturer Narrative: A4-A6: UNK. B3: UNK. D4 (EXPERIATION DATE); UNK. NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -4.00/2.0/055 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS OTHER. REPORTEDLY, "ASYMPTOMATIC" AND "IN FUTURE VISITS WE WILL SEE THE EVOLUTION." H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - BLURRED VISION WAS ALSO NOTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950559&pc=QCB | 2023826-2023-01857 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+5.0/168 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17019319&pc=QCB | 2023826-2023-01948 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H5- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. ADDITIONAL DATA: B5- REPORTEDLY "EXCHANGE IS PLANNED END OF OCTOBER." CLAIM# (B)(4). Manufacturer Narrative: ADDITIOAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LENGTH LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT IS HAPPY". D6B.- (B)(6) 2023 SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE:"2993" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17056313&pc=QCB | 2023826-2023-02124 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.0/4.0/24 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS REPOSITIONING WAS PERFORMED AFTER A LENS ROTATION OCCURED. THE LENS ROTATED AGAIN. REPORTEDLY, "THIS LENS ICL SUFFERED A ROTATION OF APPROXIMATELY 16 DEGREES. MORE INFORM THAT THE LENS IN QUESTION WAS ALREADY REPOSITIONED.". Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. ADDITIONAL DATA: B3- "(B)(6) 2023" SHOULD BE ADDED. B7- "PROGRESSIVE MYOPIA"SHOULD BE ADDED D6B. "(B)(6) 2023" SHOULD BE ADDED H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+4.0/024 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS ROTATION AND REPOSITIONING. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THAT DID NOT RESOLVE THE PROBLEM. REPORTEDLY "THIS LENS ICL SUFFERED A ROTATION OF APPROXIMATELY 16 DEGREES. MORE INFORM THAT THE LENS IN QUESTION WAS ALREADY REPOSITIONED IN THE AXIS INDICATED IN THE DIAGRAM AND POSITIONED IN THE SAME PLACE". ON (B)(6) 2023 THE LENS WAS EXPLANTED. REPORTEDLY "THE DOCTOR WANT TO IMPLANT ANOTHER ONE TO SOLVE THE PROBLEM". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FIBRE ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119692&pc=QCB | 2023826-2023-02161 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -9.0/4.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. LENS WAS DIRECTLY ON THE ANTERIOR LENS SURFACE. Manufacturer Narrative: CORRECTED DATA: B5 - THE PATIENT HAD BLURRY VISION AS A RESULT OF THE WRONG IMPLANTED LENS - SHOULD HAVE BEEN INCLUDED IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119698&pc=QCB | 2023826-2023-02168 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.0/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. TRANSLENT LOSS OF BCVA AND LOW VAULT WITH ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2023 AND DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2023. A LONGER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: 2137 ADDED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119699&pc=QCB | 2023826-2023-02167 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.0/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. TRANSLENT LOSS OF BCVA AND LOW VAULT WITH ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2023 AND DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2023. A LONGER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: 2137 ADDED FOR CORRECTION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17145853&pc=QCB | 2023826-2023-02393 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.50/2.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2023. THE SURGEON NOTES A POSSIBLE LENS EXCHANGE DUE TO LENS ROTATION. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE ON (B)(6) 2023. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17151996&pc=QCB | 2023826-2023-02323 | 2023-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+3.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AMD RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/3.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) /2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17167504&pc=QCB | 2023826-2023-02462 | 2023-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.0/3.50/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. IT WAS REPORTED THAT THIS LENS WAS INCORRECTLY IMPLANTED INTO THE PATIENT'S LEFT EYE, AS IT WAS INTENDED FOR THE PATIENT'S RIGHT EYE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE REPLACED WITH A SAME MODEL/LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE PATIENT HAD BLURRY VISION AS A RESULT OF THE WRONG IMPLANTED LENS - SHOULD HAVE BEEN INCLUDED IN PREVIOUS MDR. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213812&pc=QCB | 2023826-2023-02508 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -5.0/4.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED DUE TO REFRACTIVE SURPRISE. REPOSITIONING DID NOT RESOLVE THE ISSUE. CAUSE OF THE EVENT IS UNKNOWN. H6: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322215&pc=QCB | 2023826-2023-02772 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, "TORIC EVO HAS ROTATED." Manufacturer Narrative: B5: A 12.6MM VTICM512.6 WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING MEIBOMIAN GLAND DYSFUNCTION. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. REPORTEDLY, "PT C/O BLURRED VISON OS." THE PROBLEM HAS NOT RESOLVED. H6: OFF-LABEL - PRE-EXISTING MEIBOMIAN GLAND DYSFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM #(B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17385337&pc=QCB | 2023826-2023-03155 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.00/3.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTES THERE WAS LENS ROTATION, THE LENS HAD ROTATED TO 161 DEGREES, 16 DEGREES OFF THE DESIRED LOCATION OF 177. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.00/3.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTES THERE WAS LENS ROTATION, THE LENS HAD ROTATED TO 161 DEGREES, 16 DEGREES OFF THE DESIRED LOCATION OF 177 AND REFRACTIVE CHANGE OVER TIME WAS OBSERVED. LENS REMAINS IMPLANTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669070&pc=QCB | 2023826-2023-03736 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, -4.5/5.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION INFERIOR WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS ROTATED. THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE, LENS TOO SMALL, KEEPS ROTATING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: H6- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE ADDED. D6B. "ON (B)(6) 2023" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER +4.50/5.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. PRODUCT NONCONFORMITY WAS REPORTED. REPORTEDLY "THE ICL ROTATED INFERIORLY TWICE NOW" AND "ICL KEEPS ON ROTATING AND PREFERS TO SETTLE VERTICALLY". THE PATIENT EXPERIENCED BLURRED VISION AND INFERIOR ROTATION OF ICL X2. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THAT DID NOT RESOLVE THE PROBLEM. REPORTEDLY " ICL HAS ROTATED TWICE NOW. ONCE CLOCKWISE AND THE SECOND TIME COUNTER CLOCKWISE. ANTERIOR SEGMENT SCANS WILL BE PROVIDED". ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH THE SAME MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "THE ICL IS TOO SMALL AND HAS NOW ROTATED INFERIORLY TWICE. IT WAS ROTATED BACK ONCE BUT HAS AGAIN ROTATED INFERIORLY". CLAIM#: (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17927921&pc=QCB | 2023826-2023-04478 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -14.50/3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO THE LENS BEING TOO SMALL. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, THE LENS ROTATED 90 DEGREES WITH THE SIZE OF THE LENS BEING PROBABLE ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT MUCH HAPPIER WITH IMPROVED VISION. ICL IN PROPER POSITION." CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17927922&pc=QCB | 2023826-2023-04479 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.0/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO THE LENS BEING TOO SMALL. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, THE LENS ROTATED 90 DEGREES WITH THE SIZE OF THE LENS BEING PROBABLE ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT MUCH HAPPIER WITH IMPROVED VISION. ICL IN PROPER POSITION." CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18088953&pc=QCB | 2023826-2023-04764 | 2023-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18275560&pc=QCB | 2023826-2023-05475 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -7.0/1.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LOW VAULT AND REFRACTIVE MISS WERE REPORTED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: 1 SIMILAR COMPLAINT: CLAIM# (B)(4) (FELLOW EYE). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18296805&pc=QCB | 2023826-2023-05487 | 2023-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.5/3.5/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE OBSERVED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18296806&pc=QCB | 2023826-2023-05486 | 2023-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/50 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE OBSERVED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18430037&pc=MTA | 2023826-2024-00047 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON 11-SEPT-2023. REPORTER STATES THE WRONG LENS WAS IMPLANTED. ON 13-SEPT-2023 REPORTER STATES THE "GOOD LENS IMPLANTED." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2024-00046 FOR INITIAL LENS. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18430037&pc=QCB | 2023826-2024-00047 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON 11-SEPT-2023. REPORTER STATES THE WRONG LENS WAS IMPLANTED. ON 13-SEPT-2023 REPORTER STATES THE "GOOD LENS IMPLANTED." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2024-00046 FOR INITIAL LENS. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4803722&pc=QCB | 2023826-2015-00169 | 2015-01-16 | Injury | STAAR SURGICAL COMPANY | 2015-05-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.0/+2.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXCHANGED ON (B)(6) 2015 WITH A SAME MODEL AND SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH, DIOPTER POWER AND SPHERE WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYES MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, WRONG LENS, UPSIDE-DOWN LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4). MDR RESUBMITTED PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6486516&pc=QCB | 2023826-2017-00626 | 2016-10-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2017-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED A 12.1 MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +1.5/3/15 WAS IMPLANTED IN PATIENT'S RIGHT EYE ON (B)(6) 2016. ON (B)(6) 2017, THE REPORTER INDICATED REFRACTIVE SURPRISE AND LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT. THE LENS REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: B2-CORRECTED. B5-PREVIOUSLY LISTED A SPHERE/CYLINDER/AXIS VALUE OF +1.5/3/15. CORRECTED VALUE: +1.5/+3.5/015. ADDITIONAL INFORMATION REPORTED: THE PATIENT "HAS A LOT OF PIGMENTATION" ON HER ICL SURFACE AND PIGMENT DISPERSION. H6-PATIENT CODE 3191: NO CODE AVAILABLE (PIGMENTATION ON ICL, PIGMENT DISPERSION). CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794355&pc=QCB | 2023826-2018-01193 | 2018-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+5.5/068 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS ROTATED NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED SEVERAL TIMES BUT KEPT ROTATING AND THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: UPDATED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.50/+5.5/068 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER ORIGINALLY NOTED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED SEVERAL TIMES BUT THE LENS KEPT ROTATING. THE PROBLEM WAS NOT RESOLVED. ON 21-JAN-2019, ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE LENS ROTATED SEVERAL TIMES DUE TO LOW VAULT (100 MICRONS). IT WAS ALSO LATER REPORTED THAT THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY. ADDITIONAL INFORMATION RECEIVED ON 22-SEP-2019 STATED THAT THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS LATER REPLACED ON (B)(6) 2019 FOR A 13.7MM VTIMC5_13.7 LENS OF DIOPTER -10.0/4.5/082 (SPHERE/CYLINDER/AXIS) AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS A PATIENT-RELATED FACTOR. EXPLANT DATE UPDATED TO (B)(6) 2018 PER NEW INFORMATION RECEIVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY. CORRECTED DATA: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.50/+5.5/068 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER ORIGINALLY NOTED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED SEVERAL TIMES BUT THE LENS KEPT ROTATING. THE PROBLEM WAS NOT RESOLVED. ON 21-JAN-2019, ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE LENS ROTATED SEVERAL TIMES DUE TO LOW VAULT (100 MICRONS). THE REPORTER STATED THAT THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8067158&pc=QCB | 2023826-2018-01672 | 2018-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -2.5/5.0/41 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT REPOSITIONING OF THE LENS IS PLANNED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -2.5/5.0/41 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2018 REPOSITIONING OF THE LENS WAS PERFORMED. IT WAS REPORTED THAT ALL IS FINE AND THE PATIENT IS HAPPY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, REPOSITIONING. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182336&pc=QCB | 2023826-2018-01886 | 2018-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+6.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT HAD DECIDED NOT TO HAVE A LVC PROCEDURE TO CORRECT CURRENT REFRACTIVE ERROR AFTER TICL IMPLANTATION. THE PLAN IS TO EXPLANT THE LENS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5-THE REPORTER INDICATED THE "LENS WAS NOT STABLE" ON (B)(6) 2019. ON (B)(6) 2019 THE REPORTER STATED, "THE SURGEON TRIED TO CORRECT THE REFRACTIVE SURPRISE WITH A REPOSITIONING OF BOTH LENSES. FOR THIS HE ROTATED THE MISALIGNED TICLS TO HE AXIS OF THE IODS. AFTER THAT (1DAY POSTOP) FIRST THE REFRACTIONS WERE GOOD AND EVERYTHING WAS FINE. AT THE NEXT EXAMINATION MID OF JANUARY THE TICLS HAVE ROTATED AGAIN TO THE SAME MISALIGNED AXIS AS BEFORE. IN THIS POSITION THE LENSES WERE NOW STABLE AND SO THE SURGEON DECIDED TO MAKE A LASER TOUCH UP TO CORRECT THE RESIDUAL REFRACTION ON BOTH EYES". PATIENT CODE: 3191 SECONDARY SURGICAL INTERVENTION, LASER TOUCH UP. CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE: "2923" NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019. BECAUSE OF STRONG BLEEDING AFTER REMOVING THE LENS OUT OF THE SULCUS, THE SURGEON DECIDED TO WAIT TO IMPLANT THE EXCHANGE LENS FOR A COUPLE OF MONTHS. THE PATIENT AND EYE ARE FINE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED, STRONG BLEEDING CORRECTED DATA: D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298862&pc=QCB | 2023826-2019-00192 | 2018-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.00/+1.5/038 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR AND PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE SURGEON REPOSITIONED THE LENS ON (B)(6) 2019 AND THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307775&pc=QCB | 2023826-2019-00228 | 2019-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -0.50/6.0/052 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT AND REFRACTIVE SURPRISE HAS BEEN REPORTED, IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437900&pc=QCB | 2023826-2019-00452 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO REFRACTIVE SURPRISE, BLURRED VISION, AND LENS ROTATION. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437914&pc=QCB | 2023826-2019-00456 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+1.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT, BLURRED VISION, AND LENS ROTATION. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR -THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC. H6-METHOD CODE ADDED CLAIM #(B)(4). Manufacturer Narrative: THE STATEMENT THAT THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN IS INCORRECT. PLEASE DISREGARD. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. DATE OF EVENT-UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8438505&pc=QCB | 2023826-2019-00435 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.00/+3.5/023 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE UNCORRECTED VISUAL ACUITY (UCVA) WAS LOWER THAN EXPECTED. THE CLINICAL REFRACTION AND UCVA DID NOT MATCH THE EXPECTED REFRACTION OF TICL CALCULATION. THE PATIENT AND SURGEONS ARE EXTREMELY DISSATISFIED WITH THE LOW UCVA THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8442017&pc=QCB | 2023826-2019-00439 | 2018-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, --11.5/+3.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8451630&pc=QCB | 2023826-2019-00504 | 2019-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICH12.6, -0.50/2.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LEN INTO PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REPORTEDLY REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAIN IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: EVENT DESCRIPTION: "THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND BLURRED VISION". DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTION INSPECTION FOUND THE LENS SPHERE AND CYLINDER TO BE WITHIN SPECIFICATIONS BUT FOUND THE LENS AXIS TO BE 3° OUT OF SPECFICIATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DATE OF EVENT IS CORRECTED TO (B)(6)2019 EVENT - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS OF -0.50/2.0/170 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019 . ON (B)(6)2019 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE REPORTER INDICATES THE CAUSE OF THE EVENT IS DEVICE. PATIENT CODE (B)(4). CORRECTED DATA: COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. CLAIM #(B)(4). Manufacturer Narrative: "VTICH12.6" SHOULD BE CORRECTED TO "VTICMO12.6" IN ORIGINAL MDR AND MDR SUPPLEMENTAL #1. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8452986&pc=QCB | 2023826-2019-00488 | 2017-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.5/+0.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON REPORTS REFRACTIVE SURPRISE, THE PATIENT STATING HE HAS "BAD SIGHT". REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: EVENT DESCRIPTION: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. THE LENS WAS EXCHANGED AND PATIENT IS "VERY SATISFIED." SHE SEES 1.2 UNCORRECTED WITH REF. +0.25. EXPLANT DATE: UNK. ADDITIONAL PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8466158&pc=QCB | 2023826-2019-00512 | 2019-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/028 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8466179&pc=QCB | 2023826-2019-00513 | 2019-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.50/+4.0/012 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 OF -4.5/4.0/12 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2019 THE LENS WAS EXCHANGED INTRAOPERATIVELY AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE LENS WAS EXCHANGED ON (B)(6) 2019 WITHIN SAME SURGERY AS REMOVAL OF INITIAL LENS. THE EXCHANGE RESOLVED THE PROBLEM. EXPLANTED DATE IN SUPPLEMENTAL MDR (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: THE STATEMENT " THE LENS REMAINS IMPLANTED" IN ORIGINAL MDR IS NOT APPLICABLE. THE LENS WAS EXPLANTED ON (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. DEVICE CODE 4117: IS NOT APPLICABLE IN ORIGINAL MDR; LENS WAS EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8485090&pc=QCB | 2023826-2019-00552 | 2019-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50/+4.5/062 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (CALCULATION ERROR). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EVENT: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2019. AT A FOLLOW-UP VISIT, THE PATIENT'S VISION HAD IMPROVED AND PATIENT WAS VERY SATISFIED. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8485259&pc=QCB | 2023826-2019-00538 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM,VTICM5_13.2, -5.50/1.0/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. A CHANGE IN REFRACTION WAS OBSERVED, THE LENS REMAINS IMPLANTED. THE REPORTER STATED " THE CORNEA IS SIMPLY THIN, THERE ARE NO KC FFKC, ETC. DIAGNOSIS. THE SURGEON THOUGHT THE REFRACTIVE SURPRISE +1,5/-1,75/176 (B)(6) 2018) POST OPERATIVELY CAUSED BY THE 11 DEGREE MISALIGNMENT OF THE LENS. THEN HE REPOSITIONED THE LENS END OF NOVEMBER AND THOUGHT EVERYTHING WOULD BE GOOD. THE CURRENT REFRACTION OF +1,00/-1,75/172 WITH A CORRECT POSITIONED TICL LIE IOD OF 2° CCW HE CANNOT EXPLAIN AND DO NOT HAVE ANY IDEA ABOUT IT. HE'S ASKING TO RECALCULATE A NEW LENS BASED ON THE CURRENT RX AND TICL POSITION". PATIENT IS WELL AND THE SURGEON PLANS TO REPOSITION THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: SHOULD BE CORRECTED TO "NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT." CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492262&pc=QCB | 2023826-2019-00580 | 2019-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.5/1.5/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS REMAINS IMPLANTED AND THE SURGEON PLANS TO EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISUAL DAMAGE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: EXPLANT DATE: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.5/1.5/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6) 2019 THE LENS WAS EXCHANGE FOR A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. PATIENT CODE 2137 SHOULD BE ADDED TO INITIAL MDR. DEVICE CODE 1401 SHOULD BE ADDED TO INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552784&pc=QCB | 2023826-2019-00666 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, TMICL12.1, -9.0/3.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTER OPINION THE CAUSE OF THE EVEN WAS DUE TO LENS BEING TOO SMALL. REPORTEDLY, REPOSITIONING WAS NOT PERFORMED. HOWEVER, IT WAS REPORTED THAT THE REPOSITIONING RESOLVED THE PROBLEM. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE: 2137 - REFRACTIVE CHANGE OVER TIME. CLAIM# : (B)(4). Manufacturer Narrative: CORRECTED DATA: IF FOLLOW-UP, WHAT TYPE: "ADDITIONAL INFORMATION" IS NOT APPLICABLE AND SHOULD BE CORRECTED TO "CORRECTION" IN PREVIOUS MDR. EVENT: "ON 04 APR 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND REFRACTIVE CHANGE OVER TIME." CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE LENS WAS NOT REPOSITIONED. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. THE VIAL WAS RETURNED TAPED TO PAPER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8563068&pc=QCB | 2023826-2019-00711 | 2019-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +8.5/+2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED, THE SURGEON NOTING LENS ROTATION AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS NOT RESOLVED. CURRENTLY, THE LENS REMAINS IMPLANTED BUT AN EXCHANGE IS PLANNED. Manufacturer Narrative: DEVICE CODE: 1494: OFF-LABEL USE (ACD UNDER 3.00MM AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS " ALL IS FINE". CAUSE IS REPORTED AS PATIENT RELATED FACTOR. EXPLANT DATE IS (B)(6) 2019 PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION REQUIRED; LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579900&pc=QCB | 2023826-2019-00760 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/088 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND HAS RESIDUAL ASTIGMATISM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 - "REFRACTIVE SURPRISE" IN INITIAL MDR NOT APPLICABLE. H6 - PATIENT CODE: 3191 - RESIDUAL ASTIGMATISM. H6 - PATIENT CODE: 2137 - PREVIOUS CODE NOT APPLICABLE. H6 - DEVICE CODE: 1401 - PREVIOUS CODE NOT APPLICABLE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611957&pc=QCB | 2023826-2019-00783 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS REMAINS IMPLANTED. REPORTEDLY "PATIENT IS FINE AND WAITING FOR EXCHANGE." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADVERSE EVENT: PRODUCT PROBLEM IN ORIGINAL MDR IS NOT APPLICABLE. (B)(4). Manufacturer Narrative: THE LENS WAS REPOSITIONED ON (B)(6) 2019, PATIENT STABLE AFTER REPOSITIOING. PER REPORTER ON (B)(6) 2019, "PATIENT IS HAPPY AFTER REPOSITIONING". PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621876&pc=QCB | 2023826-2019-00803 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE DOCTOR SUSPECTED THE REFRACTION OF THE FIRST LENS WAS INCORRECT. Manufacturer Narrative: RESULT CODE: 213, SHOULD BE CORRECTED TO RESULT CODE: 170. CONCLUSION CODE: 4315, SHOULD BE CORRECTED TO CONCLUSION CODE: 25. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE LENS WIDTH WAS IN SPECIFICATION BUT LENGTH AND REFRACTION (FUNCTIONAL) VERIFICATION FAILED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8640314&pc=QCB | 2023826-2019-00905 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND ASTIGMATISM OVERCORRECTION. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/ VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8640353&pc=QCB | 2023826-2019-00910 | 2019-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, - 6.5/+3.0/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE.THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2019 AND WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652062&pc=QCB | 2023826-2019-00914 | 2019-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.5/+4.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: "REPORTEDLY, THE LENS REMAINS IMPLANTED." SHOULD BE CORRECTED TO "THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. PATIENT CODE: 3191- SECONDARY SURGERY INTERVENTION, EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: "PROBLEM PRODUCT" SHOULD BE CORRECTED TO "ADVERSE EVENT" IN THE INITIAL MDR. "06-MAY-2019" SHOULD BE CORRECTED TO "02-MAY-2019" IN THE INTIAL MDR. - "MALFUNCTION" SHOULD BE CORRECTED TO "SERIOUS INJURY" IN THE INITIAL MDR. ADDITIONAL INFORMATION: "REQUIRED INTERVENTION" SHOULD BE ADDED IN THE INTIAL MDR. "LENS ROTATION" SHOULD BE ADDED IN THE INITIAL MDR. DEVICE CODE: 2923- DEVICE DISLODGED/DISLOCATED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652063&pc=QCB | 2023826-2019-00913 | 2019-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: "REPORTEDLY, THE LENS REMAINS IMPLANTED." SHOULD BE CORRECTED TO "THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. PATIENT CODE: 3191- SECONDARY SURGICAL INTERVENTION, EXCHANGE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652121&pc=QCB | 2023826-2019-00912 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. THE SURGEON REPORTS REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, LENS REMAINS IMPLANTED. THE SURGEON STATES THE PATIENT'S VISION DECREASED GRADUALLY WITHIN THE LAST 2 MONTHS; NO PAIN, NO HISTORY OF EYE INJURY. Manufacturer Narrative: CORRECTED DATA: ALL PREVIOUS MDRS SHOULD BE CORRECTED TO SERIOUS DUE TO SECONDARY SURGERY FOR LENS REPOSITIONING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADD "LENS ROTATION WAS REPORTED." SHOULD BE ADDED TO THE INITIAL MDR. DEVICE CODE 2923: DISLODGED OR DISLOCATED; LENS ROTATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652191&pc=QCB | 2023826-2019-00902 | 2019-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, - 17.0/+1.0/083 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SAME MODEL AND LENGTH LENS AND THE PROBLEM WAS RESOLVED. THERE WAS AN ADDITION OF SUTURE. THE EYE IS QUIET AND VISION IS GOOD. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE LENS WIDTH WAS IN SPECIFICATION BUT LENGTH AND REFRACTION (FUNCTIONAL) VERIFICATION FAILED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION THE PROBABLE CAUSE OF THIS EVENT IS RELATED TO MANUFACTURING AND PROCESSING OF THIS LENS. CORRECTIVE ACTIONS ARE BEING DEVELOPED. RESULT CODE: 213 & 114 - SHOULD BE CORRECTED TO "RESULT CODE: 170." CONCLUSION CODE: 4315 - SHOULD BE CORRECTED TO "CONCLUSION CODE: 25." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8662649&pc=QCB | 2023826-2019-00865 | 2019-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA AND REFRACTIVE SURPRISE WAS OBSERVED. UPON FOLLOW-UP ON (B)(6) 2019, THE LENS WAS ROTATED/DISPLACED BY 25 DEGREES NOT RESOLVING THE PROBLEM. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS "SMALL SIZE LENS." AS OF PATIENT'S LAST VISIT ON (B)(6) 2019 "LENS STABLE AFTER EXCHANGE AND VAULT INCREASED TO 500 MICRON." Manufacturer Narrative: DEVICE CODE 1494: OFF LABEL USE: ACD<3.0MM. "UPON FOLLOW-UP ON 06/MAY? 2019" SHOULD BE CORRECTED TO "UPON FOLLOW-UP ON 06/MAY/2019" CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671657&pc=QCB | 2023826-2019-01018 | 2019-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.5/+4.5/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON REPORTS REFRACTIVE SURPRISE, WHICH WAS ATTEMPTED TO BE CORRECTED ON (B)(6) 2018 VIA REFRACTIVE TREATMENT ("LASER TOUCH UP"). THIS WAS NOT SUCCESSFUL AS CURRENTLY THERE IS AGAIN A REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: "REPORTEDLY, THE LENS REMAINS IMPLANTED." SHOULD BE CORRECTED TO "ON 08/08/2019 THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS. THE EXCHANGE RESOLVED THE PROBLEM." IN THE INITIAL MDR. PATIENT CODE 3191: SECONDARY SURGERY, EXCHANGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: RESULT CODE 213 SHOULD BE ADDED TO THE PREVIOUS MDR. CONCLUSION CODE 4310 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671659&pc=QCB | 2023826-2019-01015 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/103 (SPHERE/CYLINDAR/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE REFRACTIVE SURPRISE. THE REPORTER STATES "LENS WAS REMOVED AND WE'LL WAIT FOR ENSURE STABILITY." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8686760&pc=QCB | 2023826-2019-00869 | 2017-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.50/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692896&pc=QCB | 2023826-2019-01063 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.0/+5.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. SEE MFR# 2023826-2019-00057 FOR PREVIOUSLY IMPLANTED LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE . THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 OF -11.0/5.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 AS AN EXCHANGE LENS. THE SURGEON NOTES LOW VAULT; LENS ROTATION; AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED AGAIN ON (B)(6) 2019 FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTER STATED; "PATIENT IS HAPPY." EXPLANT DATE: (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8710343&pc=QCB | 2023826-2019-01099 | 2018-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8710345&pc=QCB | 2023826-2019-01100 | 2018-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. ON (B)(6) 2018 A LONGER LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8712615&pc=QCB | 2023826-2019-01096 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1, -10.50/2.5/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. IT WAS REPORTED THAT ON (B)(6) 2018 THE PATIENT NOTICED A SHARP REDUCTION OF VISUAL ACUITY, THE LENS HAD ROTATED 90 DEGREES. IN (B)(6) 2018 THE LENS WAS REPOSITIONED AND THE PATIENT RECOVERED VISION OF 20/20 WITHOUT GLASSES. A CONSULT ONE YEAR LATER (B)(6) 2019) AGAIN NOTICED DETERIORATION OF VISION AND THE LENS ROTATED AGAIN. THE SURGEON PLANS TO REMOVE AND REPLACE THE LENS WITH A LONGER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8741157&pc=QCB | 2023826-2019-01132 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; Material Split, Cut or Torn; Physical Resistance/Sticking | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO13.7, -15.50/4.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS "CUT AT THE END OF THE INSERTION" INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER INDICATED THE CAUSE OF THE CUT WAS DUE TO THE LENS BEING "STUCK TO FOAM TIP." THE LENS REMAINS IMPLANTED, "AS PART OF IT WAS CUT HENCE THE DIAMETER OF LENS WAS SMALLER AND THE LENS DISLOCATED." REPORTER INDICATED THE "ICL IS IN A VERTICAL POSITION" AND PATIENT'S "VISION IS POOR DUE TO ASTIGMATISM." THE REPORTER INDICATED NO PATIENT INJURY AND PLANS TO EXCHANGE THE LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: "THE LENS REMAINS IMPLANTED" SHOULD BE CORRECTED TO "THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019" IN THE INITIAL MDR. "PLANS TO EXCHANGE THE LENS" SHOULD BE REMOVED IN THE INITIAL MDR. PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION, EXPLANT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8767927&pc=QCB | 2023826-2019-01204 | 2019-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.50/3.50/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, AT PATIENTS LAST VISIT, BLURRED VISION WAS REPORTED AS STATUS OF THE EYE. NO TREATRMENT OR INTERVENTION HAS BEEN PERFORMED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (PIGMENT DISPERSION, ANISOCORIA, LENS REPOSITIONING). LENS REPOSITIONING WAS PERFORMED APPROXIMATELY FIVE WEEKS AFTER IMPLANTATION, POST OP STATUS WAS RECORDED AS " OS ANISOCORIA LIKELY DUE TO PIGMENT DISPERSION/ MILD SPHINCTER DAMAGE". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774160&pc=QCB | 2023826-2019-01193 | 2019-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/050 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). Manufacturer Narrative: EVENT UPDATED TO: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/050 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTED LOW VAULT AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED ON (B)(6) 2019FOR A LONGER LENGTH LENS, VTICM5 13.2MM -13.0/2.5/043 (SPHERE/CYLINDER/AXIS), AND THE PROBLEM WAS RESOLVED. THE PATIENT IS REPORTED TO BE "HAPPY"." (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774161&pc=QCB | 2023826-2019-01191 | 2019-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. REPORTEDLY, THE DOCTOR CHOSE THE INITIAL LENS TO TARGET LOWER CORRECTION, BUT THE ACTUAL CORRECTION WAS EVEN LOWER THAN WHAT HE'D EXPECTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: HE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774162&pc=QCB | 2023826-2019-01194 | 2019-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTED LOW VAULT AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS, VTICMO 13.2MM -15.0/3.5/138 (SPHERE/CYLINDER/AXIS), AND THE PROBLEM WAS RESOLVED. THE PATIENT IS REPORTED TO BE "HAPPY"." PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774166&pc=QCB | 2023826-2019-01192 | 2019-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, SIMILAR SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: ACD < 3.0MM NOT APPLICABLE IN INITIAL MDR (JAPAN APPROVED LENS USE FOR ACD > 2.8MM). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8780383&pc=QCB | 2023826-2019-01181 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.50/3.0/095 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION AND LOW VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY THE SURGEON PANNS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. SURGEON NOTES THAT PATIENT'S "REFRACTION CHANGES AFTER THE ROTATION HAPPENED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2019. REPORTEDLY A LONGER LENGTH LENS WAS IMPLANTED AND 'THE PATIENT IS DOING VERY GOOD.' PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION, LENS EXPLANT). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8780390&pc=QCB | 2023826-2019-01182 | 2019-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.50/3.0/075 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION AND LOW VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY THE SURGEON PANNS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. SURGEON NOTES THAT PATIENT'S "REFRACTION CHANGES AFTER THE ROTATION HAPPENED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT CODE 2137 NOT APPLICABLE IN INITIAL MDR. DEVICE CODE 1401 NOT APPLICABLE IN SUPPLEMENTAL MDR #1. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8790563&pc=QCB | 2023826-2019-01149 | 2019-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM, VTICM5_12.6, -6.5/1.0/53 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. THE DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8809463&pc=QCB | 2023826-2019-01260 | 2019-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8826240&pc=QCB | 2023826-2019-01295 | 2019-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/2.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE/MODEL LENS BUT DIFFERENT POWER AND THE PROBLEM RESOLVED. POSTOPERATIVE REPORT STATED "PATIENT IS HAPPY" H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8826253&pc=QCB | 2023826-2019-01215 | 2019-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: 'A REPOSITIONING BY 0 DEGREES WAS ALSO REPORTED POST-OPERATIVELY - THIS IS NOT CLEAR.' PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, REPOSITIONING. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK; ETHNICITY- UNK; RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8851480&pc=QCB | 2023826-2019-01292 | 2019-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.00/3.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871965&pc=QCB | 2023826-2019-01394 | 2019-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+3.0/063 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: INITIAL MDR SHOULD HAVE INCLUDED 'THE SURGEON REPORTED A LENS ROTATION.' (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS ROTATED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL WITH 6/6 VISION AND IS HAPPY. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONED (ROTATED). CLAIM#: (B)(4). Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876530&pc=QCB | 2023826-2019-01401 | 2019-07-11 | Injury | STAAR SURGICAL COMPANY | 2019-08-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE DUE TO THE WRONG LENS (LENS CALCULATED FOR LEFT EYE(OS)) WAS IMPLANTED IN THE RIGHT EYE. THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ALL IS FINE AND THE PATIENT IS HAPPY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889723&pc=QCB | 2023826-2019-01363 | 2018-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.50/4.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE CHANGE OVER TIME AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. REPORTER INDICATED THAT "REMAINING ASTIGMATISM WILL BE TREATED WITH LASER LATER." Manufacturer Narrative: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND RESIDUE ON THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8899033&pc=QCB | 2023826-2019-01421 | 2019-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 12.50/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN A BIOHAZARD BAG SHOULD ALSO HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE CODE: 1401 SHOULD HAVE ALSO BEEN INCLUDED IN INITIAL MEDWATCH REPORT. PATIENT'S AGE: SHOULD BE CORRECTED TO (B)(6) 1984 IN INITIAL MEDWATCH REPORT.(B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8902020&pc=QCB | 2023826-2019-01475 | 2019-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Decentration; Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, 15.00/2.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. REPOSITIONING WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED " THERE IS A CONSTANT DECENTRATION OF THE TICL DUE TO ITS SMALL DIAMETER. IT IS NECESSARY TO EXCHANGE IT AND REPLACE IT WITH A LARGER ONE". AT PATIENTS LAST VISIT "DECENTERED TICL/ BAD VISION" WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 1360: LENS DECENTRATION, DEVICE CODE 1494: OFF- LABEL USE. ANTERIOR CHAMBER DEPTH (ACD) OUTSIDE INDICATED RANGES (LESS THAN 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8919048&pc=QCB | 2023826-2019-01530 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICH12.6, 6.50/0.5/090, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT CODE: 2137- REFRACTIVE CHANGE OVER TIME. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS; RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941301&pc=QCB | 2023826-2019-01577 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +5.0/+2.5/163 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS WITH THE SAME POWER AND PLACED THE LENS ON THE CORRECT AXIS AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941308&pc=QCB | 2023826-2019-01578 | 2019-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +6.0/+1.5/161 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL BUT DIFFERENT POWER LENS AND THE PROBLEM RESOLVED. RESULTS WERE REPORTED AS "VERY SATISFACTORY". CLAIM#: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN WITH A STRAIGHT SCRATCH. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941312&pc=QCB | 2023826-2019-01584 | 2019-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND LOSS OF BCVA. REPORTEDLY, THE LENS ROTATED AFTER ONE WEEK, WAS REPOSITIONED, AND THEN ROTATED AGAIN AFTER ONE WEEK. THE LENS CURRENTLY REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. ADD TO PREVIOUSLY REPORTED PATIENT CODE 3191: (LENS EXCHANGE). DEVICE CODE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962475&pc=QCB | 2023826-2019-01590 | 2019-08-03 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.00/+4.50/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS WAS REPOSITIONED ON (B)(6) 2019. THE PATIENTS VISION WAS IMPROVED TO 20/20 AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966506&pc=QCB | 2023826-2019-01631 | 2019-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A LONGER AND SIMILAR POWER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL BUT DIFFERENT POWER LENS. DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966515&pc=QCB | 2023826-2019-01644 | 2018-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+3.0/044 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: RESULT CODE 213. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE /DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984516&pc=QCB | 2023826-2019-01674 | 2019-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.0/+2.0/042 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: METHOD CODE 4111 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. EVENT: PER UPDATED INFORMATION, THE LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. METHOD CODE 4117 NO LONGER APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984800&pc=QCB | 2023826-2019-01653 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM TMICL12.1 LENS, -5.00/+2.5/063 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT ,LENS ROTATION AND BLURRY VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT CODE:3191- SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT, OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) THE INITIAL LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2019. THE REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM. REFERENCE CLAIM#: (B)(4).FOR REPLACEMENT LENS. EXPLANT DATE: (B)(6) 2019. DEVICE EVALUATION: LENS RETURNED IN LIQUID IN EYE DROP BOTTLE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY, RACE: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984823&pc=QCB | 2023826-2019-01652 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM TMICL12.1 LENS, -6.5/+4.0/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT ,LENS ROTATION AND BLURRY VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). THE INITIAL LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2019. THE REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM. REFERENCE CLAIM# 430733 FOR REPLACEMENT LENS. EXPLANT DATE: (B)(6) 2019. LENS RETURNED IN LIQUID IN A EYE DROP BOTTLE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS TYPE OF REPORTABLE EVENT: SERIOUS INJURY. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE:3191- SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. DEVICE CODE:1401- BLURRED VISON. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9010339&pc=QCB | 2023826-2019-01656 | 2019-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WITH LOSS OF BCVA AND LENS ROTATION WAS REPORTED. LENS WAS REPOSITIONED ON (B)(6) 2019 BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS SMALL LENS SIZE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND AN OPTIC TEAR, THE HAPTIC TORN, AND DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE -UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9072751&pc=QCB | 2023826-2019-01724 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -10.50/1.5/082( SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE REPORTER STATED "PATIENT IS RESTLESS AND WORRIED ABOUT NOT SEEING SATISFACTORILY." LENS REMAINS IMPLANTED. SURGEON BELIEVES THE LENS WAS BADLY LABELED AT THE FACTORY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A VTICM5_12.6, -10.50/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE REPORTER STATED, "PATIENT IS RESTLESS AND WORRIED ABOUT NOT SEEING SATISFACTORILY". ON (B)(6) 2019 THE LENS WAS EXPLANTED, AND A REPLACEMENT LENS WAS IMPLANTED ON THE SAME DATE. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. DOCTOR IS SATISFIED WITH THE NEW RESULT. CLAIM#: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT ON (B)(6) 2019 A NEW LENS WAS IMPLANTED AND THE PATIENTS CONDITION IS SATISFACTORILY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK ,DATES OF REPORT: UNK, UTENSILS: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9100740&pc=QCB | 2023826-2019-01734 | 2019-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: TTHE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/013 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9100883&pc=QCB | 2023826-2019-01688 | 2019-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +1.5/+6.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT STILL HAD SOME ASTIGMATISM LEFT. THE PATIENT WILL HAVE A LASER PROCEDURE TO RESOLVE THE ASTIGMATISM. THE CAUSE OF THE EVENT WAS UNKNOWN. SEE MFR REPORT # 2023826-2019-01733 FOR EXCHANGED LENS. Manufacturer Narrative: PATIENT CODE: 3191 - RESIDUAL ASTIGMATISM SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. METHOD CODE 4111 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND BOTH HAPTICS SPLIT AND BROKEN. UNABLE TO IDENTIFY LENS MODEL OR TO PERFORM DIMENSIONAL/REFRACTIVE INSPECTION DUE TO CONDITION OF LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115524&pc=QCB | 2023826-2019-01731 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A SAME LENGTH , -10.00/2.0/091 (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2019. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE /DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115526&pc=QCB | 2023826-2019-01732 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A SAME LENGTH , -10.50/3.0/088 (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2019. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE /DEBRIS ON PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9175742&pc=QCB | 2023826-2019-01897 | 2019-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -6.00/1.5/124 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO A CALCULATION MISTAKE BY THE SURGEON AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE POST OPERATIVE DATA ARE ALL FINE BUT PATIENT IS NOT 100% HAPPY, THE REPORTER STATED " MAYBE PSYCHOLOGICAL REASONS". Manufacturer Narrative: THE LENS EXCHANGE DATE (B)(6) 2019 IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: CORRECTED DATA: DATE OF EVENT IS INITIAL MDR IS CORRECTED TO (B)(6) 2019. DATE RECEIVED BY MANUFACTURER IN INITIAL MDR IS CORRECTED TO 10-SEPT-2019. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9188695&pc=QCB | 2023826-2019-01950 | 2019-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.5/+1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE EVENT DESCRIPTION PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT POWER LENS. THE PROBLEM IS RESOLVED. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: VISUAL INSPECTION FOUND HAPTIC BROKEN AND LENS STICKY/GUMMY. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: -UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9199834&pc=QCB | 2023826-2019-01922 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+2.5/011 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO INCORRECT INPUT DATA. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS OPTIC SPLIT AND THE HAPTIC TORN AND BROKEN. UNABLE TO PERFORM DIMENSIONAL AND FUNCTIONAL INSPECTION DUE TO DAMAGE OPTIC LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9246796&pc=QCB | 2023826-2019-02028 | 2019-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+4.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ICA, LOSS OF BCVA, AND PARTIAL PUPIL DILATION. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS "EXCESSIVE SIZE." Manufacturer Narrative: REQUIRED SURGICAL INTERVENTION. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -7.50/4.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; LOSS OF BCVA AND PARTIAL PUPIL DILATION. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTER INDICATED THAT THE "AV IS OK." PATIENT CODE:3191 SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9276012&pc=QCB | 2023826-2019-02076 | 2019-01-11 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0/+1.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS HAD ROTATED AND WAS RE-POSITIONED ON (B)(6) 2019. THE LENS AGAIN ROTATED ON (B)(6) 2019 AND WAS RE-POSITIONED. TO DATE THE LENS REMAINS IMPLANTED AND IS IN A STABLE POSITION BUT THERE IS STILL A REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9276004&pc=QCB | 2023826-2019-02073 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.00/+3.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONED CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279650&pc=QCB | 2023826-2019-02060 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.0/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ON (B)(6) 2019 THE LENS WAS REPOSITIONED. THE PROBLEM IS RESOLVED. ADDITIONAL PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279651&pc=QCB | 2023826-2019-02064 | 2019-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/016 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: IT WAS REPORTED THAT ON(B)(6) 2020 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND SIZE BUT DIFFERENT POWER LENS AND THAT THE PATIENT WAS EXTREMELY HAPPY WITH 6/6 VISION. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: EXPLANTED DATE: (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279666&pc=QCB | 2023826-2019-02061 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.5/+2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2019 THE LENS WAS REPOSITIONED AND THE PROBLEM IS RESOLVED. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306655&pc=QCB | 2023826-2019-02095 | 2019-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VTICM5_12.1, -10.50/2.00/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGE FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED DRY, IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND A BENT HAPTIC. DIMENSIONAL AND FUNCTIONAL INSPECTION WAS PERFORMED AND LENS WAS FOUND TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: RESULT CODE 213 SHOULD BE IN THE PREVIOUS MDR. H6- CONCLUSION CODE 4310 SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(6). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336756&pc=QCB | 2023826-2019-02176 | 2019-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, BLURRED VISION, RESIDUAL CYLINER, A "LINE" IN THE VISUAL FIELD, AND A WHITE LIGHT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: H6 - DEVICE CODE: 1401 SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9344421&pc=QCB | 2023826-2019-02185 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.50/1.00/100 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE(OD) ON (B)(6) 2015. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. THE REPORTER STATED " WE ARE ASKING FOR AN ICL EXCHANGE BECAUSE OF TOUCHDOWN AND MYOPIZATION". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM RESOLVED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM# (B)(4). Manufacturer Narrative: SEX:UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9349221&pc=QCB | 2023826-2019-02186 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -5.50/0.5/90 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE(OS) ON (B)(6) 2015. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. THE REPORTER STATED " WE ARE ASKING FOR AN ICL EXCHANGE BECAUSE OF TOUCHDOWN AND MYOPIZATION". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXCHANGED ON 28/JAN/2020 WITH A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM RESOLVED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM# (B)(4). Manufacturer Narrative: SEX:UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9364603&pc=QCB | 2023826-2019-02211 | 2019-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/3.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 12/JUL/2019. REFRACTIVE SURPRISE WAS OBSERVED ON (B)(6) 2019, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2019, "A REPOSITIONING IS PLANNED, PATIENT'S CURRENT CONDITION 20/32 AND BCVA IS 20/20." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE). B5- THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND THE REPORTER STATES "PATIENT IS HAPPY AFTER REPOSITIONING." H6-PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; LENS REPOSITIONING. CLAIM# (B)(4). Manufacturer Narrative: DATE OF EVENT IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. DATE OF THIS REPORT IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. THE REPOSITIONING WAS PERFORMED ON (B)(6) 2020, THE REPOSITIONING SOLVED THE PROBLEM AND PATIENT IS HAPPY. DATE RECEIVED BY MANUFACTURER IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9403756&pc=QCB | 2023826-2019-02304 | 2019-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICH12.1, 8.00/3.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH, DIFFERENT DIOPTER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVERTIME. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.1MM VTICH 12.1 8.00/3.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND DIFFERENT DIOPTER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. REFERENCE MFR# 2023826-2020-00132 FOR REPLACEMENT LENS. H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9405715&pc=QCB | 2023826-2019-02305 | 2019-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, 0.00/4.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 LENS REPOSITIONING WAS PERFORMED DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. Manufacturer Narrative: LENS ROTATION WAS NOTICED ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9406617&pc=QCB | 2023826-2019-02302 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.00/1.0/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. IT WAS NOTED THAT THE AXIS USED IN CALCULATION WAS NOT ALIGNED WITH REFRACTION. ON (B)(6) 2019 LENS REPOSITIONING WAS PERFORMED ACHIEVING GOOD OUTCOMES. CAUSE OF THIS EVENT WAS REPORTED AS A USER ERROR. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.00/1.0/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED DUE TO USER ERROR (AXIS USED IN CALCULATION WAS NOT ALIGNED WITH REFRACTION). ON (B)(6) 2019 LENS REPOSITIONING WAS PERFORMED AND THE PROBLEM WAS RESOLVED. PATIENT IS IN GOOD CONDITION AND HAS NORMAL VISION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3- (B)(6) 2019" SHOULD BE CORRECTED TO (B)(6) 2019" IN THE INITIAL MDR. B4- (B)(6) 2019" SHOULD BE CORRECTED TO (B)(6) 2019" IN THE INITIALL MDR. G4- (B)(6) 2019" SHOULD BE CORRECTED TO (B)(6) 2019" IN THE INITIALL MDR. G4- (B)(6) 2019" SHOULD BE CORRECTED TO (B)(6) 2019" IN THE PREVIOUS MDR. B5- "ROTATION WAS ALSO OBSERVED." SHOULD BE ADDED TO THE PREVIOUS MDR. H6- DEVICE CODE 2993 IS NOT APPICABLE CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9417044&pc=QCB | 2023826-2019-02209 | 2019-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.00/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441185&pc=QCB | 2023826-2019-02358 | 2019-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +1.5/+1.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION (LENS REPOSITIONING) CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9461812&pc=QCB | 2023826-2019-02380 | 2019-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH THE SAME MODEL/LENGTH LENS BUT DIFFERENT DIOPTER AND THE PROBLEM RESOLVED, "PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476277&pc=QCB | 2023826-2019-02392 | 2019-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.0/+3.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. THE REPORTER STATES THE SURGEON POSITIONED THE LENS IN THE WRONG AXIS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2019 THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512599&pc=QCB | 2023826-2019-02440 | 2015-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATINET'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2020 THE LENS WAS EXCHNAGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- ADVERSE EVENT SHOULD BE MARKED IN MDR SUPPLEMENTAL #1. H10- "WITH MOISTURE" SHOULD BE REMOVED FROM PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512630&pc=QCB | 2023826-2019-02441 | 2015-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATINET'S LEFT EYE (OS) ON (B)(6) 2015. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2020 THE LENS WAS EXCHNAGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516211&pc=QCB | 2023826-2019-02483 | 2019-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -6.5/3/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6- WORK ORDER SEARCH: "NO" SHOULD BE CORRECTED TO "ONE" IN THE INITIAL MDR. H6- RESULT CODE 213 SHOULD BE CORRECTED TO 114 IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516575&pc=QCB | 2023826-2019-02451 | 2019-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -3.50/4.5/168 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. D7: EXPLANT DATE:(B)(6) 2020. H6: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION: EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LEN WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4 IS NOT APPLICABLE IN MDR SUPP1 D7 SHOULD BE 19-JAN-2020 IN MDR SUPP1 H6- INVESTIGATING FINDINGS CODE 213 SHOULD BE IN THE PREVIOUS MDR H6- INVESTIGATING CONCLUSION SODE 4310 SHOULD BE IN THE PREVIOUS MDR CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532308&pc=QCB | 2023826-2019-02470 | 2019-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -1.5/+2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION-LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND STICKY SUBSTANCE ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE) B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 ; -1.5/+2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON 24-SEP-2019. ON 17-MAR-2020 THE LENS WAS EXCHANED WITH ASAME MODEL AND LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATES PATIENT IS HAPPY. CAUSE OF THIS EVENT IS REPORTED AS UNKNOWN. D7- EXPLANT DATE: (B)(6)-2020 H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. (B)(4). Manufacturer Narrative: PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532323&pc=QCB | 2023826-2019-02457 | 2016-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - (B)(6) 2020 SHOULD BE CORRECTED TO (B)(6) 2020 IN SUPPLEMENTAL #1 MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532331&pc=QCB | 2023826-2019-02456 | 2016-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON 18/03/2020 THE LENS WAS EXCHANGED WITH A LONGER AND DIFFERENT MODEL LENS AND THE PROBLEM RESOLVED. PATIENT WAS REPORTED TO BE HAPPY. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9538844&pc=QCB | 2023826-2020-00003 | 2019-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.00/1.5/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2019. THE PROBLEM WAS RESOLVED. REPORTEDLY, "DUE TO DATA INPUT ERROR DURING CALCULATIONS, THE DEGREE WAS NOT SUITABLE." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 2137 SHOULD BE ADDED TO THE INITIAL MDR. H6- DEVICE CODE 2923 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO CODE 1401 IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9541982&pc=QCB | 2023826-2020-00024 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.5/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE AFTER MISALIGMENT WAS OBSERVED. IT WAS REPORTED THAT THE REFRACTIVE SURPRISE WAS NOTED ON (B)(6) 2019. PATIENTS CURRENT UCVA : 20/25, BCVA 20/20. LENS REMAINS IMPLANTED AND LENS REPOSITIONING IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS REPOSITIONED ON (B)(6) 2019. THE REPORTER INDICATED THE REPOSITIONING WENT WELL." H6- PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION- LENS REPOSITIONING. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- "ADVERSE EVENT" SHOULD BE MARKED IN THE PREVIOUS MDR. B3- "REQUIRED INTERVENTION" SHOULD BE MARKED IN THE PREVIOUS MDR. D11- "INJECTOR SYSTEM MODEL: FTP, LOT #: UNK" SHOULD BE ADDED IN THE INITIAL MDR. H1- "SERIOUS INJURY" SHOULD BE MARKED IN THE PREVIOUS MDR. H6- DEVICE CODE 2923 SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9604383&pc=QCB | 2023826-2020-00166 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 14.00/+1.0/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO SURGEON WAS NOT HAPPY WITH LENS POSITION. ANOTHER SAME MODEL LENS WAS IMPLANTED AS THE EXCHANGED LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED BLURRED VISION. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. THE LENS EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTICS TORN. THE LENS WAS RETURNED IN THREE PIECES. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9663126&pc=QCB | 2023826-2020-00221 | 2019-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.50/+1.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689069&pc=QCB | 2023826-2020-00227 | 2020-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH_12.6 IMPLANTABLE COLLAMER LENS, 0.00/5.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. LENS WAS REPOSITIONED ON (B)(6) 2020 BUT IT DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON 01/OCT/2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY". CLAIM#:(B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS, FUNCTIONAL INSPECTION FOUND THE LENS OUT OF SPECIFICATIONS. ADDITIONAL INFORMATION: H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: 715477 Manufacturer Narrative: WEIGHT, ETHNICITY & RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689821&pc=QCB | 2023826-2020-00228 | 2020-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +2.00/+4.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. LENS WAS REPOSITIONED ON (B)(6) 2020 BUT IT DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED LATER THAT ON (B)(6) 2020 A 2ND LENS REPOSITIONING WAS PERFORMED AND THIS RESOLVED THE PROBLEM. "PATIENT IS HAPPY". CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721916&pc=QCB | 2023826-2020-00313 | 2019-11-22 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.0/+3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY THE LENS FIRST ROTATED BEGINNING (B)(6) 2019 BY 1 DEGREE. ON (B)(6) 2019 THE LENS HAD ROTATED 90 DEGREES AND WAS REPOSITIONED. ON (B)(6) 2019 THE LENS AGAIN ROTATED 10 DEGREES. THE REPORTER STATES, "THE VISION HAS DECREASED AFTER 1 WEEK POSTOPERATION. WE REPOSITIONED THE LENS 30 DEGREE[S]. THE LENS WAS MORE FLOATING DURING I/A, COMPARING TO OS." REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9720695&pc=QCB | 2023826-2020-00337 | 2020-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.5/5.5/158 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE REPORTER STATED "FOLLOWING EXPLANTATION ON (B)(6) 2020- NO OTHER PROCEDURE PERFORMED. PT. VISITED ON (B)(6) AND THE EYE HAS SETTLED WELL. VISUAL ACUITY 0.2 LOGMAR IN HER OLD SPECTACLES." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781532&pc=QCB | 2023826-2020-00341 | 2019-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.50/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 WITH A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS INDUCED ASTIGMATISM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT WITH BLOOD STAINS AND DEBRIS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781544&pc=QCB | 2023826-2020-00423 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.50/2.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. REPORTEDLY, PATIENT IS NOT HAPPY. PATIENT AND SURGEON ARE LOOKING TO REPLACE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION B5: ON (B)(6) 2020, THE LENS WAS REMOVED AND THE NEXT MONTH ON (B)(6) 2020 A REPLACEMENT LENS OF THE SAME SIZE DIFFERENT MODEL; SIMILAR POWER WAS IMPLANTED AND PROBLEM RESOLVED. D7: EXPLANT DATE: ON (B)(6) 2020. H6: PATIENT CODE 3191, SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- CONCLUSION CODE 4310 SHOULD BE ADDED TO PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9795627&pc=QCB | 2023826-2020-00418 | 2019-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: CORRECTED DATA: H6-RESULT CODE 114 IN PREVIOUS MDR IS CORRECTED TO 213. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814234&pc=QCB | 2023826-2020-00472 | 2019-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819168&pc=QCB | 2023826-2020-00456 | 2019-03-21 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824536&pc=QCB | 2023826-2020-00359 | 2018-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.50/2.0/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THE PATIENT'S POSTOPERATIVE FEEDBACK WAS POOR VISUAL QUALITY, WITH DOUBLE SHADOW, AND BLURRED VISION." AFTER PATIENT AND SURGEON CONSULTATION, THEY DECIDED TO TAKE THE LENS OUT. PER THE REPORTER ON (B)(6) 2020, "PATIENT'S CURRENT CONDITION IS GOOD. AND THE POOR VISION, DOUBLE SHADOW, AND BLURRED VISION HAS CLEARED. THE REMOVAL OF LENS RESOLVED THE PROBLEM." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3- "VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS" NEEDS TO BE INCLUDED IN PREVIOUS MDR. H10-PATIENT CODE 3191-POOR VISUAL QUALITY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842752&pc=QCB | 2023826-2020-00519 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/044 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE REPORTER STATES THAT THERE WAS PLANNED MONOVISION FOR THE LEFT EYE HOWEVER THE RESULT WAS EMMETROPIA. Manufacturer Narrative: LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL/LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTER STATES "PATIENT IS HAPPY". PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910962&pc=QCB | 2023826-2020-00687 | 2020-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.0/+3.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL/LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910963&pc=QCB | 2023826-2020-00695 | 2019-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/132 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS OUT OF SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6-METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910968&pc=QCB | 2023826-2020-00688 | 2020-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.0/+4.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION: LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913390&pc=QCB | 2023826-2020-00741 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928366&pc=QCB | 2023826-2020-00716 | 2019-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.00/2.5/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- "THE PATIENT EXPERIENCED BLURRED VISION." SHOULD BE IN THE INITIAL MDR. H6- PATIENT CODE 2137 SHOULD BE IN THE INITIAL MDR. H6- DEVICE CODE 1401 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936282&pc=QCB | 2023826-2020-00768 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+3.5/126 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, REFRACTIVE CHANGE OVER TIME, AND PERMANENT LOSS OF BCVA. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR AND SAYS, "PROBABLY THE WTW WAS UNDER MEASURED THE 1ST TIME." Manufacturer Narrative: ADDITIONAL NARRATIVE: H6 - WORK ORDER SEARCH - NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTED DATA: B2 - CORRECTED TO 'REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE' IN INITIAL MDR. B5 - 'A REPOSITIONING WAS PERFORMED ON (B)(6) 2020, BUT THIS DID NOT RESOLVE THE PROBLEM.' SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION - LENS REPOSITIONING) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2021, AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936302&pc=QCB | 2023826-2020-00769 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.0/+2.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, REFRACTIVE CHANGE OVER TIME, AND PERMANENT LOSS OF BCVA. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR AND SAYS, "PROBABLY THE WTW WAS UNDER MEASURED THE 1ST TIME." Manufacturer Narrative: ADDITIONAL NARRATIVE: H6 - WORK ORDER SEARCH - NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTED DATA: B2 - CORRECTED TO 'REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE' IN INITIAL MDR. B5 - 'A REPOSITIONING WAS PERFORMED ON (B)(6) 2020, BUT THIS DID NOT RESOLVE THE PROBLEM.' SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION - LENS REPOSITIONING) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# : (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6)2021 AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A SCRATCH ON THE LENS OPTIC AND DEBRIS ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941318&pc=QCB | 2023826-2020-00836 | 2020-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/6.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: ADVERSE EVENT OR PRODUCT PROBLEM: ADVERSE EVENT. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. EVENT: THE REPORTER INDICATED THAT MEDICAL INTERVENTION WAS REQUIRED. REFRACTIVE TREATMENT (PRK) WAS PERFORMED AND THIS RESOLVED THE PROBLEM. THE LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. PATIENT CODE 3191: REFRACIVE TREATMENT (PRK). METHOD CODE 4117: LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941412&pc=QCB | 2023826-2020-00837 | 2020-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.0/6.0/78 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER INDICATED THAT MEDICAL INTERVENTION WAS REQUIRED. REFRACTIVE TREATMENT (PRK) WAS PERFORMED AND THIS RESOLVED THE PROBLEM. THE LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. H6-PATIENT CODE 3191: REFRACTIVE TREATMENT (PRK). METHOD CODE 4117: LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965157&pc=QCB | 2023826-2020-00894 | 2020-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9979190&pc=QCB | 2023826-2020-00925 | 2020-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -07.50/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH SPHERICAL LENS ON (B)(6) 2020. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3- "DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS." SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983361&pc=QCB | 2023826-2020-00935 | 2020-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+3.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B1 - ADVERSE EVENT. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5 - THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+3.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH SIMILAR POWER LENS; DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE EXCHANGE RESOLVED THE PROBLEM. POST EXCHANGE UCVA 20/20. "PATIENT IS HAPPY.". D7- EXPLANT DATE- (B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL INFORMATON H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10043595&pc=QCB | 2023826-2020-01033 | 2018-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.50/4.0/094 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILARE (SPHERE/ CYLINDER/ AXIS) LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, INSIDE A CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10043598&pc=QCB | 2023826-2020-01034 | 2018-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.00/2.0/075 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILARE (SPHERE/ CYLINDER/ AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10049158&pc=QCB | 2023826-2020-01060 | 2020-02-17 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 12.50/+1.0/067 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10053461&pc=QCB | 2023826-2020-01097 |  | Injury | STAAR SURGICAL COMPANY | 2020-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.50/2.5/104 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH MYOPIA ASTIGMATISM SINCE AGE 10YRS ON (B)(6) 2016 AND VISION WAS CORRECTED, VA 20/16. ON (B)(6) 2019 THE PATIENT COMPLAINED OF BLURRED VISION "AFTER PLAYING FOOTBALL (BOUNCE A BALL WITH HEAD)." THE DOCTOR FOUND VA TO BE 20/25+1 PH20/20-1, REFRACTION +0.62 = 2.25 X 162 AND ICL (IMPLANTABLE COLLAMER LENS) HAD ROTATED. ON (B)(6) 2019 THE DOCTOR REPOSITIONED THE LENS AND VISION RETURNED TO 20/15-2. ON (B)(6) 2020 THE PATIENT RETURNED WITH COMPLAINT OF BLURRY VISION AFTER WAKING UP. HE DENIED ANY TRAUMA. HIS VA WAS 20/20-2 PH 20/16-2, REFRACTION +0.75 = -2.12 X 162 AND ICL (IMPLANTABLE COLLAMER LENS) WAS ROTATED COUNTERCLOCKWISE TO 14 DEGREE POSITION. THE DOCTOR STATED THAT " A WEEK LATER THE ICL IS STILL IN WORSEN ROTATED POSITION." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "THE SURGEON REPOSITIONED THE LENS AGAIN CLOCKWISE 9 DEGREES ON (B)(6) 2020 BUT DURING A FOLLOW UP ON (B)(6) 2020 THE SURGEON FOUND THAT THE LENS WAS ALIGNED CLOCKWISE 60 DEGREES. THE LENS REMAINED IMPLANTED. THE SURGEON WOULD LIKE TO EXCHANGE THIS LENS TO 13.2." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE SURGEON REMOVED THE INITIAL LENS ON (B)(6) 2021 AND IMPLANTED THE REPLACEMENT LENS VTICMO13.2-12.50/+2.5/095 ((B)(6)) IN THE SAME OPERATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075062&pc=QCB | 2023826-2020-01094 | 2020-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+4.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PAITENT IS HAPPY. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATON - LENS RETURNED IN A CASE/VIAL; DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075970&pc=QCB | 2023826-2020-01084 | 2020-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/+1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFO: THE LENS WAS EXCHANGED FOR A SAME SIZE AND MODEL, DIFFERENT SPHERE/AXIS LENS. THE PROBLEM WAS RESOLVED. H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. H6-PATIENT CODE 3191 (SECONDARY SURGERY, LENS EXCHANGE). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094011&pc=QCB | 2023826-2020-01124 | 2018-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140609&pc=QCB | 2023826-2020-01195 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTER THAT ON 10/06/2020 THE LENS WAS EXPLANTED. LATER ON THE SAME DAY A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. "PATIENT IS HAPPY WITH VISION". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158148&pc=QCB | 2023826-2020-01242 | 2020-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158172&pc=QCB | 2023826-2020-01244 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/017 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FIBER. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163423&pc=QCB | 2023826-2020-01299 | 2018-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.0/040 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS NOT TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- TYPE OF INVESTIGATION: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163480&pc=QCB | 2023826-2020-01233 | 2018-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.5/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H6- TYPE OF INVESTIGATION: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL WIT MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS NOT TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10185146&pc=QCB | 2023826-2020-01344 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | P030016 | 1 | Event Description: IT WAS REPORTED THAT A TORIC ICL WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND ROTATED ABOUT 10 DEGREES. THE REPORTER STATED " LENS REPOSITIONING WAS PERFORMED, LENS HAS CURRENTLY REMAINED IN THE CORRECT POSITION AND THE PATIENT IS 20/20 UNCORRECTED AT THIS TIME." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.5/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION BUT IT DID NOT RESOLVE THE PROBLEM, THE LENS ROTATED BACK TO THE SAME POSITION ONE WEEK LATER. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189013&pc=QCB | 2023826-2020-01379 | 2019-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -16.0/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CORRECTION DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189480&pc=QCB | 2023826-2020-01289 | 2019-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE DIMENSIONAL VERIFICATION WAS IN SPECIFICATION BUT REFRACTION (FUNCTIONAL) VERIFICATION FAILED. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A NON-TORIC ICL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10190922&pc=QCB | 2023826-2020-01389 | 2020-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -2.50/+1.0/103 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE CHANGE OVERTIME. THE SURGEON IMPLANTED ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222184&pc=QCB | 2023826-2020-01419 | 2020-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/125 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SHORTER, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242474&pc=QCB | 2023826-2020-01425 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00/+1.50/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS REPORTED THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, NON TORIC LENS BUT SAME LENGTH LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271198&pc=QCB | 2023826-2020-01493 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -1/2.50/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BENT HAPTIC AND DEBRIS ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - "THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -1/2.50/95 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM." CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 2137 SHOULD BE ADDED TO INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE 1401 SHOULD BE ADDED TO INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10272062&pc=QCB | 2023826-2020-01471 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.0/1.5/10 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS IN A MICEO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10317988&pc=QCB | 2023826-2020-01533 | 2020-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.0/+2.5/044 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353747&pc=QCB | 2023826-2020-01635 | 2020-06-29 | Injury | STAAR SURGICAL COMPANY | 2020-07-31 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REDUCED VISION. THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401675&pc=QCB | 2023826-2020-01700 | 2020-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10432548&pc=QCB | 2023826-2020-01786 | 2018-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/1.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHAGED FOR A DIFFERENT DIOPTER LENS DUE TO TRANSIENT LOSS OF BCVA. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN / HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D10-DEVICE NOT AVAILABLE FOR EVALUATION. H1-THIS TYPE OF REPORT IS CORRECTED TO SERIOUS INJURY. H2- DEVICE EVALUATION IN PREVIOUS MDR IS NOT APPLICABLE. H3-DEVICE EVALUATED BY MANUFACTURER IN PREVIOUS MDR IS CORRECTED TO NO. H6- METHOD CODE 10 IN PREVIOUS MDR IS NOT APPLICABLE. RESULT CODE 114 IN PREVIOUS MDR IS NOT APPLICABLE. SUPPLEMENTAL #1 MEDWATCH REPORT SUBMITTED IN ERROR WITH WRONG MDR NUMBER YEAR. THE DEVICE EVALUATION IS FOR MDR# (2023826-2019-01786) CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN AMICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10432698&pc=QCB | 2023826-2020-01791 | 2018-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-7.5/1.0/038 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHAGED FOR A DIFFERENT DIOPTER LENS DUE TO TRANSIENT LOSS OF BCVA. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476366&pc=QCB | 2023826-2020-01926 | 2020-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -5.50/+2.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H3-DEVICE EVALUATION IN PREVIOUS MDR: FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476403&pc=QCB | 2023826-2020-01928 | 2020-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/158 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LEN. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476428&pc=QCB | 2023826-2020-01927 | 2020-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/025 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672562&pc=QCB | 2023826-2019-01031 | 2013-07-02 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+0.5/133 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE BUT DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. POST-OP WAS REPORTED AS PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672587&pc=QCB | 2023826-2019-01032 | 2013-07-02 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -5.0/+0.5/047 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL BUT SAME LENGTH SIZE LENS AND THE PROBLEM RESOLVED. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8819696&pc=QCB | 2023826-2019-01303 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS REMAINS IMPLANTED BUT SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: 'THE LENS WAS REPORTED AS HAVING LOW VAULTING' CORRECTED TO 'THE LENS WAS REPORTED AS HAVING LOW VAULTING, LOSS OF VISUAL ACUITY, AND LENS DISLOCATION. THE SURGEON ALSO NOTED MICROSPHEROPHAKIA WHEN THE PUPIL WAS DILATED AS THE INFERIOR ZONULES WERE VISIBLE'. PATIENT CODE 2692 NOT APPLICABLE IN INITIAL MDR. LENS INCIDENTS WERE REPORTED. DEVICE CODE 1494: OFF-LABEL USE (ACD < 2.8MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8819702&pc=QCB | 2023826-2019-01304 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS REMAINS IMPLANTED BUT SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: 'THE LENS WAS REPORTED AS HAVING LOW VAULTING' CORRECTED TO 'THE LENS WAS REPORTED AS HAVING LOW VAULTING, LOSS OF VISUAL ACUITY, AND LENS DISLOCATION. THE SURGEON ALSO NOTED MICROSPHEROPHAKIA WHEN THE PUPIL WAS DILATED AS THE INFERIOR ZONULES WERE VISIBLE'. PATIENT CODE 2692 NOT APPLICABLE IN INITIAL MDR. LENS INCIDENTS WERE REPORTED. DEVICE CODE 1494: OFF-LABEL USE (ACD < 2.8MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876501&pc=QCB | 2023826-2019-01393 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS SEVERAL MONTHS AGO IN THE PATIENTS RIGHT EYE (OD), AND NOW THE SURGEON HAS REPORTED THE LENS HAS ROTATED CLOCKWISE AND HAS INDUCED SOME ASTIGMATISM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: 'THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/168 (SPHERE/CYLINDER/AXIS) ON 19-FEB-2019 IN THE PATIENTS RIGHT EYE (OD). THE SURGEON HAS REPORTED THE LENS HAS ROTATED CLOCKWISE AND HAS INDUCED SOME ASTIGMATISM, THEREFORE CREATING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.'PATIENT CODE 2692 NOT APPLICABLE IN INITIAL MDR. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT FOUND WITHIN ASSOCIATED LOTS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876502&pc=QCB | 2023826-2019-01392 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS SEVERAL MONTHS AGO IN THE PATIENTS LEFT EYE (OS), AND NOW THE SURGEON HAS REPORTED THE LENS HAS ROTATED CLOCKWISE AND HAS INDUCED SOME ASTIGMATISM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: CORRECTED TO OTHER SERIOUS IN INITIAL MDR. 'THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.0/11 (SPHERE/CYLINDER/AXIS) ON (B)(6) 2019 IN THE PATIENTS LEFT EYE (OS). THE SURGEON HAS REPORTED THE LENS HAS ROTATED CLOCKWISE AND HAS INDUCED SOME ASTIGMATISM, THEREFORE CREATING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.' TYPE OF REPORTABLE EVENT: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. PATIENT CODE 2692 NOT APPLICABLE IN INITIAL MDR. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT FOUND WITHIN ASSOCIATED LOTS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANTED: UNK. EXPLANTED: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9272155&pc=QCB | 2023826-2019-02042 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL | Misfocusing | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6 - PATIENT CODE 2692 CORRECTED TO PATIENT CODE 2137. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK DATE OF EVENT: UNK THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6329420&pc=QCB | 2023826-2017-00238 | 2016-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-02-14 05:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, 0/5.5/179 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND LENS ROTATION. Manufacturer Narrative: ADDITIONAL DATA: THE PATIENT EXPERIENCED EXCESSIVE VAULTING, LENS ROTATION, AND REFRACTIVE SURPRISE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B1, B2-UPDATED B5-ADDITIONAL INFORMATION RECEIVED: THE REPORTER STATES THAT THE PATIENT HAS "VERY BAD VISION," CAN'T READ WHEN IT'S BRIGHT/DAYLIGHT. THE SURGEON TRIED TO ROTATE THE LENS BUT IT RETURNED TO THE SAME POSITION. THE VAULT OF THE LENS IS VARYING AND PIGMENT DISPERSION IS REPORTED ALONG WITH "A LOT OF PIGMENTATION ON THE ICL SURFACE." H6-PATIENT CODE 3191: NO CODE AVAILABLE (BAD VISION, DIFFICULTY READING IN BRIGHT/DAYLIGHT, SECONDARY SURGERY, LENS REPOSITIONING, PIGMENT DISPERSION, PIGMENTATION ON ICL) H6-RESULT CODE CORRECTED FROM SUPPLEMENTAL REPORT #1. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113436&pc=QCB | 2023826-2018-01755 | 2018-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+6.0/002 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT HAD A CORNEAL ASTIGMATISM OF 8 DIOPTER. AFTER IMPLANTATION OF THE ICL THE CORNEAL ASTIGMATISM REDUCED TO 2 DIOPTER. THEREFORE THE CURRENT OUTCOME IS PLANO CYLINDER -6.00 DIOPTER. THE SURGEON PLANS TO EXPLANT THE LENS FOR ANOTHER LENS. THE LENS TO DATE REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.00/6.0/002 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THE PATIENT EXPERIENCED REFRACTIVE SURPRISE DUE TO POST-OPERATIVE REDUCTION IN CORNEAL ASTIGMATISM FROM 8 TO 2 DIOPTRES. THE LENS CURRENTLY REMAINS IMPLANTED WITH EXPLANT AND NEW LENS IMPLANTATION PLANNED FOR THE FUTURE. PATIENT CODE 2692: PREVIOUS CODE NOT APPLICABLE. DEVICE CODE 3189: PREVIOUS CODE NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED SUCCESSFULLY ON (B)(6) 2018. THE DOCTOR WILL WAIT APPROX. 1 MONTH BEFORE IMPLANTING A NEW LENS. AFTER THE EXPLANTATION THE PATIENT IS FINE. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. CORRECTED DATA: COMMON DEVICE NAME CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS. CLAIM#: (B)(4). Manufacturer Narrative: DISTRIBUTOR CHECKED IN ERROR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9596829&pc=QCB | 2023826-2020-00132 | 2019-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, 7.50/3.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 AS A REPLACEMENT LENS. THE SURGEON OBSERVED THE LENS ROTATION. THE LENS REMAINS IMPLANTED. NO PLANNED MEDICAL TREATMENT FOR THE MOMENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2019-02304 FOR INITIAL LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH 12.6 IMPLANTABLE COLLAMER LENS; 7.50/3.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 AS A REPLACEMENT LENS. REPORTER STATES LENS ROTATED AGAIN AND REFRACTIVE CHANGE OVER TIME WAS NOTED ON (B)(6) 2019. LENS REMAINS IMPLANTED CAUSE IS REPORTED AS PATIENT RELATED FACTOR. REFERENCE MFR# 2023826-2019-02304 FOR INITIAL LENS. H6-PATIENT CODE-3191 FOR REFRACTIVE CHANGE OVER TIME. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 3191 FOR REFRACTIVE CHANGE OVER TIME SHOULD BE CORRECTED TO PATIENT CODE: 2137. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED AT A UNKNOWN LATER DATE. THE LENS WAS EXPLANTED AND REPLACED WITH A NON-TORIC SPHERICAL LENS OF THE SAME LENGTH. D7 - UNK. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10164916&pc=QCB | 2023826-2020-01277 | 2020-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -2.50/+1.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(4) 2020. THE LENS WAS EXPLANTED ON (B)(4) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURER DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17541754&pc=QCB | 2023826-2023-03429 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; Swelling/ Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, EYE SWELLING AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THAT THE SURGEON HAD REPOSITIONED THE LENS ON (B)(6) 2023 BEFORE EXCHANGING THE LENS. H3 - DEVICE EVALUATION - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND NO VISBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15008783&pc=QCB | 2023826-2022-02203 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.00/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. TASS WAS REPORTED AND PATIENT COMPLAINED OF BLURRED VISION. DIAGNOSTICS WAS PERFORMED - NON CONTACT TONOMETRY MEDICAL EXAMINATION. TREATMENT WAS PERFORMED - CLEANING, DROPPING MEDICINE IN THE EYE. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039569&pc=QCB | 2023826-2023-01883 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.0/6.0/171 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. TOXIC ANTERIOR SEGMENT SYNDROME(TASS) AND BLURRED VISION WAS OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND ADDITION OF ORAL STEROIDS AND OINTMENTS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WASREPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H6: HEALTH IMPACT CODE: "4625 - ADDITIONAL SURGERY: ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391254&pc=QCB | 2023826-2023-03035 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE AND TOXIC ANTERIOR SEGMENT SYNDROME(TASS, EYE DROP TREATMENT PERFORMED). THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS OTHER, TASS CAUSED BY INJECTOR, ASTIGMATISM CORRECTION. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -14.0/3.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME SIZE SIMILAR POWER LENS DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS, EYE DROP TREATMENT PERFORMED). THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS OTHER, TASS CAUSED BY INJECTOR. THE DOCTOR BELIEVES THAT TASS MAY HAVE BEEN INDUCED BY USING AN INJECTOR THAT WAS NOT STAAR'S INJECTOR. H6 HEALTH EFFECT CLINICAL CODE: 2137 SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391263&pc=QCB | 2023826-2023-03036 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.0/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE AND TOXIC ANTERIOR SEGMENT SYNDROME(TASS, EYE DROP TREATMENT PERFORMED). THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS OTHER, TASS CAUSED BY INJECTOR, ASTIGMATISM CORRECTION. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -12.0/2.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME SIZE SIMILAR POWER LENS DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS, EYE DROP TREATMENT PERFORMED). CAUSE OF THE EVENT IS REPORTED AS OTHER, TASS CAUSED BY INJECTOR. THE DOCTOR BELIEVES THAT TASS MAY HAVE BEEN INDUCED BY USING AN INJECTOR THAT WAS NOT STAAR'S INJECTOR. H6 HEALTH EFFECT CLINICAL CODE: 2137 SHOULD BE REMOVED. CLAIM#: 730392 Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6323359&pc=QCB | 2023826-2016-01547 | 2016-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-10 05:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Misfocusing | Blurred Vision; Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -8.0/+2.0/62 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, BLURRED VISION AND DIPLOPIA (DOUBLE VISION). THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/60. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): NA THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8445669&pc=QCB | 2023826-2019-00475 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE PATIENT REPORTED THAT HE HAD VISIAN TORIC IMPLANTABLE COLLAMER LENS IMPLANTED 2 WEEKS AGO AND HAS VERY BLURRY AND DOUBLE VISION. THE PATIENT REPORTED THE CAUSE WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703193&pc=QCB | 2023826-2020-00266 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Visual Disturbances | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+3.0/082 (SPHERE/CYLINDER/AXIS), IN THEIR LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT REPORTED HAS BLURRED VISION, ESPECIALLY AROUND LUMINOUS SIGNS AND DIFFICULTY READING. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H1- "SERIOUS" SHOULD BE CORRECTED TO "MALFUNCTION" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852522&pc=QCB | 2023826-2020-02677 | 2020-03-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -01.00/+6.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND PATIENT EXPERIENCED BLURRED VISION AND GLARE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B1 - CORRECTED TO PRODUCT PROBLEM. B2 - OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS SHOULD BE REMOVED. H1 - CORRECTED TO MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11606813&pc=QCB | 2023826-2021-00835 | 2020-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.5/2.0/088 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE AND DOUBLE VISION WAS OBSERVED. LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS THE DEVICE " SMALL DIAMETER OF ICL". THE REPORTER STATED " SURGEON HAS TRIED TO REPOSITION THE TICL TWO TIMES AND AFTER A COUPLE DAYS THE LENS WAS DECENTRED AGAIN, IN THE SAME POSITION." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE DEVICE WAS THE CAUSE OF THE EVENT BUT THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED "SMALL DIAMETER OF ICL, SULCUS ANATOMY?" CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12061900&pc=QCB | 2023826-2021-02038 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: ON (B)(6) 2021 WE RECEIVED NOTIFICATION FROM A PATIENT STATING THAT SHE HAD A 12.6MM, VTICM5_12.6, -6.50/0.5/077, IMPLANTABLE COLLAMER LENS IMPLANTED IN HER RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER STATED " AFTER ONE DAY I GOT TRAUMA IN MY HEAD AND I WENT FOR DOCTOR FOR EXAMINATION AND I COMPLAIN THAT MY VISION NOT CLEAR SHE ADVICE ME TO DO ANOTHER OPERATION FOR MY RIGHT EYE TO FIX THE AXIS AS MAY BE THE LENS ALIGNMENT MOVE FROM THE AXIS AND AFTER 2 WEEKS FROM SECOND OPERATION STILL I HAVE ASTIGMATISM IN MY RIGHT EYE." PER UPDATED INFORMATION FROM THE PATIENT SHE IS CONSULTING WITH A SURGEON FOR SECOND OPINION. THE PATIENT STATED " SHE EXAMINED MY EYES AFTER DILATING THE PUPILS AND SHE SAID TO ME THE AXIS OF THE LENS NOT ON THE CORRECT ALIGNMENT. SO SHE SAID TO ME FIRST TO FIX THE AXIS , AND IF STILL THERE IS ASTIGMATISM THAT MEAN NEED TO CHANGE THE LENS AND SHE SAID MAY BE THERE IS A NEW PRESCRIPTION FROM THE LENS COMPANY TO FIX THE AXIS IN MY CASE." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - DATE OF (B)(6) 2021, SHOULD BE CORRECTED TO (B)(6) 2021 IN INTIAL MEDWATCH REPORT. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - HEAD TRAUMA - SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17874022&pc=QCB | 2023826-2023-04415 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Patient-Device Incompatibility; Optical Problem | Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/012 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION. THE EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATED THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTORS, "PATIENT CANNOT ADAPT TO VISUAL GHOSTING AFTER IMPLANTATION OF CRYSTALS". THE REPORTER ALSO STATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473794&pc=QCB | 2023826-2024-00122 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -2.0/4.0/100 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. VISUAL ACUITY NOT IMPROVED. LENS REMAINS IMPLANTED. REPORTEDLY, VAULT HAS IMPROVED WITH GOOD RESULTS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10454795&pc=QCB | 2023826-2020-01865 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: REPORTER INDICATED THAT PATIENT HAD SURGERY 6 WEEKS AGO AND "PATIENT IS COMPLAINING OF SLIGHT BLURRY VA, GLARE, AND HALO IN HER LEFT EYE (OS). REPORTEDLY, HIGH VAULT IS ALSO OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634467&pc=QCB | 2023826-2020-02280 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT REPORTED MILD GLARE AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-02067. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED AT THE LAST VISIT THE PATIENT APPEARED TO BE STILL SYMPTOMATIC FOR GLARE, HALOS AND BLURRY VISION. A 30 DEGREE ROTATION OF THE LENS WAS MEASURED AT LAST VISIT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT REPORTED MILD GLARE/HALOS AT NIGHT AND BLURRY VISION. THE LENS REMAINED IMPLANTED. AT THE POST-OP VISIT ON (B)(6) 2020 THE PATIENT WAS DOING WELL, EXPRESSED MILD DISCOMFORT BUT DIDN'T EXPRESS ANY SYMPTOMS OF GLARE/HALOS. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-02067. H6: MEDICAL DEVICE PROBLEM CODE: 2923 - THIS CODE WAS USED IN ERROR, THE LENS DID NOT ROTATE. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6)2020. THE PATIENT REPORTED MILD GLARE AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT#: 2023826-2020-02067. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634506&pc=QCB | 2023826-2020-02282 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT REPORTED GLARE/HALOS AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-02066. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 : THE REPORTER INDICATED AT POST-OP VISIT ON (B)(6) 2020, IT SHOWED THE PATIENT FELT VISION SEEMED BETTER. AT LAST VISIT ON (B)(6) 2020, THERE WERE NO COMPLAINTS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - CHANGED TO PRODUCT PROBLEM. B4 - "OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS" SHOULD BE REMOVED AS THIS CLAIM SHOULD BE CORRECTED TO NON SERIOUS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11573285&pc=QCB | 2023826-2021-00770 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. PATIENT IS EXPERIENCING GLARE/HALOES AND BLURRED VISION. REFRACTIVE SURPRISE WAS ALSO REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE RESIDUAL REFRACTION WAS ADDRESSED WITH LASIK SURGERY, "OD BCVA 20/20-2, UCVA 20/25, VAULTING 1.0CCT, IOP-NOT MEASURED". THE LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11573286&pc=QCB | 2023826-2021-00769 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PATIENT IS EXPERIENCING GLARE/HALOES AND BLURRED VISION. REFRACTIVE SURPRISE WAS ALSO REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE RESIDUAL REFRACTION WAS ADDRESSED WITH LASIK SURGERY, "OS BCVA 20/20-1, UCVA 20/20-1, VAULTING 1.0CCT, IOP-NOT MEASURED". THE LENS REMAINS IMPLANTED. CLAIM #(B)(4). Manufacturer Narrative: A THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11743352&pc=QCB | 2023826-2021-01264 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.5/3.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION AND THE PROBLEM RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11743370&pc=QCB | 2023826-2021-01271 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/4.0/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION AND THE PROBLEM RESOLVED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12866483&pc=QCB | 2023826-2021-04509 | 2021-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, (B)(4), IMPLANTABLE (B)(6) WAS IMPLANTED INTO THE PATIENTS (B)(6) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REMOVED AND LATER REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE (B)(6). IT WAS REPORTED THAT THE PROBLEM RESOLVED. THE REPORTED STATED " GLARE HAS DECREASED AND PATIENT CAN FUNCTION BETTER." IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE, THE REPORTER STATED " LARGE VAULTING CAUSING A LOT OF (B)(6) NO CONSTRICTING NORMALLY." Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14415838&pc=QCB | 2023826-2022-01589 | 2021-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -15.5/+5.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION; REFRACTIVE SURPRISE; GLARE/HALOES; LOW VISUAL ACUITY; BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER THE SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH; BUT DIFFERENT MODEL/AXIS/POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND "OTHER". Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. H3 - DEVICE ANTICIPATED BUT NOT YET BEGUN SHOULD, SHOULD HAVE BEEN SELECTED IN THE INITIAL MEDWATCH REPORT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14710644&pc=QCB | 2023826-2022-01931 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN JOURNAL OF CATARACT AND REFRACTIVE SURGERY IN THE YEAR 2022 FROM A RETROSPECTIVE STUDY PERFORMED IN EYE AND ENT HOSPITAL OF FUDAN UNIVERSITY, CHINA. THIS ARTICLE'S PURPOSE IS TO EVALUATE THE INCIDENCE, CAUSES, AND OUTCOMES OF IMPLANTABLE COLLAMER LENS (ICL) RE-ALIGNMENT OR EXCHANGE FOLLOWING IMPLANTATION OF THE EVO ICL. THE STUDY INCLUDED A TOTAL OF 10,258 EYES THAT WERE TREATED WITH EVO ICL/TICL IMPLANTATION, 22 EYES (0.21%) OF 20 PATIENTS UNDERWENT ICL RE-ALIGNMENT OR EXCHANGE.AMONG THESE 22 EYES, 13 EYES (0.13%) OF 12 PATIENTS UNDERWENT REPOSITIONING, INCLUDING RE-ALIGNMENT FOR MISALIGNED TICL (10 EYES, 0.10%) AND VERTICAL ROTATION FOR NON-TORIC ICL WITH EXCESSIVE VAULT (3 EYES, 0.03%). NINE EYES (0.09%) OF 8 PATIENTS UNDERWENT ICL/TICL EXCHANGE. SEVEN (0.07%) OF THESE 9 EYES UNDERWENT EXCHANGE DUE TO EXCESSIVE VAULT OF THE ICL/TICL AND TWO (0.02%) EYES DUE TO TICL MISALIGNMENT COMBINED WITH LOW VAULT. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15219661&pc=QCB | 2023826-2022-02758 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS OF -2.00/1.5/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE, GLARE/HALOES AND BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND RESIDUE ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5: CORRECTED TO: THE REPORTER INDICATED THAT A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS OF +2.00/1.5/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE, GLARE/HALOES AND BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. H6 - INVESTIGATION CONCLUSION CODE 4310 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MDR#1. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15937237&pc=QCB | 2023826-2022-04284 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -11.5/2.5/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED WITH GLARE AND HALOES AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER - 11.50/2.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION, GLARES/HALOES, AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND POWER BUT LONGER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "ALL WELL, LENS IS STABLE, REFRACTION OUTCOME IS SATISFACTORY". THE CAUSE OF THE EVENT IS UNKNOWN AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6- MEDICAL DEVICE PROBLEM CODE: "2923" SHOULD BE ADDED. CLAIM# (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15952662&pc=QCB | 2023826-2022-04354 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+3.5/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15953610&pc=QCB | 2023826-2022-04353 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525763&pc=QCB | 2023826-2024-00168 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1, -6.0/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525764&pc=QCB | 2023826-2024-00167 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -6.0/1.5/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11382799&pc=QCB | 2023826-2021-00531 | 2020-12-18 05:00 | Injury | STAAR SUGICAL COMPANY | 2021-02-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -11.50/2.0/67 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE. EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. LENS REMAINS IMPLANTED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE. HALO AND GLARE BY LENS POSITION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13. 7MM, VTICM5_ 13. 7, -11.50/2.0/67 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE. HALO AND GLARE BY LENS POSITION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IT WAS REPORTED THAT THE LENS WAS REMOVED AND REPLACED AT AN UNKNOWN DATE AND THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13. 7MM, VTICM5_ 13. 7, -11.50/2.0/67 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE. HALO AND GLARE BY LENS POSITION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IT WAS REPORTED THAT THE LENS WAS REMOVED AND REPLACED AT AN UNKNOWN DATE AND THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT-RACE- UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLE, PIGMENT DISPERSION. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11382825&pc=QCB | 2023826-2021-00532 | 2020-12-18 05:00 | Injury | STAAR SUGICAL COMPANY | 2021-02-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE. EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. LENS REMAINS IMPLANTED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE." WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13. 7MM, VTICM5_ 13. 7, -7 .00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE." WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IT WAS REPORTED THAT THE LENS WAS REMOVED AND REPLACED AT AN UNKNOWN DATE AND THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13. 7MM, VTICM5_ 13. 7, -7 .00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE." WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IT WAS REPORTED THAT THE LENS WAS REMOVED AND REPLACED AT AN UNKNOWN DATE AND THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT: CLINICAL CODE (B)(4): NARROW ANGLE, PIGMENT DISPERSION. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10454806&pc=QCB | 2023826-2020-01866 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: REPORTER INDICATED THAT PATIENT HAD SURGERY 6 WEEKS AGO AND "PATIENT IS COMPLAINING OF SLIGHT BLURRY VA, GLARE, AND HALO IN HER RIGHT EYE (OD). REPORTEDLY, HIGH VAULT IS ALSO OBSERVED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - REPORTER INDICATED THAT THE PATIENT HAD SURGERY 6 WEEKS AGO IN HER RIGHT EYE (OD) AND PATIENT " HAS HIGH VAULT OU BUT THE PUPIL OD IS SLUGGISH AND DOESN'T REACT MUCH". LENS REMAINS IMPLANTED. H6 - PATIENT CODE: 2227, 2140, 2137 SHOULD BE REMOVED. H6 - HEALTH EFFECT CLINICAL CODE : 4581 - SLUGGISH PUPIL. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16555053&pc=QCB | 2023826-2023-00862 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00/+1.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16555127&pc=QCB | 2023826-2023-00863 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+0.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16555168&pc=QCB | 2023826-2023-00861 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.50/+2.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16559490&pc=QCB | 2023826-2023-00864 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.50/+1.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNKNOWN ETHNICITY: UNKNOWN RACE: UNKNOWN (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16636031&pc=QCB | 2023826-2023-01064 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD). AT 3 WEEKS POST-OP, THE PATIENT COMPLAINED OF BLURRINESS (MILD) CENTRALLY, MONOCULAR DIPLOPIA, GLARE/HALOS AND GHOSTING/DOUBLING OF HER MIDPERIPHERAL VISION. THERE IS PERMANENT LOSS OF BCVA. THE CAUSE OF THE EVENT IS THE DEVICE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS ROTATED (REPOSITIONED) ON (B)(6) 2023. THE PATIENT IS UNCORRECTED 20/25- OD AT LAST VISIT ON (B)(6) 2023. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY & RACE UNKNOWN. IMPLANTED/ EXPLANTED DATES UNKNOWN. ADVERSE EVENT PROBLEM: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE ERROR, PERMANENT LOSS OF BCVA; WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11158539&pc=QCB | 2023826-2021-00092 | 2020-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020 DUE TO BEING TOO LONG. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, BLURRED VISON AND GLARE/HALOS. THE LENS WAS WELL CENTERED, ANGLE CLOSURE, CLOSE TO SUTURE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THE PATIENT HAD AN ITCHY EYE BUT NO VISION COMPLAINTS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE; SUTURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8457907&pc=QCB | 2023826-2019-00477 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497526&pc=QCB | 2023826-2019-00568 |  | Injury | STAAR SURGICAL COMPANY | 2019-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -11.0/+1.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED A SAME SIZE, DIFFERENT MODEL LENS DUE TO LOW VAULT, LENS ROTATION, PIGMENT DISPERSION, REFRACTIVE CHANGE OVER TIME, GLARE/HALOS, BLURRED VISION, AND IRIS ATROPHY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736620&pc=QCB | 2023826-2019-01125 | 2019-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOS, AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252685&pc=QCB | 2023826-2019-01959 | 2019-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/071 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR/DEVICE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10077079&pc=QCB | 2023826-2020-01109 | 2018-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.00/1.00/019 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT WAS COMPLAINING OF EXCESSIVE GLARE/HALOES, MONOCULAR DIPLOPIA AND GHOSTING. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2018 TO CORRECT 2 DEGREES OF ERROR BUT SYMPTOMS PERSISTED DESPITE CORRECT IOL POSITION POST OP. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (GHOSTING, MONOCULAR DIPLOPIA). H6 - MEDICAL DEVICE PROBLEM CODE 2923. CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.00/+1.00/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, BLURRED VISION, GLARE AND HALOS, MONOCULAR DIPLOPIA, GHOSTING, ROTATION: REPOSITIONING ATTEMPTED. LENS REPOSITIONED BUT THAT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. STATUS OF THE EYE: "BLURRED VISION IN RIGHT EYE AND HALOS AT NIGHT". ADDITIONAL INFORMATION PROVIDED: "WANTS REPLACEMENT OF ICL IN RIGHT EYE". IT WAS ALSO REPORTED, "THIS IS STILL UNDER CONSULT SO NO FURTHER UPDATES OR INTERVENTION DECIDED UPON". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721918&pc=QCB | 2023826-2020-00310 | 2016-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Patient-Device Incompatibility; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2016. THE SURGEON REPORTS LENS ROTATION, GLARES/HALOS, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), LENS DISLOCATION/SUBLUXATION, PIGMENT DISPERSION, REFRACTIVE CHANGE OVER TIME, AND BLURRED VISION. THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. IN ADDITION REFRACTIVE TREATMENT WAS PERFORMED AFTER IMPLANT DUE TO LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR-THE SURGEON BELIEVES THE PATIENT'S SULCUS IS IRREGULAR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/SAME LENGTH LENS. THE PROBLEM WAS RESOLVED. H6- HEALTH EFFECT IMPACT CODE- 4629: SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: SYMPTOMS NOTED (B)(6) 2016. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131800&pc=QCB | 2023826-2019-01761 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | NA | 1 | Event Description: RECEIVED AN ARTICLE INDICATING A SURGEON IMPLANTED A 13.7MM TORIC ICL IMPLANTABLE COLLAMER LENS, -7.50/+6.0/054 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) IN 2017. POSTOPERATIVELY, THE ICL WAS REPORTED AS BEING REPOSITIONED THREE TIMES BECAUSE OF ROTATION. THE PATIENT REPORTED BLURRY VISION WITH DIPLOPIA. THERE WAS SOME IRIS PIGMENT SEEN ON THE ANTERIOR SURFACE OF THE ICL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9246802&pc=QCB | 2023826-2019-02026 | 2019-10-04 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -7.0/+4.0/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO BLURRED VISION, DOUBLE-VISION, AND NEAR VERTICAL ROTATION OF LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9246803&pc=QCB | 2023826-2019-02027 | 2019-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -5.5/+2.5/063 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO BLURRED VISION, DOUBLE-VISION, AND NEAR VERTICAL ROTATION OF LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5 - WOUND LEAKAGE AND ADDITION OF 3 SUTURES WERE REPORTED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE - WOUND LEAKAGE. H6 - PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION - ADDITION OF 3 SUTURES. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9406313&pc=QCB | 2023826-2019-02303 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/2.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A NON TORIC, DIFFERENT DIOPTER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BAD VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B4 - CORRECTED TO 06-NOV-2019 IN INITIAL MDR G4 - CORRECTED TO 06-NOV-2019 IN INITIAL MDR H6 - PATIENT CODE 2137 CORRECTED TO PATIENT CODE 2140 IN INITIAL MDR H6 - DEVICE CODE 1401 CORRECTED TO DEVICE CODE 2923 IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938191&pc=QCB | 2023826-2020-00814 | 2020-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/121 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT HAS DIFFICULTY WITH SEEING NEAR AND READING. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER AND THE PROBLEM RESOLVED. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CORRECTED TO (B)(6) 2019. SHOULD BE CORRECTED TO (B)(6) 2019 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119235&pc=QCB | 2023826-2021-02320 | 2021-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Impairment; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -14.0/+1.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) HALOS, GLARE, SPARKLES, PIGMENT DISPERSION, AND BLURRED VISION WAS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS DEVICE. Manufacturer Narrative: H6: INCLUDE CLINICAL CODE: 2227. CLAIM #(B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -14.0/+1.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH A HISTORY OF DIPLOPIA AND PROGRESSIVE MYOPIA ON (B)(6) 2021. EXCESSIVE VAULT; TRANSIENT LOSS OF BCVA; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA); HALOS; GLARE; SPARKLES; PIGMENT DISPERSION; AND BLURRED VISION WAS REPORTED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF A DIFFERENT MODEL. POSTOPERATIVE REPORTS THAT THE LENS WAS ROTATED 10 DEGREES AND THE PATIENT RE-ENTERED SURGERY TO ACCOMMODATE. AFTERWARDS REPORTS THAT THE PATIENT IS WELL, VISION WITHOUT HALOS OR FLASHES, VAULT OF 800, ANGLES DEGREE II, WILL CHECK UP EVERY 2 MONTHS TO SEE THE EVOLUTION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119268&pc=QCB | 2023826-2021-02319 | 2021-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Impairment; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/+0.5/126 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) HALOS, GLARE, SPARKLES, PIGMENT DISPERSION, AND BLURRED VISION WAS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS DEVICE. Manufacturer Narrative: H6 - INCLUDE CLINICAL CODE: 2227. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -12.0/+0.5/126 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH A HISTORY OF DIPLOPIA AND PROGRESSIVE MYOPIA ON (B)(6) 2021. EXCESSIVE VAULT; TRANSIENT LOSS OF BCVA; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA); HALOS; GLARE; SPARKLES; PIGMENT DISPERSION; AND BLURRED VISION WAS REPORTED. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF A DIFFERENT MODEL. POSTOPERATIVE REPORTS THAT THE PATIENT FEELS GOOD SEEING; VERY HAPPY; NORMAL IOP OF 12; VAULT OF 800; VISION WITHOUT HALOS OR FLASHES; ANGLES GRADE II-III BY GONIOSCOPY; WILL KEEP HIM IN CHECK-UPS EVERY 2 MONTHS TO SEE THE EVOLUTION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: 4581 (SIGNIFICANT REDUCTION OF ICA, PIGMENT DISPERSION, SPARKLES). INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241293&pc=QCB | 2023826-2021-02893 | 2021-06-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Capsular Bag Tear | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS PREPARING TO IMPLANT A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/076 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AND THE LENS CAPSULE WAS NICKED DURING THE PRIMARY INCISION POINT. THE ICL LENS WAS IMPLANTED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17292528&pc=QCB | 2023826-2023-02649 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Capsular Bag Tear | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/176 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE ANTERIOR CAPSULE RUPTURED AT THE TIME THE MAIN INCISION WAS MADE. PHACOEMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PHACOEMULSIFICATION (CATARACT SURGERY), IOL IMPLANTED. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10247584&pc=QCB | 2023826-2020-01427 | 2020-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/073 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). WHEN THE LENS WAS REMOVED, THE LENS CONTACTED THE CRYSTALLINE LENS AND A MILD ANTERIOR SUBCAPSULAR OPACITY DEVELOPED. THE SURGEON IMPLANTED ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS EYE MOVED WHEN THE SURGEON INSERTED THE INITIAL ICL. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR/UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014526&pc=QCB | 2023826-2020-02985 | 2020-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR) WAS OBSERVED. ANOTHER ICL LENS HAS NOT BEEN IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10990261&pc=QCB | 2023826-2020-02963 | 2020-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LOW VAULT AND LENS OPACITY ASC (OBSERVED: (B)(6) 2020). ON (B)(6) 2020 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: B5- PER NEW INFORMATION NO INTERVENTION OR TREATMENT WAS PERFORMED FOR LENS OPACITY. IT WAS REPORTED THAT THE LENS OPACITY WAS NOT PRODUCT RELATED, THE REPORTER STATED THAT THE DOCTOR WAS NOT SURE HOW THE LENS OPACITY CAME ABOUT. IT MAYBE RELATED TO HIS MANIPULATION DURING SURGERY. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT, AND DEFORMED WITH DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894548&pc=QCB | 2023826-2021-01709 | 2021-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER STATES A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS -12.0/+1.5/109 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO A POSTERIOR SUBCAPSULAR CATARACT. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. MODERATE CORNEAL EDEMA AND MODERATE ANTERIOR CHAMBER CELL ARE REPORTED AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "WHILE THE MECHANISM OF CATARACT FORMATION REMAINS SPECULATIVE, THE TIME COURSE FOR THE DEVELOPMENT OF THE CATARACT AND THE PRESENCE OF ANTERIOR CAPSULE FINDINGS IN THIS CASE STRONGLY SUGGEST AN ASSOCIATION BETWEEN THE LASER PERIPHERAL IRIDOTOMY AND THE POSTERIOR SUBCAPSULAR CATARACT." Manufacturer Narrative: B5-ADDITIONAL INFORMATION: RESOLUTION DATE WAS 07-JUN-2021. ALL SYMPTOMS RETURNED TO NORMAL. (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986324&pc=QCB | 2023826-2021-01855 | 2021-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS OPACITY-ASC. PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED. THE SURGEON REPORTS THIS WAS AN ICL-INDUCED CATARACT. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: LENS OPACITY-ASC (ANTERIOR SUBCAPSULAR), REPORTED IN INITIAL MDR IS CORRECTED TO LENS OPACITY-PSC (POSTERIOR SUBCAPSULAR). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021151&pc=QCB | 2023826-2021-01906 | 2021-02-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.5/1.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASC). IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4) OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064338&pc=QCB | 2023826-2021-02028 | 2021-05-26 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM, TMICL13.2, -10.0/1.0/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS OPACITY WAS OBSERVED. LENS REMAINS IMPLANTED. THE REPORTER STATED THAT PATIENT IS HAPPY WITH VISION AND THE SURGEON IS NOT TREATING IT AT THE TIME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872840&pc=QCB | 2023826-2021-04560 | 2021-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -11.5/1.5/051 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO SUBCAPSULAR CATARACT (ASC) WHICH WAS OBSERVED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: G3: INITIAL MDR DATE SHOULD BE CORRECTED TO 11-NOV-2021. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12957841&pc=QCB | 2023826-2021-04691 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -11.00/1.0/129 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED ANTERIOR SUBCAPSULAR CATARACT GRADE 1 (ASC). LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED (B)(6) 2021 AND THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082661&pc=QCB | 2023826-2021-04920 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED AT ONE WEEK POST-OP, A SMALL ANTERIOR CATARACT WAS OBSERVED. THE PATIENT IS 20/15 OU AND SO FAR THE CATARACT IS A NON ISSUE. THE LENS REMAINED IMPLANTED AND THE PATIENT WILL BE MONITORED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED AT ONE WEEK POST-OP, A SMALL ANTERIOR CATARACT WAS OBSERVED. THE PATIENT IS 20/15 OU AND SO FAR THE CATARACT IS A NON ISSUE. THE LENS REMAINED IMPLANTED AND THE PATIENT WILL BE MONITORED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK; WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK; DATE OF EVENT: UNK; EXPLANT DATE: UNK. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: SEX: UNK; WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK; DATE OF EVENT: UNK; EXPLANT DATE: UNK. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13563729&pc=QCB | 2023826-2022-00543 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2020. THE SURGEON HAS REPORTED THE PATIENT NOW HAS SHALLOW VAULT AND SMALL SUBCAPSULAR CATARACTS IN THE MIDPERIPHERAL ZONE SUPERIORLY. IT WAS ALSO NOTED THE PATIENT IS HAPPY VISUALLY AND REFRACTIVE HAS BEEN CONSISTENT. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER UPDATED THAT THE SURGEON ANTICIPATES TO DO AN EXCHANGE IN THE FUTURE. POSTOPERATIVE NOTES INDICATED "NO SUBJECT CHANGE IN VISION, VERY SMALL PERIPHERAL CATARACT". H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2020. THE SURGEON HAS REPORTED THE PATIENT NOW HAD SHALLOW VAULT AND SMALL SUBCAPSULAR CATARACTS IN THE MIDPERIPHERAL ZONE SUPERIORLY. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH/SIMILAR POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13563731&pc=QCB | 2023826-2022-00542 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN DEC 2020. THE SURGEON HAS REPORTED THE PATIENT NOW HAS SHALLOW VAULT AND SMALL SUBCAPSULAR CATARACTS IN THE MIDPERIPHERAL ZONE SUPERIORLY. IT WAS ALSO NOTED THE PATIENT IS HAPPY VISUALLY AND REFRACTIVE HAS BEEN CONSISTENT. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER UPDATED THAT THE SURGEON ANTICIPATES TO DO AN EXCHANGE IN THE FUTURE. POSTOPERATIVE NOTES INDICATED "NO SUBJECT CHANGE IN VISION, SMALL PERIPHERAL CATARACT". H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#:(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2020. THE SURGEON HAS REPORTED THE PATIENT HAD SHALLOW VAULT AND SMALL SUBCAPSULAR CATARACTS IN THE MIDPERIPHERAL ZONE SUPERIORLY. ON 29-JUN-2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH/SIMILAR POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. . CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14116278&pc=QCB | 2023826-2022-01207 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE AS A REPLACEMENT LENS. IT WAS INDICATED "THE CASE WENT VERY WELL BUT SEVERAL HOURS LATER THE PATIENT COULD NOT SEE WELL". IT WAS NOTED DURING SLIT LAMP THE PATIENT HAD WHAT LOOKED LIKE "VERY SIGNIFICANT CENTRAL ANTERIOR OPACITY COVERING MOST OF THE CRYSTALLINE LENS; IOP; VAULT; CORNEA WERE UNREMARKABLE". LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14197863&pc=QCB | 2023826-2022-01310 | 2021-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM12.0; -10.0/1.5/77 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2013. THE SURGEON REPORTS OPACITY - NS CATARACT WHICH WAS OBSERVED ON (B)(6) 2021. ALSO ON THIS DATE THE LENS WAS EXPLANTED AND REPLACED WITH ANOTHER IOL. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506297&pc=QCB | 2023826-2022-03329 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/016 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO A CATARACT DEVELOPED. THE LENS WAS REPLACED WITH AN IOL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH DEBRIS ON THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PHACOEMULSIFICATION (CATARACT SURGERY), IOL IMPLANTED. B3: DATE OF EVENT (B)(6) 2022. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/016 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND THE DEVELOPMENT OF A CATARACT (NUCLEAR SCLEROTIC), OBSERVED ON (B)(6) 2022. PHACOEMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. D6B: (B)(6) 2022. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15679254&pc=QCB | 2023826-2022-03845 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 -12.5/2.5/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2022 LENS OPACITY (ASC) WAS OBSERVED. IT WAS REPORTED THAT ON (B)(6) 2022 CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN IOL WAS IMPLANTED. PROBLEM RESOLVED. IN THE REPORTER OPINION THE LENS DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED "PATIENT DEVELOPED EARLY CATARACT". Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15679288&pc=QCB | 2023826-2022-03846 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 -12.0/3.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 LENS OPACITY (ASC) WAS OBSERVED. IT WAS REPORTED THAT ON (B)(6) 2022 CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN IOL WAS IMPLANTED. PROBLEM RESOLVED. IN THE REPORTER OPINION THE LENS DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED " PATIENT DEVELOPED EARLY CATARACT ". Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16018250&pc=QCB | 2023826-2022-04428 | 2022-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 . LOW VAULT WITH LENS OPACITY (CATARACT) ASC WAS REPORTED. THE LENS WAS REMOVED ON (B)(6) 2022 AND ON THE SAME DAY BUT DIFFERENT SURGERY A REPLACEMENT LENS OF DIFFERENT MODEL AND LONGER LENGTH WAS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS DEVICE; USER ERROR AND UNKNOWN. THE DEVICE FAILED TO PERFORM AS INTENDED. "EXCESSIVE LOW ICL VAULT MOST LIKELY CONTRIBUTED TO LENS OPACITY FORMATION." SEE MFR # 2023826-2022-04656 FOR EXCHANGE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WITH LENS OPACITY (CATARACT) ASC WAS REPORTED. THE LENS WAS REMOVED ON (B)(6) 2022 AND ON THE SAME DAY BUT DIFFERENT SURGERY A REPLACEMENT LENS OF DIFFERENT MODEL AND LONGER LENGTH WAS IMPLANTED. REPORTEDLY THE EXCHANGE LENS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16345841&pc=QCB | 2023826-2023-00483 | 2022-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.00/3.0/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO LENS OPACITY (CATARACT); ANTERIOR SUBCAPSULAR CATARACTS (ASC). PHACOEMULSIFICATION WAS DONE AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. "LENS TOUCH BETWEEN ICL AND ANTERIOR CAPSULE." Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16526528&pc=QCB | 2023826-2023-00873 | 2023-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6 -11.0/1.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. IT WAS REPORTED THAT LENS OPACITY (ASC) WAS OBSERVED ON THE SAME DATE OF SURGERY, THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY AND THE PROBLEM RESOLVED. IN THE REPORTER OPINION USER ERROR WAS THE CAUSE OF THE EVENT. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE DOCTOR MADE AN INCISION THAT DAMAGED THE PATIENT'S ANTERIOR CRYSTALLINE CAPSULE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16527034&pc=QCB | 2023826-2023-00874 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2 0.50/5.5/007 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. IT WAS REPORTED THAT LENS OPACITY (NS) WAS OBSERVED ON THE SAME DATE OF SURGERY, THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY AND THE PROBLEM RESOLVED. THE REPORTER STATES " THE DR INFORMS ME THAT SHE HAD AN INTRAOPERATIVE ADVERSE EVENT, SHE TOUCHES THE CRYSTALLINE LENS AT THE TIME OF THE INJECTION OF THE LENS". INTRAOPERATIVE ADVERSE EVENT WAS THE CAUSE OF THE CATARACT. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16541830&pc=QCB | 2023826-2023-00855 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND AN ANTERIOR SUBCAPSULAR CATARACT (ASC) HAD DEVELOPED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN; ETHNICITY: UNKNOWN; RACE: UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545625&pc=QCB | 2023826-2023-00856 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND AN ANTERIOR SUBCAPSULAR CATARACT (ASC) HAD DEVELOPED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFORMATION: UNK. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16569934&pc=QCB | 2023826-2023-00972 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/0.5/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING; AND LENS OPACITY (ASC) THAT WAS OBSERVED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF A DIFFERENT MODEL AND LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - (B)(4): OFF-LABEL USE (ACD < 3.0MM) (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16694265&pc=QCB | 2023826-2023-01299 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -12.0/+1.5/17 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. IT WAS INDICATED THAT THE PATIENT EXPERIENCED SUBCAPSULAR OPACITY. ON (B)(6) 2023 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "HIPERLAXITY" AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16770068&pc=QCB | 2023826-2023-01385 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13,2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO AN ANTERIOR SUBCAPSULAR OPACITY (CATARACT) WAS OBSERVED ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791232&pc=QCB | 2023826-2023-01444 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.1 -8.0/2.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT AND LENS OPACITY WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN OPTIC AND HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791528&pc=QCB | 2023826-2023-01445 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6 -7.0/1.5/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT AND LENS OPACITY WAS OBSERVED. ON 13-FEB-2023 THE LENS WAS EXPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN OPTIC AND HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989892&pc=QCB | 2023826-2023-01992 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1 -10.50/1.00/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS OPACITY (CORTICAL) WAS OBSERVED ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS REPORTED THAT CATARACT SURGERY WAS ALSO PERFORMED. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: A1: (B)(6). H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17226939&pc=QCB | 2023826-2023-02566 | 2023-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.50/2.0/052 (SPHERE / CYLINDER / AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS OPACITY (OBSERVED: ON (B)(6) 2023) AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT AN IOL WAS IMPLANTED AND PATIENT IS HAPPY. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17263291&pc=QCB | 2023826-2023-02664 | 2023-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, ASC LENS OPACITY WAS OBSERVED AND THE LENS WAS REMOVED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17308923&pc=QCB | 2023826-2023-02762 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.0/0.5/97 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO LENS OPACITY, ASC THAT WAS OBSERVED IN (B)(6) 2023. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "SURGEON TOLD: LENS OPACITY WAS NOT PRODUCT RELATED." THE PROBLEM WAS RESOLVED. THE REPORTER ALSO STATED, "THE SURGEON EXPLANTED THE ICL AND IMPLANTED A MONFOCAL IOL ALL FINE! PATIENT IS HAPPY." Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342578&pc=QCB | 2023826-2023-02887 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+1.0/019 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO AN ANTERIOR SUBCAPSULAR (ASC) OPACITY (CATARACT) WAS NOTED ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486835&pc=QCB | 2023826-2023-03271 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF CATARACT AND REFRACTIVE SURGERY IN 2023 TITLED, "CLINICAL OUTCOMES OF TORIC IMPLANTABLE COLLAMER LENS.' THE ARTICLE STATES THAT 8 EYES DEVELOPED ANTERIOR SUBCAPSULAR CATARACTS, OF THESE FOUR UNDER WENT CATARACT SUGERY. TWO EYES LENSES REQUIRED REPLACEMENT DUE TO LOW VAULT. ANOTHER LENS WAS REPOSITIONED DUE TO OCULAR TRAUMA. WHILE THE FINAL LENS WAS REPOSITIONED THREE TIMES AND ITS VAULT WAS MODERATE. CATARACT SUGER WAS PERFORMED DUE TO THE ANTERIOR SUBCAPSULAR CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: "2993", "1494- OFF-LABEL USE (ACD<3.0MM)" AND "1494- OFF LABEL USE (AGE<21)" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4629" AND "4614" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. B5- AN ARTICLE WAS PUBLISHED IN JOURNAL OF CATARACT AND REFRACTIVE SURGERY PUBLISHED AHEAD OF PRINT A CASE STUDY OF "EVALUATION OF LONG TERM CLINICAL OUTCOMES AFTER TORIC IMPLANTABLE COLLAMER LENS IMPLANTATION". THE PATIENTS EXPERIENCED ANTERIOR SUBCAPSULAR CATARACTS, LENS REPOSITIONING, AND LENS ROTATION. THIS STUDY INCLUDED "133 EYES OF 77 PATIENTS WHO UNDERWENT TICL IMPLANTATION FOR THE CORRECTION OF MYOPIA AND ASTIGMATISM BETWEEN 2005 AND 2009". REPORTEDLY, "OF THE 133 EYES, EIGHT (6.2%) DEVELOPED ANTERIOR SUBCAPSULAR CATARACTS DURING THE 8-10-YEAR FOLLOW-UP PERIOD; OF THESE, FOUR (3.1%) UNDERWENT CATARACT SURGERY. MOREOVER, ONE (0.78%) EYE REQUIRED AN ICL REPLACEMENT 7 MONTHS AFTER SURGERY DUE TO A LOW VAULT. ANOTHER EYE (0.78%) REQUIRED AN ICL REPLACEMENT 10 MONTHS AFTER SURGERY DUE TO A LOW VAULT, AND CATARACT SURGERY WAS PERFORMED 41 MONTHS AFTER THE INITIAL ICL SURGERY. ADDITIONALLY, 1 EYE (0.78%) HAD A 45° AXIAL MISALIGNMENT DUE TO AN ACCIDENTAL BLUNT OCULAR TRAUMA BY ANOTHER PERSON'S ARM AND UNDERWENT REPOSITIONING SURGERY 10 MONTHS AFTER THE INITIAL SURGERY. ANOTHER EYE (0.78%) UNDERWENT ICL REPOSITIONING SURGERY THREE TIMES AT 4, 7, AND 41 MONTHS AFTER SURGERY. ITS VAULT WAS MODERATE, BUT REPEATED ICL ROTATION OCCURRED, AND CATARACT SURGERY WAS PERFORMED 9 YEARS AFTER SURGERY DUE TO THE ANTERIOR SUBCAPSULAR CATARACT". IN CONCLUSION, THE AUTHORS STATED THAT "TICL SURGERY WAS CONSIDERED AN EFFECTIVE METHOD TO CORRECT MYOPIA AND ASTIGMATISM". (B)(4) Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17595282&pc=QCB | 2023826-2023-03530 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT CATARACTS WAS OBSERVED IN PATIENT. LENS WAS IMPLANTED IN 2021. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS OPACITY (CATARACT) ASC WAS OBSERVED ON (B)(6) 2023. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT IS TRAILING A SCL FOR THE REFRACTIVE SHIFT AND WILL REFER CATARACT SURGERY FOR NOW. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTION: G2: REPORTED AS FOREIGN - CORRECT TO REMOVE FOREIGN. CLAIM #(B)(4). Manufacturer Narrative: A1 - UNK A2 - UNK, A3 - UNK, A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY), D4 - UNK, D6A - UNK, H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18007848&pc=QCB | 2023826-2023-04808 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/+2.50/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS OPACITY (CORTICAL) ASSESSED ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH A CATARACT LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18268802&pc=QCB | 2023826-2023-05293 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/131 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE. THE LENS WAS EXPLANTED DUE TO THE DEVELOPMENT OF AN ANTERIOR CHAMBER CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18419826&pc=QCB | 2023826-2023-05835 | 2023-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -0.00/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO LENS OPACITY. PROBLEM RESOLVED. IT WAS REPORTED THAT CATARACT SURGERY AND IOL IMPLANTATION WAS PERFORMED AND THE PATIENT IS HAPPY. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960156&pc=QCB | 2023826-2020-00852 | 2020-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S LEFT (OS) EYE. REPORTEDLY, THE PATIENT DEVELOPED PSC CATARACTS ON (B)(6) 2019 THAT APPEARED TO RESOLVE BY (B)(6) 2019. THE CATARACT REAPPEARED IN (B)(6) 2019. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS (B)(4): DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 12.6MM TMICL 12.6 OF -7.50/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT (OS) EYE ON (B)(6) 2019. IT WAS REPORTED THAT THE PATIENT DEVELOPED PSC CATARACT ON (B)(6) 2019, THAT APPEARED TO RESOLVE BY SLIT LAMP AND IMAGING BY (B)(6) 2019. THE CATARACT RE-APPEARED IN (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020, AND THIS RESOLVED THE PROBLEM. STATUS OF EYE: S/P CATARACT EXTRACTION WITH PCIOL IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B3- DATE OF EVENT: (B)(6) 2020 B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENTS LEFT (OS) EYE ON (B)(6) 2019. IT WAS REPORTED THAT THE PATIENT DEVELOPED PSC CATARACT ON (B)(6) 2019 THAT APPEARED TO RESOLVE BY SLIT LAMP AND IMAGING BY (B)(6)2019. THE CATARACT RE-APPEARED IN (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 AND THIS RESOLVED THE PROBLEM. STATUS OF EYE: S/P CATARACT EXTRACTION WITH PCIOL IMPLANTATION. D6A- (B)(6) 2019, D6B- (B)(6) 2020. H6- CLINICAL CODE 4581- "TRANSIENT LOSS OF VISION". TYPE OF INVESTIGATION 4110- LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14287968&pc=QCB | 2023826-2022-01304 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS) 2 YEARS AGO. THE PATIENT NOW HAS AN ANTERIOR SUBCAPSULAR CATARACT (ASC) AND SYNECHIAS IN THE ANGLE 0° AND 180° BECAUSE OF THE UPSIDE DOWN SITUATION OF THE ICL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14287961&pc=QCB | 2023826-2022-01303 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD) 2 YEARS AGO. THE PATIENT NOW HAS AN ANTERIOR SUBCAPSULAR CATARACT (ASC) AND SYNECHIAS IN THE ANGLE 0° AND 180° BECAUSE OF THE UPSIDE DOWN SITUATION OF THE ICL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S RIGHT (OD) ON ON (B)(6) 2019. THE PATIENT EXPERIENCED LOW VAULT; ANTERIOR SUBCAPSULAR CATARACT (ASC) AND SYNECHIAS DUE TO THE LENS IMPLANTED UPSIDE DOWN. ON (B)(6) 2022 THE LENS WAS EXPLANTED; CATARACT SURGERY WAS PERFORMED AND IOL WAS IMPLANTED AND PROBLEM WAS RESOLVED." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. MODEL # EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12437079&pc=QCB | 2023826-2021-03475 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/044 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE AND THE PATIENT DEVELOPED A CATARACT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) -2021. THE PATIENT HAD BLURRED VISION. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS DEBRIS ON LENS. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - CATARACT SURGERY, IOL IMPLANTED CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4). CLAIM #:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840495&pc=QCB | 2023826-2021-04372 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS OPACITY-ASC, AND VISION LOSS.THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "SIZE ERROR, OCOS CALCULATION." REPORTEDLY, THE PROBLEM WAS NOT RESOLVED WITH EXPLANT. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY CATARACT SURGERY IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC WAS BROKEN AND BENT AND FIBERS WERE ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13146115&pc=QCB | 2023826-2021-04373 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS OPACITY-ASC, AND VISION LOSS. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "SIZE ERROR, OCOS CALCULATION." REPORTEDLY, THE PROBLEM WAS NOT RESOLVED WITH EXPLANT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: CATARACT SURGERY IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14108334&pc=QCB | 2023826-2022-01183 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Unintended Movement | Cataract; Blurred Vision | NA | 1 | Event Description: ON 14-MAR-2022 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN NATURE PORTFOLIO, ENTITLED, 'MULTICENTER CLINICAL OUTCOMES OF HOLE IMPLANTABLE COLLAMER LENS IMPLANTATION IN MIDDLE AGED PATIENTS'. THE STUDY REPORTS A TOTAL OF 118 EYES OF 65 PATIENTS AGED 45-65 YEARS WITH MYOPIC REFRACTIVE ERRORS RANGING FROM -2.13 TO -18.75 DIOPTERS (D) UNDERWENT HOLE ICL IMPLANTATION AND ROUTINE POSTOPERATIVE EXAMINATIONS. THE AVERAGE OBSERVATION PERIOD WAS 2.2 ± 1.0 YEARS. THE SAFETY AND EFFICACY INDICES WERE 1.08± 0.21 AND 0.87 ± 0.25, RESPECTIVELY. MANIFEST REFRACTION CHANGES OF 0.20 ± 0.43 D OCCURRED FROM 1 MONTH TO THE FINAL VISIT AFTER ICL IMPLANTATION. EIGHT EYES (6.8%) DEVELOPED ASYMPTOMATIC ANTERIOR SUBCAPSULAR CATARACT (ASC) IMMEDIATELY AFTER SURGERY, AND THREE EYES (2.5%) DEVELOPED CLINICALLY SIGNIFICANT SYMPTOMATIC NUCLEAR CATARACTS DURING THE FOLLOW-UP PERIOD. THIS MULTICENTER STUDY DEMONSTRATED A 6.8% (8/118) INCIDENCE OF SECONDARY SURGICAL INTERVENTION WITH AN AVERAGE FOLLOW-UP PERIOD OF 2.2 YEARS. TWO EYES REQUIRED ROTATION OF A TORIC ICL, AND THREE EYES EXCHANGED THE ICL DUE TO EXCESSIVE HIGH VAULT. THREE EYES (2.5%) DEVELOPED CLINICALLY SIGNIFICANT NUCLEAR CATARACTS POSTOPERATIVELY, AND SIMULTANEOUS LENS EXTRACTION AND PHACOEMULSIFICATION WITH IOL IMPLANTATION WERE SUCCESSFULLY PERFORMED. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15578068&pc=QCB | 2023826-2022-03617 | 2022-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -3.0/1.50/064 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT INDICATED THE PATIENT WILL BE MEASURED AGAIN IN A FEW MONTHS; THE INITIAL REFRACTIVE MUST BE RE-DETERMINED THEN A DECISION ABOUT A NEW IMPLANTATION WILL OCCUR. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+1.50/64 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. A CATARACT WAS OBSERVED. A IOL WAS IMPLANTED AND THE PROBLEM. STATUS OF THE EYE: "THE PATIENT HAS ALSO DEVELOPED A CATARACT". ADDITIONAL INFORMATION: THE ICL WAS EXCHANGED FOR AN IOL. THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+1.50/64 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND A CATARACT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND AN IOL WAS IMPLANTED. THIS RESOLVED THE PROBLEM. STATUS OF THE EYE: "THE PATIENT HAS ALSO DEVELOPED A CATARACT". ADDITIONAL INFORMATION: THE ICL WAS EXCHANGED FOR AN IOL. THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16018251&pc=QCB | 2023826-2022-04429 | 2022-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT; PERMANENT LOSS OF BCVA AND LENS OPACITY (CATARACT) ASC WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS DEVICE; USER ERROR AND UNKNOWN. THE DEVICE FAILED TO PERFORM AS INTENDED. "EXCESSIVE LOW ICL VAULT MOST LIKELY CONTRIBUTED TO LENS OPACITY PROGRESSION." SEE MFR# 2023826-2022-04657. Manufacturer Narrative: ADDITIONAL INFORMATION: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT; PERMANENT LOSS OF BCVA AND LENS OPACITY (CATARACT) ASC WAS REPORTED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. REPORTEDLY THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17547301&pc=QCB | 2023826-2023-03523 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/051 9SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND LENS OPACITY. IT WAS REPORTED THAT REFRACTIVE TREATMENT AND PHACO-EMULSIFICATION WITH IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM RESOLVED. THE REPORTER STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED, FOR CAUSE DETAILS THE REPORTER STATED " THE REFRACTIVE SURPRISE RESULTED IN A HYPEROPIC LVC (LASER VISION CORRECTION) AND UNSATISFACTORY VISION. THE ASC WAS OF LESSER ISSUE". Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/051 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL ON LENS SURFACE SPECIFICALLY RESIDUE THROUGHOUT THE LENS. CLAIM #(B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- IMPACTL CODE 4625: REFRACTIVE TREATMENT AND CATARACT SURGERY WITH IOL IMPLANTATION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741313&pc=QCB | 2023826-2023-04041 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Cataract; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.00/+0.50/153 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS IMPLANTED UPSIDE DOWN AND THE PATIENT HAD BLURRY VISION. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPOSITIONED AND IMPLANTED THE CORRECT WAY. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND DEVICE CAUSALITY IS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6- HEALTH EFFECT - CLINICAL CODE: "1766" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.00/0.5/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A CORTICAL LENS OPACITY AND BLURRED VISION. ON (B)(6) 2023, WAS EXPLANTED AND REPOSITIONED THE CORRECT WAY. THE PROBLEM WAS NOT RESOLVED, "REFRACTIVE LENS EXCHANGE/CATARACT SURGERY IS REQUIRED AS PATIENT IS +6 AND HAS A LENS OPACITY" AND " I WILL INFORM YOU OF THE RLE SURGERY DATE AND OUTCOME". STATUS OF THE EYE: "VERY BLURRY REFRACTION". THE CAUSE OF THE EVENT IS USER ERROR. CAUSE DETAILS: "LENS TWISTED ON IMPLANTATION. AIR BUBBLE CONFUSED POSITIONING HOLE." CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16019193&pc=QCB | 2023826-2022-04657 | 2022-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS AN EXCHANGE LENS. IT IS REPORTED THAT "THE ICL HAS LOW VAULT WITH SIGNIFICANT INFERIOR DECENTRATION AND CCW ROTATION. DESPITE REPLACING ICL WITH ONE STEP LARGER SIZED ICL THE VAULT REMAINS SHALLOW AND INFERIOR DECENTRATION SUPPORTS UNDERSIZING OF REPLACEMENT ICL DESPITE INCREASING SIZE ONE STEP." THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2022-04429 FOR INITIAL LENS. Manufacturer Narrative: CLINICAL CODE: 4581: SIGNIFICANT INFERIOR DECENTRATION. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: RECEIVED INFORMATION THAT THIS EXCHANGE LENS DID RESOLVE THE ISSUE OF LOW VAULT. THIS CLAIM IS NOT A COMPLAINT DUE TO NO PRODUCT DEFICIENCY AGAINST THE LENS. PLEASE DISREGARD THE INITIAL MDR. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16688954&pc=QCB | 2023826-2023-01202 | 2022-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/2.5/84(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASSESSED: (B)(6) 2022), PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL AND BLURRED VISION AND THE PROBLEM RESOLVED. THE REPORTER STATED THAT ON (B)(6) 2022 PILOCARPINE WAS USED FOR AROUND 1 MONTH, PATIENT PUPIL STILL UNREACTIVE AND STOP USING PILOCARPINE THEN. CATARACT SURGERY IS PLANNED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: UNREACTIVE/FIXED PUPIL, PIGMENT DISPERSION. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678146&pc=QCB | 2023826-2023-01164 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. A CATARACT WAS OBSERVED AND THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READING VA. THE LENS REMAINED IMPLANTED BUT THE PLAN WAS TO REMOVE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. AND LENS OPACITY (ANTERIOR SUBCAPSULAR) WAS OBSERVED. THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READY VA. THE CAUSE WAS UNKNOWN. D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EVENT DATE:UNK. EXPIRATION DATE: UNK. EXPLANT DATE: UNK. MFR DATE: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - DIFFICULTY READING VA. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16692052&pc=QCB | 2023826-2023-01165 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. A CATARACT WAS OBSERVED AND THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READING VA. THE LENS REMAINED IMPLANTED BUT THE PLAN WAS TO REMOVE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/050 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. AND LENS OPACITY (ANTERIOR SUBCAPSULAR) WAS OBSERVED. THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READY VA. THE CAUSE WAS UNKNOWN. D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: A4: WEIGHT UNK. A5: ETHNICITY UNK. A6: RACE UNK. B3:EVENT DATE UNK. D4: EXPIRATION DATE UNK. D6B: EXPLANT UNK. H4: DEVICE MANUFACTURE DATE UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DIFFICULTY READING VA. CLAIM # 432841
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17610615&pc=QCB | 2023826-2023-03705 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS OPACITY (CATARACT) ASC WAS OBSERVED ON (B)(6) 2023. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT IS TRAILING A SCL FOR THE REFRACTIVE SHIFT AND WILL REFER CATARACT SURGERY FOR NOW. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT - CLINICAL CODE: "1766" AND "2137" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "2199" SHOULD BE REMOVED. "4614 " SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE REMOVED. "1401" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.00/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED AN ASC LENS OPACITY OBSERVED ON 17/MAY/2023, REFRACTIVE CHANGE OVER TIME, AND A SPHERICAL CONTACT LENS PRESCRIBED. THE LENS REMAINS IMPLANTED. STATUS OF THE EYE:" PATIENT IS TRIALING SCL (SPHERICAL CONTACT LENS) OS FOR THE REFRACTIVE SHIFT; WILL DEFER CATARACT SURGERY FOR NOW." CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D6A - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611928&pc=QCB | 2023826-2019-00824 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -5.50/2.5/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA, LENS OPACITY (ASC), SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS, AND BLURRED VISION WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS PERFORMED. AS OF (B)(6) 2019 PATIENT'S "VISION IS GREAT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT CODE: 114 SHOULD BE CORRECTED TO 213 IN INITIAL MEDWATCH REPORT. PATIENT CODE: 1766 SHOULD BE INCLUDED IN INITIAL MEDWATCH REPORT. DATE OF THIS REPORT: 04/17/2019 SHOULD BE CORRECTED TO 04/16/2019 IN INITIAL MEDWATCH REPORT. DEVICE CODE: 1494 OFF-LABEL USE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611958&pc=QCB | 2023826-2019-00791 | 2019-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/0.5/150 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOSS OF BCVA, BLURRED VISION, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS SCHEDULED FOR (B)(6) 2019 BUT THE PATIENT PASSED AWAY ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: "REMOVED ON 10-MAR-2019" SHOULD BE CORRECTED TO "REMOVED ON (B)(6) 2019" IN INITIAL MDR. "10-MAR-2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936253&pc=QCB | 2023826-2020-00745 | 2020-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -8.0/+2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS OPACITY PSC CORTICAL. TRANSIENT LOSS OF BCVA IS ALSO REPORTED. REPORTEDLY, THE PATIENT DEVELOPED PSC CATARACTS ON (B)(6) 2019 THAT APPEARED TO RESOLVE BY (B)(6) 2019. THE CATARACT REAPPEARED IN (B)(6) 2019 AND BCVA WAS LIMITED TO 20/50. Manufacturer Narrative: H6 - METHOD CODE 3331. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID. VISUAL INSPECTION FOUND A CIRCULAR MARK ON THE OPTIC. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE PROBLEM WAS RESOLVED WITH S/P CATATARCT EXTRACTION AND A PCIOL IMPLANTATION." THIS STATEMENT SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4-DATE IN INITIAL MDR IS CORRECTED TO (B)(6) 2020. D10- DATE RECEIVED BY MANUFACTURER IN INITIAL MDR IS CORRECTED TO 13-MAR-2020. G4- DATE IN INITIAL MDR IS CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476134&pc=QCB | 2023826-2020-01935 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR OPACITY AND BLURRED VISION. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. SEE MFR REPORT # 2023826-2020-01934 FOR INITIAL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS INSPECTED UNDER 2X MAGNIFICATION AND THERE WAS NO VISIBLE DAMAGE OR RESIDUE NOTED. CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. IT IS LIKELY THAT THESE LENSES WERE INADVERTENTLY SWAPPED DURING RETURN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16904423&pc=QCB | 2023826-2023-01686 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Cataract; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS OPACITY (CATARACT, ANTERIOR SUBCAPSULAR) WAS OBSERVED ON (B)(6) 2022. THERE WAS PIGMENT DISPERSION AND THE PATIENT COMPLAINED OF BLURRED VISION AND GLARE/HALOES. THE PROBLEM WAS RESOLVED. CHAFING OF IRIS WITH RESULTING SIGNIFICANT IRIS TRANSILLUMINATION BETWEEN 1 WEEK AND 1 MONTH POST OP WITH SLOW PROGRESSION; DESPITE EXCELLENT VAULT. BELIEVED TO BE MECHANICAL ISSUE RELATED TO LENS/IRIS APPROXIMATION. THE POST-OP STATUS OF THE EYE AFTER LENS REMOVAL WAS STABLE IRIS TRANSILLUMINATION. ANISOCORIA IMPROVED BUT STILL WITH MYDRIASIS AND OVOID PUPIL. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, CHAFING OF IRIS, IRIS TRANSILLUMINATION, MYDRIASIS. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706991&pc=QCB | 2023826-2020-02407 | 2020-09-15 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Capsular Bag Tear | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/094 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY, DUE TO AN ANTERIOR SUBCAPSULAR OPACITY AND A CAPSULAR TEAR HAD OCCURRED. CATARACT SURGERY WAS PERFORMED AND AN IOL LENS WAS IMPLANTED IN THE CILIARY SULCUS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO 2 PITS WERE NOTED IN THE LENS CAPSULE (ANTERIOR) DIRECTLY BELOW THE IRIDOTOMIES. WHEN THE LENS WAS INSERTED A CAPSULAR TEAR EXTENDED FROM THE PITS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND THE HAPTICS TORN. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586886&pc=QCB | 2023826-2020-00067 | 2019-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Cataract; Corneal Decompensation; Corneal Edema; Glaucoma; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AFTER AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S EYE. "IOP WAS 27 AFTER IMPLANTATION, BUT LATER INCREASED AND REACHED 70 MM HG. MANNITOL WAS USED MULTIPLE TIMES AND PRESSURE WENT DOWN BUT THERE IS SEVERE CORNEAL EDEMA (HAZY CORNEA)." LENS REMAINS IMPLANTED. REPORTEDLY SURGEON IS "THINKING OF EXPLANTING THE LENS." REPORTER INDICATED, "PATIENT HAS ACUTE GLAUCOMA, THE SURGEON DOES NOT KNOW THE CAUSE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: A1 - (B)(6). B5 - PER UPDATED INFORMATION, THE PATIENT ALSO DEVELOPED ANTERIOR SUBCAPSULAR CATARACT, PUPIL BLOCK, CORNEAL DECOMPENSATION, AND UNREACTIVE (FIXED) PUPIL. TOPICAL STEROIDS AND GLAUCOMA MEDICATIONS WERE PRESCRIBED. 'DIFFERENCE IN COLOUR PERCEPTION BETWEEN EYES' WAS ALSO REPORTED. THE PATIENT IS REPORTEDLY 'STABLE AND SLOWLY IMPROVING.' THE LENS REMAINS IMPLANTED AND THE SURGEON DOES NOT PLAN TO REOPERATE. THE CAUSE OF THE EVENT WAS REPORTED AS PATIENT RELATED FACTOR. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (PUPIL BLOCK, UNREACTIVE (FIXED) PUPIL, DIFFERENCE IN COLOUR PERCEPTION). H6 - PATIENT CODE 3191: MEDICAL INTERVENTION (TOPICAL STEROIDS, GLAUCOMA MEDICATIONS). H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 2.8MM). H6 - METHOD CODE 4110: WORK ORDER SEARCH - NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8970000&pc=QCB | 2023826-2019-01541 | 2019-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Cataract; Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.50/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ANGLE CLOSURE WITH ELEVATED IOP (26 MMHG) AND LOW VAULT WERE ASSESSED ON (B)(6) 2019. EYE STATUS ON (B)(6) 2019, "AC QUIET WITH MINIMAL CORNEAL EDEMA. NEAR VA J8." LENS REMAINS IMPLANTED. Manufacturer Narrative: "LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SAME LENGTH, SAME MODEL LENS. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED IN THE INITIAL MDR. PATIENT CODE 3191: SECONDARY SURGERY- EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). ON A FOLLOW UP VISIT ON (B)(6) 2019, IT WAS REPORTED THAT THE PATIENT HAS "SLIGHT ANTERIOR CAPSULE OPACITY, ANTERIOR CHAMBER QUIET AND OPEN ANGLES". SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. EXPLANTED DATE: (B)(6) 2019. PATIENT CODE: 1791, PATIENT CODE: 1766, PATIENT CODE 3191: NO CODE AVAILABLE (ANGLE CLOSURE), DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), DEVICE CODE 2017: IMPROPER OR INCORRECT PROCEDURE OR METHOD. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254838&pc=QCB | 2023826-2019-00098 |  | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Endophthalmitis; Blurred Vision; Vitrectomy; No Code Available | NA | 1 | Event Description: A JOURNAL OF CATARACT REFRACTIVE SURGERY ARTICLE WAS RECEIVED ON 21-DEC-2018, ORIGINALLY PUBLISHED IN AUGUST 2018, TITLED 'ATYPICAL ENDOPHTHALMITIS AFTER INTRAOCULAR COLLAMER LENS IMPLANTATION.' THE ARTICLE DESCRIBES A CASE IN WHICH A (B)(6) WOMAN DEVELOPED ENDOPHTHALMITIS IN THE LEFT EYE FOLLOWING BILATERAL INTRAOCULAR COLLAMER LENS IMPLANTATION. BLURRED VISION, LOSS OF VISION, AND VISUAL ACUITY REDUCED TO HAND MOTIONS WERE REPORTED AT 3 DAYS POSTOP. THE LENS WAS IMMEDIATELY EXPLANTED AND VITRECTOMY COMBINED WITH INTRAVITREAL ANTIBIOTIC INJECTION WAS PERFORMED. STAPH EPIDERMIDIS WAS DETECTED IN THE VITREOUS TAP AND THE PATIENT WAS TREATED WITH INTRAVENOUS ANTIBIOTICS, ANTIBIOTIC EYEDROPS, AND PARABULBAR INJECTIONS OF DEXAMETHASONE. THE CDVA RECOVERED TO 20/20 AFTER A 6-MONTH PERIOD. POSSIBLE CAUSES INCLUDE THE PATHOGENS WERE INTRODUCED WITH THE ICL, OR DURING THE SURGERY, OR THAT AN ENDOGENOUS INFECTION UNRELATED TO THE ICL IMPLANTATION OCCURRED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: CORRECTED DATA: A JOURNAL OF CATARACT REFRACTIVE SURGERY ARTICLE TITLED 'ATYPICAL ENDOPHTHALMITIS AFTER INTRAOCULAR COLLAMER LENS IMPLANTATION' PUBLISHED IN AUGUST 2018 DESCRIBES A CASE IN WHICH A WOMAN DEVELOPED UNILATERAL ENDOPHTHALMITIS AFTER BILATERAL INTRAOCULAR COLLAMER LENS IMPLANTATION. THREE DAYS AFTER SURGERY THE PATIENT PRESENTED WITH SUDDEN LOSS OF VISION IN THE LEFT EYE. THE VISUAL ACUITY WAS REDUCED TO HAND MOTIONS. THE POSTERIOR POLE OF THE LEFT EYE WAS NOT VISIBLE BECAUSE OF THE PRESENCE OF DENSE INFILTRATES IN THE VITREOUS CAVITY. THE ICL WAS EXPLANTED AND A PARS PLANA VITRECTOMY COMBINED WITH AN INTRAVITREAL INJECTION OF CEFAZOLIN WAS PERFORMED. A REGIME OF TOPICAL ANTIBIOTICS AND STEROIDS WERE GIVEN. AN ANTIBIOTIC CULTURE SHOWED STAPH EPIDERMIDIS SENSITIVE TO CEFAZOLIN. AFTER SIX MONTHS A BARELY DETECTABLE NUCLEAR SCLEROTIC CATARACT WAS NOTED. TWO YEARS LATER THE PATIENT UNDERWENT ROUTINE CATARACT SURGERY WITH MONOFOCAL IOL IMPLANTATION. CLAIM#:(B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15211444&pc=QCB | 2023826-2022-02678 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Cataract; Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: ONE CASE OF ASYMPTOMATIC ANTERIOR SUBCAPSULAR OPACITY WAS OBSERVED IN THE T-ICL GROUP THAT PRESENTED WITH A CENTRAL VAULT LESS THAN 100UM NO VISUAL IMPAIRMENT WAS NOTICED DURING FOLLOW-UP. SIGNIFICANT AXIS MISALIGNMENT WAS REPORTED IN 6 EYES. IOL EXCHANGE WAS PERFORMED IN 4 EYES. WITH T-ICL IMPLANTATION: ONE CASE DUE TO POSTOPERATIVE ACUTE ANGLE CLOSURE GLAUCOMA TRIGGERED BY IGH VAULT AND 3 CASES DUE TO SIGNIFICANT AXIS ROTATION SECONDARY TO LOW VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9874226&pc=QCB | 2023826-2020-00634 | 2020-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, IRIS ATROPHY, LENS OPACITY ASC AND PUPIL BLOCK WERE OBSERVED. SECONDARY SURGICAL INTERVENTION (YAG AND ANTERIOR CHAMBER IRRIGATION) AND MEDICAL INTERVENTION (EYE DROPS TO REDUCE INFLAMMATION) WERE PERFORMED/PRESCRIBED. THE REPORTER STATED " THE OPACITY WAS SEEN BEFORE AC IRRIGATION AND BEFORE YAG PI. PATIENT IOP IS STABLE AND VA IS IMPROVING. THE DOCTOR IS WAITING FOR PUPIL IMPROVEMENTS SINCE IT SEEMS TO BE SLIGHTLY STUCK TO THE ICL." PATIENTS LAST VISIT DATE (B)(6) 2020 UCVA WAS REPORTED TO BE 20/30, BCVA PIN-HOLE 20/25, ASC OPACITY REMAINS THE SAME, IN CONTINUES EVALUATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, IRIS ATROPHY, LENS OPACITY ASC AND PUPIL BLOCK; ELEVATED IOP AND INFLAMMATION WERE OBSERVED. SECONDARY SURGICAL INTERVENTION (YAG AND ANTERIOR CHAMBER IRRIGATION) AND MEDICAL INTERVENTION (EYE DROPS TO REDUCE INFLAMMATION) WERE PERFORMED/PRESCRIBED. THE REPORTER STATED " THE OPACITY WAS SEEN BEFORE AC IRRIGATION AND BEFORE YAG PI. PATIENT IOP IS STABLE AND VA IS IMPROVING. THE DOCTOR IS WAITING FOR PUPIL IMPROVEMENTS SINCE IT SEEMS TO BE SLIGHTLY STUCK TO THE ICL." PATIENTS LAST VISIT DATE ON (B)(6) 2020 UCVA WAS REPORTED TO BE 20/30, BCVA PIN- HOLE 20/25, ASC OPACITY REMAINS THE SAME, IN CONTINUED EVALUATION. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- INCLUDED "ELEVATED IOP AND INFLAMMATION" TO STATEMENT. H6-PATIENT CODE 1937- ELEVATED IOP. 1932- INFLAMMATION. H6-DEVICE CODE: 1401-BLURRED VISION. H6-DEVICE CODE 2993 IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107370&pc=QCB | 2023826-2019-01751 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Cataract; Inflammation; Intraocular Pressure Increased; Pain; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -18.00/+1.0/177 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. A COUPLE OF HOURS AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED PAIN AND A NEGATIVE EFFECT ON HER VISION. A CATARACT HAD DEVELOPED. THE LENS WILL BE EXPLANTED BUT TO DATE REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 1937: INFLAMMATION. PATIENT CODE 3191: ANGLE CLOSURE. CLAIM#: (B)(4). Manufacturer Narrative: LENS IMPLANTED UPSIDE DOWN. PATIENT CODE 1937 IN PREVIOUS MDR CORRECTED TO 1932. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16261376&pc=QCB | 2023826-2023-00326 | 2022-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Cataract; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER (PATIENT CURRENT OBSERVING DOCTOR) STATES THAT A SURGEON AT A DIFFERENT CLINIC IMPLANTED A 13.7MM VTICMO 3.7 IMPLANTABLE COLLAMER LENS OF -7.00/0.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE CURRENT OBSERVING DOCTOR SAW PATIENT FOR THE FIRST TIME ON (B)(6) 2022, WHERE HE UNFORTUNATELY BECAME AWARE OF SEVERAL MEDICAL ERRORS. "TEMPORAL CORNEAL INCISION, ANTERIOR CHAMBER SHALLOWER, IRIS TEMPORALLY DEPIGMENTED AND "RAGGED", PIGMENT DISPERSION IN THE ANTERIOR CHAMBER, WIDE AND PLEGIC PUPIL, ICL SPECKLED WITH PIGMENT, ANTERIOR SUBCAPSULAR LENS OPACITY- CATARACT." COMPLAINT QUESTIONNAIRE PROVIDED REPORTS CATARACT (ASC); ANGLE CLOSURE WITH ELEVATED IOP; LENS DISLOCATION/ SUBLUXATION; PIGMENT DISPERSION; UNREACTIVE FIXED PUPIL; BLURRED VISION; AND IRIS ATROPHY WAS REPORTED. CURRENT STATUS OF THE EYE: "EYE ALREADY CALM, CORNEA CLEAR, ANTERIOR CHAMBER SHALLOW, PUPIL WIDE DILATED, ICL IN SITUATION ROTATED COUNTER CLOCKWISE TO INTENDED POSITION, ANTERIOR SUBCAPS CAT. CATARACT FORMATION." CAUSE REPORTED AS USER ERROE; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE IMPLANTING SURGEON HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10538715&pc=QCB | 2023826-2020-02049 | 2020-07-08 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Cataract; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; Vomiting; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/1.50/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE SURGEON ASSESSED AN ELEVATED IOP OF 22MMHG. THIS WAS CONTROLLED WITHIN THE FIRST 24 HOURS WITH MEDICATION. ALSO REPORTED IS PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, DILATED PUPIL, GLARE/HALOES, IRIS ATROPHY, PUPIL BLOCK, AND ANTERIOR SUBCAPSULAR OPACITIES. THE PATIENT ALSO EXPERIENCED NASEAU AND EXCESSIVE MOVEMENT DUE TO VOMITING CAUSED BY THE ANESTHESIA. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE PUPILLARY DILATION WAS "SLOWLY CORRECTED WITHIN THE FIRST MONTH, LEAVING AN ANISOCORIA OF APPROX. 2.5MM AND PRESENTING SECTORAL ATROPHY OF THE IRIS." THE PATIENT ALSO REPORTS "DECREASE IN THE INTENSITY OF THE COLORS." THE PATIENT STILL EXPERIENCES GLAUKOMFLECKEN, IRIS PIGMENT DISPERSION, AND GLARES AT NIGHT, BUT IS OVERALL "REALLY SATISFIED." THE CAUSE OF THE EVENT IS REPORTED AS "REMNANT OF OVD BEHIND THE ICL." Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE US.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8807833&pc=QCB | 2023826-2019-01275 | 2019-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -11.00/1.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE DAY AFTER OPERATION THE EYE HAD UNDERGONE MYDRIASIS, WHICH WAS IMPOSSIBLE TO NARROW. ON (B)(6) 2019 THE LENS WAS EXPLANTED. THE SURGEON STATED " THE EVENT WAS DUE TO AN UNKNOWN PROBLEM. I CAN ONLY SUPPOSE, THAT DIAMETER ICL WAS BIGGER. WE STILL DONT KNOW WHAT HAPPENED. CATARACT (BEHIND CAPSULA ANTERIOR) WAS OBSERVED AFTER IMPLANTATION. SHE STILL HAS PARALYTIC MYDRIASIS (5-6 MM).'' PATIENT CONDITION IS STILL THE SAME, ON (B)(6) 2019 THE DIAMETER OF THE PUPIL WAS 5-6 MM AND IT DIDN'T RESPOND TO THE LIGHT. PATIENT USED PILLOCARPIN 1% THE LAST 2 MONTHS. Manufacturer Narrative: DATE OF THIS REPORT: 20-JUN-2019. CONCOMITANT MEDICAL PRODUCTS: CARTRIDGE MODEL-SFC-45; LOT#-UNK SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT'S DATE OF BIRTH: (B)(6) 1996. PATIENT CODE 3191: NO CODE AVAILABLE (UNREACTIVE FIXED PUPIL) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 26 AUG 2019. OVD USED: EYEVISC PFS 2 ML (HYDROXYPROPYL METHYLCELLULOSE 2% W/V) OPHTALMIC SOLUTION USP FROM BIOTECH IOP WAS 50.0MM THE DAY AFTER SURGERY, WHEN MYDRIASIS WAS NOTICED. THE ANGLES WERE OPEN. MEDICATIONS USED TO TREAT PATIENT: DEXAMETASON-INIECTION 3 DAYS ONE TIME PER DAY, BETOPTIC AND AZOPT IN DROPS FOR 3 DAYS AND THEN THE IOP WAS OK -14 MM. PATIENT CODE 3191: MEDICAL INTERVENTION. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12978921&pc=QCB | 2023826-2021-04658 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; Pain; Pupillary Block | NA | 1 | Event Description: A JOURNAL ARTICLE WAS RECEIVED FROM THE JOURNAL OF CURRENT GLAUCOMA, ((B)(6) 2021), FOR "BILATERAL CATARACT DEVELOPMENT AND PUPILLARY BLOCK GLAUCOMA FOLLOWING IMPLANTABLE COLLAMER LENS". THE ARTICLE SITED A STUDY OF A (B)(6) MAN WITH HIGH MYOPIA FOR ROUTINE FOLLOW-UP STATUS POST BILATERAL PHAKIC ICL PLACEMENT. THE VISUAL ACUITY WAS REDUCED DUE TO AN ANTERIOR SUBCAPSULAR CATARACT AND ELEVATED IOP IN BOTH EYES WITH ADVANCE GLAUCOMATOUS VISUAL FIELD DEFECTS. THE PATIENT HAD PAIN IN THE RIGHT EYE. THE PIS IN THE RIGHT EYE WERE ENLARGED BY YAG. THE PATIENT WAS TREATED WITH TOPICAL GLAUCOMA MEDICATIONS. THE LEFT EYE UNDERWENT SAME-DAY PHAKIC ICL EXPLANTATION AND CATARACT SURGERY TO PREVENT FURTHER VISUAL FIELD LOSS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. MODEL #: EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620830&pc=QCB | 2023826-2022-03493 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/084, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR (ASC) CATARACT. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. POST-OP, THE PUPIL IN MYDRIASIS NON REACTIVE AND SLIGHT EPITHELIAL HAZE. THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - EPITHELIAL HAZE, MYDRIATIC PUPIL SHOULD HAVE BEEN INCLUDED. B5: POSTOP COMMENTS ADDITIONAL NOTED PUPIL IN MYDRISASI NON REACTIVE; SLIGHT EPITHELIAL HAZE. SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515453&pc=QCB | 2023826-2023-03415 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Cataract; Intraocular Pressure Increased; Retinal Detachment; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: AN ARTICLE WAS RECEIVED FOR "EVALUATION OF THE EVO/EVO+ SPHERE AND TORIC VISION ICL: SIX MONTH RESULTS FROM THE UNITED STATES FOOD AND DRUG ADMINISTRATION CLINIC TRIAL". THE TRIAL ENROLLED 629 EYES OF 327 PATIENTS. THE ARTICLE REPORTED THE FOLLOWING ADVERSE EVENTS - TRANSIENT INCREASE OF IOP (DUE TO RETAINED OVD); LENS REPOSITIONING AND LENS EXPLANT/EXCHANGE (2 EYES); EXCESSIVE VAULTING; BLURRED VISION (1 EYE); CATARACT DEVELOPMENT/MYOPIC SHIFT/REDUCED VISUAL ACUITY (1 EYE); HALOES/GLARE (1 EYE); RETINAL DETACHMENT/VITRECTOMY (3 EYES). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - THIS IS REGARDING A CLINICAL TRIAL THAT STAAR CONDUCTED. STAAR ALREADY REPORTS ALL CLINICAL TRIAL COMPLAINTS TO THE COMPLAINTS DEPARTMENT, THIS COMPLAINT SHOULD BE VOIDED, SINCE WE WOULD HAVE ALREADY CAPTURED ALL EVENTS LISTED IN THE ARTICLE THROUGH INTERNAL REPORTS. CLAIM # (B)(4). Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - VITRECTOMY. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366584&pc=QCB | 2023826-2021-00480 | 2021-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/010 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY, LENS OPACITY ASC WAS REPORTED; OBSERVED ON (B)(6) 2021, AFTER LENS WAS EXPLANTED. MIOL SURGERY PLANNED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO BROKEN HAPTIC. H6: 1766 SHOULD HAVE BEEN INCLUDED PREVIOUSLY. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368673&pc=QCB | 2023826-2021-00463 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -11.50/+2.5/179 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING ROTATED. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED. THE ICL WAS TOO SMALL. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: B5: THE ICL LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. H6: HEALTH EFFECT IMPACT CODE: 4625 - CATARACT SURGERY AND IOL IMPLANT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021862&pc=QCB | 2023826-2021-02006 | 2020-08-06 04:00 | Death | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Contamination; Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+6.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS DUE TO LENS OPACITY-CORTICAL. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES THAT THE LENS HAS A DEPOSIT ON ITS SURFACE. Manufacturer Narrative: B5-FINAL BCVA 0.85, UCVA 0.7. H6-METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THIS IS A CASE OF FOREIGN DEBRIS ON THE LENS THAT CAUSED THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. AND THERE WAS RESIDUE AND DEBRIS ON THE LENS SURFACE. RED PARTICULATES WERE NOTED, ON LENS OPTIC. CLAIM# (B)(4). Manufacturer Narrative: 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS; -9.5/+6.0/107. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378577&pc=QCB | 2023826-2021-03379 | 2013-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +4.0/+2.0/011 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2013. ON (B)(6) 2021 THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY THE PATIENT EXPERIENCED LENS OPACITY-OBSERVED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PHACO-EMULSIFICATION AND IOL IMPLANTATION PERFORMED. PROBLEM RESOLVED. H3-DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A CONTACT LENS CASE. THERE WAS DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND THERE WAS DEBRIS THROUGHOUT THE LENS. H6-HEALTH IMPACT CODE 4650 (PHACO-EMULSIFICATION AND IOL IMPLANTATION). CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465183&pc=QCB | 2023826-2022-03166 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; Unintended Movement | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS FLIPPED IN THE EYE. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT: ADDITIONAL SURGERY: 4625: CATARACT SURGERY, IOL IMPLANTED. H6: DEVICE CODE: 1494: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR/CORTICAL CATARACT. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED THE NEXT DAY AND TORIC IOL WAS IMPLANTED. THE PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (THE EYE WAS TOO SOFT AND PATIENT SQUEEZED WHEN THE INCISION FOR ICL WAS MADE, CUTTING THE ANTERIOR CAPSULE, WHICH INDUCED THE CATARACT). CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766105&pc=QCB | 2023826-2023-01343 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY, "PAT STILL WITH LENS OPACITY SUPRISINGLY, BUT BETTER VAULT." Manufacturer Narrative: CORRECTED DATA: B1: SHOULD BE CORRECTED TO ADVERSE EVENT. B2: OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS SHOULD BE ADDED FOR CORRECTION. H1: SERIOUS INJURY SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986861&pc=QCB | 2023826-2023-04603 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.5/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. MILD LENS OPACITY (ASC) WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT IS VERY HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED A MILD ASC. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "PATIENT VERY HAPPY." H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4614" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986877&pc=QCB | 2023826-2023-04604 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/1.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. MILD LENS OPACITY (ASC) WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT IS VERY HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/1.5/112 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED A MILD ASC. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "PATIENT VERY HAPPY". H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4614" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175024&pc=QCB | 2023826-2023-05105 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE. LENS OPACITY WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307838&pc=QCB | 2023826-2019-00213 | 2018-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED, THE SURGEON NOTING LENS OPACITY CORTICAL. REPORTEDLY, THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.5/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. A CAPSULAR TEAR OCCURRED. A CORTICAL CATARACT WAS NOTED. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND THE PATIENT UNDERWENT PHACOEMULSIFICATION AND IOL IMPLANTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN THE ORIGINAL MDR PATIENT CODE 3191: CAPSULAR TEAR, SECONDARY SURGERY, LENS EXPLANT, CATARACT SURGERY, IOL IMPLANTATION. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8426297&pc=QCB | 2023826-2019-00417 | 2018-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORT STATED "BEFORE THE SURGERY, DOCTORS HAD OBSERVED CATARACTS. AND THE DOCTOR HAD FULL COMMUNICATION WITH THE PATIENT". A VTICMO13.2, -8.5/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018 AND LATER REMOVED ON THE SAME DAY. LENS OPACITY (CORTICAL, DATE: (B)(6) 2018) WAS REPORTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO PATIENT RELATED FACTOR. NO TREATMENT OR INTERVENTION HAS BEEN PERFORMED. NON-TRAUMATIC CATARACT IN THE FELLOW EYE WAS REPORTED AS A PRE-EXISTING CONDITION. Manufacturer Narrative: THE REPORTER INDICATED THAT THAT THIS OPACITY WAS NOT PRESENT BEFORE THE ICL SURGERY. TO AVOID UNNECESSARY COMPLICATIONS, THE DOCTOR DECIDED NOT TO IMPLANT THE LENS FOR THE TIME BEING. THE PATIENT IS UNDER OBSERVATION AND LENS IMPLANTATION MAY BE CONSIDERED FOR THE FUTURE. IT WAS ALSO REPORTED THAT THE OPACITY MIGHT HAVE OCCURRED DUE TO LENS TOUCH DURING ICL IMPLANTATION. CONCLUSION CODE- CORRECTED FROM "24" TO "4315" FOR INITIAL MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8601622&pc=QCB | 2023826-2019-00811 | 2018-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-10.50/1.0/096 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: OTHER RELEVANT HISTORY: CATARACT IN THE OPERATIVE EYE. DEVICE CODE- 1494: OFF-LABEL USE (CATARACT IN OPERATIVE EYE). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8932863&pc=QCB | 2023826-2019-01545 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 14.00/+2.5/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO A CATARACT HAD DEVELOPED. THE PATIENT HAD A LASIK PROCEDURE AFTER THE ICL SURGERY. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LASIK) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. METHOD CODE: 4111 SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. IMPLANT DATE: (B)(6) 2018. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10011427&pc=QCB | 2023826-2020-00996 | 2020-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.00/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED, PATIENT IS IN VERY GOOD CONDITION. THE REPORTER STATED THAT LENS OPACITY ASC WAS NOTED BEFORE FIRST SURGERY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.00/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH AN UNSPECIFIED LENS ROTATION AND THE PROBLEM RESOLVED, PATIENT IS IN VERY GOOD CONDITION. THE REPORTER STATED THAT LENS OPACITY ASC WAS NOTED BEFORE FIRST SURGERY. D7-EXPLANT DATE-(B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. PATIENT CODE 1766-LENS OPACITY(CATARACT). DEVICE CODE 2923- ROTATION. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.00/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON (B)(6) 2020 REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED, PATIENT IS IN VERY GOOD CONDITION. THE REPORTER STATED THAT LENS OPACITY ASC WAS NOTED BEFORE FIRST SURGERY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATON : LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285662&pc=QCB | 2023826-2020-01480 |  | Injury | STAAR SURGICAL COMPANY | 2020-07-16 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/038 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. REPORTEDLY, THE PATIENT WAS OPERATED ON WITHOUT COMPLICATIONS. AT THE END OF THE SURGERY THE DOCTOR NOTICED AN OPACIFICATION IN THE CENTER OF THE LENS CORTEX BUT WASN'T SURE IF IT WAS THERE PREVIOUSLY. VAULTING INITIALLY HIGH BUT IS OK NOW. REFRACTION IS +0.75 INSTEAD OF THE EXPECTED -0.5D. ANTERIOR SUBCAPSULAR OPACITIES SEEN BUT VISION IS "TOO LOW." IT IS REPORTED THAT THE LENS WAS EXPLANTED AND CATARACT SURGERY WAS PERFORMED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10337903&pc=QCB | 2023826-2020-01604 | 2020-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/2.00/95 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. IT WAS REPORTED THAT THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED ON THE SAME DAY DUE TO LENS OPACITY (ASC). THE REPORTER STATED THAT THE OPACITY WAS OBSERVED ON THE DAY OF SURGERY, THE PATIENTS CONDITION IS WELL NOW AND UCVA IS 20/32. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429096&pc=QCB | 2023826-2020-01830 | 2019-05-15 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS OPACITY ASC. CATARACT SURGERY WAS PERFORMED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429117&pc=QCB | 2023826-2020-01846 | 2019-05-15 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS OPACITY ASC. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9742235&pc=QCB | 2023826-2020-00373 |  | Injury | STAAR SURGICAL COMPANY | 2020-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: TICL LENS ROTATION WAS REPORTED. EMAIL STATES "WE HAVE A PATIENT THAT HAS HAD HIS ICL ROTATE AND WE'LL NEED TO REPOSITION." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: A1: 231999120 TK. B5- THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.5/2.5/111 (SPHERE/CULINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION IS NOT NECESSARY. " DECREASED VISION RELATED TO CRYSTALLINE LENS CHANGED (PSC) SO NO TX UNTIL PT READY FOR CATARACT SURGERY. PRIMARY OPEN ANGLE OR NARROW GLAUCOMA WAS REPORTED AS A PRE-EXISTING CONDITION. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO OTHER SERIOUS IN INITIAL MDR. B4 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR. G4 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR & SUPPLEMENTAL MDR#1. H6 - PATIENT CODE 2692 CORRECTED TO PATIENT CODE 1766 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689304&pc=QCB | 2023826-2020-00184 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -7.50/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS EXCHANGE TO OCCUR. Manufacturer Narrative: CORRECTED DATA: B4- "16-JAN-2020" SHOULD BE CORRECTED TO "15-JAN-2020" IN THE INITIAL MDR. G4- "16-JAN-2020" SHOULD BE CORRECTED TO "15-JAN-2020" IN THE INITIAL MDR. ADDITIONAL INFORMATION: B5- "ON (B)(6) 2020 THE LENS WAS EXPLANTED WITH A PLAN FOR EXCHANGE. HOWEVER, THE PATIENT DEVELOPED AN INTRAOPERATIVE CATARACT DURING EXPLANTATION AND THE DECISION WAS MADE TO ABORT IMPLANTATION OF REPLACEMENT ICL." H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; LENS EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: PATIENT CODE 1766 SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON (B)(6) 2020 AN ICL IMPLANTATION IN THE RIGHT EYE WITH A LARGER LENS WAS PERFORMED." CLAIM# (B)(4). Manufacturer Narrative: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17844162&pc=QCB | 2023826-2023-04346 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Cataract; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS REPOSITIONING ON (B)(6) 2022 WAS PERFORMED DUE TO ROTATION AND DISPHOTOPSIA ISSUES BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED FOR A LONGER LENGTH SIMILAR POWER LENS (SPHERE/CYLINDER/AXIS) ON (B)(6) 2023. PROBLEM RESOLVED. Manufacturer Narrative: H6- HEALTH EFFECT IMPACT CODE 4581: DISPHOTOPSIA ISSUES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - (B)(6) 1983, A3 - M, B3 - 13-SEP-2022. CORRECTION: B5 - REPORTED AS: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS REPOSITIONING ON (B)(6) 2022 WAS PERFORMED DUE TO ROTATION AND DISPHOTOPSIA ISSUES BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED FOR A LONGER LENGTH SIMILAR POWER LENS (SPHERE/CYLINDER/AXIS) ON (B)(6) 2023. PROBLEM RESOLVED. CORRECT TO: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS REPOSITIONING ON (B)(6) 2022 WAS PERFORMED DUE TO ROTATION AND GLARE/HALOES BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2023 AND IN AS SECOND SURGERY ON (B)(6) 2023 A LONGER LENGTH SIMILAR POWER LENS (SPHERE/CYLINDER/AXIS) WAS IMPLANTED. PROBLEM RESOLVED. REPORTEDLY, MILD ASC AND PATIENT IS VERY HAPPY. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17844244&pc=QCB | 2023826-2023-04347 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Cataract; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXCHANGED WITH LONGER LENGTH SIMILAR POWER LENS(SPHERE/CYLINDER/AXIS) DUE TO DISPHOTOPSIA ISSUES ON (B)(6) 2023. PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - (B)(6)1983, A3 - M, B3 - 13-SEP-2022. CORRECTION B5 - REPORTED AS THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXCHANGED WITH LONGER LENGTH SIMILAR POWER LENS(SPHERE/CYLINDER/AXIS) DUE TO DISPHOTOPSIA ISSUES ON (B)(6) 2023. PROBLEM RESOLVED. CORRECT TO: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXPLANTED AND A NEW LONGER LENGTH SIMILAR POWER LENS(SPHERE/CYLINDER/AXIS) WAS LATER IMPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND GLARE/HALOES ON (B)(6) 2023. PROBLEM RESOLVED. REPORTEDLY MILD ASC, PATIENT IS VERY HAPPY. CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT IMPACT CODE 4581: DISPHOTOPSIA ISSUES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941309&pc=QCB | 2023826-2019-01571 | 2019-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -13.0/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 CATARACT SURGERY WAS PERFORMED DUE TO LENS OPACITY-CORTICAL, GLARE/HALOS, POOR VISION, AND CORTICAL CHANGES. THE SURGEON STATES, "VISION POOR WITH GLARE/HALOS, SOME DOUBLE." Manufacturer Narrative: "IT WAS LATER REPORTED AT A FOLLOW UP DATE THAT THE PATIENT STATUS IS GOOD AND STABLE WITH NO VISUAL COMPLAINTS." SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10171666&pc=QCB | 2023826-2020-01273 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+4.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2020 THE SURGEON ASSESSED LENS OPACITY ASC. ALSO REPORTED IS LOW VAULT AND LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H6- PATIENT CODE 3191- SECONDARY SURGERY: LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, EXCHANGE FOR A LONGER LENGTH LENS DID NOT RESOLVE THE PROBLEM - LOW VAULT REMAINS AND THE SURGEON WILL CONTINUE TO MONITOR MILD ASC CATARACT. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10171670&pc=QCB | 2023826-2020-01272 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+4.5/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020 THE SURGEON ASSESSED LENS OPACITY ASC. ALSO REPORTED IS LOW VAULT AND LENS ROTATION. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON 14-JUL-2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H6- PATIENT CODE 3191- SECONDARY SURGERY: LENS EXCHANGE CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, EXCHANGE FOR A LONGER LENGTH LENS DID NOT RESOLVE THE PROBLEM - LOW VAULT REMAINS AND THE SURGEON WILL CONTINUE TO MONITOR MILD ASC CATARACT. CORRECTIONS: B5 - "BLURRED VISION WAS ALSO REPORTED." SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT CLINICAL CODE 2140 SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13698235&pc=QCB | 2023826-2022-00761 | 2022-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Conjunctivitis; Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.00/+2.50/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT INTRAOCULAR PRESSURE INCREASED AND COMPLAINED OF BLURRY VISION AND PAIN ON THE DAY OF SURGERY; MEDICATION WAS PRESCRIBED AND THE SURGEON PERFORMED IRRIGATION BEFORE SENDING HER HOME. ONE DAY POST OP THE PATIENT HAD HIGH IOP AND CORNEAL EDEMA CENTRALLY WITH CONJUNCTIVAL REDNESS. PUPILLARY BLOCK 3+ CELLS/FLARES NOTED. ON THIS DAY THE SURGEON REPORTED NOTICING A THIN MEMBRANE OVER THE CENTRAL HOLE OF THE ICL; WHICH THEY PERFORMED AN ADDITIONAL IRRIGATION WITH ADDITIONAL MEDICATION PRESCRIBED. IT WAS INDICATED AFTER ANOTHER POSTOP VISIT SURGEON PERFORMED AN ADDITIONAL IRRIGATION AND THE MEMBRANE WASHED OUT DURING THIS AND THE PATIENT'S IOP BECAME STABLE. IT WAS FURTHER INDICATED 1 WEEK POSTOP THE IOP IS STABLE; NO CORNEAL EDEMA BUT PATIENT HAS PUPIL MID-DILATED DUE TO THE PUPILLARY BLOCK AND 1.2 CT VAULTING. THE PATIENT WILL CONTINUE BEING MONITORED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11937992&pc=QCB | 2023826-2021-01756 | 2021-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Corneal Clouding/Hazing | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+3.5/068 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS REMAINED IMPLANTED. THE CORNEA WAS A BIT HAZY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED ON (B)(6) 2021 AND REPLACED WITH A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE LENS HAD ROTATED AGAIN AND WAS REMOVED AND REPLACED WITH A MFIOL. SEE MFR. REPORT#: 2023826-2021-03263 FOR REPLACEMENT LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200502&pc=QCB | 2023826-2021-02790 | 2021-06-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Clouding/Hazing; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -09.00/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. POST-OP, THE VISION WAS HM AND THERE WAS SIGNIFICANT CORNEAL CLOUDING. LENS LOOKED APPROPRIATE AND PHAKIC LENS WAS CLEAR. THE LENS REMAINED IMPLANTED . THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02534. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM LENS IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY POST-OP THE VISION WAS HM AND THERE WAS SIGNIFICANT CORNEAL CLOUDING. LENS LOOKED APPROPRIATE AND PHAKIC LENS WAS CLEAR. THE LENS REMAINS IMPLANTED. THE SERIAL NUMBER FOR THIS LENS WAS NOT PROVIDED. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652056&pc=QCB | 2023826-2019-00934 | 2019-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Unintended Movement | Corneal Decompensation | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/138 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. REPORTEDLY, THE LENS WAS UPSIDE DOWN AFTER WRONG MANIPULATION. THIS OCCURRED ON (B)(6) 2019. THE LENS WAS REMOVED INTRA-OPERATIVELY. ON (B)(6) 2019 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. THE PROBLEM IS NOTED AS RESOLVED AND "NO PATIENT INJURY" IS NOTED, HOWEVER SURGEON STATES PATIENT STILL COMPLAINS OF HALOS, ESPECIALLY DURING NIGHT DRIVING. IN ADDITION, CORNEAL DECOMPENSATION IS NOTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. REFERENCE MFR REPORT# 2023826-2019-00935 FOR REPORTING REGARDING HALOS WITH REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A LENS VIAL WITH LIQUID. VISUAL INSPECTION FOUND THAT THE LENS OPTIC IS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15965364&pc=QCB | 2023826-2022-04580 | 2022-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Corneal Decompensation | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.5/1.0/007 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS THE LENS WAS IMPLANTED UPSIDE DOWN; AND THERE WAS CORNEAL DECOMPENSATION. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR; THE DEVICE AND A PATIENT RELATED FACTOR OF THE PATIENT HAD MOVED THE EYE DURING INSERTION; AND THE SURGEON LOST VISUALIZATION OF THE ICL. ICL ALSO FLIPPED DESPITE PROPER LOADING PER STANDARD PROTOCOL. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SAME LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840739&pc=QCB | 2023826-2021-04397 | 2021-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Corneal Edema; Blurred Vision; Discomfort; Eye Pain | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+0.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO CORNEAL DECOMPENSATION. REPORTEDLY, "HOURS AFTER UNEVENTFUL SURGERY, PAIN, DECREASED VISION; ON EXAMINATION: CORNEAL DECOMPENSATION, NO DETAILS VISIBLE." THE PROBLEM IS MARKED AS RESOLVED AFTER EXPLANT HOWEVER THERE WAS STILL SOME CORNEAL EDEMA AT 7H. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES THE PATIENT STARTED HAVING DISCOMFORT A FEW HOURS AFTER SURGERY AND THE LENS HAD TO BE REMOVED TWO DAYS LATER. SYMPTOMS ABATED AFTER EXPLANT. THE SURGEON STATES THE ICL LOOKED FINE WITHOUT VISIBLE DEFECTS AND HE REMAINS UNCERTAIN ABOUT THE EXACT CAUSE OF THE EVENT. CLAIM# (B)(4). Manufacturer Narrative: D4: CORRECTED MODEL#, SERIAL#, AND EXPIRATION DATE. D6A: CORRECTED TO ON (B)(6) 2021. D6B: CORRECTED TO ON (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16431957&pc=QCB | 2023826-2023-00571 | 2021-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Corneal Decompensation; Corneal Edema; Foreign Body Sensation in Eye; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A COPY OF AN ARTICLE WAS FOUND FROM IJO, DATED (B)(6) 2023. THE ARTICLE WAS IN REGARD TO "SPONTANEOUS DISLOCATION OF A TORIC IMPLANTABLE COLLAMER LENS RESULTING IN CORNEAL ENDOTHELIAL DECOMPENSATION: A CASE REPORT AND REVIEW OF THE LITERATURE". THE ARTICLE INVOLVED A 47 YEAR OLD MAN, WHO PRESENTED WITH COMPLAINTS OF FOGGY VISION, WITH FOREIGN BODY SENSATION AND PHOTOPHOBIA IN HIS RIGHT EYE (OD). THE PATIENT RECEIVED BILATERAL ICL IMPLANTATION FOR CORRECTION OF HIGH AMETROPIA IN 2013. A TICMV4 MODEL LENS, WITH POWER OF -23.00/+2.5/095 WAS IMPLANTED IN THE RIGHT EYE. POST-OP, THE ICL IN RIGHT EYE SLIGHTLY ROTATED AND PROLAPSED. THIS LENS WAS REMOVED AND ANOTHER TICL LENS WITH POWER OF -23.00/+2.5/100 WAS IMPLANTED BUT THIS LENS ALSO PROLAPSED WITH SLIGHT ROTATION. THE PATIENT REGAINED A SATISFACTORY DISTANCE UNCORRECTED VISUAL ACUITY (UCVA) OF 20/20. THE PATIENT DID NOT WANT THE LENS REMOVED. A CAPSULE TENSION RING WAS IMPLANTED. DESCEMET'S STRIPPING ENDOTHELIAL KERATOPLASTY WAS PERFORMED FOR RESTORING VISUAL ACUITY. THERE WAS CORNEAL ENDOTHELIAL DECOMPENSATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AN ARTICLE WAS PUBLISHED CITING "SPONTANEOUS DISLOCATION OF A TORIC IMPLANTABLE COLLAMER LENS RESULTING IN CORNEAL ENDOTHELIAL DECOMPENSATION: A CASE REPORT AND REVIEW OF THE LITERATURE". THE ARTICLE REPORTED THE SURGEON IMPLANTED A TICMV4 IMPLANTABLE COLLAMER LENS OF DIOPTER -23.00/+2.5/95 (SPHERE/CYLINDER/ AXIS) INTO THE PATIENTS RIGHT EYE (OD) IN 2013 TO CORRECT HIGH AMETROPIA. THE PATIENT EXPERIENCED: CORNEAL OEDEMA, LENS ROTATION, ICL PROLAPSE, FOGGY VISION, FOREIGN BODY SENSATION, PHOTOPHOBIA, REDUCED ENDOTHELIAL CELL COUNT, UNSTABLE ZONULES, PHACOEMULSIFICATION AND INTRAOCULAR LENS IMPLANTATION, CAPSULAR TENSION RING INSERTION, DESCEMET'S STRIPPING ENDOTHELIAL KERATOPLASTY, RESULTING IN CORNEAL DECOMPENSATION. AS PER ARTICLE 'POSTOPERATIVELY, THE TICL IN THE RIGHT EYE SLIGHTLY ROTATED AND PROLAPSED. WITH CONSENT, ANOTHER TICL WITH A DIAMETER OF 12.0 MM AND A POWER OF 23.00/+2.5×100 WAS LATER IMPLANTED IN THE RIGHT EYE AFTER THE FORMER ONE WAS EXPLANTED. HOWEVER, THE TICL AGAIN PROLAPSED WITH SLIGHT ROTATION POSTOPERATIVELY. SINCE THE OPTICAL ZONE OF THE TICL STILL REMAINED IN THE CENTRAL FRONT OF THE PUPIL, THE PATIENT REGAINED A SATISFACTORY DISTANCE UNCORRECTED VISUAL ACUITY (UCVA) OF 20/20. CLINICAL DECISION FOR EXPLANTATION OF THE TICL IN THE RIGHT EYE WAS MADE. HOWEVER, THE MAN REFUSED TO TAKE FURTHER SURGERY AND WAS LOST TO FOLLOW UP IN THE FOLLOWING 8Y'. REPORTEDLY 'ON PRESENTATION, UCVA WAS HAND MOTION FOR THE RIGHT EYE AND 20/200 (REFRACTION, 4.00) FOR THE LEFT EYE. IOP WAS 20 MM HG FOR THE RIGHT EYE AND 18 MM HG FOR THE LEFT EYE. SLIT-LAMP EXAMINATION AFTER PUPILLARY DILATATION SHOWED ROTATION OF APPROXIMATELY 45 AND PROLAPSE OF THE TICL IN THE RIGHT EYE WITH CORNEAL EDEMA. THE CORNEAL ENDOTHELIAL CELL DENSITY WAS 2281 CELLS/MM2 FOR THE LEFT EYE, WHILE UNOBTAINABLE FOR THE RIGHT EYE. EXAMINATION OF ULTRASOUND BIOMICROSCOPY (UBM) SHOWED THAT THE TICL WAS STILL BEHIND THE IRIS, BUT THE EDGES OF THE TICL ARCHED FORWARD INTO THE ANTERIOR CHAMBER. THE SYMPTOMS 'FINALLY RESULTED IN CORNEAL ENDOTHELIAL DECOMPENSATION'. ADDITIONALLY 'PHACOEMULSIFICATION AND INTRAOCULAR LENS IMPLANTATION AFTER EXPLANTATION OF THE TICL WAS PERFORMED ON HIS RIGHT EYE, AND A CAPSULAR TENSION RING WAS ALSO IMPLANTED DUE TO DETECTED UNSTABLE ZONULES. THREE MONTHS LATER, DESCEMET'S STRIPPING ENDOTHELIAL KERATOPLASTY WAS PERFORMED FOR RESTORING VISUAL ACUITY'. H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE>45). H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL USE (ACD<3.0MM. H6- HEALTH EFFECT- CLINICAL CODE: 4581- REDUCED ENDOTHELIAL CELL COUNT, CAPSULAR TENSION RING INSERTION, DESCEMET'S STRIPPING ENDOTHELIAL KERATOPLASTY, UNSTABLE ZONULES. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPIRATION DATE: EXPIRATION DATE UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - FOGGY, PROLAPSE, PHOTOPHOBIA. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634551&pc=QCB | 2023826-2020-02295 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+5.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION (CORNEAL EDEMA) AND UNREACTIVE (FIXED) PUPIL. THE IOP WAS CONTROLLED WITH AGGRESSIVE TREATMENT. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO AN OVERSIZED LENS. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON 12-OCT-2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS ASYMPTOMATIC, THE ANTERIOR SEGMENT WAS OK, THERE WAS SOME PIGMENT DISPERSION AND THE PUPIL WAS MYDRIATIC. THE PATIENT WAS HAPPY WITH THE RESULTS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558781&pc=QCB | 2023826-2021-00704 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -15.0/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS IOP CONTROLLED WITH ALPHAGAM, HIGH VAULT, SHALLOW AC, "CRAPPY" VISION AT 20/200 AND PUPIL IS DILATED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: SIGNIFICANT REDUCTION OF ICA, CORNEAL DECOMPENSATION, CORNEAL EDEMA ALSO REPORTED. THE CAUSE IS REPORTED AS RETAINED VISCOELASTIC BEHIND THE ICL AND "ICL STILL A BIT OVERSIZED." REPORTEDLY THE PATIENT IS RECOVERED, VISION IS GOOD. H6-CLINICAL CODE 4581 (SIGNIFICANT REDUCTION OF ICA). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558806&pc=QCB | 2023826-2021-00699 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS -15.5/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTEDLY EXPLANTED THE LENS AND REPLACED IT WITH A SHORTER LENS BECAUSE THE LENS WAS "TOO BIG." Manufacturer Narrative: B5: ADDITIONAL INFORMATION: EXCESSIVE VAULT-TOUCHING THE CORNEA, ELEVATED IOP, 4+ CORNEAL EDEMA, CORNEAL DECOMPENSATION, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION REPORTED. CAUSE IS REPORTED AS DEVICE: ICL TOO LARGER PER CALCULATOR. PROBLEM MARKED AS RESOLVED. HOWEVER, SEE MFR REP# 2023826-2021-00704 FOR REPLACEMENT LENS. ORIGINAL LENS IMPLANTED, REMOVED, AND REPLACED WITHIN THE SAME SURGERY. CLAIM #(B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUIED IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11143969&pc=QCB | 2023826-2021-00113 | 2020-10-14 04:00 | Injury | STAAR SUGICAL COMPANY | 2021-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Toxic Anterior Segment Syndrome (TASS); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: ON 11 DEC 2020 WE RECEIVED NOTIFICATION OF AN ARTICLE ENTITLED 'OCCLUSION OF AQUAPORT FLOW IN A CASE OF TOXIC ANTERIOR SEGMENT SYNDROME FOLLOWING IMPLANTED COLLAMER LENS SURGERY CAUSING SEVERE PUPILLARY BLOCK.' THE ARTICLE REPORTS " A V4C TORIC ICL (STAAR SURGICAL) (VTICMO) WAS IMPLANTED IN THE LEFT EYE OF A (B)(6) YEAR-OLD WOMAN. AFTER 1 WEEK, THE LENS WAS 15° OFF AXIS AND UNEVENTFUL LENS ALIGNMENT CORRECTION (LENS REPOSITIONING) WAS PERFORMED. AT POSTOPERATIVE DAY 1, AN INTRAOCULAR PRESSURE (IOP) OF 11 MM HG AND MILD CORNEAL EDEMA WERE OBSERVED. AT POSTOPERATIVE DAY 7, THERE WAS AN UNCORRECTED DISTANCE VISUAL ACUITY OF 20/800, AN IOP OF 44 MM HG, DIFFUSE CORNEAL EDEMA, AND FIBRIN STRANDS IN THE ANTERIOR CHAMBER. TASS WAS DIAGNOSED AND TOPICAL STEROIDS, CYCLOPLEGIA, ANTIGLAUCOMA DROPS, AND ORAL ACETAZOLAMIDE WERE PRESCRIBED. AT POSTOPERATIVE DAY 9, DISPERSED ENDOTHELIAL PIGMENT WITH A FIXED MID-DILATED PUPIL WERE IDENTIFIED. ANTERIOR SEGMENT OPTICAL COHERENCE TOMOGRAPHY SHOWED STRANDS OF FIBRIN BLOCKING THE CENTRAL ICL HOLE AND ANGLE CLOSURE. PUPILLARY BLOCK RELATED TO THE FIBRIN OCCLUDING THE AQUAPORT HOLE CONSECUTIVE TO TASS WAS DIAGNOSED. THE PATIENT UNDERWENT ICL REMOVAL, BUT HAD A PERSISTENT ATONIC, HYPOREFLEXIVE PUPIL AS A COMPLICATION." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM # : (B)(4). Manufacturer Narrative: D4: CORRECTED MODEL # FROM VTICMO TO VTICM5. ADDITIONAL INFORMATION: B5: PATIENT IS STABLE. IOL EXPLANTED. WAITING FOR CATARACT SURGERY AND IRIS RECONSTRUCTION DUE TO URRETS- ZAVALIA SYNDROME. PROBLEM RESOLVED BUT WITH COPLICATIONS (CORNEAL DECOMPENSATION, GONIO SYNECHIAE 360- OCULAR HYPERTENSION). CAUSE OF EVENT TASS. METHYLCELLULOSE USED IN THE EYE INTRA-OPERATIVELY. LAST KNOWN BCVA: FINGER COUNTING AT 1 METER. CAUSE OF TASS IS UNKNOWN. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION, FIXED DILATED PUPIL. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18337518&pc=QCB | 2023826-2023-05553 | 2023-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Corneal Edema; Toxic Anterior Segment Syndrome (TASS) | UNK | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +04.00/+4.0/085 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED. DIFFUSE CORNEAL EDEMA DIAGNOSED WITHIN 1 -2 DAYS. TREATMENT OF TOPICAL CORTICOSTEROIDS AND ORAL ANTIBIOTICS WERE PRESCRIBED. THERE WAS A DIAGNOSIS OF CORNEAL DECOMPENSATION. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: G4: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. H6: HEALTH EFFECT IMPACT CODE: 4644 - TOPICAL CORTICOSTEROIDS, ORAL ANTIBIOTICS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12741585&pc=QCB | 2023826-2021-04219 | 2021-09-10 | Injury | STAAR SURGICAL COMPANY | 2021-11-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -16.0/2.5/99 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED AND AT A LATER DATE REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND CORNEAL DECOMPENSATION. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699716&pc=QCB | 2023826-2019-00980 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Corneal Decompensation; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/5.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PUPIL BLOCK WITH ELEVATED IOP (50MMHG), CORNEAL DECOMPENSATION, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, AND IRIS ATROPHY WAS OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, INTRAOCULAR INJECTION AND THE LENS WAS REMOVED ON (B)(6) 2019. PER THE REPORTER, ON (B)(6) 2019 PATIENT'S EYE STATUS, "IRIS ATROPHYIC AND PUPIL DILATED. REPORTER NOTED CORNEA BECAME CLEAR AFTER 1 MONTH. EYE IS CALM. CAUSE OF THE EVENT IS REPORTED AS "URETS-ZAVALIA SYNDROME DUE TO OCCLUSION OF THE AQUAPORT FIRST WITH VISCOELASTIC THEN WITH IRIS PIGMENTS." Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE 1937: ELEVATED IOP. DEVICE CODE: 1401: MISFOCUSING; PIGMENT DISPERSION, FIXED PUPIL, IRIS ATROPHY. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 1401 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO "DEVICE CODE 3191: PIGMENT DISPERSION, FIXED PUPIL, IRIS ATROPHY" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984533&pc=QCB | 2023826-2019-01677 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.5/+3.0/094 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE, CORNEAL DECOMPENSATION AND UNREACTIVE (FIXED) PUPIL. THERE WAS AN ADDITION OF NEW PERIPHERAL IRIDECTOMY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE PATIENTS POST-OP CONDITION WAS NORMAL CORNEAL THICKNESS, NO EDEMA, IOP 10MMHG, CELL COUNT (B)(6). Manufacturer Narrative: METHOD CODE: 4111 SHOULD HAVE BEEN INCLUDED IN INITAL MEDWATCH REPORT. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED. DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9132032&pc=QCB | 2023826-2019-01799 | 2019-05-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Corneal Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -5.5/3/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 AS A REPLACEMENT LENS TO CORRECT EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. IT WAS REPORTED THAT THE LENS RESOLVED THE PROBLEM BUT CORNEAL EDEMA WAS OBSERVED, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: ADVERSE EVENT OR PRODUCT PROBLEM- PRODUCT PROBLEM. DATE OF EVENT- (B)(6) 2019. EXPLANT DATE- (B)(6) 2019. TYPE OF REPORTABLE EVENT- PREVIOUS MDR SERIOUS IS CORRECTED TO MALFUNCTION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10590854&pc=QCB | 2023826-2020-02254 | 2020-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1/0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, THE REPORTER STATED " LARGE EXUDATION IS SEEN IN THE ANTERIOR CHAMBER; BACTERIAL CULTURE IN INTRAOCULAR FLUID: STERILE GROWTH; EYE B ULTRASOUND SUGGESTED NO OBVIOUS ABNORMALITIES IN THE VITREOUS CAVITY. ON (B)(6) YU UNDERWENT ANTERIOR CHAMBER IRRIGATION OF THE LEFT EYE +ICL INTRAOCULAR LENS REMOVAL, AND RECEIVED ANTI-INFLAMMATORY AND ANTI-INFECTIVE TREATMENT AFTER SURGERY. THE PATIENT DID NOT COMPLAIN OF ANY SPECIAL DISCOMFORT. PHYSICAL EXAMINATION: VOS:0.02; CORRECTION: OS: -9.25/-2.00×175 0.5; NCT: OS: 16.7MMHG; LEFT EYE CONJUNCTIVAL CONGESTION (+), INCISION AIRTIGHT, MILD EDEMA OF THE CORNEA, THE FORMER ROOM, AQUEOUS HUMOR, ROOM (-), THE IRIS TEXTURE CLEAR, PUPILS ROUND SHAPE, BULK DRUGS SEX, 6 MM IN DIAM SIZE, LIGHT REFLEX SLOW, TRANSPARENT, LENS CAPSULE MEMBRANE BEFORE A LARGE NUMBER OF PIGMENTATION, FUNDUS VISIBLE DEPENDING ON THE NIPPLE BOUNDARY CLEAR, VISIBLE CONTRACTION ARC, RETINAL FLAT ABOVE, BELOW THE RETINA FUZZY VISIBLE FLAT. THE PATIENT'S VISUAL ACUITY RETURNED TO THAT BEFORE OPERATION. AFTER FINDING THE SYMPTOMS, THE PATIENT WAS GIVEN ANTERIOR CHAMBER IRRIGATION AND INTRAVITREAL INJECTION OF VANCOMYCIN, SYSTEMIC MEDICATION: AMPICILLIN, SULBACTAM SODIUM, TOPICAL ATROPINE, PREDNISOLONE ACETATE EYE DROPS, DIANBISHU, PRANOPROFEN, LEVOFLOXACIN." IT WAS REPORTED THAT THE PROBLEM HAS RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060189&pc=QCB | 2023826-2020-03154 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Corneal Edema | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. RIGHT INFERIOR CORNEAL EDEMA IS REPORTED. Manufacturer Narrative: B5-ADDITIONAL INFO: PER THE REPORTER THE PROBLEM IS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1. ADVERSE EVENT CORRECTED TO PRODUCT PROBLEM. B2. OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS IS NOT APPLICABLE. H1. SERIOUS INJURY CORRECTED TO MALFUNCTION. H6. MEDICAL DEVICE PROBLEM CODE 2993 IS CORRECTED TO 3189. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563283&pc=QCB | 2023826-2021-00754 | 2021-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, -10.00/2.0/050 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, "STILL LETTLE EDEMA OF THE CORNEA, IMPROVING.". PER THE REPORTER ON (B)(6)2021, "ACTUALLY THE PATIENT CAN SEE 20/20 WITHOUT CORRECTION AND VAULT IS CORRECT. HE PREVIOUSLY HAD A LITTLE BIT OF A HIPEROPIA AND A VERY UNSATISFACTORY CLOSE VISION. MUCH BETTER NOW." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM VTICH 13.2; +10.00/2.0/050 (SPHERE/CYLINDER/AXIS); IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY; "STILL LITTLE EDEMA OF CORNEA; IMPROVING." CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. DATE RECEIVED BY MANUFACTURER - THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725276&pc=QCB | 2023826-2021-04182 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.50/5.0/122, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT EXCESSIVE VAULT. FOR STATUS OF THE EYE (SIGNS / SYMPTOMS) THE REPORTER STATED "STILL HAS EDEMA AND STITCH DUE TO WOUND LEAKAGE INTRA-OP. AC IS NOT SHALLOW ANYMORE. BUT IT WILL RECOVER SOON." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. D4 - CORRECTED TO "SERIAL #: (B)(6), EXPIRATION DATE: 30 - SEP - 2022" IN INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION IT WAS CLARIFIED THAT THIS CASE OF EDEMA WAS NO PRODUCT RELATED, BUT SURGERY RELATED, AND IT RESOLVED. THE EDEMA DID NOT REQUIRE TREATMENT BEYOND WHAT IS STANDARD AFTER OCULAR SURGERY. THERE WAS NO PATIENT INJURY. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840505&pc=QCB | 2023826-2021-04378 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | Corneal Edema | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/080 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS MILD CORNEAL EDEMA. Manufacturer Narrative: B5: MILD CORNEAL EDEMA REPORTED AS AN EXPECTED OCCURENCE. THERE WAS NO BEYOND STANDARD CARE TO TREAT THE EDEMA. THE EDEMA HAS RESOLVED. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256532&pc=QCB | 2023826-2022-00159 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.00/4.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021; AS A REPLACEMENT LENS TO A PREVIOUS LENS IMPLANTED UPSIDE DOWN ISSUE. IT WAS NOTED THE PATIENT NOW EXPERIENCES SEVERE CORNEAL EDEMA. THE PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI) # (B)(4) SHOULD BE REMOVED FROM INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B5- " REPORTEDLY, 'OS HAD SEVERE CORNEAL ODEMA BUT HAS SETTLED WITH SATISFACTORY OUTCOME IN BOTH EYE.' THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16065223&pc=QCB | 2023826-2022-04761 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM_12.6; -9.5/1.0/007 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2022. THE SURGEON REPORTS THE PATIENT HAD MASSIVE POSTOPERATIVE CORNEAL EDEMA (4TH CENTRAL EDEMA) AND ADDITIONALLY NOTED; "CORNEAL EDEMA OUT OF PROPORTION TO AMOUNT ANTICIPATED WILL CONTINUE TO CLOSELY MONITOR FOR RESOLUTION. AT A LATER DATE IT WAS NOTED THAT THE PATIENT WAS DOING WELL AND AT WEEK 1.5 SHE WAS 20/30 UNCORRECTED. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545693&pc=QCB | 2023826-2023-00851 | 2022-12-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/123 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON 13-DEC-2022. POST-OP THERE WAS MINOR CORNEAL EDEMA AT NASAL SIDE. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00850. POST-OP THE PATIENT HAS STABILIZED AND FEELS BETTER. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253007&pc=QCB | 2023826-2023-02655 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE REPORTER INDICATED THERE WAS TRACE CORNEAL EDEMA POST-OP. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK . D6B: NA. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: AT THE LAST VISIT ON (B)(6) 2023, THE SLE WAS ALL NORMAL. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17261690&pc=QCB | 2023826-2023-02573 | 2023-06-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/090 (SPHERE/CYLINDER/AXIS), ON (B)(6)2023. POST-OP, CORNEAL EDEMA WAS RESOLVING. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-02572. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE CORNEAL EDEMA HAS BEEN RESOLVED AND THE PATIENT IS DOING WELL. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18007851&pc=QCB | 2023826-2023-04806 | 2023-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH_12.6; +7.50/+4.50/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS IMPLANTED UPSIDE DOWN. INTRAOPERATIVELY THE LENS WAS REMOVED AND PATIENT INJURY WAS NOTED - EDEMA. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS LENS BREAKAGE DURING EXTRACTION AND UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT - IMPACT CODE: "2199" SHOULD BE REMOVED AND "4614" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18419305&pc=QCB | 2023826-2023-05832 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -8.50/1.0/147 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT WITH ROTATION. FOR STATUS OF THE EYE THE REPORTER STATED THAT THE EYE WAS FIRST DAY POSTOPERATIVE WITH CORNEAL EDEMA AND MIOSIS FROM SURGERY. IT WAS REPORTED THAT AT THE NEXT APPOINTMENT ON (B)(6) 2023 PUPIL WAS STILL IN MIOSIS FROM MYOSTST. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: MIOSIS. H6 TYPE OF INVESTIGATION LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11448803&pc=QCB | 2023826-2021-00609 | 2021-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/160 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), IT TORE. IT WAS REMOVED INTRAOPERATIVELY AND THERE WAS NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2021. THE STATUS OF THE EYE IS REPORTED AS AC CELL GR. 1, 0 FLARE, CORNEAL EDEMA. NO INFECTION OR NON-STANDARD POSTOP CARE WAS PERFORMED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER-POSSIBLE OVD IN INJECTOR. Manufacturer Narrative: B5-PREVIOUSLY STATED A -13.0/+2.0/160 (SPHERE/CYLINDER/AXIS). CORRECTED VALUE IS -11.5/+2.0/146. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A PLASTIC CONTAINER WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11845967&pc=QCB | 2023826-2021-01547 | 2021-03-08 | Injury | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN AND THE LENS TORE. THERE WAS A LOT OF EDEMA AFTER THE SURGERY. THERE WAS SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR CELL LOSS. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR - PROBABLY THE LENS WAS NOT LOADED PROPERLY (A BIT "TWISTED"). THE ICL UNFOLDED WRONG IN THE EYE. PROBABLY THE ENDOTHELIUM WAS DAMAGED WHILE REMOVING THE ICL. THE SURGEON DECIDED NOT TO IMPLANT ANOTHER ICL LENS, THE PATIENT WILL CONTINUE WITH A CL IN THIS EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14006701&pc=QCB | 2023826-2022-01052 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -12.5/1.0/146 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND CORNEAL DECOMPENSATION. THE EXCHANGE RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "LOOKS GOOD, PATIENT HAPPY. RESOLUTION OF INFERIOR CORNEAL EDEMA WHICH HAD DEVELOPED FROM INTERMITTENT IK TOUCH." THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16413617&pc=QCB | 2023826-2023-00685 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10/1.5/108 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023; AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS; AND CORNEAL ODEMA DUE TO THE PREVIOUS TRICKY EXCHANGE (CONSTRICTION OF PUPIL DURING PROCEDURE). IT WAS ADDITIONALLY NOTED THAT "VISION IS EXCELLENT THEREFORE SURGEON DOESN'T WANT TO REMOVE LENS SO IS GOING TO LEAVE IN SITU AND MONITOR PATIENT". PROBLEM IS REPORTED AS NOT RESOLVED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS. TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17054420&pc=QCB | 2023826-2023-02174 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -10.0/+2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE SURGEON REPORTS THE PATIENT HAS CORNEAL EDEMA, AND THE POST OP REFRACTION WAS +0.75/-1.75/165. SURGEON ADVISED THE NEXT REVIEW AFTER 6 WEEKS. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER 10.0/+2.0/90 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED CORNEAL OEDEMA AND RESIDUAL REFRACTION. REPORTEDLY "PATIENT IS HAVING RESIDUAL OF +1.75/-3@175. CHECKED POST-OP ON 25TH MAY WITH CORNEAL EDEMA, AND THE POST-OP REFRACTION WAS +0.75/-1.75/165. SURGEON ADVISED THE NEXT REVIEW AFTER 6 WEEKS". THE LENS REMAINS IMPLANTED. H6-HEALTH EFFECT - CLINICAL CODE: "4581" RESIDUAL REFRACTION. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17222444&pc=QCB | 2023826-2023-02560 | 2023-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -11.50/3.5/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, DISCOMFORT AND ANGLE CLOSURE (IN/OUT), IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO THE DEVICE "LENS TOO LARGE". A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT EDEMA WAS NOTICED AFTER IMPLANTATION OF THE REPLACEMENT LENS, HOWEVER THE EDEMA RESOLVED AND IT DID NOT REQUIRE TREATMENT BEYOND WHAT IS STANDARD AFTER OCULAR SURGERY. THE REPORTER STATED THAT THE CAUSE OF THE EDEMA WAS DUE TO PATIENTS REACTION TO SURGERY. Manufacturer Narrative: H1 - MALFUNCTION CORRECTED TO SERIOUS IN INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE; DISCOMFORT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685454&pc=QCB | 2023826-2021-04058 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Corneal Edema; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER STATES A TORIC IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS A 1-DAY POSTOP MRX SPHERICAL EQUIVALENT OF +2.25D; POSSIBLE SUBTLE CORNEAL EDEMA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5 ADDITIONAL INFORMATION: PER THE REPORTER, "THIS SPONTANEOUSLY RESOLVED AS HYPEROPIA WAS FROM TEMPORARY CORNEA EDEMA. NO ISSUES NOW." (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER, AGE, SEX, WEIGHT, ETHNICITY, RACE- UNK. DATE OF EVENT- UNK. IMPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13716395&pc=QCB | 2023826-2022-00768 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Corneal Edema; Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.50/4.0/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS IMPLANTED UPSIDE DOWN AND WOULDN'T OPEN; WHICH THE SURGEON MANAGED TO EVENTUALLY OPEN IT AND LEFT IT IN PLACE. THERE WAS THEN CORNEAL EDEMA AND REFRACTIVE SURPRISE; WHICH THE SURGEON NOTED IT WAS UPSIDE DOWN. THE LENS WAS THEN EXPLANTED ON THE SAME DATE. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. POSTOPERATIVE REPORT INDICATED "EXCHANGED AS SUCCESSFUL. SURGEON NOW UNDERSTANDS HOW TO ADAPT TECHNIQUE TO ENSURE LENS UNFOLDS CORRECTLY" THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17273272&pc=QCB | 2023826-2023-02723 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.00/3.5/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) AS A REPLACEMENT LENS, ON 16-JUN-2023. IT WAS NOTED THE PATIENT HAS GRADE 1 EDEMA AND BLURRY VISUAL ACUITY BUT IMPROVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17275327&pc=QCB | 2023826-2023-02810 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.00/4.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS, ON (B)(6) 2023. IT WAS NOTED THE PATIENT HAS GRADE 1 EDEMA AND BLURRY VISUAL ACUITY BUT IMPROVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18295971&pc=QCB | 2023826-2023-05524 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/2.0/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS A REFRACTIVE ERROR DUE TO NEGATIVE SIGN NOT ENTERED INTO CYLINDER PART OF REFRACTION IN OCOS. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16689487&pc=QCB | 2023826-2023-01207 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Edema; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -3.5/5.0/098 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, BLURRED VISION AND CORNEAL EDEMA. CAUSE OF EVENT WAS UNKNOWN." IRIS ATROPHY IN HIGH SECTOR" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). PER UPDATED INFORMATION THE PATIENT IS IN PSEUDOPHAKIA. MEDICAL INTERVENTION ( INTENSIVE FREQUENCY OF CORTICOSTEROIDES DROPS AND ANTIEDEMA DROPS) WAS PRESCRIBED FOR CORNEAL EDEMA. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BROKEN HAPTIC. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS HAPTIC DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: ADVERSE EVENT PROBLEM - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). ADVERSE EVENT PROBLEM - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16746796&pc=QCB | 2023826-2023-01372 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Corneal Edema; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.0/4.0/038 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION; REFRACTIVE CHANGE OVER TIME AFTER THE LENS HAD ROTATED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND AN INCISION ENLARGEMENT WAS ALSO REPORTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+4.0/38 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULT WITH ROTATION, REFRACTIVE CHANGE OVER TIME, INCISION ENLARGEMENT AND A MILD CORNEAL OEDEMA. REPORTEDLY "AFTER CONSULTATION WITH THE SURGEON -MILD CORNEAL EDEMA WAS EXPECTED AT IT WAS AN INITIAL PATIENT SPECIFIC OF CORNEA TENDENCY OF CORNEA". ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIMILAR POWER; LONGER LENGTH AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE WAS REPORTED AS "MILD CORNEAL EDEMA - TEMPORAL PERIPHERAL PART. OS REFRACTION +0.75/-1.75 AX 5". THE CAUSE OF THE EVENT AS UNKNOWN. H6- HEALTH EFFECT- CLINICAL CODE: 4581- " INCISION ENLARGEMENT." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12242256&pc=QCB | 2023826-2021-02943 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Glaucoma; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -15.5/2.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTED ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK; (IN (B)(6)) HYPHEMA ON THE ANTERIOR AND POSTERIOR SURFACE OF THE ICL AND WAS PRESCRIBED MEDICATION WHICH THE EYE RECOVERED TO 1.5 CORRECTED VISUAL ACUITY AND 1.2 NAKED EYE VISION. AT A FOLLOW UP DATE HYPHEMA AND MILD CORNEAL EDEMA WAS NOTED WITH ADDITIONAL MEDICATION; AS WELL AS CONDITION OF SECONDARY GLAUCOMA ASSOCIATED WITH HYPHEMA CONTINUED. THE PATIENT IS BEING OBSERVED. THE PROBLEM WAS REPORTED AS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17935878&pc=QCB | 2023826-2023-04537 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Corneal Edema; Headache; Blurred Vision; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -12.5/+1.00/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED LENS ROTATION, THE SURGEON ALSO NOTED HEADACHES, CORNEAL EDEMA, BLURRY VISION AND GLARES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR - PREOP AXIS WAS NOT ACCURATE BASED ON THE TOPOGRAPHY AND AUTO-REFRACT AXIS OF ABOUT 49 DEGREES AND IT WAS A PRE-OP MEASUREMENT ERROR NOT AN ISSUE WITH THE LENS ITSELF. Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THAT THE PATIENT HAD EYE PAIN AND MEDICATION WAS USED, TIMOLOL AND DIAMOX. CLAIM#:(B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17895044&pc=QCB | 2023826-2023-04425 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Inflammation | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+4.0/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT HAD SIGNIFICANT CORNEAL EDEMA AND MARKED INFLAMMATION. PATIENT WAS PRESCRIBED HIGH DOSE OF STEROIDS AND FINISHING TOPICAL ABX. THE SUTURE WAS REMOVED ON (B)(6) 2023 WITHOUT COMPLICATION. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-04389. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6B: UNK . H6: HEALTH EFFECT IMPACT CODE: 4644 - STEROIDS, TOPICAL ABX H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16390681&pc=QCB | 2023826-2023-00595 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Corneal Edema; Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/021 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. CAUSE WAS "INFLAMMATORY REACTION DUE TO IRIS CHAFING AND PIGMENT DISPERSION." STATUS OF THE EYE IS PSEUDOPHAKIA. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- ANGLE CLOSURE; PIGMENT DISPERSION; IRIS CHAFING. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/21 INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED INFLAMMATION; IRIS CHAFING; PIGMENT DISPERSION; ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. REPORTEDLY "INFLAMMATORY REACTION DUE TO IRIS CHAFING AND PIGMENT DISPERSION". ON (B)(6) 2022 THE LENS WAS EXPLANTED. PHACOEMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PSEUDOPHAKIA". CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL DATA: H6-HEALTH EFFECT- CLINICAL CODE: "1791-MILD CORNEAL EDEMA" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/21 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 INFLAMMATION, IRIS CHAFING, PIGMENT DISPERSION, ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED. REPORTEDLY "INFLAMMATORY REACTION DUE TO IRIS CHAFING AND PIGMENT DISPERSION". ON (B)(6) 2022 THE LENS WAS EXPLANTED. PHACOEMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "MILD CORNEAL EDEMA". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9612453&pc=QCB | 2023826-2020-00172 | 2019-12-24 | Injury | STAAR SURGICAL COMPANY | 2020-01-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Inflammation; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.5/3.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. "TASS, AC CELLS 3+" WAS OBSERVED ON (B)(6) 2019. MEDICAL INTERVENTION WITH TOPICAL STEROIDS AND TOPICAL AND ORAL ANTIBIOTICS WAS INITIATED. ON (B)(6) 2019 BCVA WAS 20/25 " CORNEAL EDEMA, INFLAMMATION AND AC CELLS FLARE BUT REDUCED FROM DAY 1" AND " STILL CLOSE MONITORING THE PATIENT BY DAILY BASIS, SUSPECT TASS AND THE INFLAMMATION IS REDUCING" WAS NOTED. THE LENS REMAINS IMPLANTED AND AS OF (B)(6) 2020 THE REPORTER REPORTED " THE PATIENTS EYES ARE QUIET DOWN AS OF YESTERDAY. PATIENTS IS SEEING 20/20. THE SURGEON DOES NOT PERFORM ANY CULTURE TO CONFIRM IT IS TASS. JUST A DIAGNOSTIC BASED ON SIGNS AND SYMPTOMS." Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9612460&pc=QCB | 2023826-2020-00171 | 2019-12-24 | Injury | STAAR SURGICAL COMPANY | 2020-01-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Inflammation; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.00/3.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. "TASS, AC CELLS 3+" WAS OBSERVED ON (B)(6) 2019. MEDICAL INTERVENTION WITH TOPICAL STEROIDS AND TOPICAL AND ORAL ANTIBIOTICS WAS INITIATED. ON (B)(6) 2019 BCVA WAS 20/40 " CORNEAL EDEMA, INFLAMMATION AND AC CELLS FLARE BUT REDUCED FROM DAY 1" AND " STILL CLOSE MONITORING THE PATIENT BY DAILY BASIS, SUSPECT TASS AND THE INFLAMMATION IS REDUCING" WAS NOTED. THE LENS REMAINS IMPLANTED AND AS OF (B)(6) 2020 THE REPORTER REPORTED " THE PATIENTS EYES ARE QUIET DOWN AS OF YESTERDAY. PATIENTS IS SEEING 20/20. THE SURGEON DOES NOT PERFORM ANY CULTURE TO CONFIRM IT IS TASS. JUST A DIAGNOSTIC BASED ON SIGNS AND SYMPTOMS." Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9278980&pc=QCB | 2023826-2019-01993 |  | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A LENS INTO PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PER REPORTER, PATIENT "HAVING PROGRESSIVELY EDEMATOUS CORNEAS." SURGEON "THINKS VAULT MAY BE NEAR 1000 BUT DIFFICULT TO TELL THROUGH CORNEA, BUT VERY CERTAIN THERE IS NO INFLAMMATORY CELLS (I.E. AS IN TASS OR ENDOPHTHALMITIS) THROUGH SMALL WINDOW OF CLARITY NEAR INCISION. REPORTEDLY SURGEON "STATED SURGERY VERY SMOOTH AND UNEVENTFUL." THE SURGEON NOTES THIS IS HIS FIRST TIME USING TIME NUVISC (BY BVI, HIGHER WEIGHT COHESIVE SODIUM HA). "IOP WAS 16 AND 17, MOVING UP TO 26 WITH INCREASINGLY EDEMATOUS CORNEAS." INTRAOCULAR MOXIFLOXACIN WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B4 - CORRECTED TO 06-OCT-2019 IN INITIAL MDR .G4 - CORRECTED TO 06-OCT-2019 IN INTIAL MDR CLAIM#(B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9278986&pc=QCB | 2023826-2019-01994 |  | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A LENS INTO PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PER REPORTER, PATIENT "HAVING PROGRESSIVELY EDEMATOUS CORNEAS." SURGEON "THINKS VAULT MAY BE NEAR 1000 BUT DIFFICULT TO TELL THROUGH CORNEA, BUT VERY CERTAIN THERE IS NO INFLAMMATORY CELLS (I.E. AS IN TASS OR ENDOPHTHALMITIS) THROUGH SMALL WINDOW OF CLARITY NEAR INCISION. REPORTEDLY SURGEON "STATED SURGERY VERY SMOOTH AND UNEVENTFUL." FIRST TIME NUVISC (BY BVI, HIGHER WEIGHT COHESIVE SODIUM HA). IOP WAS 16 AND17, MOVING UP TO 26 WITH INCREASINGLY EDEMATOUS CORNEAS. INTRAOCULAR MOXIFLOXACIN WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B4 - CORRECTED TO 06-OCT-2019 IN INITIAL MDR. G4 - CORRECTED TO 06-OCT-2019 IN INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: G4 - CORRECTED TO 07-JAN-2020 IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). MODEL #: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334443&pc=QCB | 2023826-2021-00358 | 2021-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ANOTHER LENS WAS NOT IMPLANTED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. POST-OP MILD CORNEAL EDEMA. THE LENS WAS DISCARDED. Manufacturer Narrative: B5 - PATIENT IS USING ARTIFICIAL TEARS EVERY ONCE IN A WHILE AS NEEDED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005314&pc=QCB | 2023826-2021-01926 | 2021-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). ANOTHER LENS WAS NOT IMPLANTED. THE IOP WAS FIRST LOWERED WITH PARACENTESIS AND TAMPONADE ICL/REFILL A/C-FLUID WITH IOP INCREASING AFTER 1 HOUR. POST-OP CORNEAL EDEMA, DILATED PUPIL, OPTIC NERVE NOT PALE AT THIS TIME BUT CONCERNED AT LONG TERN CORNEAL AND OPTIC NERVE HEALTH. PATIENT WAS TAKING KRATOM-CHOLINERGENIC BLOCKER ALPHA AGONIST; POSSIBLE OPIOD EFFECTS. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND WITH RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14645301&pc=QCB | 2023826-2022-01780 | 2022-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXPLANTED DUE TO SIZING ISSUES. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. FOR CAUSE DETAILS THE REPORTER STATED "AT THE DAY 1 POSTOPERATIVE VISIT, THE UNCORRECTED VISUAL ACUITY MEASURED 20/320 OS. INTRAOCULAR PRESSURE WAS 45 MMHG OS. EXAMINATION DEMONSTRATED SEVERE CORNEAL EDEMA OS WITH NARROW ANGLES OU." A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AS REMOVAL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLES. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - PER EMAIL ON (B)(6) 2023, PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: IMPLANT DATE, EXPLANT DATE:UNK TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473796&pc=QCB | 2023826-2024-00131 | 2023-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129041&pc=QCB | 2023826-2018-01773 | 2018-11-09 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -12.5/+2.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. REPORTEDLY, THE PATIENT PRESENTED WITH BLURRY VISION, CORNEAL EDEMA, HIGH IOP AND EXCESSIVE VAULT. THE SURGEON PERFORMED A YAG PI AND PATIENT WAS PLACED ON IOP MEDICATIONS. CURRENTLY, THE PATIENT'S SYMPTOMS ARE RESOLVING AND VISION HAS IMPROVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9709011&pc=QCB | 2023826-2020-00272 | 2019-12-17 | Injury | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.00/+1.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. POST-OP, THERE WAS CORNEAL EDEMA AND THE INTRAOCULAR PRESSURE (IOP) INCREASED, WHICH WAS MANAGED WITH WOUND BURPING, PO DIAMOX, PREDNEFRINE FORTE AND TOPICAL G COMBIGAN. THE CORNEAL EDEMA AND THE IOP HAD SETTLED DOWN IN LESS THAN 2 WEEKS. POST-OP, THE PUPIL WAS REPORTED AS OVOID (OVAL SHAPED). THE LENS REMAINS IMPLANTED AND THE SURGEON AND PATIENT ARE WAITING FOR THE CORNEA TO SETTLE DOWN FOR A COUPLE OF MONTHS BEFORE ANY DECISIONS ARE MADE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10498203&pc=QCB | 2023826-2020-01980 | 2020-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +6.00/+1.5/032 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE VAULT WAS MUCH IMPROVED, THE ANGLE WAS OPEN AND IOP WAS NORMAL. THERE WAS CORNEAL EDEMA, WHICH WAS CAUSING REDUCTION OF VISION. SEE MFR REPORT # 2023826-2020-02050 FOR REPLACEMENT LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 1791 AND 2137 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. H6 - DEVICE CODE: 1401 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - WORK ORDER SEARCH RESULTS AND INFORMATION SHOULD BE REMOVED AS IT'S NOT APPLICABLE FOR THIS CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS SPHERE/CYLINDER/AXIS SHOULD BE CHANGED TO +5.00/1.5/049. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930510&pc=QCB | 2023826-2020-02837 | 2020-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Corneal Edema; Intraocular Pressure Increased; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS THAT IMMEDIATELY POSTOP THERE WAS ELEVATED IOP, PAIN, CORNEAL EDEMA, AND A VAULT OF 1200. DIAMOX WAS ADMINISTERED. DAY 21 POSTOP THE VAULT IS 1000, THE PUPIL DOESN'T TIGHTEN WELL, NO PAIN OR EDEMA, IOP 18. Manufacturer Narrative: B5-UPDATED INFORMATION: LENS WAS EXPLANTED. CLAIM# (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: ON 18-NOV-2020, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. VAULTING WAS RESOLVED WITH EXCHANGE, HOWEVER SEE MFR CLAIM#: (B)(4) FOR RESIDUAL SYMPTOMS AFTER EXCHANGE. CAUSE OF THE EVENT IS REPORTED AS A PATIENT RELATED FACTOR: "THE MEASUREMENT OF THE WTW WAS CORRECT AND ICL CALCULATION WAS COHERENT WITH IT. BUT APPARENTLY, THE ICL OF 13.2MM WAS TOO BIG AND DIDN'T FOUND THE GREAT POSITION OF IMPLANTATION IN THE SULCUS. WE HAD TO REMOVE IT FOR A SMALLER ONE." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473806&pc=QCB | 2023826-2024-00132 | 2023-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Corneal Edema; Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. ELEVATED IOP AND CORNEAL EDEMA WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK A2 - UNK A3 - UNK A4 - UNK A5 - UNK A6 - UNK B3 - UNK D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY) D4 - UNK D6A - UNK H4 - UNK CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120422&pc=QCB | 2023826-2019-01744 | 2012-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -2.5/+0.5/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, AND CORNEAL EDEMA. AN ENLARGEMENT OF THE PI-YAG WAS PERFORMED. THE PROBLEM IS RESOLVED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER INFORMATION RECEIVED ALSO STATES THAT THE ACD WAS NOT MEASURED PER DFU INSTRUCTIONS (ACD ASSUMED TO BE 3.2MM). INFORMATION INITIALLY RECEIVED VIA AN ARTICLE PUBLISHED IN CLINICAL OPHTHALMOLOGY ENTITLED 'A PROSPECTIVE PILOT STUDY USING A LOW POWER PIGGY-BACK TORIC IMPLANTABLE COLLAMER LENS TO CORRECT RESIDUAL REFRACTIVE ERROR AFTER MULTIFOCAL IOL IMPLANTATION.' Manufacturer Narrative: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120423&pc=QCB | 2023826-2019-01743 | 2013-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -1.0/+1.0/142 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, AND CORNEAL EDEMA. AN ENLARGEMENT OF THE PI-YAG WAS PERFORMED. THE PROBLEM IS RESOLVED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER INFORMATION RECEIVED ALSO STATES THAT THE ACD WAS NOT MEASURED PER DFU INSTRUCTIONS (ACD ASSUMED TO BE 3.2MM). INFORMATION INITIALLY RECEIVED VIA AN ARTICLE PUBLISHED IN CLINICAL OPHTHALMOLOGY ENTITLED 'A PROSPECTIVE PILOT STUDY USING A LOW POWER PIGGY-BACK TORIC IMPLANTABLE COLLAMER LENS TO CORRECT RESIDUAL REFRACTIVE ERROR AFTER MULTIFOCAL IOL IMPLANTATION.' Manufacturer Narrative: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8780421&pc=QCB | 2023826-2019-01173 | 2019-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/4.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA (PRE-OP BCVA = 20/80 POST-OP BCVA = 20/500) AND ELEVATED IOP (40 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2019. LENS WAS REMOVED ON (B)(6) 2019. PER REPORTER, PATIENTS CONDITION ON LAST VISIT, (B)(6) 2019, IOP 38 MMHG, CORNEAL EDEMA, MYDRIASIS, PHOTOPHOBIA, BLURRED VISION, PAIN. REPORTER INDICATED ON (B)(6) 2019, THAT PATIENT RECEIVED TREATMENT FOR HIGH IOP, DILATING DROPS WITH STEROIDS AND NON-STEROIDS. "NO MORE CORNEAL EDEMA AFTER" LENS REMOVAL. AFTER LENS REMOVAL, SURGEON CHOSE CLE (CLEAR LENS EXTRACTION) TO RESOLVE THE REFRACTIVE ISSUE. THE PATIENT IS FINE. THE REPORTER INDICATED THAT CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. CORRECTED TO NO, DEVICE EVALUATION ANTICIPATED BUT NOT YET BEGUN, IN INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13546404&pc=QCB | 2023826-2022-00529 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -8.00/+1.00/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THERE WAS REPORT OF PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE; ON (B)(6) 2022 IT WAS ALSO NOTED PATIENT EXPERIENCED MID DILATED PUPIL; UNCORRECTED VISUAL ACUITY; CORNEAL EDEMA AND A YAG PI SURGERY WAS PERFORMED WHICH RELIEVED THE PRESSURE AND WHEN THE PATIENT RETURNED (B)(6) 2022 THEY HAD IMPROVED VISUAL ACUITY/REDUCED INTRAOCULAR PRESSURE AND MID-DILATED PUPIL. THE PROBLEM WAS MARKED AS RESOLVED BUT NOTED PATIENT IS SCHEDULED TO COME IN FOR A 1 MONTH FOLLOW UP. Manufacturer Narrative: B5 - IT WAS LATER CLARIFIED THAT "EXAMINATION ON (B)(6) 2022- DEMONSTRATED REDUCED UNCORRECTED VISUAL ACUITY OF 20/32; MILD CORNEAL EDEMA; A MID-DILATED PUPIL AND NON-PATENT PERIPHERAL IRIDOTOMY OD. REPEAT LASER IRIDOTOMY OD IMMEDIATELY RELIEVED THE PRESSURE; AND THE SUBJECT RETURNED THE FOLLOWING DAY (B)(6) 2022, WITH IMPROVED VISUAL ACUITY; REDUCED INTRAOCULAR PRESSURE AND PERSISTENT MID-DILATED PUPIL." THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION :LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334355&pc=QCB | 2023826-2021-00359 | 2021-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ANOTHER LENS WAS NOT IMPLANTED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. POST-OP MILD CORNEAL EDEMA. THE LENS WAS DISCARDED. Manufacturer Narrative: B5 - THE PATIENT STATED SEEING A GLARE THE LEFT EYE (OS) WHILE NIGHT DRIVING AND USES ARTIFICIAL TEARS OCCASIONALLY AS NEEDED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11738953&pc=QCB | 2023826-2021-01207 | 2021-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | Corneal Edema; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN AND REMOVED A 13.2MM VTICM5_13.2, -8.50/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2021. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, "LACK OF EXPERIENCE (3RD LENS IMPLANTED), ADDED VISCOS IN 1/2 INFERIOR JUST BEFORE IMPLANTATION, WHICH PROBABLY LED A MOMENTUM FOR LENS REVERSAL. Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 OF -8.50/2.5/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS IMPLANTED; REMOVED AND RE-IMPLANTED INTRA-OP WITHIN THE SAME SURGERY DUE TO THE LENS IMPLANTED UPSIDE DOWN. ENLARGEMENT OF INCISION WAS REPORTED. "3 HOURS POST OP A SMALL CENTRAL AREA OF CORNEAL EDEMA. AT 10 DAY POST OP REFRACTION AT RIGHT EYE IS ABSOLUTE PERFECT." H6-HEALTH IMPLACT: 1791 CORNEAL EDEMA. H6-HEALTH IMPACT: 4625 ENLARGEMENT OF INCISION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253006&pc=QCB | 2023826-2023-02572 | 2023-05-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Corneal Edema; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/090 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN IN THE PATIENTS LEFT EYE (OS). THE LENS WAS DISCARDED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP, CORNEAL EDEMA WAS RESOLVING. SEE MFR. REPORT # 2023826-2023-02573 FOR REPLACEMENT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK; A5: UNK; A6: UNK; H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "1791" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) UPSIDE DOWN ON (B)(6) 2023. REPORTEDLY "THE LENS FLIPPED INSIDE THE EYE". THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A LENS OF THE SAME PARAMETERS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "CORNEA EDEMA STILL RESOLVING". THE CAUSE OF THE EVENT IS USER ERROR AND ITS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: " ICL LOADED LIKE NORMAL ROTATED DURING UNFOLDING AND WAS UPSIDE DOWN. ICL EXPLANTED AND REPLACED." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6549625&pc=QCB | 2023826-2016-01508 | 2016-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-05-05 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Corneal Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -11.5/+1.0/92 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT EXPERIENCED CORNEAL EDEMA POST-OP. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S WEIGHT: UNK. (B)(4). WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM # (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8438196&pc=QCB | 2023826-2019-00368 | 2019-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.00/0.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS IMPLANTED UPSIDE DOWN, SURGEON ENLARGED THE INCISION TO REMOVE THE LENS DURING THE INITIAL SURGERY. IT WAS REPORTED THE IRIS WAS PARTIALLY INJURED. THE CORNEAL INCISION WAS CLOSED AND AN ANTI-INFLAMATORY AND ANTIBIOTIC EYE DROPS WERE "INDICATED". SURGEON DID A MACULAR OCT FOR SMALL LOSS OF IRIS PIGMENT IN THE TEMPORAL AREA AND IT IS NORMAL. PATIEN'S CURRENT BCVA IS 20/20 AND IS ONLY USING "TEARS". SURGEON STATES THE CAUSE OF THE EVENT IS THAT THE "PROCEDURE IS COMPLICATED." CURRENTLY, THERE IS NO CORNEAL EDEMA, LENS IS CLEAR, PUPIL IS NORMAL. Manufacturer Narrative: "THE CORNEAL INCISION WAS CLOSED" SHOULD BE CORRECTED TO "A SUTURE WAS PLACED IN THE INCISION" EVENT DESCRIPTION: 'INDICATED' SHOULD BE CORRECTED TO 'PRESCRIBED' (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: INPLANT/EXPLANT DATE IN PREVIOUS MDR SHOULD BE CORRECTED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: DATE IN PREVIOUS MDR SHOULD BE CORRECTED TO 19-JUN-2019. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120433&pc=QCB | 2023826-2019-01745 | 2012-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Edema; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -0.5/+2.0/146 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE SURGEON REPORTS EXCESSIVE VAULT, ANGLE NARROWING, ANTERIOR BULGING OF THE IRIS, AND CORNEAL EDEMA. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER INFORMATION RECEIVED ALSO STATES THAT THE ACD WAS NOT MEASURED PER DFU INSTRUCTIONS (ACD ASSUMED TO BE 3.2MM REGARDLESS OF THE ACTUAL INDIVIDUAL ACD VALUE). INFORMATION INITIALLY RECEIVED VIA AN ARTICLE PUBLISHED IN CLINICAL OPHTHALMOLOGY ENTITLED 'A PROSPECTIVE PILOT STUDY USING A LOW POWER PIGGY-BACK TORIC IMPLANTABLE COLLAMER LENS TO CORRECT RESIDUAL REFRACTIVE ERROR AFTER MULTIFOCAL IOL IMPLANTATION.' Manufacturer Narrative: EVENT LENS DIOPTER CORRECTED TO -0.5/2.0/138 (SPHERE/CYLINDER/AXIS) IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: SEX: MALE. DATE OF EVENT: (B)(6) 2012. EVENT: PER UPDATED INFORMATION RECEIVED: THE LENS WAS IMPLANTED ON (B)(6) 2012. THE LENS WAS EXCHANGED ON (B)(6) 2012. IMPLANT DATE: (B)(6) 2012. EXPLANT DATE: (B)(6) 2012. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN ASSOCIATED LOTS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312692&pc=QCB | 2023826-2019-02106 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. POST-OP, THE PATIENT HAD CORNEAL EDEMA. Manufacturer Narrative: THE REPORTER INDICATED THAT NOW, ALL IS FINE WITH THE PATIENT. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824673&pc=QCB | 2023826-2020-00425 | 2020-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Corneal Edema; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -14.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT AND CORNEAL EDEMA WERE OBSERVED. THE LENS WAS INTRAOPERATIVELY REMOVED WITH ENLARGEMENT OF PI AND YAG. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR.ON PATIENT'S (B)(6) 2020 "MILD PERSISTENT CORNEAL EDEMA OBSERVED AND PATIENT CONSIDERING PLACING SMALL ICL." PER THE SURGEON ON (B)(6) 2020, "PATIENT PREFERS TO USE GLASSES AND CONTACT, NOT TO ATTEMPT ANOTHER ICL SURGERY." Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-ADDITIONAL INFORMATION- ORIGINAL PI PERFORMED 1 WEEK PRIOR TO SURGERY; AFTER THE ICL WAS PLACED THE VAULT WAS VERY HIGH AND THE ANTERIOR CHAMBERS WERE VERY SHALLOW. AT 1-2 HRS POST-OP SURGEON ENLARGED THE PI TO INSURE THEY WERE PATENT BUT THIS DID NOT LEAD TO DEEPENING OF THE ANTERIOR CHAMBER OVER THE NEXT HOUR. SURGEON THEN DECIDED TO EXPLANT THE LENS. THIS WAS DONE SAME DAY DIFFERENT SURGERY. H3-DEVICE EVALUATION-LENS RETURNED DRY IN POUCH WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CORRECTED DATA: H1-MALFUNCTION IN INITIAL MDR IS CORRECTED TO SERIOUS INJURY. (B)(4). Manufacturer Narrative: CORRECTED DATA: B4- IN INITIAL MDR IS CORRECTED TO (B)(6) 2020. B5- PATIENT ALSO EXPERIENCED "EXTREMELY NARROW ANGLES" SHOULD BE INCLUDED IN INITIAL MDR. D10- LENS RETURNED DATE IN PREVIOUS MDR IS CORRECTED TO (B)(6) 2020. G4- IN INITIAL MDR IS CORRECTED TO (B)(6) 2020. H6- PATIENT CODE 3191-NARROW ANGLES. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193768&pc=QCB | 2023826-2015-01441 | 2015-09-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-01 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Off-Label Use; Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; No Information; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +1.50/0.5/X090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE SURGEON IS PLANNING ON EXPLANTING AND EXCHANGING THE LENS FOR A SHORTER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED/ADDITIONAL INFORMATION: ON (B)(4) 2015 THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM IS RESOLVED. PATIENT ALSO EXPERIENCED ANTERIOR BULGING OF THE IRIS WITH ANGLE NARROWING AND INTERMITTENT CORNEAL EDEMA. THE LENS WAS IMPLANTED PIGGYBACK WITH AN MIOL. ADDITIONAL INFORMATION RECEIVED IN AN ARTICLE PUBLISHED IN CLINICAL OPHTHALMOLOGY, ENTITLED 'A PROSPECTIVE PILOT STUDY USING A LOW POWER PIGGY-BACK TORIC IMPLANTABLE COLLAMER LENS TO CORRECT RESIDUAL REFRACTIVE ERROR AFTER MULTIFOCAL IOL IMPLANTATION.' THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER THE LENS SIZE WAS CALCULATED USING AN ACD OF 3.2MM REGARDLESS OF THE ACTUAL ACD VALUE. ADDED REPORT SOURCE. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: PER UPDATED INFORMATION, THE LENS WAS IMPLANTED ON (B)(6) 2015. THE LENS WAS EXCHANGED ON (B)(6) 2016. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN ASSOCIATED LOTS. CORRECTED DATA: CORRECTED TO QCB IN INITIAL MDR. CORRECTED TO 07-AUG-2012. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498160&pc=QCB | 2023826-2019-02408 | 2019-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.50/+3.5/065 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND THE PATIENT FELT PAIN. THE INCISION WAS ENLARGED TO REMOVE THE LENS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR AND DEVICE. Manufacturer Narrative: POSTOPERATIVE REPORT STATED PATIENT'S EYE PRESSURE RETURNED TO NORMAL AND IS HEALING WITH CORNEAL EDEMA. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8860009&pc=QCB | 2023826-2019-01291 | 2019-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Edema; Pain; Red Eye(s); Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.00/3.5/007 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. BLURRED VISION, "PAINFUL, AND RED EYE" WAS OBSERVED. LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS "THE SIZES OF THE LENS PROBABLY TOO BIG." PER PATIENTS LAST VISIT ON (B)(6) 2019, "REDUCED VISION LEFT EYE STILL RIGID BIG PUPIL, IRIS DEFECT AS 13.00, CORNEAL EDEMA 201090410, RIGID PUPIL HAS NOT RETURNED TO NORMAL SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191 - IRIS DEFECT AND FIXED PUPIL. DEVICE CODES 3001 AND 1583 THAT WERE PREVIOUSLY REPORTED ARE NOT APPLICABLE. COMMON DEVICE NAME/ PRODUCT CODE CORRECTED. MTA PRODUCT CODE IN INITIAL MDR NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10114934&pc=QCB | 2023826-2020-01179 | 2020-05-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema; Swelling/ Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -8.00/2.0/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, "LITTLE CORNEAL EDEMA" WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H10 - H6 - PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION LENS EXCHANGE, SHOULD BE REMOVED IN INITIAL MEDWATCH REPORT AS IT IS NOT APPLICABLE. B5 - IT WAS ALSO INDICATED THE PATIENT EXPERIENCED EYELID SWELLING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10600060&pc=QCB | 2023826-2020-02050 | 2020-08-20 | Injury | STAAR SURGICAL COMPANY | 2020-09-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Visual Disturbances | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +5.00/+1.5/010 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THIS WAS THE REPLACEMENT LENS FOR MFR REPORT # 2023826-2020-01980. THE PATIENTS POST-OP STATUS WAS THE VAULT WAS MUCH IMPROVED, THE ANGLE WAS OPEN AND IOP WAS NORMAL. THERE WAS CORNEAL EDEMA, WHICH WAS CAUSING REDUCTION OF VISION. THE REPORTER HAS INDICATED THE EVENT HAS NOW RESOLVED AND OUTCOME IS SATISFACTORY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9126212&pc=QCB | 2023826-2019-01791 | 2019-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Edema; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.50/3.0/175 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-ANGLES AND GLARE/HALOES. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM BUT FOR STATUS OF THE EYE CORNEAL EDEMA WAS REPORTED. Manufacturer Narrative: PATIENT CODE: 1791 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310673&pc=QCB | 2023826-2020-01551 | 2016-04-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Touch; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.50/+1.5/148 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED HAPTICS CONTACTED THE PERIPHERAL CORNEAL ENDOTHELIUM. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. THE REPORTER INDICATED THEY DIDN'T HAVE EXPERIENCE IN THE EARLY TIME WHEN ICL FLOATED IN THE ANTERIOR CHAMBER WITH THE "OVD FREE" TECHNIQUE. THIS PATIENT WAS PART OF A STUDY COMPARING THE SAFETY OF IMPLANTABLE COLLAMER LENS (ICL) WITH AND WITHOUT OPHTHALMIC VISCOSURGICAL DEVICE (OVD). Manufacturer Narrative: CORRECTED DATA: B1 - CHANGED TO PRODUCT PROBLEM. B2 - OTHER SERIOUS (IMPORTANT MEDICAL EVENTS) SHOULD BE REMOVED AS EVENT IS NON SERIOUS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9350551&pc=QCB | 2023826-2019-02189 | 2019-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Ulcer; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE INTRAOCULAR PRESSURE WAS REPORTED AS BEING HIGH. THE PATIENT COMPLAINED OF POOR VISION AND SEES DISTORTED LIGHTS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/081 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE INTRAOCULAR PRESSURE WAS REPORTED AS BEING HIGH. THE PATIENT COMPLAINED OF POOR VISION AND SEES DISTORTED LIGHTS. TREATMENT WITH STEROID, ANTIBIOTIC LUBRICANT AND ACETAZOLAMIDE WAS INITIATED AND A THERAPEUTIC CONTACT LENS INSERTED. PAIN WAS REPORTED AND CORNEAL ULCER. THE LENS REMAINS IMPLANTED. PATIENT CODE: 3191 MEDICATION TREATMENT WITH STEROID, ANTIBIOTIC LUBRICANT AND ACETAZOLAMIDE PRESCRIBED, THERAPEUTIC CONTACT LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12914251&pc=QCB | 2023826-2021-04626 | 2021-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-13.50/+2.00/3, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO UNCOMFORTABILITY. IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM AND THAT THERE IS NO PLAN TO IMPLANT ANOTHER LENS.Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-13.50/+2.00/3, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO UNCOMFORTABILITY. IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM AND THAT THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12914252&pc=QCB | 2023826-2021-04625 | 2021-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-14.50/+2.50/175, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO UNCOMFORTABILITY. IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM AND THAT THERE IS NO PLAN TO IMPLANT ANOTHER LENS.Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-14.50/+2.50/175, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO UNCOMFORTABILITY. IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM AND THAT THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16649780&pc=QCB | 2023826-2023-01077 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO POSTOPERATIVE BINOCULAR VISION BALANCE, ELIMINATE DISCOMFORT AND ACHIEVE THE DESIRED CORRECTION PURPOSE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT : UNKNOWN; ETHNICITY: UNKNOWN; RACE: UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453736&pc=QCB | 2023826-2023-03240 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Discomfort; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 OF -11.00/3.5/078 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTER OPINION THE CAUSE FOR THIS EVENT WAS THE DEVICE. " LENS TOO LARGE, CAUSING DISCOMFORT" WAS REPORTED FOR CAUSE DETAILS. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/3.5/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND DISCOMFORT. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH THE SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS DEVICE- "LENS TOO LARGE, CAUSING DISCOMFORT'. CLAIM# (B)(6). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9038497&pc=QCB | 2023826-2019-01660 | 2019-05-10 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND DISCOMFORT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881043&pc=QCB | 2023826-2020-00668 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT DISCOMFORT (DUE TO PRESBYOPIA). THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: D4 - CORRECTED TO (B)(6) 2018. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10092802&pc=QCB | 2023826-2020-01146 | 2019-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/2.50/82, IMPLANTABLE COLLAMER LENS TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND REPLACED, AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. THE REPORTER STATED "BEFORE IMPLANTING THE LENS, I DECIDED IT WOULD BE OKAY EVEN THOUGH I RECOGNIZED THERE WAS A CRACK ON THE LENS. A FEW DAYS LATER, THE PATIENT FELT UNCOMFORTABLE SO, I SUSPECTED THE CAUSE OF THE CRACK AND REPLACED IT WITH A NEW LENS. COMPLETED THE SURGERY ALTHOUGH A CRACK WAS SEEN AROUND A FOOTPLATE AFTER INSERTING IN THE ANTERIOR CHAMBER. AT 3DAY POST-OP, DECIDED TO REPLACE IT WITH NEW LENS BECAUSE THE PATIENT COMPLAINED OF THE UNSTABLE VISION." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. PATIENT CODE 3191: NO CODE AVAILABLE (UNSTABLE VISION). PATIENT CODE: 2330 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364450&pc=QCB | 2023826-2020-01645 | 2020-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO "DISCOMFORT." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS STRETCHED OUT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364455&pc=QCB | 2023826-2020-01644 | 2020-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO "DISCOMFORT." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12020655&pc=QCB | 2023826-2021-01913 | 2021-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Dry Eye(s); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.5/1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PERMANENT FEELING OF PRESSURE AND DRY EYE. IT WAS REPORTED THAT THE PROBLEM RESOLVED AND LENS WILL STAY IN THE EYE. THE REPORTER ALSO STATED THAT THE PATIENT WILL GET PILOCARPINE DROPS AND PUNCTUM PLUGS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND SUBJECTIVE FEELING OF EYE PRESSURE WITHOUT HIGH IOP (10 MMHG)." SHOULD BE ADDED TO THE INITIAL MDR. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4). Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON (B)(6) 2023 WITH A SHORTER LENGTHL LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ALL GOOD. PATIENT IS HAPPY." STATUS OF THE EYE IS "LENS WILL STAY IN THE EYE!" ADDITIONAL INFORMATION STATES "PATIENT WILL GET PILOCARPINE DROPS AND PUNCTUL PLUGS." CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12020847&pc=QCB | 2023826-2021-01914 | 2021-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Dry Eye(s); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.5/1.0/132(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PERMANENT FEELING OF PRESSURE AND DRY EYE. IT WAS REPORTED THAT THE PROBLEM RESOLVED AND LENS WILL STAY IN THE EYE. THE REPORTER ALSO STATED THAT THE PATIENT WILL GET PILOCARPINE DROPS AND PUNCTUM PLUGS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND SUBJECTIVE FEELING OF EYE PRESSURE WITHOUT HIGH IOP (12 MMHG)." SHOULD BE ADDED TO THE INITIAL MDR. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4). Manufacturer Narrative: B5 - LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON 23 JUN 2023 AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE (B)(4): PERMANENT FEELING OF PRESSURE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16507962&pc=QCB | 2023826-2023-00837 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Dry Eye(s); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8/+2.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS. IT WAS INDICATED THAT THE PATIENT HAS MILD DRY EYES WITH IRRITATION. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - IRRITATION. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16507993&pc=QCB | 2023826-2023-00838 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Dry Eye(s); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7/+3/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. IT WAS INDICATED THAT THE PATIENT HAS MILD DRY EYES WITH IRRITATION. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - IRRITATION. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12267071&pc=QCB | 2023826-2021-03055 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Dry Eye(s); Blurred Vision; Halo |  | 1 | Event Description: AN ARTICLE PUBLISHED, TITLED "TORIC IMPLANTABLE COLLAMER LENS FOR THE TREATMENT OF MYOPIC ASTIGMATISM" WHICH COVERED A STUDY FOR TREATMENT OF MYOPIC ASTIGMATISM. IT INDICATED THAT OF 96 EYES, OF THE MULTIPLE PATIENTS13 EYES REQUIRED SECONDARY INTERVENTION WHICH INCLUDED: PRK SURGERY, LASIK, POSTOPERATIVE LIMBAL RELAXING INCISIONS, LENS REPOSITIONING. IT WAS ALSO NOTED DRY EYE SYMPTOMS WERE PRESENT IN 13% OF EYES WHICH PEAKED AT 29%, 39% OF EYES HAD GLARE/HALOES AT ONE MONTH WHICH IMPROVED TO 15% OF EYES AT TWELVE MONTHS. THERE WERE THREE EYES IN TWO PATIENTS WHICH REQUIRED ENLARGMENT OF PI (YAG) TO AVOID IMPENDING ANGLE CLOSURE GLAUCOMA. IT WAS NOTED " IN SUMMARY TORIC ICL IS SAFE AND EFFECTIVE OPTION FOR PATIENT'S WITH MYOPIC ASTIGMATISM SEEKING SURGICAL CORRECTION". LENSES REMAIN IMPLANTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11464544&pc=QCB | 2023826-2021-00625 | 2020-11-28 05:00 | Malfunction | STAAR SUGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Dry Eye(s); Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/1.0/124 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. PATIENT EXPERIENCED GLARE/HALOES. LENS REMAINS IMPLANTED,BEING MONITORED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) "MILD DRYNESS INFERIOR CORNEA" WAS REPORTED. THE REPORTER STATED "FEELS LIKE SEES DIPLOPIA IN DIMMER LIGHT AT DISTANCE AND NEAR." SURGEON IS WAITING TO SEE IF NEUROADAPTATION OCCURS OVER TIME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: PRODUCT PROBLEM; H1: MALFUNCTION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11464664&pc=QCB | 2023826-2021-00626 | 2020-11-28 05:00 | Malfunction | STAAR SUGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Dry Eye(s); Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.0/1.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PATIENT EXPERIENCED GLARE/HALOES. LENS REMAINS IMPLANTED,BEING MONITORED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) "MILD DRYNESS INFERIOR CORNEA" WAS REPORTED. THE REPORTER STATED "FEELS LIKE SEES DIPLOPIA IN DIMMER LIGHT AT DISTANCE AND NEAR." SURGEON IS WAITING TO SEE IF NEUROADAPTATION OCCURS OVER TIME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: PRODUCT PROBLEM. H1: MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983283&pc=QCB | 2023826-2022-01033 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Dry Eye(s); Inflammation; Intraocular Pressure Increased; Blurred Vision; Discomfort; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -12.5/+2.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS LOW VAULT, ELEVATED IOP, PIGMENT DISPERSION, AC INFLAMMATION CELLS +, SPK, HEAVINESS AND EYE DISCOMFORT, PHOTOPHOBIA, EYE PAIN, UNABLE TO OPEN EYE, DRY EYE. REPORTEDLY MEDICATION WAS ALSO ADMINISTERED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: BLURRED VISION AND ROUND PUPIL ALSO REPORTED. H6: CLINICAL CODE 4581 (ROUND PUPIL). CLAIM #(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983288&pc=QCB | 2023826-2022-01032 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Dry Eye(s); Inflammation; Intraocular Pressure Increased; Blurred Vision; Discomfort; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -10.5/+1.0/055 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS LOW VAULT, ELEVATED IOP, PIGMENT DISPERSION, AC INFLAMMATION CELLS +, SPK, HEAVINESS AND EYE DISCOMFORT, PHOTOPHOBIA, EYE PAIN, UNABLE TO OPEN EYE, DRY EYE. REPORTEDLY MEDICATION WAS ALSO ADMINISTERED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: BLURRED VISION AND ROUND PUPIL ALSO REPORTED. H6-CLINICAL CODE 4581 (ROUND PUPIL). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15583155&pc=QCB | 2023826-2022-03588 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Dry Eye(s); Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6, -14.5/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AS A REPLACEMENT LENS. IT WAS REPORTER THAT IT RESOLVED THE PROBLEM BUT FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) DRYNESS AND FLOATERS WERE REPORTED. FOR PATIENTS CURRENT CONDITION IT WAS REPORTED THAT THE EYE IS S/P ICL AND DOING WELL, NO INTERVENTION WAS NECESSARY AND IT HAS IMPROVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9278988&pc=QCB | 2023826-2019-01995 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Edema; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A LENS INTO PATIENT'S RIGHT EYE (OS) ON (B)(6) 2019. PER REPORTER, "ORIGINAL VAULT DAY ONE WAS 700-800 WITH 1+ EDEMA AT ONE WEEK WITH INCREASING VAULT TO 1000-1200. IOPS 23/22 AT ONE-TWO HOUR VISIT, 21/21 AT ONE WEEK, 19/17 AT ONE MONTH." REPORTEDLY PATIENTS VISION IS 20/20 BUT HALOS AND VAULT 1100-1200 WITH REPORTEDLY MID-DILATED, POORLY REACTIVE PUPILS ARE BEING OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: G4: IN THE INITIAL MDR IS CORRECTED TO 06-OCT-2019. H6-PATIENT CODE 1937 IN INITIAL MDR IS NOT APPLICABLE. PATIENT CODE 2140 IN INITIAL MDR IS NOT APPLICABLE. ADDITIONAL INFORMATION: H6-PATIENT CODE 2227-"HALOES". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4- IN INITIAL MDR DATE 07-OCT-2019 CORRECTED TO 06-OCT-2019. CLAIM# (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK. MODEL #: UNK. DEVICE MANUFACTURER DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9278996&pc=QCB | 2023826-2019-01996 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Edema; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A LENS INTO PATIENT'S LEFT EYE (OD) ON (B)(6) 2019. PER REPORTER, "ORIGINAL VAULT DAY ONE WAS 700-800 WITH 1+ EDEMA AT ONE WEEK WITH INCREASING VAULT TO 1000-1200. IOPS 23/22 AT ONE-TWO HOUR VISIT, 21/21 AT ONE WEEK, 19/17 AT ONE MONTH." REPORTEDLY PATIENTS VISION IS 20/20 BUT HALOS AND VAULT 1100-1200 WITH REPORTEDLY MID-DILATED, POORLY REACTIVE PUPILS ARE BEING OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B4- IN INITIAL MDR DATE (B)(6) 2019 CORRECTED TO (B)(6) 2019. CLAIM# : (B)(4). Manufacturer Narrative: CORRECTED DATA: G4- IN THE INITIAL MDR IS CORRECTED TO 06-OCT-2019. H6-PATIENT CODE 1937 IN INITIAL MDR IS NOT APPLICABLE. PATIENT CODE 2140 IN INITIAL MDR IS NOT APPLICABLE. ADDITIONAL INFORMATION: H6-PATIENT CODE 2227-HAOLES. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK. MODEL #: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10868449&pc=QCB | 2023826-2020-02719 | 2020-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/+2.00/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ENDOPHTHALMITIS, ADDITIONAL INTERVENTION WAS PERFORMED (ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND INTRAOCULAR INJECTION-CEFRAZIDIME). IT WAS REPORTED THAT A CULTURE WAS NOT PERFORMED, DOCTOR SUSPECTS CAUSE IS DEVICE RELATED, PATIENTS CURRENT CONDITION IS STABLE AND BCVA WAS 20/20.ENDOPHTHALMITIS WAS RESOLVED THROUGH TREATMENT AND WILL CONTINUE TO BE MONITORED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION- LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(6). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938962&pc=QCB | 2023826-2021-01613 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -14.50/3.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. A 1MM TEAR IN DESCEMET'S MEMBRANE. LENS REMAINS IMPLANTED. REPORTEDLY, 'AS OF NOW WE ARE WAITING TO SEE IF HER CORNEA WILL CONTINUE TO CLEAR." WHEN THE CORNEA CLEARS ENOUGH, ASSESS THE STATUS OF THE ENDOTHELIUM TO SEE HOW MUCH DAMAGE PATIENT HAS SUSTAINED BY THE THREE PROCEDURES ON THIS EYE. CURRENTLY PATIENT'S ANGLES ARE ADEQUATE AND THE IOP IS ACCEPTABLE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. MFR # 2023826-2021-01652. Manufacturer Narrative: B4: CORRECTED DATE TO 2021-MAY-23. G3: INITIAL MDR DATE SHOULD BE CORRECTED TO 2021-MAY-23. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK, DATE OF EVENT - UNK. (B)(4): DESCEMET MEMBRANE TEAR. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023217&pc=QCB | 2023826-2021-04898 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision | N/A | 1 | Event Description: A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -10.5/+2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS ENDOPHTHALMITIS AFTER ONE WEEK OF SURGERY. TREATMENT WAS PROVIDED, THE LENS WAS EXPLANTED. VITREORETINAL INJECTION AND OTHER MEDICATION WAS GIVEN. THE PROBLEM IS RESOLVED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS INITIALLY IMPLANTED (B)(6) 2021. THE LENS WAS EXPLANTED (B)(6) 2021. ADDITIONAL SYMPTOM OF BLURRED VISION REPORTED. H6-DEVICE CODE 1494: (AGE<21YRS) . CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9180458&pc=QCB | 2023826-2019-01904 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. IT WAS REPORTED THE PATIENT HAD BLURRED VISION AND THE SURGEON SUSPECTED ASEPTIC ENDOPHTHALMITIS. FIBRIN DEPOSITION AND LENS REAR SURFACE DEPOSITION WERE NOTED. THE SURGEON PERFORMED AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND DECADRON INJECTION TO THE SUBCONJUNCTIVAL WAS PERFORMED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: DEVICE CODE 2993 CORRECTED TO DEVICE CODE 1401 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE INFLAMMATION HAD SUBSIDED AFTER 12 DAYS OF TREATMENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136571&pc=QCB | 2023826-2023-04906 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Hypopyon; Inflammation; Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT WAS FOUND TO HAVE ENDOPHTHALMITIS AND PERMANENT LOSS OF BCVA. INTRAOCULAR PRESSURE (IOP) CHECK AND ANTERIOR SEGMENT IMAGING WERE PERFORMED ON (B)(6) 2023. PATIENT WAS PRESCRIBED MEDICATIONS FOR INCREASED IOP. PATIENT HAD HYPOPYON, FIBRIM NET, CELLS IN AC, WITH REDNESS, SEVERE LIGHT SENSITIVITY AND PAIN, INITIALLY DIAGNOSED AS TASS DUE TO FAST COURSE OF SIGNS AND CLEAR B-SCAN BY DOCTOR AND RETINAL SPECIALIST. PATIENT HAD VITREOUS BIOPSY WITH INTRAVITREAL INJECTION OF VANCOMYCIN AND CEFTAZIDIME TO RULE OUT AND TREAT ENDOPHTHALMITIS AS WELL AS SUBTENON KENALOG DUE TO INITIAL FAST COURSE OF SIGNS OF TASS. THE NEXT DAY, DUE TO CELLS ON VITREOUS B-SCAN, THE PATIENT HAD BILATERAL PPV (PARS PLANA VITRECTOMY) WITH INTRAVITREAL INJECTION OF CEFTAZIDIME, VANCOMYCIN AND VITREOUS CULTURE. VITREOUS CULTURE GREW PSEUDOMONA AERUGINOSA AND CURRENTLY THE PATIENT IS HOSPITALIZED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE PATIENT IS RECEIVING IV MEROPENEM, TOPICAL ANTIBIOTICS AND GLAUCOMA MEDS. IN THE LAST EXAMINATION BY RETINAL SPECIALIST, SHE WAS STILL HAVING PAIN IN BOTH EYES AND PERIORBITAL INFLAMMATION THAT HAD BEEN IMPROVED SINCE INITIAL HOSPITALIZATION EVALUATED BY CT AND MRI OF THE BRAIN AND ORBITS. CURRENT UCVA IS NLP OU (B)(6) 2023). H6: HEALTH EFFECT IMPACT CODE: 4644 - MEROPENEM. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK, D6B: UNK, H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PERMANENT LOSS OF BCVA, REDNESS, LIGHT SENSITIVITY. H6: HEALTH EFFECT IMPACT CODE: 4644 - CEFTAZIDIME, VANCOMYCIN. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PARS PLANA VITRECTOMY. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136578&pc=QCB | 2023826-2023-04905 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Hypopyon; Inflammation; Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT WAS FOUND TO HAVE ENDOPHTHALMITIS AND PERMANENT LOSS OF BCVA. INTRAOCULAR PRESSURE (IOP) CHECK AND ANTERIOR SEGMENT IMAGING WERE PERFORMED ON (B)(6) 2023. PATIENT WAS PRESCRIBED MEDICATIONS FOR INCREASED IOP. PATIENT HAD HYPOPYON, FIBRIM NET, CELLS IN AC, WITH REDNESS, SEVERE LIGHT SENSITIVITY AND PAIN, INITIALLY DIAGNOSED AS TASS DUE TO FAST COURSE OF SIGNS AND CLEAR B-SCAN BY DOCTOR AND RETINAL SPECIALIST. PATIENT HAD VITREOUS BIOPSY WITH INTRAVITREAL INJECTION OF VANCOMYCIN AND CEFTAZIDIME TO RULE OUT AND TREAT ENDOPHTHALMITIS AS WELL AS SUBTENON KENALOG DUE TO INITIAL FAST COURSE OF SIGNS OF TASS. THE NEXT DAY, DUE TO CELLS ON VITREOUS B-SCAN, THE PATIENT HAD BILATERAL PPV (PARS PLANA VITRECTOMY) WITH INTRAVITREAL INJECTION OF CEFTAZIDIME, VANCOMYCIN AND VITREOUS CULTURE. VITREOUS CULTURE GREW PSEUDOMONA AERUGINOSA AND CURRENTLY THE PATIENT IS HOSPITALIZED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE PATIENT IS RECEIVING IV MEROPENEM, TOPICAL ANTIBIOTICS AND GLAUCOMA MEDS. IN THE LAST EXAMINATION BY RETINAL SPECIALIST, SHE WAS STILL HAVING PAIN IN BOTH EYES AND PERIORBITAL INFLAMMATION THAT HAD BEEN IMPROVED SINCE INITIAL HOSPITALIZATION EVALUATED BY CT AND MRI OF THE BRAIN AND ORBITS. CURRENT UCVA IS NLP OU ((B)(6) 2023). H6: HEALTH EFFECT IMPACT CODE: 4644 - MEROPENEM. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK, D6B: UNK, H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PERMANENT LOSS OF BCVA, REDNESS, LIGHT SENSITIVITY. H6: HEALTH EFFECT IMPACT CODE: 4644 - CEFTAZIDIME, VANCOMYCIN. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PARS PLANA VITRECTOMY. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND/OR PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9583802&pc=QCB | 2023826-2020-00105 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Endophthalmitis; Hypopyon; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT DEVELOPED ENDOPHTHALMITIS AND HYPOPYON. THE SURGEON RINSED THE PATIENT'S EYE AND GAVE HER MEDICINE. THE SITUATION IS UNDER CONTROL AND PATIENT HAS RECOVERED WELL. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR, THE PATIENT WAS NOT PAYING ATTENTION TO HYGIENE. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11092007&pc=QCB | 2023826-2020-03236 | 2020-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Inflammation; Red Eye(s); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/062 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. ON THE 9TH DAY AFTER SURGERY, THE PATIENT COMPLAINED OF MILD REDNESS, A LITTLE PHOTOPHOBIA, MILD CILIARY HYPEREMIA, AND A FEW PLANKTON CELLS IN ANTERIOR CHAMBER. MEDICATION WAS PRESCRIBED. ON 10TH DAY, VISUAL ACUITY WAS SLIGHTLY DECREASED AND PLANKTONIC CELLS HAD INCREASED FROM PREVIOUS DAY. SOME LITTLE FIBROUS MEMBRANE EXUDATE ON SURFACE OF LENS. THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS AND AQUEOUS HUMOR WERE SENT FOR TESTING AND THE BACTERIAL CULTURE TESTING RESULT WAS NEGATIVE. MEDICATION WAS PRESCRIBED, THE INFLAMMATION HAS DISAPPEARED. THE PATIENT STATUS IS STABLE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT THE PATIENT EXPERIENCED ENDOPHTHALMITIS. ALSO ON (B)(6) 2021 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. PATIENT'S VISUAL ACUITY HAS RECOVERED TO PRE OPERATIVE UCVA AND BCVA IN THE LAST VISIT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10463780&pc=QCB | 2023826-2020-01914 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Loss of Vision; Vitrectomy; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/1.50/165 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ENDOPHTHALMITIS AND LOSS OF VISION WERE OBSERVED. DIAGNOSTICS (PRP OF VITREOUS SAMPLE) AND MEDICAL INTERVENTION (VITREOUS BIOPSY, CORE VITRECTOMY, INTRAVITREAL VANCOMYCIN, INTAVITREAL PIPERACILLIN TAZOBACTAM) WERE PERFORMED. VITREOUS BIOPSY SAMPLE REVEALED PSEUDOMONAS AERUGINOSA. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND ADDITIONAL INTERVENTION/TREATMENT WAS PERFORMED (ANTERIOR CHAMBER WASH, VITREOUS BIOPSY, VITRECTOMY, FLUID- AIR EXCHANGE, SILICONE OIL INJECTION, INTRAVITREAL VANCOMYCIN, INTRAVITREAL PIPERACILLIN TAZOBACTAM, INTRAVITREAL DEXAMETHSONE). VITREOUS BIOPSY WAS SENT FOR REPEAT GRAMS STAIN, KOH MOUNT, CULTURE AND SENSITIVITY FOR BACTERIA AND FUNGUS. ON (B)(6) 2020 VITREOUS CULTURE SHOWED NO GROWTH. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INFORMATION RECEIVED ON 21 SEP 2020: PATIENT HAD NO GENERAL HEALTH CONDITIONS. PATIENT WAS PREPPED WITH POVIDONE IODINE. A PROPHYLACTIC INTRACAMERAL (MOXIFLOXACIN) ANTIBIOTIC WAS USED AT THE TIME OF IMPLANTAION. PATIENT HAS PL+VE VISION IN THE AFFECTED EYE. VISUAL PROGNOSIS IS NOT GOOD. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9772189&pc=QCB | 2023826-2020-00389 | 2020-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Vitrectomy; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.50/+6.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE DEVELOPMENT OF ENDOPHTHALMITIS. DIAGNOSTIC TESTS WERE PERFORMED WHICH FOUND SERRATIA. A VITRECTOMY WAS PERFORMED AND PATIENT WAS ADMINISTERED INTRAOCULAR INJECTION. THE LENS WAS DISCARDED. THE PATIENT DID NOT RETURN TO THE CLINIC BUT WENT TO A HOSPITAL, WHERE THE LENS WAS EXPLANTED, AFTER TESTS AND TREATMENT. CURRENTLY, THE INFLAMMATORY REACTION WAS WEAKEN AND THE PATIENT'S CONDITION IS IMPROVING. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708898&pc=QCB | 2023826-2020-02435 | 2019-09-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Eye Injury; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. THE LENS WAS SUCCESSFULLY EXCHANGED INTRA-OPERATIVELY. PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY STATED PATIENT INJURY OCCURRED. CORRECTED STATEMENT: NO PATIENT INJURY OCCURRED. H6-CLINICAL AND HEALTH IMPACT CODES CORRECTED. DEVICE PROBLEM CODE 2993 CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B2 - OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708903&pc=QCB | 2023826-2020-02434 | 2019-09-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Eye Injury; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. THE LENS WAS SUCCESSFULLY EXCHANGED INTRA-OPERATIVELY. PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY STATED THAT PATIENT INJURY OCCURRED. CORRECTED STATEMENT: NO PATIENT INJURY OCCURRED. H6-CLINICAL AND HEALTH IMPACT CODES CORRECTED. DEVICE PROBLEM 2992 CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B2 - OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS SHOULD BE REMOVED AS IT'S NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17222220&pc=QCB | 2023826-2023-02559 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE REPORTER STATES " LENS IS TOO LARGE. PATIENT EXPERIENCING INTERMITTENT PAIN. ASSUMING CAUSING PATIENT TO GO IN AND OUT OF ANGLE CLOSURE LESS THAN LEFT EYE". LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE REPORTER STATES " PATIENT WHO'S LENS IS TOO LARGE AND CAUSING HER INTERMITTENT PAIN, ASSUMING CAUSING HER TO GO IN AND OUT OF ANGLE CLOSURE OS>OD". CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980501&pc=QCB | 2023826-2020-02926 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), PAIN, PHOTOPHOBIA, AND FIXED PUPIL. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED/ADDITIONAL INFORMATION: AN ADDITIONAL SUTURE WAS ALSO PERFORMED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE VAULT WAS RESOLVED. REFERENCE MFR FILE# (B)(4) FOR ADDITIONAL SYMPTOMS WITH THE REPLACEMENT LENS. H6-HEALTH IMPACT CODE UPDATED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136765&pc=QCB | 2023826-2022-02524 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. INTERMITTENT EYE PAIN IS REPORTED. THE LENS REMAINS IMPLANTED. THE REPORTER STATES CAUSE WAS DUE TO THE DEVICE. CAUSE DETAILS PROVIDED BY THE REPORTER DESCRIBE MULTIPLE PATIENT RELATED FACTORS THAT POTENTIALLY CONTRIBUTE TO THE INTERMITTENT EYE PAIN. THE REPORTER STATES, "THIS SUBJECT'S SYMPTOMS ARE COMPATIBLE WITH INTERMITTENT SUBACUTE ANGLE CLOSURE GLAUCOMA PRECIPITATED BY ACCOMMODATIVE MIOSIS DURING COMPUTER WORK. HOWEVER, THERE ARE FEATURES WHICH ARGUE AGAINST THIS DIAGNOSIS, INCLUDING THE ABSENCE OF ANY SIGNS OR SYMPTOMS FOLLOWING PHARMACOLOGIC DILATION, PATENT PIS AND NORMAL ICL VAULT. THE HISTORY OF MIGRAINE AND THE CONSTELLATION OF SYMPTOMS IN THIS SUBJECT SUGGESTS THAT A DIAGNOSIS OF BENIGN EPISODIC UNILATERAL MYDRIASIS MAY ALSO BE ENTERTAINED. NEVERTHELESS, THIS DIAGNOSIS MUST ONLY BE MADE IF INTERMITTENT ANGLE CLOSURE IS FIRST RULED OUT." Manufacturer Narrative: (B)(4). Manufacturer Narrative: B1: CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. B5: "REPORTEDLY, AFTER ORIGINAL TMICL SURGERY SUBJECT BEGAN EXPERIENCING EYE PAIN, FEELING OF PRESSURE, NAUSEA, AND PULSATING VISION WITH HEAR BEAT, INTERMITTENTLY OVER THE PAST SEVERAL MONTHS. EPISODES WOULD HAPPEN WHEN READING AND WHEN IN DIM LIGHTING. SUBJECT HAD OS TMICL REPLACED ON (B)(6) 2023 WHICH DECREASED THE VAULT OF THE ICL OVER THE CRYSTALLINE LENS. SUBJECT WILL BE FOLLOWED UNTIL HER EYE PAIN HAS BEEN SHOWN TO BE RESOLVED. " SHOULD BE ADDED TO THE INITIAL MDR. H1: CORRECTED TO: SERIOUS INJURY IN THE INITIAL MDR. H6: HEALTH EFFECT - IMPACT CODE: 2199 SHOULD BE REMOVED FROM THE INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: B5- IT WAS ALSO REPORTED, "SEVERITY: GRADE 3/SEVERE, OUTCOME: UNKNOWN AND RELATIONSHIP TO DEVICE: UNKNOWN". H6-HEALTH EFFECT- IMPACT CODE: 4629 SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- HEALTH EFFECT- CLINICAL CODE: 4581- INTERMITTENT MYDRIASIS. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED INTERMITTENT EYE PAIN AND INTERMITTENT MYDRIASIS. REPORTEDLY, "AFTER ORIGINAL TMICL SURGERY THE SUBJECT BEGAN EXPERIENCING EYE PAIN, FEELING OF PRESSURE, NAUSEA, AND PULSATING VISION WITH HEARTBEAT, INTERMITTENTLY OVER THE PAST SEVERAL MONTHS. EPISODES WOULD HAPPEN WHEN READING AND WHEN IN DIM LIGHTING". IT WAS ALSO REPORTED, "SEVERITY: GRADE 3/SEVERE, OUTCOME: UNKNOWN AND RELATIONSHIP TO DEVICE: UNKNOWN". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: " "(B)(6) 2022 SUBJECT COMPLAINED OF INTERMITTENT EPISODES OVER THE PAST SEVERAL WEEKS OF "FEELING EYE PRESSURE" OS WITH ACCOMPANYING OCULAR PAIN, HAZY VISION, OCCASIONAL NAUSEA AND A LARGER PUPIL THAN FELLOW EYE. UNCORRECTED VISUAL ACUITY MEASURED 20/20 AND INTRAOCULAR PRESSURE WAS 14 MMHG, AND THE REMAINDER OF THE EXAMINATION WAS UNREMARKABLE. THIS ADVERSE EVENT WAS NOTED AS RESOLVED 06APR2022. ON (B)(6) 2022 THE SUBJECT PRESENTED FOR AN UNSCHEDULED VISIT REPORTING EYE PAIN WHICH STARTED (B)(6) 2022. VISUAL ACUITY MEASURED 20/16, INTRAOCULAR PRESSURE WAS 16 AND LENS VAULT MEASURED 616 [?]M OS. THE TMICL LENS WAS EXCHANGED ON (B)(6) 2023 FOR VTICM5 12.1, -12.5/1.0/093 EVO+ ICL LENS. THE POSTOPERATIVE COURSE WAS UNEVENTFUL. ON 03MAR2023, THE SUBJECT REPORTED THERE HAD BEEN NO FURTHER EPISODES OF PAIN." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6: HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16942788&pc=QCB | 2023826-2023-01890 | 2022-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/1.5/011 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT COMPLAINT OF PAIN AROUND THE EYE AND THE SURGEON NOTES THE PATIENT HAD MUSCAE VOLITANTES (EYE FLOATERS). THE LENS WAS EXPLANTED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - MUSCAE VOLITANTES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387537&pc=QCB | 2023826-2023-05730 | 2023-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Eye Pain; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/115 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT WAS EXPERIENCING ORBITAL PAIN DUE TO PSYCHOLOGICAL EFFECTS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18389058&pc=QCB | 2023826-2023-05729 | 2023-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Eye Pain; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/067 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT WAS EXPERIENCING ORBITAL PAIN DUE TO PSYCHOLOGICAL EFFECTS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778992&pc=QCB | 2023826-2021-01327 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Eye Pain; Swelling/ Edema | N/A | 1 | Event Description: HE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE SUSPENSORY LIGAMENT OF THE EYE LENS THAT WAS FOUND TO BE RELAXED, AND PATIENT REPORTED SEVERE EYE PAIN, SWELLING AND DISTENSION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE INTRAOCULAR LENS WAS IMPLANTED INTO THE EYE TO SINK. AFTER THE OPERATION, THE PATIENT HAD SEVERE EYE PAIN, SWELLING AND DISTENSION, WHICH AFFECTED HIS LIFE. THEREFORE, THE LENS WAS STRONGLY REQUIRED TO BE REMOVED." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779023&pc=QCB | 2023826-2021-01328 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Eye Pain; Swelling/ Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. SEVERE EYE PAIN, SWELLING AND DISTENSION WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. REPORTEDLY, "DURING THE OPERATION, THE SUSPENSORY LIGAMENT OF THE LENS WAS FOUND TO BE RELAXED, AND THE LENS WAS IMPLANTED INTO THE EYE TO SINK. AFTER THE OPERATION, THE PATIENT HAD SEVERE EYE PAIN, SWELLING AND DISTENSION, WHICH AFFECTED HIS LIFE. THEREFORE, THE LENS WAS STRONGLY REQUIRED TO BE REMOVED." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16622823&pc=QCB | 2023826-2023-01088 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Fatigue; Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT IS HYPEROPIC AND SHOWING RESIDUAL CYL; HAS HEADACHES AND FATIGUE; STEROIDS WERE PRESCRIBED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT IS HYPEROPIC AND SHOWING RESIDUAL CYL; HAS HEADACHES AND FATIGUE; STEROIDS WERE PRESCRIBED. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SAME MODEL AND SAME LENGTH BUT DIFFERENT DIOPTER LENS. IT HAS NOT BEEN REPORTED IF THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6A: SHOULD BE CORRECTED TO 09-FEB-2023. H6: 4629 SHOULD BE CORRECTED TO 4627. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT IS HYPEROPIC AND SHOWING RESIDUAL CYL; HAS HEADACHES AND FATIGUE; STEROIDS WERE PRESCRIBED. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SAME MODEL AND SAME LENGTH BUT DIFFERENT DIOPTER LENS.IT IS UNCLEAR WHETHER LENS REMOVAL AND REPLACEMENT OCCURRED INTRAOPERATIVELY. RESIDUAL CYLINDER WAS PRESENT AFTER THE EXCHANGE. IT HAS NOT BEEN REPORTED IF THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 HYPEROPIC/SHOWING CYL. TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715562&pc=QCB | 2023826-2023-03867 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Fatigue; Nausea; Visual Disturbances; Dizziness; Halo; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES, DIZZINESS, NAUSEA, EYE FATIGUE AND PAIN. THE PROBLEM WAS RESOLVED. THE INTRAOCULAR PRESSURE WAS NORMAL AND GOOD NEAR AND FAR VISION. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715631&pc=QCB | 2023826-2023-03868 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Fatigue; Nausea; Visual Disturbances; Dizziness; Halo; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+3.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES, DIZZINESS, NAUSEA, EYE FATIGUE AND PAIN. THE PROBLEM WAS RESOLVED. THE INTRAOCULAR PRESSURE WAS NORMAL AND GOOD NEAR AND FAR VISION. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685354&pc=QCB | 2023826-2021-04060 | 2021-09-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Foreign Body Sensation in Eye; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT GOT EXTREMELY ANXIOUS, FELT LIKE SOMETHING WAS IN THE EYE AND COULD NOT "GO TO BED." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. REPORTEDLY, THE PATIENT COULD NOT ACCEPT THE LENS PSYCHOLOGICALLY. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685388&pc=QCB | 2023826-2021-04059 | 2021-09-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Foreign Body Sensation in Eye; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO THE PATIENT GOT EXTREMELY ANXIOUS, FELT LIKE SOMETHING WAS IN THE EYE AND COULD NOT "GO TO BED." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. REPORTEDLY, THE PATIENT COULD NOT ACCEPT THE LENS PSYCHOLOGICALLY. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025900&pc=QCB | 2023826-2020-01016 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Foreign Body Sensation in Eye; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT FELT UNCOMFORTABLE AND HAD FOREIGN BODY SENSATION. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025905&pc=QCB | 2023826-2020-01017 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Foreign Body Sensation in Eye; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT FELT UNCOMFORTABLE AND HAD FOREIGN BODY SENSATION. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10868572&pc=QCB | 2023826-2020-02721 | 2020-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.00/+2.50/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND GLAUCOMA. IT WAS REPORTED THAT PATIENT'S CURRENT CONDITION IS GOOD AND GLAUCOMA WAS RESOLVED THROUGH MEDICATION. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10868645&pc=QCB | 2023826-2020-02720 | 2020-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.00/+2.50/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND GLAUCOMA. IT WAS REPORTED THAT PATIENT'S CURRENT CONDITION IS GOOD AND GLAUCOMA WAS RESOLVED THROUGH MEDICATION. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: UNK. UNK. UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453104&pc=QCB | 2023826-2021-00590 | 2021-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -4.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE PATIENT'S EYE DUE TO CLOSED ANGLE GLAUCOMA BEING OBSERVED. NO PATIENT INJURY. ADDITIONAL SUTURE TO CLOSE THE WOUND. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN CONTACT LENS CASE, DRY, RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND STRETCHED OUT, WITH RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16142126&pc=QCB | 2023826-2023-00099 | 2022-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.50/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. POST-OP, THERE WAS REMAINING PUPIL DISSIPATION AFTER AN ACUTE GLAUCOMA ATTACK. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00098. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4)..
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766002&pc=QCB | 2023826-2023-01455 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -10.00+1.50/103 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED PIGMENT DISPERSION; GLAUCOMA AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ADDITIONALLY NOTED THE PATIENT IS ON ANTI-GLAUCOMA DROPS FOR NOW. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION . TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773437&pc=QCB | 2023826-2023-01461 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -12.00/+1.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED PIGMENT DISPERSION; GLAUCOMA AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ADDITIONALLY NOTED THE PATIENT IS ON ANTI-GLAUCOMA DROPS FOR NOW. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8592286&pc=QCB | 2023826-2019-00776 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma; Headache; Intraocular Pressure Increased; Pain | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2019. ON (B)(6) 2019 THE PATIENT WENT TO ER BECAUSE SHE HAD A VERY BAD HEADACHE AND HER EYE IN PAIN. THE GLAUCOMA SURGEON IN ER PERFORMED OCT AND IOP MEASUREMENT, ACUTE IOP AND HIGH VAULT WERE OBSERVED. NO SECONDARY SURGICAL INTERVENTION OR MEDICAL TREATMENT WERE REPORTED TO BE PERFORMED. IT WAS REPORTED THAT THE PATIENT FELT DOWN AND UNCOMFORTABLE IN THE FIRST WEEK, HAD WRONG PERCEPTION ABOUT THE POST OPERATION EFFECT. AFTER POST OP 5 DAYS, 6 DAYS AND 1 WEEK, PATIENT CAN ACCEPT ICL. IOP IS STABLE 9-10 MMHG. AT THE MOMENT PATIENT HAS GOOD VA 20/20, ANGLES ARE OPEN AND 10 MMHG. HOWEVER, DUE TO EXCESSIVE VAULT, THE SURGEON WILL EXCHANGE TO A SMALLER SIZE LENS( IT WAS REPORTED THAT THE SURGEON FORGOT TO DEDUCT THE CT FROM ACD WHEN CALCULATING FOR INITIAL IMPLANT). LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CHANGED TO ADVERSE EVENT. CHANGED TO SERIOUS INJURY. THE GLAUCOMA SURGEON ALSO OBSERVED ACUTE GLAUCOMA DURING OCT AND IOP MEASUREMENT.(B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766109&pc=QCB | 2023826-2023-01337 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Glaucoma; Headache; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)-2023, YAG WAS PERFORMED DUE TO TRANSIENT LOSS OF BCVA AND ELEVATED IOP. REPORTEDLY, "PATIENT HAD INTERMITTENT TRANSIENT VISION LOSS RESOLVED WITH LASER PI. HER LATEST IOP IS 12MMHG. POSSIBLE INTERMITTIENT RESULTS OF PUPILARY BLOCK PRIOR TO LASER PI IN THIS CASE. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(6). Manufacturer Narrative: CORRECTED DATA: B5- THE PATIENT ALSO EXPERIENCED GRAYISH VISION, VISUAL FIELD DEFECT, ISCHAEMIC ATTACK, AND PUNCTATE EPITHELIAL EROSIONS. H6- HEALTH EFFECT- CLINICAL CODE: 4581- GRAYISH VISION, VISUAL FIELD DEFECT, ISCHAEMIC ATTACK, AND PUNCTATE EPITHELIAL EROSIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED TRANSIENT LOSS OF BCVA, YAG WAS PERFORMED, ELEVATED IOP: 24 MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, PERIORBITAL EYE PAIN, MEDICATION WAS PRESCRIBED, GLAUCOMA, INTERMITTENT PUPILLARY BLOCK, AND HEADACHES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. REPORTEDLY "THE PATIENT DEVELOPED OS SECONDARY GLAUCOMA POST ICL, PATIENT DECIDED TO REMOVE BOTH EYE ICL". CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153715&pc=QCB | 2023826-2022-02531 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Glaucoma; Inflammation; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -12.00/+6.0/005 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS GRADUAL LOSS OF VISION AND DIAGNOSIS OF BLINDNESS ((B)(6) 2022). BLINDNESS DUE TO GLAUCOMA RESULTING FROM ICL IMPLANTATION. THE SPECIFIC ETIOLOGY IS UNKNOWN. THE GLAUCOMA PROGRESSED RELATIVELY QUICKLY WITH REPEATED MILD ELEVATIONS IN IOP AND SEDATION WITH MEDICATION (ANTI GLAUCOMA AND ANTI INFLAMMATION) . THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK, SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- " DATE OF (B)(6) 2014 SHOULD BE CORRECTED TO (B)(6) 2014." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM 12.5 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/+6.0/005 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT EXPERIENCED LOSS OF VISION; BLINDNESS DUE TO GLAUCOMA (DIAGNOSED ON (B)(6) 2022; ELEVATED IOP; SECONDARY YAG PERFORMED; LENS REPOSITIONING, LENS ROTATION; INFLAMMATION; PIGMENT; AND MEDICATION WAS PRESCRIBED. REPORTEDLY "THE LENS REMAINS IMPLANTED. THE IMPLANTED LENSES ARE NOT LENSES SOLD BY STAAR JAPAN. THE LENSES WERE PURCHASED BY PHYSICIANS THROUGH PRIVATE IMPORTATION." THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE WAS REPORTED AS "BLINDNESS DUE TO GLAUCOMA RESULTING FROM ICL IMPLANTATION. THE SPECIFIC ETIOLOGY IS UNKNOWN. THE GLAUCOMA PROGRESSED RELATIVELY QUICKLY WITH REPEATED MILD ELEVATIONS IN IOP AND SEDATION WITH MEDICATION." THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN." H6- HEALTH IMPACT-CLINICAL CODE: 4581- PIGMENT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15434301&pc=QCB | 2023826-2022-03078 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient Device Interaction Problem | Glaucoma; Intraocular Pressure Increased; Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EYE PAIN; ANGLE CLOSURE WITH ELEVATED IOP AND BLURRED VISION. "PAIN IN RIGHT SIDE AFTER IMPLANTATION. PUSHING FORWARD ON IRIS. GLAUCOMA SUSPECT POSSIBLE." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP; PIGMENT DISPERSION; BLURRED VISION; ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. STATUS OF EYE IS "OCCASIONAL CLOUDINESS RIGHT EYE ;MGD." CAUSE IS REPORTED AS PATIENT RELATED FACTOR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6: CLINICAL CODE 4581: PIGMENT DISPERSION. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15434626&pc=QCB | 2023826-2022-03079 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; Patient Device Interaction Problem | Glaucoma; Intraocular Pressure Increased; Blurred Vision; Halo; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.50/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EYE PAIN; ANGLE CLOSURE WITH ELEVATED IOP AND GLARE/HALOES. "PAIN IN LEFT SIDE AFTER IMPLANTATION. PUSHING FORWARD ON IRIS. GLAUCOMA SUSPECT POSSIBLE; PHOTOSENSITIVITY. PIGMENTATION." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.50/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE ELEVATED IOP; PIGMENT DISPERSION; BLURRED VISION; ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. STATUS OF EYE IS "MGD." CAUSE IS REPORTED AS PATIENT RELATED FACTOR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6-CLINICAL CODE 4581- PIGMENT DISPERSION. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317868&pc=QCB | 2023826-2019-02130 | 2019-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +7.5/+1.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO A "POST-OP GLAUCOMA PROBLEM" AND ANGLE CLOSURE WITH ELEVATED IOP. AN ADDITION OF NEW PI WAS ALSO PERFORMED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR, PATIENT-RELATED FACTOR, AND UNKNOWN. THE SURGEON REPORTS THE MOST LIKELY CAUSE IS THAT THE PATIENT POSTOPERATIVELY INSTILLED MIDRINP EYEDROPS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317889&pc=QCB | 2023826-2019-02135 | 2019-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +6.5/+1.0/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO A "POST-OP GLAUCOMA PROBLEM" AND ANGLE CLOSURE WITH ELEVATED IOP. AN ADDITION OF NEW PI, IRIS SUTURE, AND ADDITIONAL SUTURE WAS ALSO PERFORMED. THE CAUSE OF THE DEVICE IS REPORTED AS USER ERROR, PATIENT-RELATED FACTOR, AND UNKNOWN. THE SURGEON REPORTS THE MOST LIKELY CAUSE IS THAT THE PATIENT POSTOPERATIVELY INSTILLED MIDRINP EYEDROPS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THERE WAS A TEAR IN THE OPTIC AND THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9672545&pc=QCB | 2023826-2020-00246 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; Pain; Vomiting; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.0/+1.0/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP. THE REPORTER STATED THAT THE PATIENT DEVELOPED DENSE AC FIBRINOUS INFLAMMATORY REACTION, SECONDARY GLAUCOMA AND RIGHT FROM IMPLANTATION COMPLAINED OF SEVERE PAIN, WATERING AND VOMITING. PATIENT WAS UNDER STEROIDS. IT WAS REPORTED THAT HE PROBLEM HAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC, RESIDUE AND DEBRIS ON THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (ANGLE CLOSURE, STEROIDS). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9672785&pc=QCB | 2023826-2020-00247 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; Pain; Vomiting; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.0/+1.5/053 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP. THE REPORTER STATED THAT THE PATIENT DEVELOPED DENSE AC FIBRINOUS INFLAMMATORY REACTION, SECONDARY GLAUCOMA AND RIGHT FROM IMPLANTATION COMPLAINED OF SEVERE PAIN, WATERING AND VOMITING. PATIENT WAS UNDER STEROIDS. IT WAS REPORTED THAT HE PROBLEM HAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC AND RESIDUE ON THE LENS. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (ANGLE CLOSURE, STEROIDS). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652117&pc=QCB | 2023826-2019-00935 | 2019-05-05 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-29 | QCB | IMPLANTABLE COLLAMER LENS | Optical Problem | Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/154 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THOUGH SURGEON REPORTS THAT THE STATUS OF THE EYE IS "ALL RIGHT," IT IS ALSO NOTED THAT THE PATIENT COMPLAINTS ABOUT HALOS, ESPECIALLY DURING NIGHT DRIVING. REFERENCE MFR REPORT# 2023826-2019-00934 FOR INITIALLY IMPLANTED LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368707&pc=QCB | 2023826-2021-00468 | 2021-01-19 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/092, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. # 2023826-2021-00097 BUT THE PROBLEM WAS NOT RESOLVED, THE PATIENT WAS NOT SATISFIED DUE TO HALOS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. EXPLANT DATE: NA. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12161974&pc=QCB | 2023826-2021-02573 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA | 1 | Event Description: THE PATIENT REPRTED THAT HE SEES LIGHT RINGS AFTER TICL IMPLANT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12162012&pc=QCB | 2023826-2021-02574 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA | 1 | Event Description: THE PATIENT REPORTED THAT HE SEES LIGHT RINGS AFTER TICL IMPLANT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378632&pc=QCB | 2023826-2021-03367 | 2021-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS HALOS AROUND LIGHT IN THE EVENING AND NIGHT IN THE CENTRAL VISUAL FIELD. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402700&pc=QCB | 2023826-2022-01550 | 2022-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -3.5/+1.5/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES AND LENS ROTATION. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THE PATIENT STILL HAD HALOES. THEN ON (B)(6) 2022 THE LENS WAS EXPLANTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299128&pc=QCB | 2023826-2022-02800 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM 13.2 IMPLANTABLE COLLAMER LENS OF -5.50/3.5/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). THE REPORTER STATES THE PATIENT IS HAVING PROBLEMS WITH HALOES. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT: UNK. IMPLANT DATE: UNK. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15495834&pc=QCB | 2023826-2022-02684 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 03-AUG-2021. EXCESSIVE VAULT WAS OBSERVED WITH GLARE AND HALOES. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: LENS WAS EXCHANGED ON (B)(6) 2022 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ALL FINE! PATIENT IS VERY HAPPY! NO MORE HALOS". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15643334&pc=QCB | 2023826-2022-03625 | 2022-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -7.0/1.00/101 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, LENS OPACITY (CATARACT) WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS REMOVED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057494&pc=QCB | 2023826-2022-04732 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- 1494 OFF-LABEL HISTORY OF AMBYLOPIA, BINOCULAR NYSTAGMUS. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057499&pc=QCB | 2023826-2022-04731 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- OFF-LABEL:1494 SHOULD BE ADDED FOR CORRECTION; HISTORY OF BINOCULAR AMBLYOPIA, BINOCULAR NYSTAGMUS. CLAIM#(B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178026&pc=QCB | 2023826-2023-00155 | 2022-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON 18-OCT-2022. THE PATIENT EXPERIENCED HALOES AT NIGHT BY CAR DASHBOARD. THE LENS REMAINED IMPLANTED. THE VA HAD IMPROVED AND PATIENT WAS DOING WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: RACE: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. DEVICE CODE: (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178121&pc=QCB | 2023826-2023-00156 | 2022-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/073 (SPHERE,CYLINDER,AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES AT NIGHT BY CAR DASHBOARD. THE LENS REMAINED IMPLANTED. THE VA HAD IMPROVED AND PATIENT WAS DOING WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16250870&pc=QCB | 2023826-2023-00339 | 2022-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.0/1.50/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO HALOS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. THE PROBLEM WAS REPORTED AS RESOLVED AND NOT RESOLVED; IT WAS INDICATED THE LENS WAS REMOVED DUE TO THE PATIENT'S WISH; REFRACTIVE OUTCOME UA UNCHANGED TO PRE-OP. PATIENT IS CONSIDERING LASER REFRACTIVE SURGERY IN ONE YEAR. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269395&pc=QCB | 2023826-2023-00312 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.5/2.5/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON 28-(B)(6) 2022, THE LENS WAS EXPLANTED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269396&pc=QCB | 2023826-2023-00309 | 2022-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY. "PATIENT REQUEST TO TAKE OUT TICL." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16328066&pc=QCB | 2023826-2023-00381 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo |  | 1 | Event Description: IN THE ARTICLE, "CORNEAL ENDOTHELIAL RING FOLLOWING THE IMPLANTATION OF TORIC IMPLANTABLE COLLAMER LENSES WITH A CENTRAL HOLE: A CASE REPORT" THE ARTICLE LOOKED INTO A CASE OF A CORNEAL ENDOTHELIAL RING AFTER TORIC IMPLANTABLE COLLAMER LENS (TICL, V4C) IMPLANTATON IN THE RIGHT EYE OF A PATIENT. REPORTEDLY, "ANTERIOR SEGMENT EXAMINATION REVEAL AN ENDOTHELIAL ANNULAR LESION OF APPROXIAMTELY 0.4 MM DIAMETER IN THE CENTRAL PART OF THE CORNEA, GRAY-WHITE IN COLOR." TOBRAMYCIN AND DEXMETHASONE EYE DROPS WERE APPLIED. AFTER 10 DAYS, DEXAMENTHASONE TOBRAMYCIN WAS REPLACED WITH FLUOROMENTHOLONE, AND TAPERED SLOWLY OVER THE NEXT WEEKS. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16499869&pc=QCB | 2023826-2023-00784 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023, THE LENS WAS REMOVED AND LATER REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, GLARE, AND HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. REPORTEDLY, "DOING WELL HAPPY WITH VISION.". Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500068&pc=QCB | 2023826-2023-00783 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND LATER REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, GLARE, AND HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. REPORTEDLY, "DOING WELL HAPPY WITH VISION." Manufacturer Narrative: PATIENT INFORMATION:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16506867&pc=QCB | 2023826-2023-00628 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00/+0.5/174 ((B)(6) 2023 DUE TO THE PATIENT EXPERIENCED HALOS AND THE PROBLEM WAS RESOLVED. HYALURONIC ACID WAS USED DURING THE EXPLANTATION, WHICH INDUCED A CATARACT. THE PATIENT IS BEING CLOSELY MONITORED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16506879&pc=QCB | 2023826-2023-00627 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00/+0.5/002 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT EXPERIENCED HALOS. HYALURONIC ACID WAS USED DURING THE EXPLANTATION, WHICH INDUCED A CATARACT. THE PATIENT IS BEING CLOSELY MONITORED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792346&pc=QCB | 2023826-2023-01442 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO HIS RIGHT EYE (OD) IN (B)(6) 2022. THE PATIENT IS DISSASTISFIED AND COMPLAINS OF HALOES. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922247&pc=QCB | 2023826-2023-01774 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/169 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF HALOS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922381&pc=QCB | 2023826-2023-01773 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF HALOS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253193&pc=QCB | 2023826-2023-02659 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -1.50/6.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "HALOS BECAUSE OF THR REMAINING CYLINDER." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. IT WAS ALSO REPORTED, "THE DOCTOR IMPLANTED AN IPCL WITH A HIGHER CYLINDER +1/+8/A80. Manufacturer Narrative: ADDITIONAL DATA: H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.50/6.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED GLARES/HALOES. REPORTEDLY "HALOES BECAUSE OF THE REMAINING CYLINDER". ON (B)(6) 2022 WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "THE DOCTOR IMPLANTED AN IPCL WITH A HIGHER CYLINDER +1/+8/A80. THE PATIENT IS HAPPY!". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: OFF LABEL- AGE <21. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253194&pc=QCB | 2023826-2023-02658 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -1.50/6.0/097 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "HALOS BECAUSE OF THR REMAINING CYLINDER." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. IT WAS ALSO REPORTED, "THE DOCTOR IMPLANTED AN IPCL WITH A HIGHER CYLINDER +1/+8/A80. Manufacturer Narrative: A4-A6:UNK. H3: OFF LABEL- AGE <21. CLAIM# (B)(4). Manufacturer Narrative: H6: HEALTH EFFECT: IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6: MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.50/6.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED GLARES/HALOES AND RESIDUAL ASTIGMATISM. REPORTEDLY "HALOES BECAUSE OF THE REMAINING CYLINDER". ON (B)(6) 2022 WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "THE DOCTOR IMPLANTED AN IPCL WITH A HIGHER CYLINDER +1/+8/A80. THE PATIENT IS HAPPY!". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# 729917.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18088946&pc=QCB | 2023826-2023-04903 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT COMPLAINED OF LIGHT RINGS/HALOS. CYCLOGEL WAS USED. AT POST-OP VISIT THE LIGHT RINGS/HALOS HAD RESOLVED AND THE PATIENT WAS 20/15 AND HAPPY. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: A2: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6B: NA. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18088947&pc=QCB | 2023826-2023-04904 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT COMPLAINED OF LIGHT RINGS/HALOS. CYCLOGEL WAS USED. AT POST-OP VISIT THE LIGHT RINGS/HALOS HAD RESOLVED AND THE PATIENT WAS 20/15 AND HAPPY. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: A2: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6B: NA. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18322977&pc=QCB | 2023826-2023-05582 | 2023-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS A PATIENT-RELATED-FACTOR WITH THE DEVICE NOT PERFORMING AS INTENDED. STATUS OF THE EYE IS REPORTED AS, "OK." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. PATIENT-RELATED-FACTOR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18322978&pc=QCB | 2023826-2023-05583 | 2023-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS A PATIENT-RELATED-FACTOR WITH THE DEVICE NOT PERFORMING AS INTENDED. STATUS OF THE EYE IS REPORTED AS, "OK." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. PATIENT-RELATED-FACTOR. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18337475&pc=QCB | 2023826-2023-05532 | 2023-08-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD DISLOCATED/SUBLUXED AND PATIENT COMPLAINED OF HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18340659&pc=QCB | 2023826-2023-05533 | 2023-08-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD DISLOCATED/SUBLUXED AND PATIENT COMPLAINED OF HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387277&pc=QCB | 2023826-2023-05785 | 2023-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING AND GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18389150&pc=QCB | 2023826-2023-05784 | 2023-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/1.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING AND GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15384508&pc=QCB | 2023826-2022-03045 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL.13.7; -10.0/+1.0/153 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAD EXPERIENCED GLARE/HALOES; WHICH THEN WAS TREATED WITH LUMIFY EYE DROPS WHICH STARTED ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED AND PATIENT REMAINS ON STUDY. Manufacturer Narrative: EXPIRATION DATE: 11-30-2023. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15384832&pc=QCB | 2023826-2022-03046 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL.13.7; -11.5/+1.0/025 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT HAD EXPERIENCED GLARE/HALOES; WHICH THEN WAS TREATED WITH LUMIFY EYE DROPS WHICH STARTED ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED AND PATIENT REMAINS ON STUDY. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: EXPIRATION DATE: 11-30-2023. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17680112&pc=QCB | 2023826-2023-03649 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+3.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE SURGEON NOTED A RING STRUCTURE ON THE LENS SURFACE AFTER THE LENS WAS IMPLANTED. THE LENS REMAINED IMPLANTED. THE PATIENTS REFRACTIONS AND VISION MEASUREMENTS WERE NORMAL BUT THE PATIENT MENTIONED FEELING UNCOMFORTABLE DUE TO SEEING HALOS AND A GRAY SPOT IN FRONT OF LEFT EYE. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581 - VISUAL DYSPHOTOPSIA. CLAIM#: (B)(4). Manufacturer Narrative: B3: UNK. D6A: UNK. D6B: NA . H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16519714&pc=QCB | 2023826-2023-00810 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -12.00/2.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2227" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE:"3001" SHOULD BE ADDED. B5- THE PATIENT ALSO EXPERIENCED GLARE/HALOS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES" SHOULD BE ADDED. ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/71 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. REPORTEDLY "THE NEW LENS WAS NEVER INSTALLED AGAIN AND WAS THEREFORE RETURNED. CLIENT CANCELED THE TREATMENT. HE HAD NO CONFIDENCE THAT THE NEW LENS WOULD BE ANY GOOD." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE- "WRONG WTW". CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17051835&pc=QCB | 2023826-2023-02126 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF 13.5/3.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO PATIENT DISCOMFORT/GLARE AND HALOES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2330" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE<21). B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6MPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED DISCOMFORT AND GLARES/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FROM ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17051836&pc=QCB | 2023826-2023-02127 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF 15.0/3.5/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO PATIENT DISCOMFORT/GLARE AND HALOES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2330" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE<21). B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED DISCOMFORT AND GLARES/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FROM ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831603&pc=QCB | 2023826-2022-00850 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Discomfort; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -5.50/1.0/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED OF SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: B5- PATIENT EXPERIENCED DISCOMFORT AND HALOS (RESIDUAL REFRACTIVE ERROR)- SHOULD HAVE BEEN INCLUDED ON INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED DRY IN WRAPPED CLOTH WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT. DEBRIS THROUGHOUT THE LENS AND ONE HAPTIC IS FOLDED IN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577870&pc=QCB | 2023826-2022-00513 | 2022-01-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Insufficient Information | N/A | 1 | Event Description: THE REPORTER INDICATED THAT EXCESSIVE VAULT WAS OBSERVED IN THE PATIENT'S RIGHT EYE (OD).Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -10.00/1.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT AND GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 -10.00/+1.5/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 15 JUN 2021. EXCESSIVE VAULT AND GLARE/HALOES WERE OBSERVED. LENS WAS EXCHANGED ON 17 MAY 2022 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. PATIENT IS HAPPY AND ALL FINE. CLAIM (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH - UNK. SEX - UNK. WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. DATE OF EVENT - UNK. PRODUCT CODE: UNK. MODEL#/ SERIAL#/ EXP DATE/ CATALOG #/ LOT#/ UDI#- UNK. IMPLANT DATE - UNK. EXPLANT DATE - UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577873&pc=QCB | 2023826-2022-00514 | 2022-01-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Insufficient Information | N/A | 1 | Event Description: THE REPORTER INDICATED THAT EXCESSIVE VAULT WAS OBSERVED IN THE PATIENT'S LEFT EYE (OS).Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -8.50/1.5/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. EXCESSIVE VAULT AND GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 -8.50/+1.5/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(3) 2021. EXCESSIVE VAULT AND GLARE/HALOES WERE OBSERVED. LENS WAS EXCHANGED ON (B)(3) 2022 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. ALL FINE AND PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNKNOWN, SEX: UNKNOWN, WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, DATE OF EVENT: UNKNOWN, COMMON DEVICE NAME: PRODUCT CODE: UNK. MANUFACTURER NAME, CITY AND STATE - UNKNOWN, IMPLANT DATE: UNKNOWN. EXPLANT DATE:UNKNOWN, PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE: UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830517&pc=QCB | 2023826-2022-03969 | 2022-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE IS REPORTED AS A PATIENT RELATED FACTOR. REPORTEDLY, "VISUAL GHOSTING," AND "TAKE OUT THE CRYSTAL." Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4) 4B LOW VALT EXCHANGE FOR SAME SIZE LENS, PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 20-DEC-2022. H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC STRETCHED, BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECTS - CLINICAL SIGNS AND SYMPTOMS-2227 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015475&pc=QCB | 2023826-2023-02011 | 2023-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; No Apparent Adverse Event | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B5: LENS WAS REMOVED DUE TO GLARE AND A STRONG APERTURE. REPORTEDLY, "THE POSTOPERATIVE GLARE AND APERTURE WERE STRONG, AND THE PATIENT COULD NOT ADAPT, SO THE CRYSTAL WAS REQUIRED TO BE REMOVED." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18107825&pc=QCB | 2023826-2023-04970 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem; No Apparent Adverse Event | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/145(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT IS OBSERVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/145(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT, AND HALOES WERE OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. IT WAS REPORTED THE HALOS DIMINISHED BECOMING BEARABLE FOR THE PATIENT. THE SURGEON IS WAITING FOR A POSTOP EXAM MAYBE (B)(6) 2024. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND INDICATED LENS TOO BIG. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586880&pc=QCB | 2023826-2020-00065 | 2019-12-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.50/3.0/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE, EXCESSIVE VAULT, AND HEADACHE WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "HEADACHE RELIEVED BY WEARING GLASSES PRESCRIBED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586884&pc=QCB | 2023826-2020-00064 | 2019-12-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.00/3.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE, EXCESSIVE VAULT WAS OBSERVED, WITH HEADACHE AND DIFFICULTY DRIVING. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "FIRST CASE TO DEVELOP HYPEROPIA POST-OP WHEN SUPPOSED TO BE, THE PATIENT SHOULD BE SLIGHTLY MYOPIC POST-OP. HEADACHE WITHOUT GLASSES. PATIENT GIVEN GLASSES TO EASE HEADACHE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13958485&pc=QCB | 2023826-2022-01028 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache |  | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND HEADACHES. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039558&pc=QCB | 2023826-2023-01879 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | Headache | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCING SEVERE AND PERSISTENT HEADACHE AND MIGRAINE. PROBLEM RESOLVED. REPORTEDLY, NO NEW LENS WILL BE IMPLANTED. PATIENT USED HER NORMAL PAINKILLERS USED FOR MIGRAINE(NAPROXEN, SUMITRAPTAN). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039559&pc=QCB | 2023826-2023-01880 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | Headache | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. LENS WAS EXPLANTED ON (B)(6)2023 DUE TO PATIENT EXPERIENCING SEVERE AND PERSISTENT HEADACHE AND MIGRAINE. PROBLEM RESOLVED. REPORTEDLY, NO NEW LENS WILL BE IMPLANTED. PATIENT USED HER NORMAL PAINKILLERS USED FOR MIGRAINE(NAPROXEN, SUMITRAPTAN). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12980949&pc=QCB | 2023826-2021-04799 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -1.00/0.50/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND HEADACHES IN THE NIGHT ON THE PATIENT'S LEFT SIDE. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13641115&pc=QCB | 2023826-2022-00645 | 2022-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -01.50/+1.0/018 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PATIENT HAD A HEADACHE AT NIGHT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14294131&pc=QCB | 2023826-2022-01424 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -17.50/3.0/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. IT REPORTED THE CAUSE OF THE EVENT WAS "OTHER" AND NOTED "ICL TOO LARGE - PATIENT WITH HIGH HTO NON-MEASURABLE) AT D+1 POST OP VISIT WITH VOMITING; SEVERE HEADACHE". THE LENS WAS EXPLANTED ON (B)(6) 2022. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2022, A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15316632&pc=QCB | 2023826-2022-02946 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.50/1.50/064 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCPLANTED DUE TO ANISEIKONIA AND HEADACHE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). EXPIRATION DATE: 04/30/2024. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18086700&pc=QCB | 2023826-2023-05008 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Headache; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6; 3/5/125 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS LOW VAULTING WITH LENS ROTATION, BLURRED VISION, HEADACHES, THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18086711&pc=QCB | 2023826-2023-05009 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Headache; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6; 4/3.5/79 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS LOW VAULTING WITH LENS ROTATION, BLURRED VISION, HEADACHES, THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749935&pc=QCB | 2023826-2020-02521 | 2020-06-17 | Injury | STAAR SURGICAL COMPANY | 2020-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), HEADACHES, INTERMITTENT ANGLE CLOSURE, AND BLURRED VISION. AN ENLARGEMENT OF PI-YAG WAS ALSO PERFORMED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749941&pc=QCB | 2023826-2020-02520 | 2020-06-17 | Injury | STAAR SURGICAL COMPANY | 2020-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), HEADACHES, INTERMITTENT ANGLE CLOSURE, AND BLURRED VISION. AN ENLARGEMENT OF PI-YAG WAS ALSO PERFORMED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882131&pc=QCB | 2023826-2023-04444 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2, +2/+1.5/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THE PATIENT NOW HAS LOSS OF BCVA AND BLURRED VISION, HEADACHES AND REDUCED CONTRAST AT NIGHT. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - ON DAY ONE POST-OP THE UCVA WAS REPORTED AS 0.9 AND 'PRESSURE INCREASE LE' WAS REPORTED WHICH WAS MANAGED BY'"EXTRA YAG IRIDOTOMY LE AFTER WHICH PRESSURE NORMAL'.EARLY ELEVATED IOP WAS REPORTED. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - LOSS OF BCVA, REDUCED CONTRASTS AT NIGHT. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE - AMBLYOPIA. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882138&pc=QCB | 2023826-2023-04445 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2, +1.00/4.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS THE PATIENT NOW HAS LOSS OF BCVA AND BLURRED VISION, HEADACHES AND REDUCED CONTRAST AT NIGHT. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - LOSS OF BCVA, REDUCED CONTRASTS AT NIGHT. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE - AMBLYOPIA. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17701923&pc=QCB | 2023826-2023-03896 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2019. REPORTEDLY, "OVER THE PAST 4 YEARS HAS DEVELOPED AN INCREASED VAULT AND NARROWING ANGLES." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THERE WAS BLURRY VISION AND HEADACHES. H6 - 4581 - NARROW ANGLES, 2137, 1880 CODE SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17701924&pc=QCB | 2023826-2023-03895 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN 2019. REPORTEDLY, "OVER THE PAST 4 YEARS HAS DEVELOPED AN INCREASED VAULT AND NARROWING ANGLES." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THERE WAS BLURRY VISION AND HEADACHES. H6 - 4581 - NARROW ANGLES, 2137, 1880 CODE SHOULD HAVE BEEN INCLUDED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708897&pc=QCB | 2023826-2020-02431 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-20 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS EXCESSIVE VAULT, AC ANGLE OF 12MM, HEADACHES, BLURRED VISION AND TUNNEL VISION. ALPHAGAN WAS ADMINISTERED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708900&pc=QCB | 2023826-2020-02430 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-20 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS EXCESSIVE VAULT, AC ANGLE OF 11MM, HEADACHES, BLURRED VISION AND TUNNEL VISION. ALPHAGAN WAS ADMINISTERED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916646&pc=QCB | 2023826-2023-01725 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Inflammation; Intraocular Pressure Increased; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0/3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT, INFLAMMATION, ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED, AND STEROIDS WERE PROVIDED. IT WAS ALSO STATED, "AS THE PATIENT WAS TOLD BY HER DOCTOR IN USA: 'PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP."" CAUSE OF THE EVENT WAS BOTH UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY" SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS." THE PROBLEM WAS NOT RESOLVED. THE REPORTER ALSO STATED, "AS PER THE PATIENT: INFLAMMATION BEING TREATED WITH STEROIDS BY HER CURRENT DOCTOR, (B)(6)," AND "PATIENT SEEMS TO HAVE A NARROW SULCUS, MEASUREMENTS WAS DONE AND REVIEWED SEVERAL TIMES BY (B)(6), THE PATIENT IMPLANTED THE LENS IN JORDAN BY (B)(6) ON (B)(6) 2022, AND IN (B)(6) 2022 THE PATIENT CONTACTED HIM INFORMING HIM THAT SHE DID AN EXAMINATION IN USA AND HER DOCTOR (B)(6) INFORMED THAT SHE HAS AN EXCESSIVE VAULT. AS THE PATIENT IS IN THE UNITED STATES AND COULD NOT COME TO JORDAN FOR REPLACEMENT SURGERY, SO SHE REQUESTED THAT SHE REPLACES THE ICL IN THE STATES, AS PER HER CURRENT DOCTOR, HE ADVISED THAT HE EXPLANTS AND THEN MONITOR THE PATIENT FOR 3 MONTHS AND THEN WE PROCEED ACCORDINGLY AS SHE SEEMS TO HAVE ELEVATED IOP.". Manufacturer Narrative: CORRECTED DATA: TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. HEALTH EFFECT- CLINICAL CODE: "1880" AND "2137" SHOULD BE ADDED. ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/95 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, INFLAMMATION, HEADACHE, BLURRED VISION AND MEDICATION WAS PRESCRIBED. REPORTEDLY "PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP". THE LENS WAS EXPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY "HE DOESN'T INTEND TO IMPLANT NEW LENSES ANY TIME IN THE IMMEDIATE FUTURE ". THE REPORTER INDICATED THE CAUSE OF THE EVENT AS A PATIENT RELATED FACTOR AND UNKNOWN. - "SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS". (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: PATIENT RELATED FACTOR. CLAIM# (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916648&pc=QCB | 2023826-2023-01726 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Inflammation; Intraocular Pressure Increased; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND INFLAMATION WITH ELEVATED IOP WERE OBSERVED, AND STEROIDS WERE PROVIDED. IT WAS ALSO STATED, "AS THE PATIENT WAS TOLD BY HER DOCTOR IN USA: 'PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP."" CAUSE OF THE EVENT WAS BOTH UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY" SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS." THE PROBLEM WAS NOT RESOLVED. THE REPORTER ALSO STATED, "AS PER THE PATIENT: INFLAMMATION BEING TREATED WITH STEROIDS BY HER CURRENT DOCTOR, (B)(6)," AND "PATIENT SEEMS TO HAVE A NARROW SULCUS, MEASUREMENTS WAS DONE AND REVIEWED SEVERAL TIMES BY (B)(6), TE PATIENT IMPLANTED THE LENS IN (B)(6) BY (B)(6) IN (B)(6) 2022, AND IN (B)(6) 2022 THE PATIENT CONTACTED HIM INFORMING HIM THAT SHE DID AN EXAMINATION IN USA AND HER DOCTOR SAMER MULKI INFORMED THAT SHE HAS AN EXCESSIVE VAULT. AS THE PATIENT IS IN THE UNITED STATES AND COULD NOT COME TO (B)(6) FOR REPLACEMENT SURGERY, SO SHE REQUESTED THAT SHE REPLACES THE ICL IN THE STATES, AS PER HER CURRENT DOCTOR, HE ADVISED THAT HE EXPLANTS AND THEN MONITOR THE PATIENT FOR 3 MONTHS AND THEN WE PROCEED ACCORDINGLY AS SHE SEEMS TO HAVE ELEVATED IOP." Manufacturer Narrative: A4-A6:UNK H6: PATIENT RELATED FACTOR. CLAIM#(B)(4). Manufacturer Narrative: H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6- HEALTH EFFECT- CLINICAL CODE: "1880" AND "2137" SHOULD BE ADDED. ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.5/78 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, INFLAMMATION, HEADACHE, BLURRED VISION AND MEDICATION WAS PRESCRIBED. REPORTEDLY "PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP". THE LENS WAS EXPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY "HE DOESN'T INTEND TO IMPLANT NEW LENSES ANY TIME IN THE IMMEDIATE FUTURE ". THE REPORTER INDICATED THE CAUSE OF THE EVENT AS A PATIENT RELATED FACTOR AND UNKNOWN. - "SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS". CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13173083&pc=QCB | 2023826-2022-00057 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/172, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON 2021 DUE TO EXCESSIVE VAULTING, SHALLOW CHAMBER, SIGNIFICANT REDUCTION OF ANGLE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT COMPLAINED OF A HEADACHE, WHICH WAS RESOLVED WITH MEDICATION (DIAMOX AND TOPICAL DRUGS). IOP REMAINS STABLE UNDER TOPICAL BETA-BLOCKER. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH ANOTHER ICL LENS WITHIN THE SAME SURGERY ON (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT- CLINICAL CODE: (B)(4). HEALTH EFFECT IMPACT CODE: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14274802&pc=QCB | 2023826-2022-01399 | 2022-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13.0/+1.0/14 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; UNREACTIVE FIXED PUPIL. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS ADDITIONALLY REPORTED THAT THERE IS NO PLANS TO REPLACE THE LENS AND THE PATIENT EXPERIENCED "GLAUCOMA SYMPTHOMATIC LIKE A HEADACHE AND VOMITING; HE DID NOT CONTACT THE CLINIC". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4) ANGLE CLOSURE, FIXED PUPIL, VOMITING. TYPE OF INVESTIGATION: (B)(4) LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433012&pc=QCB | 2023826-2023-00707 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -16.5/1.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; HEADACHES; SHALLOW ANTERIOR CHAMBER. ON (B)(6) 2023 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND A PATIENT RELATED FACTOR AND NOTED "PATIENT HAS EXTREME FLOPPY IRIS WHICH IS CAUSING EXTREME BOWING OF IRIS TISSUE FORWARDS LEADING TO REDUCTION OF IRIDOCORNEAL ANGLE; LEFT EYE WORSE THAN RIGHT. IOP SPIKES PRESENT CAUSING HEADACHE. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE, SHALLOW ANTERIOR CHAMBER. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433013&pc=QCB | 2023826-2023-00708 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -17.0/2.0/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; HEADACHES; SHALLOW ANTERIOR CHAMBER. THE PROBLEM IS NOT RESOLVED; AS THE SURGEON NOTED THE PATIENT STILL HAS HIGH VAULT WITH EXPLANT OR YAG PI IN CONSIDERATION. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND A PATIENT RELATED FACTOR AND NOTED "PATIENT HAS EXTREME FLOPPY IRIS WHICH IS CAUSING EXTREME BOWING OF IRIS TISSUE FORWARDS LEADING TO REDUCTION OF IRIDOCORNEAL ANGLE; LEFT EYE WORSE THAN RIGHT. IOP SPIKES PRESENT CAUSING HEADACHE. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE, SHALLOW ANTERIOR CHAMBER. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16947904&pc=QCB | 2023826-2023-01916 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.50/6.0/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPOSITIONED/ROTATED, PER REPORTER THE PATIENT PRESENTED WITH ELEVATED INTRAOCULAR PRESSURE, HEADACHE AND NAUSEA FOLLOWING REPOSITIONING. FIBRIN PLUG OVER CENTRAFLOW HOLE WAS OBSERVED, ONCE RELEASED, INTRAOCULAR PRESSURE RETURNED TO NORMAL AND PROBLEM RESOLVED. Manufacturer Narrative: H6: HEALTH EFFECT- CLINICAL CODE 4581: NAUSEA, FIBRIN. H6: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781744&pc=QCB | 2023826-2020-00390 |  | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -12.0/+2.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, BLURRED VISION, AND HEADACHES. THE PROBLEM IS RESOLVED. REPORTEDLY THE PATIENT IS "SEEING WELL, NO MORE HEADACHES. COMPLIANT WITH DROPS." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11200458&pc=QCB | 2023826-2021-00183 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.00/1.5/78 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED PAIN, HEADACHE, LIGHT SENSITIVITY, AND BLURRY VISION. PUPIL BLOCK WITH ELEVATED IOP WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION (YAG PI) AND MEDICAL INTERVENTION (TROPICAMIDE, PHENYLEPHRINE, COSOPT PF, DIAMOX, BROMINIDINE) WERE PERFORMED/PRESCRIBED. IT WAS REPORTED THAT THE PROBLEM RESOLVED, STATUS OF EYE "NORMAL". IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE. Manufacturer Narrative: B3 - (B)(6) 2020. B5 - EVALUATION ON 19-DEC-2020: EXAMINATION DEMONSTRATED CORNEAL EDEMA, A MID-DILATED PUPIL, ELEVATED IOP, ELEVATED ICL VAULT AND IOP. YAG PERFORMED. AQUEOUS TAP DONE (B)(6) 2020. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14145766&pc=QCB | 2023826-2022-01205 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Headache; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Eye Pain | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -11.5/+1.5/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 A PI-YAG WAS PERFORMED, THE SURGEON REPORTING TRANSIENT LOSS OF BCVA, PERIORBITAL EYE PAIN, HEADACHE, ELEVATED IOP, AND POSSIBLE INTERMITTENT BOUTS OF PUPILLARY BLOCK PRIOR TO LASER PI. THE LENS REMAINS IMPLANTED. REPORTEDLY THE TRANSIENT VISION LOSS WAS RESOLVED WITH THE LASER PI; LATENT IOP IS 12MMHG. Manufacturer Narrative: B5 LENS WAS EXPLANTED ON (B)(6) 2023. STATUS OF EYE IS REPORTED AS "PATIENT DEVELOPED LEFT EYE SECONDARY GLAUCOMA POST ICL IMPLANTATION ASSOCIATED WITH HEADACHE. CAUSE IS REPORTED AS PATIENT-RELATED FACTOR. "PATIENT COMPLAINTED HEADACHE ASSOCIATED WITH EYE PAIN." CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED EMAIL, THIS CLAIM IS FOR THE RIGHT EYE (OD). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9677189&pc=QCB | 2023826-2020-00248 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Vomiting; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_13.7, -7.50/2.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION WERE REPORTED. THE REPORTER ALSO STATED THAT "IRIS HAS BEEN DAMAGED AT THE LIMBAL REGION, WHERE YOU CAN SEE THAT THE PIGMENT LAYER HAS BEEN DISSOLVED." SECONDARY SURGICAL INTERVENTION (YAG) WAS PERFORMED ON POST-OP DAY 1 "BECAUSE THE CONDITION OF THE EYE AND PATIENT'S RAPPORT (HEADACHE AND VOMITING DEBUTING CA 8 H AFTER SURGERY) SUGGESTED THAT IOP HAD BEEN HIGH." IN THE SURGEONS OPINION THE CAUSE OF THE EVENT WAS DUE TO ICL TOO LARGE/ ACTUAL SULCUS DIAMETER SMALLER THAN EXPECTED FROM PREOP MEASUREMENTS. FOR PATIENTS' CURRENT CONDITION THE SURGEON REPORTED " PATIENT IS VERY PLEASED WITH THE OUTCOME, WEARS NO LENSES OR GLASSES AND WOULD LIKE TO KEEP EVERYTHING AS IT CURRENTLY IS; SO CLOSE TO EMMETROPIA OA. HOWEVER, THERE IS RATHER LARGE AMOUNTS OF PIGMENTATION ON ICL, IRIS ATROPHY AND THE VAULT IS STILL CA 4 OA". THE SURGEON PLANS TO REMOVE THE ICL WITH NO ICL IMPLANTATION." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT CLINICAL CODE 1937- INTRAOCULAR PRESSURE INCREASE SHOULD BE IN THE INITAL MDR. H6- INVESTIGATION CONCLUCIONS CODE 4310 SHOULD BE IN THE PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9677374&pc=QCB | 2023826-2020-00249 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Vomiting; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_13.7, -7.00/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION WERE REPORTED. SECONDARY SURGICAL INTERVENTION (YAG) WAS PERFORMED ON POST-OP DAY 1 "BECAUSE THE CONDITION OF THE EYE AND PATIENT'S RAPPORT (HEADACHE AND VOMITING DEBUTING CA 8 H AFTER SURGERY) SUGGESTED THAT IOP HAD BEEN HIGH." IN THE SURGEONS OPINION THE CAUSE OF THE EVENT WAS DUE TO ICL TOO LARGE/ACTUAL SULCUS DIAMETER SMALLER THAN EXPECTED FROM PREOP MEASUREMENTS. FOR PATIENTS' CURRENT CONDITION THE SURGEON REPORTED " PATIENT IS VERY PLEASED WITH THE OUTCOME, WEARS NO LENSES OR GLASSES AND WOULD LIKE TO KEEP EVERYTHING AS IT CURRENTLY IS; SO CLOSE TO EMMETROPIA OA. HOWEVER, THERE IS RATHER LARGE AMOUNTS OF PIGMENTATION ON ICL, IRIS ATROPHY AND THE VAULT IS STILL CA 4 OA". THE SURGEON STATED "I PLAN TO EXCHANGE THE ICL FOR ONE TWO SIZES DOWN." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT CLINICAL CODE 1937- INTRAOCULAR PRESSURE INCREASED SHOULD BE IN THE PREVIOUS MDR. H6- HEALTH EFFECT CLINICAL CODE 4851- IRIS ATROPHY SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616940&pc=QCB | 2023826-2021-00926 | 2021-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Headache; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A LENS WITH THE WRONG AXIS WAS IMPLANTED. ANOTHER LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR/UNKNOWN. Manufacturer Narrative: CORRECTED DATA: A1: PATIENT IDENTIFIER - (B)(6). B3: DATE OF EVENT - (B)(6) 2021. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND A HEADACHE. THERE WAS NO IOP ELEVATION AND NORMAL VAULT. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS STABLE AND VISION IS GOOD. THE REPORTER INDICATED THE MANIFEST REFRACTION AXIS WAS TYPED WRONG ON THE ORIGINAL ICL FORM (15 INSTEAD OF 115). D6A: (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13958504&pc=QCB | 2023826-2022-01029 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THE PATIENT IS EXPERIENCING EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED HEADACHES - SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941304&pc=QCB | 2023826-2019-01572 | 2018-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO "HEADACHE." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE OPTICE AND HAPTIC WERE TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941310&pc=QCB | 2023826-2019-01573 | 2018-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO "HEADACHE." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275294&pc=QCB | 2023826-2023-00263 | 2022-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+3.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON AUG-2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE PATIENT ALSO COMPLAINED OF PAIN AND PRESSURE SENSATION IN THE EYE, SO THE LENS WAS REMOVED. POST-OP, THE PATIENT HAS BEEN SEEING THE SURGEON EVERY 2 WEEKS, SOMETIMES COMPLAINING OF PAIN IN THE EYE, SOMETIMES OF PRESSURE AND HEADACHES. SHE HAS ALREADY VISITED THE EMERGENCY ROOM SEVERAL TIMES (WITHOUT RESULTS). H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PRESSURE SENSATION. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: UNK. PROCODE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275298&pc=QCB | 2023826-2023-00262 | 2022-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+2.5/087(SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON AUG-2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT IDENTIFIER: UNK. PROCODE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE PATIENT ALSO COMPLAINED OF PAIN AND PRESSURE SENSATION IN THE EYE, SO THE LENS WAS REMOVED. POST-OP, THE PATIENT HAS BEEN SEEING THE SURGEON EVERY 2 WEEKS, SOMETIMES COMPLAINING OF PAIN IN THE EYE, SOMETIMES OF PRESSURE AND HEADACHES. SHE HAS ALREADY VISITED THE EMERGENCY ROOM SEVERAL TIMES (WITHOUT RESULTS). H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PRESSURE SENSATION. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182322&pc=QCB | 2023826-2018-01885 | 2018-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.00/+4.0/033 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED SEVERE HEADACHE. THE REPORTER INDICATED THE RELATED SYMPTOM MAY BE DUE TO SHALLOW ANTERIOR CHAMBER (MID PERIPHERAL IRIDOCORNEAL TOUCH IN TURN CAUSING HEADACHE AND EYE PAIN). THE LENS REMAINS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT IS THE DEVICE. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE EXCHANGE LENS DID NOT RESOLVE THE ISSUE. REFERENCE MFR REPORT # 2023826-2019-00435 FOR FURTHER DETAILS ON THE EXCHANGE LENS. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND BLUE RESIDUE ON LENS SURFACE AND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070995&pc=QCB | 2023826-2022-02369 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Headache; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES AND COMPLAINED OF HEADACHES; AND EXAMINATION SHOWED NO ABNORMALITIES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070996&pc=QCB | 2023826-2022-02370 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Headache; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/1.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES AND COMPLAINED OF HEADACHES; AND EXAMINATION SHOWED NO ABNORMALITIES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986875&pc=QCB | 2023826-2023-04601 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Headache; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.0/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS, HEADACHE, AND PATIENT DISSTATISFACTION. PROBLEM RESOLVED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548122&pc=QCB | 2023826-2020-02091 | 2020-08-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Headache; Vomiting | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PRESSURE 16MMHG, OPEN ANGLE, WERE OBSERVED. HEADACHE AND VOMIT WERE OCCSSIONALLY FELT BY PATIENT. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, "SERGEON IS GOING TO REVIEW PATIENT IN 1 MONTH. NO SPECIFICT ACTION, ONLY OBSERVATION FOR NOW." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - "REPORTEDLY, THE SURGEON THINKS THE CAUSE OF THE HEADACHE AND VOMITING IS DUE TO THE PATIENT'S ADAPTATION OF THE VISION." SHOULD BE ADDED TO THE INITIAL MDR. CORRECTED DATA: B5 - "...IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020." CORRECTED TO "...IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020." IN THE INITIAL MDR. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548144&pc=QCB | 2023826-2020-02092 | 2020-08-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Headache; Vomiting | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.00/1.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. PRESSURE 16MMHG, OPEN ANGLE, WERE OBSERVED. HEADACHE AND VOMIT WERE OCCASIONALLY FELT BY PATIENT. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, "SURGEON IS GOING TO REVIEW PATIENT IN 1 MONTH. NO SPECIFIC ACTION, ONLY OBSERVATION FOR NOW." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5 - "REPORTEDLY, THE SURGEON THINKS THE CAUSE OF THE HEADACHE AND VOMITING IS DUE TO THE PATIENT'S ADAPTATION OF THE VISION." SHOULD BE ADDED TO THE INITIAL MDR. B5 - "IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2020." CORRECTED TO "IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2020." IN THE INITIAL MDR. CLAIM#: 717258 Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216366&pc=QCB | 2023826-2023-02429 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Hyphema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -9.5/1.0/011 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED VERTICALLY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. HYPHEMA WAS OBSERVED. IRRIGATION WAS PERFORMED. THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH IMPACT CODE: "4625- ANTERIOR CHAMBER IRRIGATION" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON VERTICALLY IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/11 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED HYPHAEMA. ANTERIOR CHAMBER IRRIGATION WAS PERFORMED BUT THAT DID NOT RESOLVE THE PROBLEM. REPORTEDLY "IRRIGATED THE HYPHEMA, AND THERE WERE NO ADDITIONAL INTERVENTIONS. THE CAUSE OF THE BLEEDING WAS COULDN'T BE DETERMINED, SO THE ICL LENS WERE REMOVED". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17870418&pc=QCB | 2023826-2023-04414 | 2023-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Hyphema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/000 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE REPORTER REPORTS A NON TRAUMATIC HYPHEMA THREE MONTHS POST OP. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FUIDS WAS PERFORMED AND THE PROBLEM RESOLVED. REPORTER STATES " ALL FINE, PATIENT IS HAPPY". CAUSE OF EVENT WAS UNKNOWN. . Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15901047&pc=QCB | 2023826-2022-04233 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO CHRONIC INFLAMMATION. THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT DECIDED NOT TO RE-IMPLANT ICL." Manufacturer Narrative: CORRECTED DATA: B5-25-OCT-2022 SHOULD BE CORRECTED TO 26-OCT-2022. D6B: EXPLANT DATE SHOULD BE CORRECTED TO 26-OCT-2022. H4-DEVICE MANUFACTURE DATE:SHOULD BE CORRECTED TO 21-JUN-2022. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15902174&pc=QCB | 2023826-2022-04232 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/74 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO CHRONIC INFLAMMATION. THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT DECIDED NOT TO RE-IMPLANT ICL." Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16479532&pc=QCB | 2023826-2023-00726 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). PATIENT DEVELOPED UNCONTROLLED INFLAMMATION AND THE LENS WAS EXPLANTED. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18183258&pc=QCB | 2023826-2023-05185 | 2023-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -6.50/2.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REBOUND POST-OP INFLAMMATION WITH MID CORNEAL STROMAL INFLAMMATION. ADDITIONAL POSTOPERATIVE STEROID DROPS WERE GIVEN. LENS REMAINS IMPLANTED. CONDITIONS ARE IMPROVING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-HEALTH IMPACT:4644: STEROID DROPS. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18183259&pc=QCB | 2023826-2023-05186 | 2023-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. REBOUND POST-OP INFLAMMATION WITH MID CORNEAL STROMAL INFLAMMATION. ADDITIONAL POSTOPERATIVE STEROID DROPS WERE GIVEN. LENS REMAINS IMPLANTED. CONDITIONS ARE IMPROVING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-HEALTH IMPACT:4644: STEROID DROPS. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12920224&pc=QCB | 2023826-2021-04630 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS), MARKED INFLAMMATION, AC ++++, WAS NOTED 6 WEEKS POST OP. MEDICAL INTERVENTION ( PILOCARPINE) WAS PRESCRIBED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED INFLAMMATION/ IRITIS; MEDICATION WAS PRESCRIBED; PUPIL ANOMALY AND PIGMENTS IN THE ANTERIOR CHAMBER. THE LENS REMAINS IMPLANTED. REPORTEDLY "PATIENT HAS NO SYMPTOMS AND 1 MONTH POST OP AFTER REPOSITIONING THE LENS I STILL SEE SOME PIGMENTS IN THE ANTERIOR CHAMBER. BUT NO INTERVENTIONS ARE PLANNED FOR NOW. I WILL JUST MONITOR THE IOP AND GIVE TOPICAL ANTI GLAUCOMA MEDS IF IOP INCREASES". H6- HEALTH EFFECT- CLINICAL CODE: 4581- PUPIL ANOMALY. H6- HEALTH EFFECT- CLINICAL CODE: 4581- PIGMENT IN ANTERIOR CHAMBER. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE OF EVENT: UNKNOWN. (EXPIRATION DATE): UNKNOWN, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): NO SERIAL NUMBERS REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15351312&pc=QCB | 2023826-2022-03010 | 2022-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF 11.0/2.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. INFLAMMATION WAS OBSERVED. IT WAS REPORTED THAT ANTERIOR CHAMBER IRRIGATION AND MEDICAL INTERVENTION WERE PERFORMED/PRESCRIBED. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO INFLAMMATION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- ANTERIOR CHAMBER IRRIGATION CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 11/30/2024. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465780&pc=QCB | 2023826-2022-03248 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT HIS PATIENT HAS INFLAMMAITON AND PUPIL HAS GOTTEN BIGGER AND NON-REACTIVE IN SPITE OF GOOD VISION AND GOOD ANATOMICAL PLACEMENT OF THE ICL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14649917&pc=QCB | 2023826-2022-01840 | 2022-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Inflammation; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -8.00/+1.50/76 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT HAD INFLAMMATION BEHIND THE ICL; VA POSTOP WAS 6/6 BUT BLURRY; AND IS CURRENTLY ON STEROIDS. PATIENT IS STILL BEING MONITORED. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8642522&pc=QCB | 2023826-2019-00911 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Inflammation; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS, -13.5/+3.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), WITHOUT ANY COMPLICATIONS. ELEVEN MONTHS AFTER THE LENS WAS IMPLANTED, THE PATIENT EXPERIENCED BLURRED VISION AFTER BLUNT OCULAR TRAUMA. EXAMINATION REVEALED MILD INFLAMMATION IN THE ANTERIOR CHAMBER AND BOTH SUPERIOR AND INFERIOR TEMPORAL HAPTICS OF THE ICL WERE DISLODGED INTO THE ANTERIOR CHAMBER AND ENTRAPPED WITHIN THE PUPIL, WITHOUT TOUCHING THE CORNEAL ENDOTHELIUM. THE LENS WAS REPOSITIONED AND THE PATIENT WAS ON MEDICATION. ONE MONTH AFTER THE REPOSITIONING SURGERY, THE BLURRED VISION WAS RESOLVED AND THERE WAS NO DAMAGE TO THE EYE. TWENTY SIX MONTHS LATER, THE PATIENT AGAIN EXPERIENCED BLUNT OCULAR TRAUMA TO THE RIGHT EYE. THE LENS WAS REPOSITIONED AND ONE MONTH AFTER THERE WERE NO COMPLICATIONS ASSOCIATED TO THE REPOSITIONING SURGERY AND NO TRAUMA OBSERVED. PIGMENT DEPOSITION ON THE POSTERIOR SURFACE OF THE ICL WAS OBSERVED AFTER SURGERY BUT THE AMOUNT OF THE PIGMENT HAD NOT CHANGED ONE YEAR AFTER THE INJURY. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: EXCESSIVE VAULTING SHOULD BE ADDED TO THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9173237&pc=QCB | 2023826-2019-01901 | 2019-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. IT WAS REPORTED THE PATIENT HAD BLURRED VISION AND TASS (TOXIC ANTERIOR SEGMENT SYNDROME). THE SURGEON PERFORMED AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE FOLLOWING MEDICATIONS WERE PRESCRIBED FOR ABOUT 10 DAYS AFTER SURGERY - RINDERON OPHTHALMIC SOLUTION, MYDRIN-P OPHTHALMIC SOLUTION AND PREDNISOLONE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THE INFLAMMATION HAD SUBSIDED AFTER 10 DAYS OF TREATMENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16981458&pc=QCB | 2023826-2023-01954 | 2023-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Optical Problem | Inflammation; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTES THE LENS HAS SOME DIFFUSE WHITISH PLAQUES/FILM ON THE LENS, WHICH LOOKS LIKE IT'S ON THE POSTERIOR SURFACE OF THE EVO LENS AND NOT THE LENS OR CORNEA. THE SURGEON ALSO INDICATED THERE IS A BIT MORE INFLAMMATION THEN TYPICAL, THE WHITISH FILM/PLAQUES APPEAR STABLE, NEITHER INCREASING OR DECREASING AND THERE SEEMS TO BE NO CONTINUED INFLAMMATION. THE SURGEON ALSO REPORTS GLARE/HALOES, BLURRED VISION, AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY, A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT MODEL/POWER WAS IMPLANTED. IT WAS NOTED "VISION IMPROVED SIGNIFICANTLY WITH EXCHANGE". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "SUSPECT THIS WAS INFLAMMATORY DEBRIS/FIBRING BUILDING UP ON THE LENS". Manufacturer Narrative: ADDITIONAL INFORMATION: INITIAL MDR IS NOT APPLICABLE AS THERE WAS NO DEVICE PROBLEM, REPORTER LATER CLARIFIED THE ISSUE WAS DIRECTLY PATIENT RELATED. CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - WHITISH PLAQUE/FILM. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704497&pc=QCB | 2023826-2023-03871 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. SHALLOW ANTERIOR CHAMBER (AC) AND HIGH VAULT WAS NOTED. THE LENS REMAINED IMPLANTED. A COURSE OF STEROIDS WAS ADMINISTERED. POST-OP, THE PATIENTS OCULAR DISCOMFORT HAS MOSTLY SUBSIDED ALREADY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. HEALTH EFFECT IMPACT CODE: 4644 - STEROIDS WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715559&pc=QCB | 2023826-2023-03872 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. SHALLOW ANTERIOR CHAMBER (AC) AND EXCESSIVE VAULTING WAS NOTED. THERE WAS PIGMENT ON ICL FROM PUPIL MARGIN, SO THERE MAY BE IRIS RUBBING. THE PATIENT COMPLAINED OF ACHE FROM TIME TO TIME. THE LENS REMAINED IMPLANTED. A COURSE OF STEROIDS WAS ADMINISTERED. POST-OP, THE PATIENTS OCULAR DISCOMFORT HAS MOSTLY SUBSIDED ALREADY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6B: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - IRIS RUBBING; ACHE. H6: HEALTH EFFECT IMPACT CODE: 4644 - STEROIDS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147542&pc=QCB | 2023826-2021-02274 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Inflammation; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 13.2MM TMICL13.2, -11.50/3.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S EYE ON (B)(6) 2021. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ELEVATED IOP, INFLAMMATION AND PIGMENT DEBRIS OBSERVED. THE LENS WAS REMOVED AND CATARACT SURGERY WAS PERFORMED AND A NEW LENS AS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE ICL WAS IMPLANTED UPSIDE DOWN. THE PATIENT ALSO DISPLAYED A TEAR IN THE ANTERIOR CAPSULE AND CATARACT FORMATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14058808&pc=QCB | 2023826-2022-01151 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/+1.5/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. ON (B)(6) 2022 SECONDARY YAG WAS PERFORMED. PER UPDATED INFORMATION AS OF (B)(6) 2022 INTRAOCULAR PRESSURE WAS 42MMHG, PATIENT IS BEING MONITORED WITH MEDICATION. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2022. THE REPORTER STATED, "PERSISTENT INFLAMMATION, HIGH IOP, VAULTING ½ CORNEA THICKNESS. EXPLANTED TO CONTROL INFLAMMATION. AND TO BETTER MANAGE IOP". CLAIM (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14062269&pc=QCB | 2023826-2022-01150 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.1, -8.00/+1.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, INFLAMMATION AND PIGMENT DISPERSION. MEDICAL INTERVENTION WAS PRESCRIBED, IT WAS REPORTED THAT THE CONDITION OF THE PATIENT IS IMPROVING. PATIENT IS BEING MONITORED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741439&pc=QCB | 2023826-2023-03979 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND PIGMENT DISPERSION. PATIENT DEVELOPED CHRONIC INFLAMMATION WITH HIGH IOP. MEDICATION WAS ADMINISTERED - ANTIGLAUCOMA; COMBIGAN; AZOPT; XALATAN. POST-OP, THE ANTERIOR CHAMBER HAD MILD INFLAMMATION WITH PIGMENT CELL. THE PATIENT STILL CONTINUES WITH 2 TYPES OF ANTIGLAUCOMA - TIMOLOL BD AND AZOPT BD. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION. H6: HEALTH EFFECT IMPACT CODE: 4644 - ANTIGLAUCOMA MEDS - COMBIGAN, AZOPT, XALATAN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549694&pc=QCB | 2023826-2023-03497 | 2023-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Inflammation; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A,12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -18.0/1.00/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, AND BLURRED VISION. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "UP TO DATE THE EYE CONTINUES WITH THE INFLAMMATORY PROCESS." CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. THE REPORTER STATED, "DUE TO A PATIENT INFLAMMATORY RESPONSE." Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984492&pc=QCB | 2023826-2019-01675 | 2018-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Iritis; Red Eye(s); No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/059 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED 2 WEEKS AFTER THE INITIAL SURGERY DUE TO POSSIBLE INFLAMMATION. ELEVATED INTRAOCULAR PRESSURE WAS ALSO REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS HAD LOW VAULTING AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE PATIENT'S LEFT EYE (OS) WAS RED AND HE CAME TO THE HOSPITAL FOR DIAGNOSIS ON (B)(6) 2018. THE DIAGNOSIS WAS IRITIS. THE PATIENT WAS TREATED WITH ANTI-INFLAMMATION AND INTRAOCULAR PRESSURE LOWERING MEDS. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE PATIENT'S CONDITION IS STABLE. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15336964&pc=QCB | 2023826-2022-02893 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | Inflammation; Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. NARROW ANGLES AND IRIS ELEVATED REPORTED DUE TO LENS IMPLANTED UPSIDE DOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED UPSIDE DOWN. ON (B)(6) 2022 THE LENS WAS REMOVED AND THE SAME LENS WAS RE-IMPLANTED IN THE CORRECT ORIENTATION. AFTER THE LENS WAS RE-IMPLANTED;. SUBSEQUENTLY, IOP INCREASED, INFLAMMATION-LIKE REACTION, MYDRIASIS. SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. PUPIL BLOCK WITH ELEVATED IOP NOT MEASURABLE. ANGLE CLOSURE WITH ELEVATED IOP NOT MEASURABLE. ANTERIOR CHAMBER CONTENTS PRESENT, 8/6 IOP DIFFICULT TO MEASURE. THE STATUS OF THE PATIENT'S EYE WAS REPORTED AS " ON AUGUST 6, IOP ROSE (UNMEASURABLE) AND ANTERIOR CHAMBER CONTENTS APPEARED, SO THE ANTERIOR CHAMBER WAS RINSED. WITH EYE DROPS AND ORAL MEDICATION, THE IOP DECREASED TO 10 ON (B)(6), BUT THERE WAS RESIDUAL CONTENT, POOR PUPIL SHAPE, AND CORNEAL EDEMA. STEROID EYE DROPS WERE CONTINUED AND THE PATIENT'S SYMPTOMS IMPROVED. CURRENTLY, SHE CONTINUES TO USE STEROIDS AND PILOCARPINE IN THE LEFT EYE. PUPIL SHAPE HAS IMPROVED, BUT THE LEFT EYE TENDS TO CONTRACT, AND THERE IS CURRENTLY A DIFFERENCE IN PUPIL DIAMETER BETWEEN THE LEFT AND RIGHT EYES. THE PROBLEM WAS RESOLVED. CAUSE IS USER ERROR; AND DEVICE DID NOT FAIL TO PERFORM AS INTENDED. D6B- (B)(6) 2022. H6- CLINICAL CODE: 4581: SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ANGLE CLOSURE. ANTERIOR CHAMBER CONTENTS PRESENT. POOR PUPIL SHAPE. DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022." D6A - CORRECTED TO: IMPLANT DATE (B)(6) 2022 . CLAIM # (B)(4). Manufacturer Narrative: EXPIRATION DATE: 30-APR-2025, CLINICAL SIGNS CODE (B)(4) - NARROW ANGLES; IRIS ELEVATED. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12833486&pc=QCB | 2023826-2021-04292 | 2021-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Inflammation; Iritis; Uveitis | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2021. AT 6 MONTHS POST-OP, THE LENS WAS REPORTED AS HAVING LOW VAULTING WHICH SUBSEQUENTLY CAUSED INFLAMMATION AND ACUTE ANTERIOR UVEITIS/ IRITIS HAD DEVELOPED. ONE DAY AFTER TAKING INTENSIVE TOPICAL STEROIDS, THE PATIENTS SYMPTOMS HAD IMPROVED DRAMATICALLY. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). Manufacturer Narrative: THE LENS WAS A -12.00/+2.0/096 DIOPTER. THE PATIENT WAS REPORTED AS DOING WELL AT LAST POST-OP VISIT. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S RIGHT EYE (OD), IN (B)(6) 2021. AT 6 MONTHS POST-OP, THE PATIENT EXPERIENCED INFLAMMATION AND ACUTE ANTERIOR UVEITIS/IRITIS. ONE DAY AFTER TAKING INTENSIVE TOPICAL STEROIDS, THE PATIENTS SYMPTOMS HAD IMPROVED DRAMATICALLY. THE LENS REMAINED IMPLANTED. H6: MEDICAL DEVICE PROBLEM CODE: 2993. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840726&pc=QCB | 2023826-2021-04293 | 2021-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Iritis; Uveitis | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2021. AT 6 MONTHS POST-OP, THE LENS WAS REPORTED AS HAVING LOW VAULTING WHICH SUBSEQUENTLY CAUSED INFLAMMATION AND ACUTE ANTERIOR UVEITIS/IRITIS HAD DEVELOPED. ONE DAY AFTER TAKING INTENSIVE TOPICAL STEROIDS, THE PATIENTS SYMPTOMS HAD IMPROVED DRAMATICALLY. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THE LENS WAS A 12.50/+1.0/068 DIOPTER. THE PATIENT WAS REPORTED AS DOING WELL AT LAST POST-OP VISIT. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPIRATION DATE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763823&pc=QCB | 2023826-2021-04230 | 2021-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Iritis; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.5/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, INFLAMMATION AND IRITIS. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT INITIALLY THE DOCTOR WANTED TO EXPLANT THE LENS BUT HAS DECIDED TO WAIT AWHILE. MEDICAL INTERVENTION (PRED FORTE, PILOCARPINE, PREDNISONE) WAS PRESCRIBED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- CORRECTION: -11.50/2.5/100 SHOULD BE CORRECTED TO -11.50/2.5/108 IN THE ORIGINAL MEDWATCH REPORT. CLAIM: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512604&pc=QCB | 2023826-2019-02425 | 2019-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Inflammation; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.0/+2.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS DEPOSITS ON THE POSTERIOR SURFACE OF THE CORNEA (INFLAMMATION) AND PIGMENT DISPERSION. THE PATIENT WAS GIVEN ADDITIONAL ANTIBIOTIC, NSAID, STEROID, AND PUPIL DILATOR MEDICATIONS TO TREAT THIS ISSUE. REPORTEDLY, LENS REMAINS IMPLANTED AND THE DEPOSITS HAVE SUBSIDED. "SMALL PIGMENT DISPERSION ON THE ICL+" IS ALSO STATED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (DEPOSITS ON POSTER SURFACE OF CORNEA) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512615&pc=QCB | 2023826-2019-02424 | 2019-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Inflammation; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.5/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS DEPOSITS ON THE POSTERIOR SURFACE OF THE CORNEA (INFLAMMATION), PIGMENT DISPERSION, AND LENS ROTATION. THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND (B)(6) 2019. THE PATIENT WAS GIVEN ADDITIONAL ANTIBIOTIC, NSAID, STEROID, AND PUPIL DILATOR MEDICATIONS TO TREAT THIS ISSUE. REPORTEDLY, LENS REMAINS IMPLANTED AND THE DEPOSITS HAVE SUBSIDED. "SMALL PIGMENT DISPERSION ON THE ICL+" IS ALSO STATED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (DEPOSITS ON POSTER SURFACE OF CORNEA) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136620&pc=QCB | 2023826-2023-05112 | 2023-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.50/+1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED. IT WAS NOTED THAT "OD PAIN EYELID EDEMA - INFLAMMATION DESPITE ANTI-INFLAMMATORY EYE DROPS FOR 3 DAYS - TOBRADEX = DEXAMETHASONE AND TOBRAMYCINE. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - EYELID EDEMA. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017451&pc=QCB | 2023826-2023-04787 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Inflammation; Red Eye(s) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.50/2.0/045 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. PIGMENT DISPERSION AND ANTERIOR CHAMBER INFLAMMATION WAS OBSERVED. ORAL STEROIDS AND TOPICAL STEROIDS WERE USED TO REDUCE THE INFLAMMATION. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.50/2.0/045 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2023. PIGMENT DISPERSION AND ANTERIOR CHAMBER INFLAMMATION WAS OBSERVED. ORAL STEROIDS AND TOPICAL STEROIDS WERE USED TO REDUCE THE INFLAMMATION. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: H6-HEALTH IMPACT:4644- ORAL STEROID AND TOPICAL STEROID EYE DROPS. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017491&pc=QCB | 2023826-2023-04788 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Inflammation; Red Eye(s) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -12.50/2.5/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. PIGMENT DISPERSION AND ANTERIOR CHAMBER INFLAMMATION WAS OBSERVED. ORAL STEROIDS AND TOPICAL STEROIDS WERE USED TO REDUCE THE INFLAMMATION. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -12.50/2.5/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2023. PIGMENT DISPERSION AND ANTERIOR CHAMBER INFLAMMATION WAS OBSERVED. ORAL STEROIDS AND TOPICAL STEROIDS WERE USED TO REDUCE THE INFLAMMATION. ON (B)(6) 2023, THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: H6-HEALTH IMPACT:4644- ORAL STEROID AND TOPICAL STEROID EYE DROPS. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18275479&pc=QCB | 2023826-2023-05522 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Red Eye(s) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -18.00/+4.00/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THAT THE PATIENT HAD LOW VAULTING, CELL +3 INFLAMMATION AND RED EYE. THE SURGEON ALSO REPORTED THE LENS WAS EXCHANGED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387531&pc=QCB | 2023826-2023-05725 | 2023-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. AT 5 DAYS POST-OP, TOXIC ANTERIOR SEGMENT SYNDROME (TASS) (FIBRIN AND 3+ CELL) WAS OBSERVED. TAP FOR CULTURE WAS PERFORMED AND WAS NEGATIVE. PREDNISOLONE, MOXIFLOXACIN INJECTION AND BROMFENAC COMPOUND DROPS WERE ADMINISTERED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED (A/C WASHOUT). PATIENT WAS STARTED ON MEDROL DOSE PACK. POST-OP THE PATIENT HAS EXCELLENT VISION AND IOP, WITH ONLY TRACE RESIDUAL INFLAMMATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK D6B: UNK H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS (A/C WASHOUT). H6: HEALTH EFFECT IMPACT CODE: 4644 - PREDNISOLONE, MOXIFLOXACIN, BROMFENAC COMPOUND DROPS, MEDROL DOSE PACK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465648&pc=QCB | 2023826-2022-03247 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. PIGMENT DISPERSION, UNEACTIVE (FIXED) PUPIL, ANTERIOR SEGMENT INFLAMMATION, AND GLARE/HALOES WAS OBSERVED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION IS NOT NECESSARY. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "MILD CELL/FLARE. 6.0MM UNREACTIVE PUPIL." UNKNOWN WAS REPORTED FOR CAUSE OF EVENT AND THE REPORTER STATED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5: PER UPDATED INFORMATION PATIENT IS SEEING WELL BUT HAS A DILATED PUPIL AND HAS BEEN USING STEROID EYE DROPS TO CONTROL INFLAMMATION. INFLAMMATION HAS BEEN IMPROVING. SURGEON WILL BE TAPERING STEROID EYE DROPS WITH NO PLANS OF SURGICAL INTERVENTION AT THE MOMENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659759&pc=QCB | 2023826-2020-02349 |  | Malfunction | STAAR SUGICAL COMPANY | 2020-10-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Insufficient Information | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -5.50/1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS MAY BE REMOVED AND REPLACED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659773&pc=QCB | 2023826-2020-02348 |  | Malfunction | STAAR SUGICAL COMPANY | 2020-10-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Insufficient Information | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -5.50/1.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS MAY BE REMOVED AND REPLACED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10817112&pc=QCB | 2023826-2020-02641 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Insufficient Information | NA | 1 | Event Description: A DAMAGED LENS WAS RETURNED BACK TO US. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION IN PREVIOUS MDR IS CORRECTED TO YES; H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE, WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND LENS DAMAGED. CLAIM#(B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. CLAIM #:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685382&pc=QCB | 2023826-2021-04091 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated; No Apparent Adverse Event | Insufficient Information | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THIS LENS WAS IMPLANTED AS A REPLACEMENT LENS AND THE PROBLEM WAS NOT RESOLVED. SEE MFR REP# 2023826-2021-03958 FOR INITIAL LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, IN LIQUID WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN WITH FIBER ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. DATE OF EVET- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16820477&pc=QCB | 2023826-2023-01524 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS ROTATED/REPOSITIONED AND THE ISSUE RESOLVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10759202&pc=QCB | 2023826-2020-02535 |  | Injury | STAAR SUGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: IT WAS REPORTED THAT A TORIC ICL PATIENT EXPERIENCED LENS ROTATION. SURGEON ATTEMPTED TO ROTATE BACK INTO POSITION BUT LENS ROTATED AGAIN. DOCTOR PLANS TO EXPLANT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5:THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -8..00/1.5/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PATIENT EXPERIENCED LENS ROTATION. SURGEON ATTEMPTED TO ROTATE LENS BACK INTO POSITION, BUT LENS ROTATED AGAIN. LENS WAS EXPLANTED. NO PATIENT INJURY WAS REPORTED. ADDITIONAL INFORMATION BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -8.0/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM # (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10759221&pc=QCB | 2023826-2020-02534 |  | Injury | STAAR SUGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: IT WAS REPORTED THAT A TORIC ICL PATIENT EXPERIENCED LENS ROTATION. SURGEON ATTEMPTED TO ROTATE BACK INTO POSITION BUT LENS ROTATED AGAIN. DOCTOR PLANS TO EXPLANT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -8.00/2.0/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. PATIENT EXPERIENCED LENS ROTATION. SURGEON ATTEMPTED TO ROTATE LENS BACK INTO POSITION, BUT LENS ROTATED AGAIN. LENS WAS EXPLANTED. NO PATIENT INJURY WAS REPORTED. ADDITIONAL INFORMATION BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -8.0/2.0/96 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION : LENS WAS RETURNED DRY WITH RESIDUE, IN SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. RESIDUE ON LENS WAS OBSERVED. (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10556172&pc=QCB | 2023826-2020-02123 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Decreased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE SURGEON REPORTS A HIGH VAULT OF 1006UM AND AN IOP<5MMHG. HOWEVER, BECAUSE OF HYPOTONY AND IRIS INCISION AND AFTER CORNEAL SUTURE, THE VAULT IS 720UM. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD- BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK WORD(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." SHOULD HAVE BEEN IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/080(SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT, IRIS PROLAPSE AND ADDITION OF SUTURE, GLARE AND HALOES WERE OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2020. A LENS OF A SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WERE "NO COMPLAINTS" FOR THE STATUS OF THE EYE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10619462&pc=QCB | 2023826-2020-02273 | 2020-09-01 | Malfunction | STAAR SUGICAL COMPANY | 2020-10-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.0/2.0/73 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. IT WAS REPORTER THAT THE PROBLEM RESOLVED, REPORTER STATED " IOP WENT DOWN. LENS REMAINS IN THE EYE. PATIENT IS HAPPY." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10619517&pc=QCB | 2023826-2020-02272 | 2020-09-01 | Malfunction | STAAR SUGICAL COMPANY | 2020-10-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/2.0/102 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. IT WAS REPORTER THAT THE PROBLEM RESOLVED, REPORTER STATED " IOP WENT DOWN. LENS REMAINS IN THE EYE. PATIENT IS HAPPY." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138575&pc=QCB | 2023826-2020-01183 | 2020-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, -3.00/2.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ELEVATED IOP (40 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR. PER THE REPORTER ON (B)(6) 2020, "PREDISPOSITION TO SUFFER FROM OPEN ANGLE GLAUCOMA. THE ANGLE IS CONFIRMED OPEN BY OCT OF THE ANTERIOR POLE. CURRENTLY INTRAOCULAR PRESSURE IS CONTROLLED BY AZARGA AND MONOPROST." Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - VTICH12.6 DIOPTER CORRECTED TO "+3.00/2.5/082" IN INITIAL MDR. H6 - DEVICE CODE: 1494, OFF-LABEL USE, ACD < 3.0MM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515444&pc=QCB | 2023826-2020-02020 | 2020-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.50/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO ELEVATED IOP (40MMHG) WITHOUR PUPIL BLOCK AND ANGLE CLOSURE OBSERVED ON (B)(6) 2020, AS WELL AS EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537754&pc=QCB | 2023826-2020-01978 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS LOST AND WILL NOT BE RETURNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537776&pc=QCB | 2023826-2020-01977 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/076 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND WITH RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10538983&pc=QCB | 2023826-2020-02161 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/2.00/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10539564&pc=QCB | 2023826-2020-02162 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/2.00/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10601013&pc=QCB | 2023826-2020-02236 | 2020-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.50/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (18.5 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE BEING OBSERVED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "BIGGER SIZE." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635478&pc=QCB | 2023826-2020-02299 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.00/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. RESIDUAL REFTRACTIVE ERROR WAS OBSERVED AND ROTATION. LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, "IT DOES NOT SEEM LIKE THE LENS ROTATED FAR FROM PLANNED ORIENTATION, AND FURTHER ROTATION IS NOT ADVISED." PATIENT WILL BE UNDER OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635480&pc=QCB | 2023826-2020-02298 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.00/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. RESIDUAL REFTRACTIVE ERROR WAS OBSERVED AND ROTATION. LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, "IT DOES NOT SEEM LIKE THE LENS ROTATED FAR FROM PLANNED ORIENTATION, AND FURTHER ROTATION IS NOT ADVISED." PATIENT WILL BE UNDER OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10701086&pc=QCB | 2023826-2020-02428 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. "HIGH IOP WAS OBSERVED 2 HOURS AFTER SURGERY. SURGEON TREATED BY RELEASING SOME VISCOELASTIC THROUGH THE PARACENTESIS AT THE SLIT LAMP AND GIVING MEDICATION." LENS REMAINS IMPLANTED. PER REPORTER ON (B)(6) 2020, SURGEON WAS NOT CONCERNED AND NOT LISTING THIS AS A PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10744517&pc=QCB | 2023826-2020-02532 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND THE PATIENT EXPERIENCED INCREASED IOP (INTRAOCULAR PRESSURE) THAT REQUIRED A CHANGE IN STANDARD POSTOPERATIVE MEDICATION REGIMEN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10744529&pc=QCB | 2023826-2020-02531 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND THE PATIENT EXPERIENCED INCREASED IOP (INTRAOCULAR PRESSURE). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060109&pc=QCB | 2023826-2020-03151 | 2020-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT CYLINDER AND AXIS LENS DUE TO ELEVATED IOP. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11013551&pc=QCB | 2023826-2020-02984 | 2020-09-28 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. A WEEK AFTER LENS WAS IMPLANTED, THERE WAS AN INTRAOCULAR PRESSURE SPIKE AND PATIENT WAS PRESCRIBED ALPHAGAN DROPS. TO DATE THE MEDICATION HAS BEEN TAPERED OFF AND LENS REMAINED IMPLANTED. THE PATIENT IS HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11013716&pc=QCB | 2023826-2020-02983 | 2020-09-28 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. A WEEK AFTER LENS WAS IMPLANTED, THERE WAS AN INTRAOCULAR PRESSURE SPIKE AND PATIENT WAS PRESCRIBED ALPHAGAN DROPS. TO DATE THE MEDICATION HAS BEEN TAPERED OFF AND LENS REMAINED IMPLANTED. THE PATIENT IS HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851811&pc=QCB | 2023826-2020-02665 | 2020-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.50/4.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ELEVATED IOP (35MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2020, THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. PER THE REPORTER ON (B)(6) 2020, SURGEON SAID PATIENT MIGHT HAVE HAD SYNECHIAE RETRO-IRIS, THEREFORE LENS DID NOT SIT PROPERLY. "SITUATION CALMED DOWN IMMEDIATELY AFTER THE EXPLANT, IOP BECAME NORMAL AND EXCEPT FEW DAYS AFTER EXPLANT SURGERY, CORNEA BECAME ALSO PERFECTLY TRANSPARENT. NO SIGNIFICANT EC LOS, NO SECONDARY CATARACT." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON PRODUCT LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(6) 2020 SHOULD BE CORRECTED TO (B)(6) 2020 IN THE PREVIOUS MDR. "LENS IS CURVED." SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135907&pc=QCB | 2023826-2021-00040 | 2020-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT AND ELEVATED IOP, ENHANCED PIS AND ADMINISTRATION OF IOP DROPS AND DIAMOX. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN A PLASTIC BAG WITH LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -15.5/+4.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON 2020 (B)(6). THE SURGEON REPORTS ELEVATED IOP, ENHANCED PIS AND ADMINISTRATION OF IOP DROPS AND DIAMOX. ON 2021 (B)(6) THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11169226&pc=QCB | 2023826-2021-00069 | 2020-03-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM, TICML12.1, 14.00/3.0/086, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD). ACCOMMODATIVE SPASMS, REFRACTIVE CHANGE OVERTIME, AND FOGGY VISION WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "AT (B)(6) 2020 VISIT, PATIENT VOICED COMPLAINT OF BLURRY VISION ONLY AFTER WORK WHICH WOULD PERSIST FOR MINUTES TO HOURS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM TICML12.1 IMPLANTABLE COLLAMER OF DIOPTER -14.00/3.0/086 (SPHERE/CYLINDER/AXIS) LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ACCOMMODATIVE SPASMS, REFRACTIVE CHANGE OVERTIME, AND FOGGY VISION WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY ON (B)(6) 2020, "THE PATIENT VOICED A COMPLAINT OF BLURRY VISION ONLY AFTER WORK WHICH WOULD PERSIST FOR MINUTES TO HOURS." THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11175631&pc=QCB | 2023826-2021-00141 | 2020-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.00/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE " LENS TOO LARGE". Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206965&pc=QCB | 2023826-2021-00109 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/2.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (29 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS." Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/2.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULT AND ELEVATED IOP (29 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS." MANUFACTURER NARRATIVE: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206967&pc=QCB | 2023826-2021-00110 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.00/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (22 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS." Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.00/2.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULT AND ELEVATED IOP (22 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS." MANUFACTURER NARRATIVE: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405310&pc=QCB | 2023826-2021-00547 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A HIGH PRESSURE SPIKE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405316&pc=QCB | 2023826-2021-00548 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.50/+2.0/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A HIGH PRESSURE SPIKE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11467925&pc=QCB | 2023826-2021-00630 | 2021-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -15.50/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. "LENGTH OF IMPLANT" WAS REPORTED AS CAUSE OF EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11467943&pc=QCB | 2023826-2021-00631 | 2021-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -14.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. "LENGTH OF IMPLANT" WAS REPORTED AS CAUSE OF EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -14.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. "LENGTH OF IMPLANT" WAS REPORTER AS CAUSE OF EVENT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11548140&pc=QCB | 2023826-2021-00736 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S EYE IS STABLE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE - THE LENS WAS TOO LARGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. EXPIRATION DATE: UNKNOWN. DEVICE MANUFACTURE DATE : UNKNOWN. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654112&pc=QCB | 2023826-2021-00967 | 2021-02-12 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "OVER SIZE LENS." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721500&pc=QCB | 2023826-2021-01160 | 2020-10-13 | Injury | STAAR SURGICAL COMPANY | 2021-04-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # ((B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721925&pc=QCB | 2023826-2021-01038 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2, -12/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. A NEW PI AND YAG WAS PERFORMED ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP(34MMHG ASSESSED ON (B)(6) 2021) ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE, "PATIENTS SULCUS IS AN ANATOMICAL VARIANET FROM THE NORMAL WHICH RESULTED IN HIGHER VAULT THAN EXPECTED WITH THE FDA NOMOGRAM." Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2, -12/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. A NEW PI AND YAG WAS PERFORMED ON (B)(6) 2021.THE LENS WAS EXPLANTED ON (B)(6) DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP(34MMHG ASSESSED ON (B)(6) 2021. A SHORTER LENGTH LENS WAS IMPLANTED (B)(6) 2021, IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE, "PATIENTS SULCUS IS AN ANATOMICAL VARIANET FROM THE NORMAL WHICH RESULTED IN HIGHER VAULT THAN EXPECTED WITH THE FDA NOMOGRAM." CORRECT INFORMATION: H6 - HEALTH EFFECT-IMPACT CODE - REPORTED AS 4629 - CORRECT TO 4627. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE PUPIL ISSUE WAS A RESULT OF THE INITIAL SURGERY NOT THE REPLACEMENT. THE PATIENT IS DOING WELL WITH ALL OTHER SIDE EFFECTS BESIDES THE PUPIL RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11822933&pc=QCB | 2023826-2021-01437 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VTICMO13.7, -14.50/0.50/146 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. A SHORTER LENGTH LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE OTHER: " WRONG LENGTH OF LENS. YES BECAUSE THE INITAL LENS WAS TOO LONG AND THE VAULT TOO HIGH". Manufacturer Narrative: (B)(6). PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12018124&pc=QCB | 2023826-2021-01921 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -14.5/2/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED IOP. ADDITIONAL MEDICAL INTERVENTION (BROMINIDINE AND TIMOLOL) WAS NECESSARY. IT WAS REPORTED THAT THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS CAUSED BY BOTH PROCEDURE AND DEVICE. Manufacturer Narrative: B5- PER UPDATED INFORMATION "CAUSED BY BOTH PROCEDURE AND DEVICE. INCREASED IOP SAME DAY POSTOPERATIVE VISIT DUE TO RETAINED VISCOELASTIC." CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12058246&pc=QCB | 2023826-2021-02040 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -13.00/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT EXPERIENCED ELEVATED IOP. REPORTER STATED " SHE WAS SEEN IN THE ER MULTIPLE TIMES AND WAS PLACED ON GLAUCOMA DROPS. IT WAS REPORTED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND A SHORTER LENGTH LENS WAS LATER IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, INSIDE A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12059229&pc=QCB | 2023826-2021-02039 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -12.00/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT EXPERIENCED ELEVATED IOP. REPORTER STATED " SHE WAS SEEN IN THE ER MULTIPLE TIMES AND WAS PLACED ON GLAUCOMA DROPS. IT WAS REPORTED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND A SHORTER LENGTH LENS WAS LATER IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, INSIDE A VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072258&pc=QCB | 2023826-2021-02069 | 2021-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, IT WAS REMOVED DUE TO ELEVATED IOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: A SHORTER LENS WAS IMPLANTED INTO THE PATIENT'S EYE AFTER THE INITIAL LENS WAS REMOVED. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A CONTACT LENS CASE. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND DEBRIS ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12097194&pc=QCB | 2023826-2021-02208 | 2021-05-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, -6.5/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND THE PROBLEM RESOLVED. USER ERROR WAS THE CAUSE OF THIS EVENT, IT WAS REPORTED THAT PERIPHERAL IRRIDECTOMIES ARE MISSING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102434&pc=QCB | 2023826-2021-02115 | 2021-06-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/6.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. TRANSIENT LOSS OF BCVA AND ELEVATED IOP (30MMHG) WITHOUT PUPIL BLOCK,ANGLE CLOSURE, AND PIGMENT DISPERSION, WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PRESENCE OF INFLAMMATION COMPOSED OF CELLS AND IRIS PIGMENTS IN THE ANTERIOR CHAMBER. OCT WILL BE DONE TO SEE IF THERE IS ADHERENCE OF LENS TO IRIS ANTI INFLAMMATORY DROPS GIVEN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12150435&pc=QCB | 2023826-2021-02532 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/118 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON DID SOME ADDITIONAL LASER SURGERY TO THE IRIDOTOMY (ENLARGED) AND THAT RESOLVED THE ISSUE. THE PATIENT IS FINE AND THERE WAS NO NEED TO EXPLANT THE LENS, THE LENS REMAINED IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12203583&pc=QCB | 2023826-2021-02557 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -12.5/1.0/131 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE SURGEON REPORTS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE AND NOTED "EXCESSIVE VAULTING CAUSING ANGLES TO NARROW". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE PATIENT WAS SEEN FOR AN EVALUATION, PATIENT HAD HIGH VAULTING BUT NORMAL IOP ON HYPOTENSIVE MEDICATION AND WAS INSTRUCTED TO DISCONTINUE MEDICATION. THEN ON (B)(6) 2021 ANOTHER EVALUATION WAS PERFORMED AND THE PATIENT HAD INCREASED IOP OFF MEDICATION AND A DECISION WAS MADE TO REPEAT/ENLARGE LPIS TO ASSESS IF ANY IMPROVEMENT IN IOP. IT WAS ALSO INDICATED IF NO IMPROVEMENT WITH IOP OCCURS A LENS EXCHANGE IS PLANNED. H6 - HEALTH EFFECT IMPACT CODE: 4625 ADDITIONAL SURGERY - LPI REPEAT/ENLARGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT'S DIURNAL IOP MEASUREMENT AND IOP IS WITHIN NORMAL LIMITS/STABLE AFTER REPEAT LPI AND OFF HYPOTENSIVE MEDICATIONS, AND THEY WOULD NOT BE PROCEEDING WITH A LENS EXCHANGE. PATIENT WILL CONTINUED BEING MONITORED. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12204759&pc=QCB | 2023826-2021-02558 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -17.0/2.5/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE SURGEON REPORTS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE AND NOTED "EXCESSIVE VAULTING CAUSING ANGLES TO NARROW". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE PATIENT WAS SEEN FOR AN EVALUATION, PATIENT HAD HIGH VAULTING BUT NORMAL IOP ON HYPOTENSIVE MEDICATION AND WAS INSTRUCTED TO DISCONTINUE MEDICATION. THEN ON (B)(6) 2021, ANOTHER EVALUATION WAS PERFORMED AND THE PATIENT HAD INCREASED IOP OFF MEDICATION AND A DECISION WAS MADE TO REPEAT/ENLARGE LPIS TO ASSESS IF ANY IMPROVEMENT IN IOP. IT WAS ALSO INDICATED IF NO IMPROVEMENT WITH IOP OCCURS A LENS EXCHANGE IS PLANNED. H6 - HEALTH EFFECT IMPACT CODE: 4625 ADDITIONAL SURGERY - LPI REPEAT/ENLARGE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT'S DIURNAL IOP MEASUREMENT AND IOP IS WITHIN NORMAL LIMITS/STABLE AFTER REPEAT LPI AND OFF HYPOTENSIVE MEDICATIONS, AND THEY WOULD NOT BE PROCEEDING WITH A LENS EXCHANGE. PATIENT WILL CONTINUED BEING MONITORED. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214302&pc=QCB | 2023826-2021-02675 | 2021-02-25 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -12.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP (45 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2021. CONCOMITANT MEDICATIONS WERE PERSCRIBED. LENS REMAINS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12250513&pc=QCB | 2023826-2021-02972 | 2021-05-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/094 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE PATIENTS INTRAOCULAR PRESSURE WAS ELEVATED DURING THE SURGERY. THE LENS WAS REPLACED FOR ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308398&pc=QCB | 2023826-2021-03083 | 2021-07-06 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.00/2.0/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (54MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308417&pc=QCB | 2023826-2021-03084 | 2021-07-06 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.5/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (60MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12407186&pc=QCB | 2023826-2021-03328 | 2021-06-20 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(32MMHG) (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12510969&pc=QCB | 2023826-2021-03705 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.00/1.00/22 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. IT WAS REPORTED THAT A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS OTHER AND IT WAS REPORTED THAT THE DEVICE FAIL TO PERFORM AS INTENDED. CAUSE DETAILS " HIGH WTW MEASUREMENT " WAS REPORTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587047&pc=QCB | 2023826-2021-03864 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+3.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND EXCESSIVE VAULTING. IOP WAS STILL HIGH ON THE 1 AND 3 DAY FOLLOW-UP POST LENS REMOVAL. PATIENT IS RESPONDING TO ANTI-GLAUCOMA MEDICATIONS AND THE IOP IS CONTROLLED. THE PATIENT WILL BE CLOSELY FOLLOWED UP. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587050&pc=QCB | 2023826-2021-03863 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+4.0/109 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND LARGE VAULT. IOP WAS STILL HIGH ON THE 1 AND 3 DAY FOLLOW-UP POST LENS REMOVAL. PATIENT IS RESPONDING TO ANTI-GLAUCOMA MEDICATIONS AND THE IOP IS CONTROLLED. THE PATIENT WILL BE CLOSELY FOLLOWED UP. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5- AS PER THE COMPLAINT QUESTIONNAIRE, THE REPORTER MARKED THE CAUSE OF THE EVENT AS UNKNOWN, HOWEVER SUBSEQUENT COMMUNICATION VIA EMAIL ON (B)(6) 2021 STATED THAT "THE DOCTOR TOLD ME TODAY THAT IN HIS OPINION THIS IS STEROID INDUCED IOP". CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS ARE REQUIRED BUT CANNOT BE PERFORMED DUE TO THE LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685440&pc=QCB | 2023826-2021-04063 | 2021-09-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+0.5/005 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED AFTER EXPLANT, NO FURTHER SURGICAL INTERVENTION IS PLANNED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "PREOPERATIVELY IOP WAS 23MMHG, BCVA 20/32; AFTER IMPLANTATION IOP WAS 60MMHG (AT 5PM), PRESSURE VIA PARACENTESIS REDUCED IOP TO 20MMHG (18 MINUTES AFTER 5PM); AT 6PM IOP INCREASED TO 50MMHG AGAIN SO EXPLANTATION WAS PERFORMED IMMEDIATELY." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685432&pc=QCB | 2023826-2021-04062 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THERE WAS NO REDUCTION OF IOP UNDER ORAL MEDICATION BY GLAUPAX. THE PROBLEM WAS RESOLVED AFTER EXPLANT. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685468&pc=QCB | 2023826-2021-04061 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+0.5/034 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THERE WAS NO REDUCTION OF IOP UNDER ORAL MEDICATION BY GLAUPAX. THE PROBLEM WAS RESOLVED AFTER EXPLANT. SEE MFR REP# 2023826-2021-04063 FOR REPLACEMENT LENS. Manufacturer Narrative: H3: DEVICE EVALUATON - LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13014953&pc=QCB | 2023826-2021-04810 | 2018-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00/+5.0/056 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT -RELATED FACTOR.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00/+5.0/056 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT -RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13049602&pc=QCB | 2023826-2021-04918 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.2MM VTICH13.2; +9.0/+1.5/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2021 THE SURGEON PERFORMED AN ADDITION OF NEW PI (YAG). THIS RESOLVED THE PROBLEM AND "ALL IS FINE". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13054368&pc=QCB | 2023826-2021-04738 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13065424&pc=QCB | 2023826-2021-04743 | 2021-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/013 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: D6: (B)(6) 2024. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13115357&pc=QCB | 2023826-2021-04768 | 2021-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/088 (SPHERE/CYLINDER/AXIS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256067&pc=QCB | 2023826-2022-00117 | 2021-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.5/3.0/091 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256668&pc=QCB | 2023826-2022-00119 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297269&pc=QCB | 2023826-2022-00186 | 2021-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL SURGERY WAS PERFORMED, WITH THE ADDITION OF NEW PIS BY YAG BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. AT LAST POST-OP VISIT THE IOP WAS 24MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-00185. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL SURGERY WAS PERFORMED, WITH THE ADDITION OF NEW PIS BY YAG BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. AT LAST POST-OP VISIT THE IOP WAS 24MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-00185. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). CLAIM # (B)(6). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). CLAIM # (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297379&pc=QCB | 2023826-2022-00185 | 2021-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE IOP INCREASED CONTINUOUSLY AFTER OPERATION. SEE MFR. REPORT # 2023826-2022-00186 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297459&pc=QCB | 2023826-2022-00184 | 2021-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL SURGERY WAS PERFORMED ON (B)(6) 2021, WITH THE ADDITION OF NEW PIS BY YAG BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. AT LAST POST-OP VISIT THE IOP WAS 29MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-00183. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL SURGERY WAS PERFORMED ON (B)(6) 2021, WITH THE ADDITION OF NEW PIS BY YAG BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. AT LAST POST-OP VISIT THE IOP WAS 29MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-00183. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297565&pc=QCB | 2023826-2022-00183 | 2021-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE IOP INCREASED CONTINUOUSLY AFTER OPERATION. SEE MFR. REPORT # 2023826-2022-00184 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE IOP INCREASED CONTINUOUSLY AFTER OPERATION. SEE MFR. REPORT # 2023826-2022-00184 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13361414&pc=QCB | 2023826-2022-00300 | 2021-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.0/6.0/135 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT A 12.1 VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.0/6.0/135 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423349&pc=QCB | 2023826-2022-00253 | 2022-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -11.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (28MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED 03/JAN/2022) EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. REPORTEDLY EYE QUIET AND PATIENT HAPPY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13481675&pc=QCB | 2023826-2022-00320 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.50/2.0/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13509882&pc=QCB | 2023826-2022-00333 | 2021-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.50/3.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND ON THE SAME DAY DIFFERENT SURGERY A SHORTER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577875&pc=QCB | 2023826-2022-00499 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -8.00/2.50/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (22MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -8.00/2.50/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (22MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577879&pc=QCB | 2023826-2022-00500 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -6.00/1.50/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (47MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13726672&pc=QCB | 2023826-2022-00758 | 2021-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE IS VERY GOOD. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13905945&pc=QCB | 2023826-2022-00907 | 2022-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.6; -6.00/+2.0/86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXPLANTED ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(4) 2022 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14116520&pc=QCB | 2023826-2022-01194 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.00/+1.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE AND PATIENT-RELATED FACTOR. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14282382&pc=QCB | 2023826-2022-01213 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS EXPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS PATIENT RELATED FACTOR. "PATIENT WITH ABNORMAL CILIARY BODY MORPHOLOGY AND NARROWING OF THE ATRIAL ANGLE ARE NOT SUITABLE FOR SURGERY." Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14282387&pc=QCB | 2023826-2022-01212 | 2022-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS EXPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS PATIENT RELATED FACTOR. "PATIENT WITH ABNORMAL CILIARY BODY MORPHOLOGY AND NARROWING OF THE ATRIAL ANGLE ARE NOT SUITABLE FOR SURGERY." Manufacturer Narrative: WEIGHT-RACE:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14403295&pc=QCB | 2023826-2022-01524 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY A REPLACEMENT LENS OF SPHERICAL AND SHORTER LENGTH WAS IMPLANTED. THE REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM. SEE MFR # 2023826-2022-01552 FOR REPLACEMENT LENS. Manufacturer Narrative: CORRECTED DATA: H6-TYPE OF INVESTIGATION CODE : 10 IN PREVIOUS MDR SHOULD BE CORRECTED TO 4110. ADDITIONAL INFORMATION: H6- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14403301&pc=QCB | 2023826-2022-01525 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/111(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED. A REPLACEMENT LENS WAS LATER IMPLANTED WITH A SPHERICAL MODEL AND SHORTER LENGTH. THE REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM. SEE MFR # 2023826-2022-01553 FOR REPLACEMENT LENS. Manufacturer Narrative: H6-TYPE OF INVESTIGATION CODE : 10 IN PREVIOUS MDR SHOULD BE CORRECTED TO 4110. ADDITIONAL INFORMATION: H6- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632891&pc=QCB | 2023826-2022-01661 | 2022-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2, -9.0/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, LENS WAS NOT PROPERLY SIZED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14687899&pc=QCB | 2023826-2022-01855 | 2022-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/1.5/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. LATER ON THE SAME DATE IN A SECOND SURGERY THE LENS WAS EXPLANTED AND NOTED "THE PATIENT WOULDN'T IMPLANTED WITH THE LENS". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14829910&pc=QCB | 2023826-2022-01877 | 2022-01-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6, -5.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14829913&pc=QCB | 2023826-2022-01878 | 2022-01-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6, -5.50/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14983249&pc=QCB | 2023826-2022-02206 | 2022-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/3.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED. IT WAS NOTED "THE PATIENT IS UNDER CONSIDERATION; AND AT THE TIME OF THE SECOND OPERATION IS UNCERTAIN". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4): CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042197&pc=QCB | 2023826-2022-02165 | 2022-02-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6, -10.5/3.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117753&pc=QCB | 2023826-2022-02455 | 2022-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/129 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP) (STEROID RESPONDER). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE PATIENT IS DOING WELL AND FEELS GOOD. THE IOP IS GOOD. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136362&pc=QCB | 2023826-2022-02502 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -5.0/3.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS HIGH IOP. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15178260&pc=QCB | 2023826-2022-02602 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: RECEIVED A COPY OF AN ARTICLE FROM THE OPEN OPHTHALMOLOGY JOURNAL. THE ARTICLE WAS IN REGARD TO "CLINICAL OUTCOMES OF TORIC IMPLANTABLE COLLAMER LENS (T-ICL) AND TORIC IMPLANTABLE PHAKIC CONTACT LENS (IPCL) FOR HIGH MYOPIA WITH ASTIGMATISM: A COMPARATIVE STUDY". THE STUDY INCLUDED 60 MYOPIC EYES DIVIDED INTO 2 GROUPS - GROUP A (30 EYES) IMPLANTED WITH T-ICL AND GROUP B (30 EYES) IMPLANTED WITH TORIC IPCL. THE RESULTS FOUND AN IOP INCREASE IN THE FIRST MONTH POSTOPERATIVELY WITH THE T-ICL, MANAGED BY TOPICAL ANTIGLAUCOMA MEDICATIONS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: A POSSIBLE CAUSE OF THE ELEVATED IOP WAS DUE TO POST-OPERATIVE RETAINED VISCOELASTIC MATERIAL, MILD POSTOPERATIVE INFLAMMATION OR STEROID RESPONDERS IN SOME CASES. BOTH GROUPS HAVE THE SAME CONDITION OF EARLY IOP INCREASE, AND WE DID NOT REFER THIS IOP INCREASE TO LENS DESIGN OR MATERIAL OR SPECIFY IT TO T-ICL. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: AGE AT TIME OF EVENT, DATE OF BIRTH: UNKNOWN, SEX: UNKNOWN, WEIGHT: UNKNOWN, ETHINICITY:UNKNOWN, RACE: UNKNOWN, DATE OF EVENT: UNKNOWN. D4: EXPIRATION DATE UNK IMPLANTED DATE: UNKNOWN, EXPLANTED DATE: UNKNOWN, DEVICE MANUFACTURE DATE: UNKNOWN. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334665&pc=QCB | 2023826-2022-02949 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS OF 5.00/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATED SIZE ICL TORIC + WTW PATIENT. Manufacturer Narrative: PATIENT IDENTIFIER:(B)(6). EXPIRATION DATE: 04/30/2023. CLAIM# 725631.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567405&pc=QCB | 2023826-2022-03565 | 2022-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.0/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT . THE LENS WAS LATER REPLACED WITH A SPHERICAL LENS OF THE SAME LENGTH DIFFERENT POWER AND THE PROBLEM RESOLVED. FOR ADDITIONAL INFORMATION THE REPORTER STATED " BECAUSE OF THE HIGH INTRAOCULAR PRESSURE AT THE INITIAL STAGE AFTER SURGERY, THE DOCTOR ORIGINALLY DECIDED TO CUSTOMIZE THE ASTIGMATIC LENS WITH SPECIAL AXIAL POSITION FOR REPLACEMENT AFTER FACE-TO-FACE CONSULTATION. AFTER REPEATED EXAMINATIONS, THE PATIENT AGAIN DECIDED TO CUSTOMIZE THE NON-ASTIGMATIC LENS FOR REPLACEMENT, AND THE ASTIGMATISM DEGREE WAS CORRECTED BY LASER SURGERY." CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6: IMPACT CODE: 4628 ADDED. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801805&pc=QCB | 2023826-2022-04085 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE INTRAOCULAR PRESSURE (IOP) WAS REPORTED AS BEING ELEVATED AND COMBIGAN BID AND PRED FORTE BID WERE INITIATED. THE ELEVATED IOP WAS RESOLVED AND WAS MEASURED WITHIN NORMAL RANGE ON (B)(6) 2022. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15811021&pc=QCB | 2023826-2022-03961 | 2022-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -18.00/2.5/057(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. CAUSE OF THE EVENT IS THE DEVICE. THE REPORTER STATED THAT THE LENS WAS, "OVERSIZED." THE PROBLEM IS NOT RESOLVED. REPORTEDLY, "EXPLANTED ICL, IOP NOT CONTROLLED AFTGER ICL REMOVAL. 30MMHG 2 WEEKS POST-OP." Manufacturer Narrative: D4: CORRECTED TO UNIQUE IDENTIFIER (UDI) # 00841542131032 IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15903005&pc=QCB | 2023826-2022-04235 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.0/076 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "2 MONTHS AFTER THE OPERATION, THE IOP HAS BEEN BETWEEN 26-29 MMHG, THE PATIENT WAS ANXIOUS AND ASKED TO REMOVE THE LENS." Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: EXPIRATION DATE: SHOULD BE CORRECTED TO 31-MAY-2025. H4-DEVICE MANUFACTURE DATE:SHOULD BE CORRECTED TO 13-JUN-2022. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15903054&pc=QCB | 2023826-2022-04234 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PRIMARY OPEN ANGLE OR NARROW ANGLE GLAUCOMA ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "2 MONTHS AFTER THE OPERATION, THE IOP HAS BEEN BETWEEN 26-29 MMHG, THE PATIENT WAS ANXIOUS AND ASKED TO REMOVE THE LENS." Manufacturer Narrative: OFF LABEL AGE<21, PRE-EXISTING CONDITION: PRIMARY OPEN ANGLE OR NARROW ANGLE GLAUCOMA. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000610&pc=QCB | 2023826-2022-04626 | 2022-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/+2.0/86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000620&pc=QCB | 2023826-2022-04625 | 2022-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/+2.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269394&pc=QCB | 2023826-2023-00310 | 2022-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. REPORTEDLY, "REMOVE AND NO LONGER IMPLANT." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269398&pc=QCB | 2023826-2023-00311 | 2022-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. REPORTEDLY, "REMOVE AND NO LONGER IMPLANT." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.5/3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. REPORTEDLY, "REMOVE AND NO LONGER IMPLANT." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN." IN THE INITAL MDR. CLAIM #(B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275247&pc=QCB | 2023826-2023-00258 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275257&pc=QCB | 2023826-2023-00257 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+2.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500236&pc=QCB | 2023826-2023-00819 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500270&pc=QCB | 2023826-2023-00818 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.00/+0.5/100 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500271&pc=QCB | 2023826-2023-00817 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.00/+0.5/045 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500966&pc=QCB | 2023826-2023-00820 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587673&pc=QCB | 2023826-2023-00952 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/109 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). AN ADDITION OF NEW PI WAS PERFORMED BY YAG ON (B)(6) 2022 BUT THE PRESSURE DID NOT GO DOWN. THE PATIENT ENDED UP HAVING A CATARACT LENS IMPLANTED. Manufacturer Narrative: PATIENT INFO: UNK. HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ADDITION OF NEW PI. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16600700&pc=QCB | 2023826-2023-01008 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.0/1.5/003 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 DUE TO REFRACTIVE SUPRISE. REPORTEDLY, "POSTOPERATIVE DAY 1 IOP AND I/A IS PERFORMED. CURRENTLY, DOCTOR IS TAKING A WAIT-AND-SEE APPROACH." THE CAUSE OF THE EVENT WAS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16721439&pc=QCB | 2023826-2023-01332 | 2022-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF --14.0/1.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPANTED AND REMOVED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: UNK. ACD <3.0. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792688&pc=QCB | 2023826-2023-01434 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT, ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WERE OBSERVED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND PATIENT RELATED FACTOR. REPORTEDLY, "HIGH IOP WITH MEDICATIONS. PLAN TO EXCHANGE TO A SMALLER SIZE (12.6). Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON 21 APR 2023 WITH A SHORTER LENGTH LENS BUT PROBLEM WAS NOT RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. LENS RETURNED IS NOT THE SIZE STATED IN THE CLAIM. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792797&pc=QCB | 2023826-2023-01433 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/147 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND ANGLE CLOSURE WITH ELEVATED IOP WERE OBSERVED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND PATIENT RELATED FACTOR. REPORTEDLY, "HIGH IOP WITH MEDICATIONS. PLAN TO EXCHANGE TO A SMALLER SIZE (12.6)." Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON (B)(6) 2023 WITH A SHORTER LENGTH LENS BUT PROBLEM WAS NOT RESOLVED. CLAIM# (B)(6). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PATIENT RELATED FACTOR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16917117&pc=QCB | 2023826-2023-01760 | 2021-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -13.00/+1.0/39 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216365&pc=QCB | 2023826-2023-02549 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). ELEVATED IOP AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT TO BE OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216368&pc=QCB | 2023826-2023-02550 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/1.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). ELEVATED IOP AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT TO BE OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK. D6B - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253137&pc=QCB | 2023826-2023-02657 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.00/1.5/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE PROBLEM WAS RESOLVED. DEVICE WAS MENTIONED AS THE CAUSE OF THE EVENT. REPORTEDLY, "ANTERIOR CATARACT. FIXED BLOWN PUPIL WITH SYNECHAIE." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322467&pc=QCB | 2023826-2023-02901 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/129 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. IT WAS REPORTED THAT THE PATIENT DECIDED TO REMOVE IT AND IS AFRAID TO IMPLANT IT AGAIN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES-UNABLE TO EVALUATE. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322480&pc=QCB | 2023826-2023-02848 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/126 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. IT WAS REPORTED THAT THE PATIENT DECIDED TO REMOVE IT AND IS AFRAID TO IMPLANT IT AGAIN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES-UNABLE TO EVALUATE. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17465253&pc=QCB | 2023826-2023-03310 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.5/1.5/004 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LENS ROTATION AND IOP SPIKES. IT WAS INDICATED THERE WILL BE NOT REPLACEMENT LENS. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17465254&pc=QCB | 2023826-2023-03311 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED IMPLANTED A 12.6MM VTICMO12.6; -16.5/1.0/002 (SPHERE / CYLINDER / AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION AND IOP SPIKES. IT WAS INDICATED THERE WILL BE NOT REPLACEMENT LENS. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17914527&pc=QCB | 2023826-2023-04273 | 2023-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO ELEVATED IOP AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136876&pc=QCB | 2023826-2023-05060 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/0.5/015 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136885&pc=QCB | 2023826-2023-05061 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/170 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18223147&pc=QCB | 2023826-2023-05376 | 2023-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -5.00/3.5/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258193&pc=QCB | 2023826-2023-05318 | 2023-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -7.5/1.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ELEVATED IOP 34MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS OBSERVED. OPHTHALMIC DROPS PRESCRIBED(BROMFENAC SODIUM HYDRATE AND LATANOPROST). NO OTHER SYSTEMS WERE OBSERVED. PROBLEM RESOLVED. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18287965&pc=QCB | 2023826-2023-05442 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC EVO ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS EYE. AT ONE MONTH POST-OP, THE PATIENT HAD PERSISTENT ELEVATED INTRAOCULAR PRESSURE (IOP), WHICH WAS BEING TREATED WITH IOP LOWERING MEDICATIONS. THE PATIENT HAD NO SYMPTOMS OF ANGLE CLOSURE AND NO SIGNS OF INFLAMMATION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK. H6: HEALTH EFFECT IMPACT CODE: 4644 - IOP LOWERING MEDICATIONS. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/071 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD). AT ONE MONTH POST-OP, THE PATIENT HAD PERSISTENT ELEVATED INTRAOCULAR PRESSURE (IOP), WHICH WAS BEING TREATED WITH IOP LOWERING MEDICATIONS. THE PATIENT HAD NO SYMPTOMS OF ANGLE CLOSURE AND NO SIGNS OF INFLAMMATION. THE LENS REMAINED IMPLANTED. AT THE MOST RECENT POST-OP, PRESSURE HAS STABILIZED ON MEDICATIONS AND THE PATIENT IS BEING MONITORED CLOSELY. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18298068&pc=QCB | 2023826-2023-05435 | 2023-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18312966&pc=QCB | 2023826-2023-05510 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. ELEVATED INTRAOCULAR PRESSURE AFTER ICL IMPLANTATION WAS REPORTED. PATIENT IS ON STEROID DROPS. VISION IS 20/40 WITH LESS THAN A HALF DIOPTER OF RESIDUAL PRESCRIPTION FOR CYL AND SPHERE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18313031&pc=QCB | 2023826-2023-05511 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. ELEVATED INTRAOCULAR PRESSURE AFTER ICL IMPLANTATION WAS REPORTED. PATIENT IS ON STEROID DROPS. VISION IS 20/30 WITH LESS THAN A HALF DIOPTER OF RESIDUAL PRESCRIPTION FOR CYL AND SPHERE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18346344&pc=QCB | 2023826-2023-05691 | 2023-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18.00/2.5/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER AND INDICATED THAT THE DEVICE PERFORMED AS INTENDED BUT THE LENS WAS TOO LARGE. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238020&pc=QCB | 2023826-2021-00168 | 2020-12-18 | Malfunction | STAAR SURGICAL COMPANY AG | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/0.5/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2013. REFRACTIVE CHANGE OVERTIME AND LOW VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANT DATE - (B)(6) 2013. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537728&pc=QCB | 2023826-2020-02067 | 2020-08-20 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON ENLARGED/ADDITION OF PERIPHERAL IRIDOTOMY, BY YAG. A SUTURE WAS REQUIRED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS VISION IS GOOD BUT HAS MILD GLARE/HALO AT NIGHT. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566551&pc=QCB | 2023826-2020-02177 | 2020-07-18 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.50/+1.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING A HUGE VAULT AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON PERFORMED A COLOBOMA ON (B)(6) 2020, WHICH CORRECTED THE SITUATION A LITTLE BIT, BUT ANGLES WERE STILL CLOSED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK, RACE: UNK, IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566577&pc=QCB | 2023826-2020-02176 | 2020-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING A HUGE VAULT AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON PERFORMED A COLOBOMA ON (B)(6) 2020, WHICH CORRECTED THE SITUATION A LITTLE BIT, BUT ANGLES WERE STILL CLOSED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE LENS HAS BEEN EXPLANTED ON (B)(6) 2020. A NEW LENS WAS LATER IMPLANTED. THE PATIENT WAS SATISFIED WITH OPERATION; AND THE DOCTORS WAS SATISFIED WITH RESULTS. D6B. (B)(6) 2020 CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE LENS HAS BEEN EXPLANTED ON AN UNKNOWN DATE. FURTHER INFORMATION HAS BEEN REQUESTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. D6B. UNK H3- DEVICE EVALUATION: LENS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT. HAPTIC TORN/BROKEN. VISUAL OBSERVATION CONCLUDED NO DIMENSIONAL INSPECTION CAN BE PERFORMED DUE LENS RETURNED IN PIECES. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +6.00/+2.0/132 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS HAS BEEN EXPLANTED ON (B)(6) 2020 AND ON THE SAME DAY A NEW LENS OF SAME MODEL BUT LONGER LENGTH WAS IMPLANTED. THIS RESOLVED THE PROBLEM. REPORTER STATES CAUSE AS "THE DOCTOR SUPPOSE THAT THE SIZE WAS SMALLER THAN THEY NEED." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634500&pc=QCB | 2023826-2020-02281 | 2020-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND INTERMITTENT ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS IMPLANTED ON (B)(6) 2020. ADDITIONAL SURGERY WAS PERFORMED, ENLARGEMENT OF PI BY YAG PROCEDURE AND THERE WAS AN ADDITIONAL SUTURE. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. H6: HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY; ENLARGEMENT OF PI BY YAG, SUTURE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634519&pc=QCB | 2023826-2020-02294 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+5.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND UNREACTIVE (FIXED) PUPIL. THE IOP WAS CONTROLLED WITH AGGRESSIVE TREATMENT. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO AN OVERSIZED LENS. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ANTERIOR SEGMENT WAS OK AND THERE WAS A LITTLE BIT OF PIGMENT DISPERSION BUT STILL OK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT IMPACT CODE: 4629 SHOULD BE CORRECTED TO 4627. B5 - IT WAS INDICATED THAT THE LENS WAS EXPLANTED ON (B)(6) 2020, AND LATER ON THE SAME DAY IN A SEPERATE SURGERY A REPLACEMENT LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. POSTOPERATIVE REPORT INDICATED THAT THE ANTERIOR SEGMENT WAS OK AND THERE WAS A LITTLE BIT OF PIGMENT DISPERSION BUT STILL OK. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635452&pc=QCB | 2023826-2020-02242 | 2020-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/2.0/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 DUE TO EXCESSIVE VAULT AND ELEVATED IOP (22MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE THAT WAS ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "OVER SIZE LENS." REPORTEDLY, "PATIENT HAPPY." Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THE PATIENT EXPERIENCED SHALLOW CHAMBER. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SHALLOW CHAMBER. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALAUTION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11099005&pc=QCB | 2023826-2020-03218 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, AND SHALLOW CHAMBER. THERE WAS AQUEOUS MISDIRECTION, WHICH WAS RELIEVED WITH VITREOUS TAPS. THE LENS WAS EXCHANGED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED, THE LENS STILL HAD AN EXCESSIVE VAULT. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. MEDICATION WAS PRESCRIBED, THE CORNEA WAS CLEAR, ICL WAS CENTERED, WITH IRIS PIGMENTS ON LENS. SEE MFR. REPORT # 2023826-2020-03220 FOR REPLACEMENT LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - HEALTH EFFECT CLINICAL CODE : 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CORRECTED DATA: B5 - "THE CORNEA WAS CLEAR, CENTERED ICL WITH NOTE OF IRIS PIGMENTS ON LENS" SHOULD BE REMOVED AS IT IS FOR THE REPLACEMENT LENS. H6 - HEALTH IMPACT CLINICAL CODE: 4581 IN H10 DESCRIPTION - SHOULD BE CORRECTED TO "HEALTH EFFECT CLINICAL CODE: 4581", AND IRIS PIGMENT SHOULD BE REMOVED AS IT IS FOR REPLACEMENT LENS. H6 - HEALTH EFFECT CLINICAL CODE : 4581 - MEDICATION, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091484&pc=QCB | 2023826-2020-03221 | 2020-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00/+3.0/062 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091943&pc=QCB | 2023826-2020-03217 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND SHALLOW CHAMBER. THERE WAS AQUEOUS MISDIRECTION, WHICH WAS RELIEVED WITH VITREOUS TAPS. THE LENS WAS EXCHANGED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED, THE LENS STILL HAD AN EXCESSIVE VAULT. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. MEDICATION WAS PRESCRIBED, THE CORNEA WAS CLEAR, ICL WAS CENTERED, WITH IRIS PIGMENTS ON LENS. SEE MFR. REPORT # 2023826-2020-03219 FOR REPLACEMENT LENS. Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH IMPACT - CLINICAL CODE: 4581 - IRIS PIGMENTS SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS OPTIC DEFORMED, THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICIATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SHALLOW CHAMBER, IRIS PIGMENTS HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY - AQUEOUS MISDIRECTION/VITREOUS TAPS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091947&pc=QCB | 2023826-2020-03220 | 2020-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING HIGH VAULTING. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. MEDICATION WAS PRESCRIBED, THE CORNEA WAS CLEAR, ICL WAS CENTERED, WITH IRIS PIGMENTS ON LENS. THIS WAS THE REPLACEMENT LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2020-03218. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT PATIENT ALSO EXPERIENCED ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2020, AND THE PROBLEM RESOLVED AND NOTED "PATIENT STILL ON AZARGA AND LUMIGAN". H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, MEDICATION. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11092035&pc=QCB | 2023826-2020-03219 | 2020-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING A HIGH VAULT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. MEDICATION WAS PRESCRIBED, CLEAR CORNEA, CENTERED ICL, WITH IRIS PIGMENTS ON LENS. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-03217. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT ALSO EXPERIENCED ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. IT WAS NOTED THE PATIENT WAS STILL ON AZARGA AND LUMIGAN. HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, BENT, WITH DEBRIS ON THE LENS. CORRECTED DATA: G6 - DISREGARD PREVIOUS SUPPLEMENTAL #2. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135881&pc=QCB | 2023826-2021-00039 | 2020-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.0/+4.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT AND ELEVATED IOP, WHICH RESPONDED TO ENHANCED PIS. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFO: UNREACTIVE (FIXED) PUPIL AND ADDITION OF NEW PI-YAG ALSO REPORTED. ON (B)(6) 2020, THE LENS WAS EXPLANTED. AFTER EXPLANT THE EYE STATUS IS LISTED AS DEEPENED AC, GOOD IOP, PUPIL STILL SLUGGISH AT 5.5MM. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: SULCUS DIAMETER IN VIVO APPEARS MUCH SMALLER THAN THE NOMOGRAM-RECOMMENDED 13.2MM TICL. H6-CLINICAL CODE 4581: [UNREACTIVE (FIXED) PUPIL, DEEPENED AC]. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343216&pc=QCB | 2023826-2021-00407 | 2021-01-15 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2, -8.50/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IIDO-CORNEAL ANGLES AND AC SHALLOW IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS SHALLOW AC. PER THE REPORTER ON (B)(6) 2021, THERE WILL NOT BE ANOTHER LENS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343218&pc=QCB | 2023826-2021-00406 | 2021-01-15 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2, -7.50/2.0/126 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IIDO-CORNEAL ANGLES AND AC SHALLOW IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS SHALLOW AC. PER THE REPORTER ON (B)(6) 2021, THERE WILL NOT BE ANOTHER LENS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366548&pc=QCB | 2023826-2021-00478 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "TOO BIG LENGTH OF ICL 13.7 FOR THIS EYE. DISSOCIATION BETWEEN THE REAL SIZE OF POSTERIOR CHAMBER AND PREDICTED BY CALCULATOR." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11389451&pc=QCB | 2023826-2021-00545 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF ICA. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405262&pc=QCB | 2023826-2021-00546 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS. THE PATIENT ENDED UP WITH A SUBSTANTIAL RESIDUAL SPHERE AND CYLINDER. THE PATIENT ALSO HAD A PRESSURE SPIKE AND WAS TAKING PRESSURE MEDICATION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478839&pc=QCB | 2023826-2021-00639 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BUT SAME LENGTH LENS ON (B)(6) 2020, AND THE PROBLEM WAS RESOLVED. THE SURGEON IMPLANTED THE EXCHANGED ICL LENS VERTICALLY TO REDUCE VAULT AND AN INCISION WAS USED TO REDUCE ASTIGMATISM. THE PATIENT WAS SATISFIED WITH THE RESULT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616998&pc=QCB | 2023826-2021-00922 | 2021-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. TOPICAL ANTIHYPERTENSIVE TREATMENT (KRYTANTEK) WAS ADMINISTERED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2021. POST-OP VISUAL ACUITY WAS 20/30 UVA; 20/25 CVA. PIO: 18MMHG WITH NO ANTI GLAUCOMA TREATMENT. SLIT LAMP - NO INFLAMMATORY FINDINGS. VAULT - 386UM, ANGLE:20 DEGREES. THE PATIENT AND DOCTOR ARE HAPPY WITH THE RESULTS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11624856&pc=QCB | 2023826-2021-00934 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM012.6, -17.50/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PIGMENT DISPERSION WAS OBSERVED. ON PATIENTS LAST VISIT DATE (B)(6) 2021 FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) " VISION IS EXCELLENT, IOP 12MM/HG WITH MEDS BUT THERE IS A IRIS DISPERSION FOR 3 MONTHS POST-OP ASSOCIATED WITH TRANSPARENT IRIS AT THE SIDE OF LENS HAPTICS" WAS REPORTED. FOR ADDITIONAL INFORMATION "HOWEVER THE LENS VAULT IS VERY LOW, THERE IS A LIGHT TRANSPARANCY AT THE SIDE OF THE HAPTICS" WAS REPORTED. ON (B)(6) 2021 THE REPORTER STATED " I LIKE TO ACKNOWLEDGE THAT PATIENT IS GETTING BETTER WEEK AFTER WEEK. DISPERSION HAS ALREADY STOPPED BUT THE ILLUMINATION IS EXISTING". LENS REMAINS IMPLANTED IT WAS REPORTED THAT NO INTERVENTION HAS BEEN PERFORMED FOR THE EYE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER STATED THAT "2 WEEKS AFTER SX HE COME WITH HIGH IOP OF 37. AND ANT.CHAMBER REACTION OF 2+. GIVEN TOPICAL ANTI GLAUCOMA AND REDUCED STEROID". THIS SHOULD HAVE BEEN ADDED IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11648783&pc=QCB | 2023826-2021-00923 | 2021-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+3.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. TOPICAL ANTIHYPERTENSIVE TREATMENT (KRYTANTEK) WAS ADMINISTERED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2021. POST-OP VISUAL ACUITY WAS 20/30 UVA; 20/25 CVA. PIO: 18MMHG WITH NO ANTI GLAUCOMA TREATMENT. SLIT LAMP: NO INFLAMMATORY FINDINGS. VAULT 439UM; ANGLE 21 DEGREES. THE PATIENT AND THE DOCTOR ARE HAPPY WITH THE RESULTS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK EXPLANT DATE: UNK. REPORT SOURCE: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11722521&pc=QCB | 2023826-2021-01209 | 2021-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/+0.5/164 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH, -11.50 DIOPTER, SPHERICAL LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP AND UNREACTIVE PUPIL (PUPIL IS REACTIVE TO LIGHT BUT VERY DILATED). FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): "AWAITING FOLLOW UP" WAS REPORTED." DEVICE" WAS REPORTED TO BE THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED " SURGEON INTERPRETATION IS THAT THE ANGLE IS OPEN BUT THAT THE CONFIGURATION AND VAULT OF ICL IS SUCH THAT THE PUPIL CANNOT CONSTICT OVER THE LIP OF THE OPTIC RESULTING IN THE PUPIL BEING MECHANICALLY PREVENTED FROM RETURNING TO NORMAL SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT IT WAS AN IRIS CONFIGURATION ISSUE AND NOT AN ISSUE WITH VAULT OR THE ICL. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11780620&pc=QCB | 2023826-2021-01371 | 2021-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.50/+6.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (THE IRIS WAS TOUCHING THE CORNEA). ANOTHER ICL LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS CALCULATION ERROR IN SUGGESTED LENS LENGTH. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11780627&pc=QCB | 2023826-2021-01370 | 2021-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00/+6.0/116 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (THE IRIS WAS TOUCHING THE CORNEA). ANOTHER ICL LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS CALCULATION ERROR IN SUGGESTED LENS LENGTH. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11796745&pc=QCB | 2023826-2021-01390 | 2021-04-05 | Injury | STAAR SURGICAL COMPANY | 2021-05-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -16/2/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND ANGLE CLOSURE. ADDITIONAL SURGICAL INTERVENTION (AQUEOUS TAP AND ENLARGEMENT OF PI YAG LASER) AND MEDICAL INTERVENTION (IOP LOWERING MEDICATION) WERE PERFORMED/PRESCRIBED PRIOR TO ICL EXPLANTATION BUT DIDN`T RESOLVE THE ISSUE. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER EXPLANT " IOP AND ANGLE GRADE RETURNED TO NORMAL RANGE." IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. Manufacturer Narrative: PATIENT INFORMATION :UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11796748&pc=QCB | 2023826-2021-01391 | 2021-04-05 | Injury | STAAR SURGICAL COMPANY | 2021-05-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -17.5/3.0/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND ANGLE CLOSURE. ADDITIONAL SURGICAL INTERVENTION (AQUEOUS TAP AND ENLARGEMENT OF PI YAG LASER) AND MEDICAL INTERVENTION (IOP LOWERING MEDICATION) WERE PERFORMED/PRESCRIBED PRIOR TO ICL EXPLANTATION BUT DIDN`T RESOLVE THE ISSUE. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER EXPLANT " IOP AND ANGLE GRADE RETURNED TO NORMAL RANGE." IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. Manufacturer Narrative: PATIENT INFORMATION :UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855173&pc=QCB | 2023826-2021-01511 | 2021-04-21 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP(24MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE, "LENS TOO BIG FOR EYE." REPORTEDLY, SURGEON TO REPLACE WITH A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855239&pc=QCB | 2023826-2021-01522 | 2021-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.00/1.0/066 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP(57MMHG), ANGLE CLOSURE WITH ELEVATED IOP(57MMHG), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, TRANSIENT LOSS OF BCVA, EXCESSIVE VAULT, IRIS ATROPHY, BLURRED VISION, AND UNREACTIVE(FIXED) PUPIL BEING OBSERVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED STUCK IN A GAUZE, WITH MOISTURE, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, WITH RESIDUE ON SURFACE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894540&pc=QCB | 2023826-2021-01710 |  | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; Unintended Movement | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -12.0/+4.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON REPORTS ELEVATED IOP. REPORTEDLY, INCREASED LASER ENERGY WAS APPLIED AND A 2ND MID PERIPHERAL PI WAS PERFORMED. PROBLEM RESOLVED. THE ELEVATED IOP WAS RELATED TO A NON PATENT SUPER PERIPHERAL PI. Manufacturer Narrative: B5: UPDATED INFO: REPORTEDLY THE PRESSURE IS UP AND THE ANGLE IS SHALLOW. THE DOCTOR BELIEVES THAT THE LENS MAY BE UPSIDE DOWN. HIGH VAULT IS REPORTED. THE DOCTOR WOULD LIKE TO EXCHANGE THE LENS. H6-HEALTH CLINICAL CODE 4581 (SHALLOW ANGLE). CLAIM# (B)(4). Manufacturer Narrative: LENS WAS EXPLANTED AND EXCHANGED AND TO THE REPORTER'S BEST KNOWLEDGE, PATIENT HAS DONE QUITE WELL. #(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938946&pc=QCB | 2023826-2021-01644 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO PUPIL BLOCK WITH ELEVATED IOP(75MMHG,ASSESSED ON (B)(6) 2021), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, NOT ACCURATE MEASUREMENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938953&pc=QCB | 2023826-2021-01645 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO PUPIL BLOCK WITH ELEVATED IOP (75MMHG,ASSESSED ON (B)(6) 2021), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE, NOT ACCURATE MEASUREMENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966819&pc=QCB | 2023826-2021-01819 | 2021-05-12 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH3.2, -10.00/3.0/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ANGLE CLOSURE WITH ELEVATED IOP(64MMHG ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT WERE OBSERVED. ADDITION OF NEW PI, AND YAG PERFORMED ON (B)(6) 2021. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE AND PATIENT-RELATED FACTOR, LENS SLIGHTLY TOO LARGE FOR PATIENT'S EYE. PATIENT IS HIGHLY HYPEROPIC HENCE SMALLER EYE. REPORTEDLY, LENS REPLACEMENT IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966836&pc=QCB | 2023826-2021-01818 | 2021-05-12 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH3.2, -10.00/2.0/024 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ANGLE CLOSURE WITH ELEVATED IOP(76MMHG ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT WERE OBSERVED. ENLARGEMENT OF PI, ADDITION OF NEW PI, AND YAG PERFORMED ON (B)(6) 2021. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE AND PATIENT-RELATED FACTOR, LENS SLIGHTLY TOO LARGE FOR PATIENT'S EYE. PATIENT IS HIGHLY HYPEROPIC HENCE SMALLER EYE. REPORTEDLY, LENS REPLACEMENT IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119272&pc=QCB | 2023826-2021-02281 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS ELEVATED PRESSURE AND POSSIBLE AQUEOUS MISDIRECTION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170834&pc=QCB | 2023826-2021-02603 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.5/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170835&pc=QCB | 2023826-2021-02602 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/1.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178448&pc=QCB | 2023826-2021-02701 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -16.0/+1.0/029 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "DEVICE BEING TOO LARGE CAUSED ANTERIOR VAULT, RESULTING IN NARROWING ANTERIOR CHAMBER AND INCREASED IOP. THE PROBLEM IS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: PER EMAIL ON (B)(6) 2021, THE REPORTER STATED THAT "REPLACEMENT LENS WERE IMPLANTED ON (B)(6) 2021. RIGHT EYE: (B)(6) & LEFT EYE: (B)(6). ALL PROBLEMS HAVE BEEN RESOLVED AT THIS TIME." CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214185&pc=QCB | 2023826-2021-02636 | 2021-05-24 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP(26.7MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2021. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT CLINICAL CODE (B)(4): SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214348&pc=QCB | 2023826-2021-02679 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.50/3.5/044 (SPHERE/ CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. PUPIL BLOCK WITH ELEVATED IOP (65MMHG) WAS ASSESSED ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. MEDICATION PRESCRIBED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE AND PATIENT RELATED FACTOR. REPORTEDLY, NON-REACTIVE PUPIL IS RESPONDING TO MYOSIS WITH THE ADMINISTRATION OF EYE DROPS (PILOCARPINE). SURGEON IS CONSIDERING DISCONTINUING DROPS AND ASSESS IF THE PUPIL REACTS TO LIGHT WITHOUT MEDICATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -2.50/3.5/044 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PUPIL BLOCK; ANGLE CLOSURE AND ELEVATED IOP REPORTED. A WEEK FOLLOWING TILTED LENS; HIGH VAULT; AND FIBRIN IN AC, AND REFRACTIVE SURPRISE WAS OBSERVED. MEDICATION WAS ADMINISTERED. IRIS ATROPHY REPORTED FOLLOWINGLY. LENS REMAINS IMPLANTED. UPDATE ABOUT PATIENT: RESIDUAL ASTIGMATISM 0.75 D; NON-REAGENT PUPIL; PUPIL DIAMETER 4,5/5MM; EYE IS QUIET; AND IOP AS 24MMHG BUT FOLLOWINGLY WAS 34 MMHG. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12228684&pc=QCB | 2023826-2021-02916 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.00/+3.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND IT WAS NOTED VISCOELASTIC WAS DRAINED IN A SEPARATE SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND INDICATED "SINCE LENS WAS TORIC AND WTW WAS IN THE LIMIT FOR BOTH SIZE; WE CHOSE THE BIGGER SIZE". Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE VAULT STILL BEING HIGH. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12264140&pc=QCB | 2023826-2021-03041 | 2021-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13/+1.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)-2021. THE SURGEON REPORTS EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; AND TOPICAL TREATMENT (PROSTAGLANDIN DROPS) WERE USED. THE SURGEON NOTED A REPLACEMENT LENS WAS ORDERED AS A SOLUTION. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406884&pc=QCB | 2023826-2021-03327 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.50/1.0/095 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (35MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2021), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443375&pc=QCB | 2023826-2021-03536 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE PATIENT EXPERIENCED ANGLE CLOSURE AND ELEVATED IOP AFTER ICL IMPLANTATION. PER UPDATED INFORMATION INTRAOCULAR PRESSURE RETURNED TO NORMAL AND PRESSURE IS BEING MONITORED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT:UNK (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478337&pc=QCB | 2023826-2021-03586 | 2021-08-13 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.3MM, TMICL13.2, -6.5/2.5/91, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTOT THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP WAS OBSERVED. IT WAS REPPORTED THAT MEDICAL OR SURGICAL INTERVENTION WAS REQUIRED. AC WASHOUT, VRT TAP AND SUTURE WOUND WAS REPORTED TO BE PERFORMED ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED. A LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND IT WAS REPORTED THAT IT IS UNSURE YET IF THE PROBLEM RESOLVED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): THE REPORTER STATED " 2+ CELL, FORMED AC, LENS EDGE VISIBLE SUP AND INF. ANGLE NARROW AT NASAL & TEMP. OPEN BUT SHALOW, DILATED IRIS." IT WAS ALSO REPORTED THAT LENSES MAY NEED TO BE REPLACED WITH SMALLER SIZE. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12502723&pc=QCB | 2023826-2021-03578 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/088, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. POST-OP, AFTER LENS WAS EXPLANTED, PATIENT HAD FIXED DILATED PUPIL WITH IRIS ATROPHY (URRETS ZAVALIA). THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558563&pc=QCB | 2023826-2021-03805 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE SURGEON REPORTS INTERMITTENT ANGLE CLOSURE WITH INCREASED IOP, PLATEAU IRIS, AND LOW-DOSE PILO ADMINISTRATION. REPORTEDLY THE VISION IS GOOD AND PATIENT DOES NOT WANT TO REMOVE THE LENS; LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12910019&pc=QCB | 2023826-2021-04438 | 2021-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM513.2, -12.50/1.50/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT; ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS DEVICE RELATED TO HIGH VAULT WITH 13.2 LENS SIZE. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12993610&pc=QCB | 2023826-2021-04882 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -10.5/2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING; PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13362153&pc=QCB | 2023826-2022-00348 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -11.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; AND ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13385990&pc=QCB | 2023826-2022-00337 | 2022-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -15.0/+1.0/036 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND NARROW ANGLES. THE PROBLEM IS RESOLVED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -15.0/+1.0/036 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND NARROW ANGLES. THE PROBLEM IS RESOLVED. Manufacturer Narrative: REPORTEDLY THE LENS WAS TOO LONG, CAUSING A SHALLOW ANGLE AND IOP ISSUES. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13385991&pc=QCB | 2023826-2022-00339 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/113 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. REPORTEDLY, "THE PATIENT HAS EYE PRESSURE IN SPITE OF ALL OCULAR TREATMENT." AFTER REMOVAL, THE EYE PRESSURE IS STABLE BUT THE PUPIL IS DILATED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC BROKEN. PRODUCT EVALUATION PERFORMED EVEN THOUGH INSPECTION WAS NOT REQUIRED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423368&pc=QCB | 2023826-2022-00265 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/2.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2021 DUE TO ELEVATED IOP (21.3MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13518185&pc=QCB | 2023826-2022-00479 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.00/+1.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5- PER ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2022: THE PROBLEM WAS REPORTED AS RESOLVED. "PATIENT DOESN'T WISH FURTHER SURGERY PATIENT WEAR GLASSES/CONTACTLENSES". CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13538841&pc=QCB | 2023826-2022-00512 | 2022-01-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.50/+2.0/057 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE; ELEVATED IOP; AND ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.50/+2.0/057 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE; ELEVATED IOP; AND ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H3 DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND WITH FIBERS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552706&pc=QCB | 2023826-2022-00437 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13562126&pc=QCB | 2023826-2022-00486 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, 07.50/+2.0/084 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THERE WAS AN ADDITION OF NEW PI BY YAG. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. VAULT AND IOP AT NORMAL POST ICL TARGET WITH SMALLER ICL. AT ONE WEEK POST-OP, THERE WAS TRACE CELL AND SPK. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. HEALTH IMPACT- CLINICAL CODE: (B)(4). HEALTH IMPACT: (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13562158&pc=QCB | 2023826-2022-00485 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, 07.50/+1.5/095 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THERE WAS AN ADDITION OF NEW PI BY YAG. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. VAULT AND IOP AT NORMAL POST ICL TARGET WITH SMALLER ICL. AT ONE WEEK POST-OP, THERE WAS TRACE CELL AND SPK. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13581966&pc=QCB | 2023826-2022-00593 | 2021-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING; ANGLE CLOSURE AND INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PRODUCT EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588934&pc=QCB | 2023826-2022-00598 | 2021-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/2.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995866&pc=QCB | 2023826-2022-00933 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP (23.3 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14058325&pc=QCB | 2023826-2022-01149 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/+3.5/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. IT WAS REPORTER THAT INTRAOCULAR PRESSURE WAS TAPER DOWN TO 12MMHG. PER UPDATED INFORMATION INTRAOCULAR PRESSURE IS CONTROLLING AND IMPROVING AND DOCTOR CONTINUES TO MONITOR PATIENT. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- "IT WAS ALSO REPORTED HTA THE PATIENT HAD A GRAND MAL SEIZURE IN THE IMMEDIATE POT OPERATIVE PERIOD ON A BACKGROUND OF A HISTORY OF SEIZURE FIVE YEARS PREVIOUSLY." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161436&pc=QCB | 2023826-2022-01083 | 2022-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.5/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(331MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2022), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439471&pc=QCB | 2023826-2022-01368 | 2022-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -9.00/5.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT; ELEVATED IOP AND SIGNIFICANT REDUCTION OF ICA WERE REPORTED. ANTI GLAUCOMA MEDICATIONS ADMINISTERED TO MAINTAIN NORMAL IOP. ENLARGED PUPIL REACTIVE TO LIGHT; IRIS BOWING; PIGMENTS IN ANTERIOR CHAMBER. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697153&pc=QCB | 2023826-2022-01766 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -15.00/3.5/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP (45MMHG ASSESSED ON (B)(6) 2022). ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT IS VERY HAPPY WITH THE NEW VISION. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15156548&pc=QCB | 2023826-2022-02583 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, 5.00/3.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15157019&pc=QCB | 2023826-2022-02584 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2 1.50/5.5/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15168658&pc=QCB | 2023826-2022-02589 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.5/+2.5/74 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE/FIXED PUPIL AND IRIS ATROPHY. IT WAS REPORTED THAT THE EXPLANT DID NOT RESOLVED THE ISSUE. THE REPORTED STATED " AFTER REMOVAL OF THE LENS, THE PUPIL DID NOT CONSTRICT. NO PLAN TO IMPLANT ANOTHER LENS." Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: MEDICAL DEVICE PROBLEM CODE: 1494- OFF LABEL USE: UNDER 21 YRS OF AGE AT TIME OF INITIAL IMPLANTATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15178277&pc=QCB | 2023826-2022-02601 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.50/+2.0/039 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. AT 3 WEEKS POST-OP IOP WAS ELEVATED TO MID 20'S AND PATIENT WAS STARTED ON TRAVATAN EYEDROP. AT 5 WEEKS POST-OP PRESSURE NORMALISED BUT PATIENT HAD OCULAR SURFACE SIDE EFFECTS FROM EYEDROP. PATIENT STARTED ON TREATMENT FOR OCULAR SURFACE DISEASE. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT WAS DIAGNOSED WITH PERIPHERAL ANTERIOR SYNECHIAE. PHACO WAS PERFORMED AFTER REMOVAL OF THE ICL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15179016&pc=QCB | 2023826-2022-02575 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -16.0/1.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282214&pc=QCB | 2023826-2022-02651 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED IOP(51 MMHG) ASSESSED ON (B)(6) 2022, AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334794&pc=QCB | 2023826-2022-02950 | 2022-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS OF 4.00/4.0/013 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(6). EXPIRATION DATE: 11/30/2024. (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15477821&pc=QCB | 2023826-2022-03263 | 2022-08-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -8.0/0.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS INRAOPERATIVELY DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506720&pc=QCB | 2023826-2022-03359 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -4.0/1.0/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO: SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PUPIL BLOCK WITH ELEVATED IOP, ANGLE CLOSURE WITH ELEVATED IOP, PIGMENT DISPERSION, AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15630190&pc=QCB | 2023826-2022-03698 | 2021-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.00/4.50/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. IN THE REPORTER'S OPINION THE CAUSE OF THE EVENT WAS THE DEVICE, THE REPORTER STATED " THE TICL WAS LARGE IN SIZE". FOR ADDITIONAL INFORMATION " EYE IS MONITORED FOR ELEVATED IOP BUT ANGLE IS STILL CLOSED BY THE IRIS AND THE TICL SIZE IS LARGE." WAS REPORTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15659600&pc=QCB | 2023826-2022-03713 | 2022-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/+2.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS OF THE SAME LENGTH DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ELEVATED IOP WITHOUT PUPIL BLOCK/ANGLE CLOSURE WAS ALSO REPORTED. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN,BROKEN AND BENT TO THE LENS. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO DAMAGED STATE OF LENS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15764262&pc=QCB | 2023826-2022-04078 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.50/2.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE SURGEON REPORTS EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE; PIGMENT DISPERSION; UNREACTIVE FIXED PUPIL; AND ADDITIONALLY NOTED "QUIET EYE; UNREACTIVE FIXED DILATED PUPIL/BORDERLINE IOP; SLIGHTLY DISCORIC PUPIL/PATIENT COMPLAINS ALSO FROM THE PUPIL SIZE; THE EXCESSIVE VAULTING VAULT CAUSED IRIS CHAFFING WITH SEVERE PIGMENT DISPERSION AND PIGMENT DEPOSITS ON THE ICL'S OPTIC. ADDITIONALLY IT CAUSED DISCORIC PUPIL AND THE PUPIL BECAME MID-DILATED AND UNRESPONSIVE; ALSO IRIS ATROPHY CAUSE BY IRIS CHAFING". ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE PROBLEM IS REPORTED AS NOTE RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED EXCESSIVE VAULTING USING THE MANUFACTURED SUGGESTED ICL SIZE; THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15910184&pc=QCB | 2023826-2022-04341 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50/+5.5/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE EYE WAS BURNING. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: SHOULD BE CORRECTED TO ADVERSE EVENT: B2: SHOULD BE CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. D4: UNIQUE IDENTIFIER NUMBER: (B)(4). H1: SHOULD BE CORRECTED TO SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50/+5.5/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE EYE WAS BURNING. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SMALLER LENGTH LENS. PROBLEM RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993011&pc=QCB | 2023826-2022-04420 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -4.00/2.5/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. "PATIENT DID NOT WANT TO UNDERGO ANOTHER LENS IMPLANTATION." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993013&pc=QCB | 2023826-2022-04419 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.00/1.5/004 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. "PATIENT DID NOT WANT TO UNDERGO ANOTHER LENS IMPLANTATION." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000862&pc=QCB | 2023826-2022-04423 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. THE LENS REMAINS IMPLANTED AND THE PATIENT DOES NOT WANT THE LENS TO BE REMOVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT 13.2MM VTICM5 13.2 IMPLANTABLE COLLAR LENS OF -13.5/1.0/129 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. THE LENS REMAINS IMPLANTED AND THE PATIENT DOES NOT WANT THE LENS TO BE REMOVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000874&pc=QCB | 2023826-2022-04424 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. THE LENS REMAINS IMPLANTED AND THE PATIENT DOES NOT WANT THE LENS TO BE REMOVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT 13.2MM VTICM5 13.2 IMPLANTABLE COLLAR LENS OF -13.0/1.0/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. REPORTEDLY THE LENS HAS BEEN EXPLANTED AT AN UNKNOWN DATE. H6:4627 ADDED AND 2199 OMITTED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT 13.2MM VTICM5 13.2 IMPLANTABLE COLLAR LENS OF -13.0/1.0/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. REPORTEDLY THE LENS HAS BEEN EXPLANTED ON (B)(6)2023. THE LENS WAS REPLACED WITH A SMALLER LENGTH LENS. PROBLEM RESOLVED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16309076&pc=QCB | 2023826-2023-00446 | 2023-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.5/2.0/050 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE SURGEON REPORTS PIGMENT DISPERSION AND ELEVATED INTRAOCULAR PRESSURE AND MEDICATION IS BEING USED TO LOWER IOP. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED AND NOTED "ICL-IRIS CHAFE". Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: (B)(4) - PIGMENT DISPERSION. CLAIM# (B)(4). Manufacturer Narrative: B5 - "THE LENS WAS EXPLANTED AND REPLACED WITH A VTICM5 12.6 WITH A DIOPTER OF -7.5/2.0/070. REPORTEDLY, " I WAS WITH DR. SKINNER FOR HIS EXPLANT RELATED TO THIS COMPLAINT. HE IMPLANTED A SMALLER SIZE AND IS WATCHING THE PATIENT."" SHOULD BE ADDED TO THE INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16343041&pc=QCB | 2023826-2023-00480 | 2022-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/+1.5/27 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND LENS REPOSITIONING WAS PERFORMED BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL USE (ACD<3.0MM). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16439865&pc=QCB | 2023826-2023-00603 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.50/005 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED DUE TO EXCESSIVE VAULT; ELEVATED IOP AND UNREACTIVE FIXED PUPIL. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: B5 - CORRECTED TO : "THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.50/008 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED DUE TO EXCESSIVE VAULT; ELEVATED IOP AND UNREACTIVE FIXED PUPIL. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE." IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: CLINICAL CODE 4581: UNREACTIVE (FIXED) PUPIL. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773438&pc=QCB | 2023826-2023-01462 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -8.5/1.0/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE SURGEON REPORTS ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ADDITIONALLY NOTED "REMOVED THE LENS AS THE IOP WENT UP WITH LOT OF PIGMENT DISPERSION". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773439&pc=QCB | 2023826-2023-01463 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.0/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ADDITIONALLY NOTED "REMOVED THE LENS AS THE IOP WENT UP WITH LOT OF PIGMENT DISPERSION". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791590&pc=QCB | 2023826-2023-01471 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/+2.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP; 30MMHG. ON (B)(6) 2023 A REPLACEMENT LENS TWO SIZES SHORTER IN LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN AND THE REPORTER STATED IT WAS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT CLINICAL CODE 4581- ANGLE CLOSURE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16873134&pc=QCB | 2023826-2023-01681 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.50/2.5/002 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THERE IS ICL DEPOSITS ON THE LENS, BUT NO REFRACTIVE OUTCOMES OR LOSS OF VISION ACUITY OR VISION QUALITY. BUT LENS TRANSPARENCY SEEMS TO BE AFFECTED; THERE IS A RISE IN IOP AND MEDICATION HAS USED AS TREATMENT. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION. B5 - IT WAS LATER INDICATED THAT THE SURGEON EXPLANTED/REPLACED THE LENS AND THE PATIENT'S POSTOP IS INDICATED AS EVERYTHING IS FINE BUT A 1 MONTH POSTOP IS SCHEDULED SOON. CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - DEPOSITS ON LENS, LENS TRANSPARENCY. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16962964&pc=QCB | 2023826-2023-01910 | 2023-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+0.5/003 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+0.5/003 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. H6 - HEALTH EFFECT CLINICAL CODE: 1937 (INTRAOCULAR PRESSURE RISE). CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16976342&pc=QCB | 2023826-2023-01871 | 2023-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT COMPLAINED OF DOUBLE VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 14-JUN-2023. G3: 14/AUG/2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DOUBLE VISION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16988828&pc=QCB | 2023826-2023-01872 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6)2022. THE LENS WAS EXPLANTED ON (B)(6)2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT COMPLAINED OF DOUBLE VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED AND THE OPTIC TORN. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DOUBLE VISION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17014968&pc=QCB | 2023826-2023-02065 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -11.0/+3.0/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE. CLAIM#:(B)(4). Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT ON 06-JUN-2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17014969&pc=QCB | 2023826-2023-02066 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -11.5/+2.5/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17091758&pc=QCB | 2023826-2023-02289 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+0.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. ANGLE CLOSURE WITH ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE ALSO NOTED. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT CLINICAL CODE 4581- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ANGLE CLOSURE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342540&pc=QCB | 2023826-2023-02889 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP) . THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE MYDRIASIS CONTINUED AFTER GLAUCOMA ATTACK. SEE MFR. REPORT # 2023826-2023-02890 FOR REPLACEMENT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON 27-APR-2023. THE LENS WAS REMOVED ON 01-JUN-2023 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE MYDRIASIS CONTINUED AFTER GLAUCOMA ATTACK. SEE MFR. REPORT # 2023826-2023-02890 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. D6B: 01-JUN-2023. CLAIM # (B)(4). Manufacturer Narrative: A3: UNK. A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17492163&pc=QCB | 2023826-2023-03432 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.50/2.5/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. OCULAR HYPER TENSION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.5/2.5/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. A SHORTER LENGTH DIFFERENT POWERED ( SPHERE/CYLINDER/AXIS) LENS WAS EXCHANGED ON (B)(6) 2023 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP 21MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2023) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK. H4 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: OCULAR HYPER TENSION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED WITHIN REPORT. CORRECTION: H4 REPORTED AS BLANK - CORRECTED TO 16/JUN/2022. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17502541&pc=QCB | 2023826-2023-03460 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.50/+1.0/75 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK, SIGNIFICANT REDUCTION OF IDIRO-CORNEAL ANGLES, ADDITIONALLY MEDICAL INTERVENTION OF COMBIGAN AND COSOPT 3 TIMES A DAY. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE SURGEON NOTED THERE IS A PLANNED LENS EXCHANGE. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS EXPLANTED AND REPLACED IN A DIFFERENT SURGERY ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLEAR CORNEA, FORMED AC, ICL IN PLACE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17502604&pc=QCB | 2023826-2023-03461 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.00/+0.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK, SIGNIFICANT REDUCTION OF IDIRO-CORNEAL ANGLES, ADDITIONALLY MEDICAL INTERVENTION OF COMBIGAN AND COSOPT 3 TIMES A DAY. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE SURGEON NOTED THERE IS A PLANNED LENS EXCHANGE. Manufacturer Narrative: B5 - LENS EXPLANTED AND REPLACED IN A DIFFERENT SURGERY ON (B)(6)2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLEAR CORNEA, FORMED AC, ICL IN PLACE. (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619513&pc=QCB | 2023826-2023-03717 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2; +5.00/+3.0/120 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE SURGEON REPORTS THERE WAS ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK, A YAG-PI WAS PERFORMED ON (B)(6) 2023 WHICH RESOLVED THE PRESSURE, THE SURGEON FELT THERE MAY BE RETAINED VISCOELASTIC AND WAS TREATED WITH DIAMOX AND DILATE THE PUPIL. BUT IT WAS LATER UPDATED THE PUPIL REFUSED TO DILATE BEYOND THE ICL AND MID DILATED PUPIL SHUT THE PI, ADDITIONAL MEDICATION WAS USED. IT WAS FURTHER UPDATED THAT PATIENT IOP IS 10 BUT STILL ON MEDS. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PUPIL DILATION ISSUES, YAG - PI. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669068&pc=QCB | 2023826-2023-03734 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.0/3.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP 21MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2023). WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. H6 - WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17913515&pc=QCB | 2023826-2023-04272 | 2023-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -7.50/1.0/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL DUE TO ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17916405&pc=QCB | 2023826-2023-04520 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/+1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND ON (B)(6) 2023 AN ADDITION OF NEW PI (YAG) WAS PERFORMED. LENS REMAINS IMPLANTED. A FOLLOW UP WILL BE PERFORMED WAS NOTED BY THE SURGEON. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/+1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2023. THIS RESOVED THE PROBLEM. . H6 - CODE 4581 - ANGLE CLOUSRE, SHOULD BE REMOVED. H6- CODE 4625 - YAG, SHOULD BE REMOVED. CLAIM #: (B)(4). Manufacturer Narrative: B5 - IT WAS LATER CLARIFIED THAT THE PROBLEM WAS NOT RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17916406&pc=QCB | 2023826-2023-04521 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.50/+2.5/121 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND ON (B)(6) 2023 AN ADDITION OF NEW PI (YAG) WAS PERFORMED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948886&pc=QCB | 2023826-2023-04441 | 2023-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6, -1.0/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP(30MMHG)(ASSESED ON (B)(6) 2023), AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, NARROW ANGEL IN MYOPIC PATIENT. EPISODES OF HIGH PREESURE WHEN DILATED PUPIL. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL USE (ACD<3.0MM) CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER +4.0/+1.5/89 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, ANGLE CLOSURE AND ELEVATED IOP. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT WAS OFFERED AN ICL EXCHANGE FOR SMALLER LENS AND REFUSED AT THE MOMENT". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS USER ERROR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "NARROW ANGLE IN HYPEROPIC PATIENT. EPISODES OF HIGH PRESSURE WHEN DILATED PUPIL". CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - HEALTH EFFECT CLINICAL CODE 4581: ANGLE CLOSURE, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948893&pc=QCB | 2023826-2023-04442 | 2023-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6, -3.0/3.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP(30MMHG)(ASSESED ON (B)(6) 2023), AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, NARROW ANGEL IN MYOPIC PATIENT. EPISODES OF HIGH PREESURE WHEN DILATED PUPIL. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: ANGLE CLOSURE H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- "-3.0/3.0/101" SHOULD BE CORRECTED TO "+3.0/3.0/101". H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL USE (ACD<3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17970807&pc=QCB | 2023826-2023-04632 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE AND PIGMENT DISPERSION. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) IT WAS REPORTED THAT THE IOP (INTRAOCULAR PRESSURE) IS NORMAL WITH SINGLE DRUG THERAPY. THE REPORTER ALSO STATED " WAITING FOR IOP TO LOWER FURTHER. THE ANGLES ARE OPEN NOW PIGMENT DISPERSION STILL PRESENT". Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4) .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17998693&pc=QCB | 2023826-2023-04775 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -15.50/+2.5/129 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS TO A PRIOR ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. IT WAS REPORTED THAT THE PROBLEM WAS NOT RESOLVED. IT IS IMPLIED THE PREVIOUS PROBLEMS STILL PERSIST. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999332&pc=QCB | 2023826-2023-04699 | 2023-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THERE WAS A SMALL TEAR IN THE ICL DURING LOADING. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE. H6 - WORK ORDER SEARCH: TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048081&pc=QCB | 2023826-2023-04728 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -7.0/1.0/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ELEVATED IOP(69MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: LENS TOO BIG. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 (IRIS COMPROMISE). H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097675&pc=QCB | 2023826-2023-05045 | 2023-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -15.00/+1.5/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, ON (B)(6) 2023 AN ADDITIONAL PI WAS PERFORMED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS NOTED THE PATIENT HAD HIGH IOP AND WAS ON GLAUCOMA DRUGS. PROBLEM IS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18337478&pc=QCB | 2023826-2023-05678 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, 10.00/+1.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). AT ONE MONTH POST-OP, THE PATIENT HAD PERSISTENT ELEVATED INTRAOCULAR PRESSURE (IOP), WHICH WAS BEING TREATED WITH IOP LOWERING MEDICATIONS. THE PATIENT HAD NO SYMPTOMS OF ANGLE CLOSURE AND NO SIGNS OF INFLAMMATION. CURRENT POST-OP FINDS THE PRESSURE HAS STABILIZED ON MEDICATION. THE PATIENT IS BEING WATCHED CLOSELY. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4644 - IOP LOWERING MEDICATIONS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537395&pc=QCB | 2023826-2020-02066 | 2020-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS BEING TOO LARGE AND WILL BE REMOVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: ADDITIONAL SURGERY WAS PERFORMED - ENLARGEMENT OF PI BY YAG. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP PATIENT EXPERIENCED HALO, GLARE AND BLURRED VISION. SEE MFR#: 2023826-2020-02282. THE CAUSE WAS DUE TO USER ERROR. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6: HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY; YAG - ENLARGEMENT OF PI. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE BLA NUMBER: NA. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135926&pc=QCB | 2023826-2021-00050 | 2020-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON THE SAME DATE, IN A SEPARATE SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H10 - CODE 4581: ANGLE CLOSURE, SHOULD HAVE BEEN SPECIFIED IN PREVIOUS SUPPLEMENTAL. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED WITH RESIDUE ON THE LENS. CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE - SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616797&pc=QCB | 2023826-2021-00840 | 2020-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -06.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE ON (B)(6) 2020. THE LENS WILL BE EXPLANTED AND EXCHANGED DUE TO OVER VAULT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -06.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PIS WERE ENLARGED, WITH AN ADDITION OF PI BY YAG ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ENLARGEMENT/ADDITION OF PIS. CORRECTED DATA: B3: (B)(6) 2020. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060666&pc=QCB | 2023826-2021-01987 | 2021-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, EYE IS QUIET. AC WAS FINE AFTER LENS REMOVAL. PATIENT NOT WILLING TO EXCHANGE AT PRESENT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. CLAIM# (B)(4) Manufacturer Narrative: H3: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT AND ELEVATED IOP WITH ANGLE CLOSURE. CLAIM# (B)(4). Manufacturer Narrative: H6: HEALTH EFFECT CLINAL CODE 1937 - ELEVATED IOP. H6: HEALTH EFFECT CLINAL CODE 4581 - ANGLE CLOSURE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161445&pc=QCB | 2023826-2022-01085 | 2022-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.0/2.5/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14647638&pc=QCB | 2023826-2022-01781 | 2022-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.50/3.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXPLANTED DUE TO SIZING ISSUES. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.5/+3.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. FOR CAUSE DETAILS THE REPORTER STATED "AT THE DAY 1 POSTOPERATIVE VISIT, THE UNCORRECTED VISUAL ACUITY MEASURED 20/32 OD. INTRAOCULAR PRESSURE WAS 18 MMHG OD. NARROW ANGLES OU." A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AS REMOVAL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLES. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTEDLY, INCREASED IOP >= 30 MMHG OR > 10 MMHG ABOVE PREOP IOP. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AS REMOVAL. PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER EMAIL ON 21 JUN 2023, PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REPORTEDLY, INCREASED IOP >= 30 MMHG OR > 10 MMHG ABOVE PREOP IOP. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AS REMOVAL. PROBLEM RESOLVED." SHOULD BE OMITTED FOR CORRECTION. H6: 1937 SHOULD BE OMITTED FOR CORRECTION. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17885076&pc=QCB | 2023826-2023-04316 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS), IN (B)(6) 2020. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND HAD ROTATED. THE LENS WAS REPOSITIONED IN (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6B: UNK. H4: UNK. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "1937" SHOULD BE REMOVED. "4581-LENS MISALIGNMENT, LENS CENTRATION ISSUES" AND "4582" SHOULD BE ADDED. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS OF DIOPTER INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2020. THE PATIENT EXPERIENCED "VERY LOW ICL VAULT FROM BEGINNING", LENS MISALIGNMENT, LENS CENTRATION ISSUES, AND LENS REPOSITIONING. THE LENS WAS REPOSITIONED DUE TO "CYLINDER ORIENTATION AND LENS CENTRATION". THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308771&pc=QCB | 2023826-2021-00318 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.00/2.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. AN ADDITIONAL OF NEW PI, YAG AND THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS, DUE TO PUPIL BLOCK WITH ELEVATED IOP (32MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (32MMHG) OBSERVED ON (B)(6) 2020, REFRACTIVE SURPRISE, AND LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "SIZING DIFFICULT, MAY BE BECAUSE OF ASIAN EYE, OR ABNORMAL SULCUS. REPORTEDLY, PATIENT "STABLE VAULT, STABLE IOP, CLEAR CORNEA. PER THE REPORTER ON (B)(6) 2021, "PATIENT HAS SUCCESSFULLY TAKEN THE REPLACEMENT LENS AND IS VERY HAPPY WITH THE RESULTS WITH NO FURTHER PROBLEMS REPORTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT, WITH RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 -15.00/2.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WITH ROTATION; REFRACTIVE SURPRISE; PUPIL BLOCK; ELEVATED IOP; AND ANGLE CLOSURE WERE REPORTED. THE LENS WAS EXPLANTED AND LATER REPLACED FOR A LONGER LENGTH LENS ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. THE REPORTER ALSO NOTES "STABLE ICL VAULT, IOP, VISION, TEMPORAL SUTURE IN SITU." CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11787674&pc=QCB | 2023826-2021-01388 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/2.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND ELEVATED IOP. LENS REMAINS IMPLANTED. FOR STATUS OF THE EYE " PATIENT CANNOT READ ANYMORE " WAS REPORTED. FOR ADDITIONAL INFORMATION THE REPORTER REPORTED " HIGH IOP OS: PROBLEM RESOLVED WITH DIAMOX AND PRESSURE WAS AFTER USING DIAMOX 20 MM HG. PROBABLY SOME VISCO REMAINED IN THE EYE. BUT THE PUPIL IN THIS EYE IS STILL ASYMMETRIC? THEY DONT KNOW WHAT IS THE CAUSE OF THIS REFRACTIVE SURPRISE. DOCTOR WANTS TO DO A LASER TO SOLVE THE REFRACTIVE SURPRISE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14507171&pc=QCB | 2023826-2022-01692 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -18.00/2.5/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ONE HAPTIC WAS NOTED TO BE TORN AFTER IMPLANTATION. PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, SUBLUXATION AND LESS VISION. MANAGEMENT WAS GIVEN FOR ELEVATED INTRAOCULAR PRESSURE. LENS REMAINS IMPLANTED. FOR ADDITIONAL INFORMATION THE REPORTER STATED, "CORNEA CLEAR, VAULT- TEMPORARILY SUBLAXATED DUE TO LESS SUPPORT (HAPTIC TEAR)". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED. EVERYTHING WENT WELL, PATIENT AND THE SURGEON ARE HAPPY WITH RESULTS. VA 6/9, VAULT OK, IOP 13MMHG. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17962122&pc=QCB | 2023826-2023-04570 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Patient-Device Incompatibility | Intraocular Pressure Increased; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/6.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING KERATOCONUS ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SUPRISE. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE IS "STABLE." CAUSE OF THE EVENT IS A PATIENT-RELATED-FACTOR WITH THE LENS NOT PERFORMING AS INTENDED. REPORTEDLY, "PATIENT HAS KERATOCONUS WHICH MAKES THE RESULT UNPREDICTABLE." Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/6.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND ELEVATED IOP. ON (B)(6) 2023, THE LENS WAS EXPLANTED BUT THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "STABLE". REPORTEDLY "THE ICL WAS REMOVED BECAUSE IT WAS TO BIG FOR THE EYE OF THE PATIENT SO THE EYE PRESSURE INCREASED". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "PATIENT HAS KERATOCONUS WHICH MAKES THE RESULT UNPREDICTABLE". H6: HEALTH EFFECT- CLINICAL CODE: "1937 SHOULD BE ADDED. CLAIM#: (B)(6). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL/PATIENT RELATED FACTOR: PRE-EXISTING KERATOCONUS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10875424&pc=QCB | 2023826-2020-02715 | 2020-07-18 04:00 | Injury | STAAR SUGICAL COMPANY | 2020-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A12.6MM, VTICMO12.6, -16.0/+1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO "TILTED ICL", ELEVATED IOP, PUPIL BLOCK, LENS DISLOCATION/SUBLUXATION AND BLURRED VISION. ADDITION OF NEW PI WAS PERFORMED AND A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON (B)(6) 2020. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11070811&pc=QCB | 2023826-2020-03177 | 2020-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED HIGH VAULT AND IOP (INTRAOCULAR PRESSURE), UNRESPONSIVE TO MAXIMUM MEDICAL THERAPY. DOCTOR EXPLANTED THE LENS AS HE COULD NOT REDUCE THE IOP ADEQUATELY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -15.50/1,0/118, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO TRANSIENT LOSS OF BCVA AND ANGLE CLOSURE WITH ELEVATED IOP. IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. H3- DEVICE EVALUATION : LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN A SPECIMEN CUP. VISULA INSPECTION FOUND ONE OF THE HAPTIC TORN AND RESIDUE AND DEBRIS ON LENS. H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308736&pc=QCB | 2023826-2021-00317 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. AN ADDITIONAL OF NEW PI, YAG AND THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS, DUE TO PUPIL BLOCK WITH ELEVATED IOP (32MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP ((32MMHG) OBSERVED ON (B)(6) 2020), REFRACTIVE SURPRISE, AND LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "SIZING DIFFICULT, MAY BE BECAUSE OF ASIAN EYE, OR ABNORMAL SULCUS. REPORTEDLY, PATIENT "STABLE VAULT, STABLE IOP, CLEAR CORNEA. PER THE REPORTER ON (B)(6) 2021, "PATIENT HAS SUCCESSFULLY TAKEN THE REPLACEMENT LENS AND IS VERY HAPPY WITH THE RESULTS WITH NO FURTHER PROBLEMS REPORTED." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT, WITH RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BENT AND DEFORMED WITH RESIDUE AND DEBRIS ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. AN ADDITIONAL OF NEW PI, YAG AND THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS, DUE TO PUPIL BLOCK WITH ELEVATED IOP (32MMHG), LENS ROTATION AND ANGLE CLOSURE WITH ELEVATED IOP ((32MMHG) OBSERVED ON (B)(6) 2020), REFRACTIVE SURPRISE, AND LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "SIZING DIFFICULT, MAY BE BECAUSE OF ASIAN EYE, OR ABNORMAL SULCUS. REPORTEDLY, PATIENT HAS "STABLE VAULT, STABLE IOP, CLEAR CORNEA, AND TEMPORAL SUTURE IN SITU." PER THE REPORTER ON (B)(6) 2021, "PATIENT HAS SUCCESSFULLY TAKEN THE REPLACEMENT LENS AND IS VERY HAPPY WITH THE RESULTS WITH NO FURTHER PROBLEMS REPORTED." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366547&pc=QCB | 2023826-2021-00485 | 2020-12-18 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -8.5/+1.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), PERIPHERAL ANTERIOR SYNECHIA (PAS), AND A "LITTLE" MYDRIASIS. AN ADDITION OF PI-YAG WAS PERFORMED, PILOCARPINE EYE DROPS ADMINISTERED, IRRIGATION AND ASPIRATION (IA) PERFORMED. THE PAS WAS RELEASED BY PEELING OFF THE ADHESIONS OF THE ANGLES BY A ROUND NEEDLE. THE PROBLEM IS NOTED AS RESOLVED THOUGH THE PUPIL IS REPORTED AS SLIGHTLY DILATED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939022&pc=QCB | 2023826-2021-01669 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL3.2, -16.00/2.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP(32MMHG, ASSESSED ON (B)(6) 2021) ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: ELEVATED IOP AND BLURRINESS WAS NOTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939023&pc=QCB | 2023826-2021-01670 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL3.2, -16.50/2.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP (25MMHG, ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: BLURRY VISION AND ANGLE CLOSURE OBSERVED. H10: CODE 4581 - ANGLE CLOSURE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222320&pc=QCB | 2023826-2021-02802 | 2021-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SECTORAL IRIS ATROPY. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY THE PROBLEM IS NOT RESOLVED THOUGH THE LENS HAS BEEN EXPLANTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS . CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/3.50/95 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2021. A REPLACEMENT LENS WAS LATER IMPLANTED OF SHORTER LENGTH LENS. "THE SPECIALIST REPORTS THAT THE AGREEMENT IS FINE, WITH A LITTLE BLURRED VISION IN THE IMMEDIATE POST-OPERATIVE CONTROL, IT HAS OPEN ANGLES OF 2 PER GONIOSCOPY, VAULT OF 700 MICRONS, STABLE IOP OF 10/10 MMHG. THE PATIENT FINDS HIMSELF BETTER WITH THE EXCHANGE, SURGERY WITHOUT COMPLICATIONS." H6- HEALTH EFFCET CLINICAL CODE; 4581-OPEN ANGLES OF 2 PER GONIOSCOPY. CLAIM# (B)(4) Manufacturer Narrative: WEIGHT, ETHNICITY, RACE -UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12874494&pc=QCB | 2023826-2021-04515 | 2021-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -14.5/2.5/62, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REMOVED AND LATER REPLACED WITH A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO ANGLE CLOSURE WITH ELEVATED IOP. IT WAS REPORTED THAT THE PROBLEM RESOLVED. FOR STATUS OF THE EYE "BLURRY VISION" WAS REPORTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B3: 05-OCT-2021. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061327&pc=QCB | 2023826-2022-01130 | 2022-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -13.0/+2.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO BLURRED VISION, IRIS SPHINCTER LOSS AT 5 O'CLOCK, HEAVILY PIGMENTS TM, KIV EXPLANT, AND IRREGULAR PUPIL. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- ELEVATED IOP SHOULD HAVE BEEN REPORTED AS AN ADVERSE EVENT IN INITIAL MDR. H6- INTRAOCULAR PRESSURE INCREASED . CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15221446&pc=QCB | 2023826-2022-02760 | 2020-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/1.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15556104&pc=QCB | 2023826-2022-03511 | 2022-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.00/6.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO HIGH INTRAOCULAR PRESSURE WITH ANGLE CLOSURE AND CHANGE IN REFRACTION. CLOUDY CORNEA WAS ALSO REPORTED, HOWEVER THE REPORTER STATED, " REGARDING THE SYMPTOMS RELATED TO THE EXPLANTED LENS, EVERYTHING RESOLVED." IN THE REPORTER'S OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17600761&pc=QCB | 2023826-2023-03684 | 2023-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -3.00/+3.00/126 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZE". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), KERATOCONUS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17608279&pc=QCB | 2023826-2023-03689 | 2023-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.00/+5.50/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZE". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), KERATOCONUS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17701985&pc=QCB | 2023826-2023-03775 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.00/0.5/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT PRESENTED WITH BLURRED VISION, PIGMENT DISPERSION AND ELEVATED IOP IN THE EARLY POST-OP PERIOD. IT WAS REPORTED THAT THE PROBLEM RESOLVED WITH MEDICAL MANAGEMENT. THE REPORTER STATES THAT THE DISPERSED PIGMENTS DECREASED WITH NORMALIZATION OF IOP AFTER CONSERVATIVE MANAGEMENT WITH NO BLURRING OF VISION. HOWEVER, THE PUPIL REMAINED SLIGHTLY DILATED AND IRREGULAR. PATIENT WAS LOST TO FOLLOW-UP. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17705120&pc=QCB | 2023826-2023-03951 | 2021-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT PRESENTED BLURRED VISION, PIGMENT DISPERSION AND ELEVATED IOP IN THE EARLY POST-OP PERIOD. IT WAS REPORTED THAT THE PROBLEM RESOLVED WITH MEDICAL MANAGEMENT. THE REPORTER STATES THAT THE DISPERSED PIGMENTS DECREASED WITH NORMALIZATION OF IOP AFTER CONSERVATIVE MANAGEMENT WITH NO BLURRING OF VISION. HOWEVER, THE PUPIL REMAINED SLIGHTLY DILATED AND IRREGULAR. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE - 4581 - PUPIL ANOMALY. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17885010&pc=QCB | 2023826-2023-04315 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+0.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS HAD A LOW VAULT. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP), PIGMENT DISPERSION AND BLURRED VISION. OCT PERFORMED ON (B)(6) 2023 REVEALED RNFL DAMAGE AND TEMPORAL SCOTOMA. MEDICATION WAS PRESCRIBED. THE LENS REMAINED IMPLANTED. POST-OP, GOOD IOP DURING 6 HOURS AFTER DROPS. PROSTAGLANDIN INHIBITOR WAS ADDED AT NIGHT. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, RETINAL NERVE DAMAGE, TEMPORAL SCOTOMA H6: HEALTH EFFECT IMPACT CODE: 4644 - HYPOTENSIVE DROPS - S. COSOPT (DORZOLAMIDE/TIMOLOL) X2, S. LUXFEN (BIMATOPROST) X2 H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+0.5/69 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED LOW VAULT, ELEVATED IOP, PIGMENT DISPERSION, OPTIC NERVE DAMAGE WITH VISUAL FIELD LOSS, FOGGY VISION, BLURRED VISION, LASER YAG PIS PERFORMED AND MEDICATION WAS PRESCRIBED. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. H6-HEALTH EFFECT- CLINICAL CODE: 4581- "FOGGY VISION" AND "OPTIC NERVE DAMAGE WITH VISUAL FIELD LOSS" SHOULD BE ADDED. H6-HEALTH IMPACT CODE: 4625- "'LASER YAG PI' WAS PERFORMED" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4614" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525781&pc=QCB | 2023826-2024-00170 | 2023-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM512.6, -11.50/3.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ANTERIOR CHAMBER IRRIGATION/VACUATION OF VISCO/FLUIDS AND ANTERIOR CHAMBER WASH ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ELEVATED IOP(40MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(DATE ASSESSED (B)(6) 2023 & (B)(6) 2023), PIGMENT DISPERSION, AND BLURRED VISION. THIS DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: PIGMENT DISPERSION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12643683&pc=QCB | 2023826-2021-04014 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Halo; Eye Pain | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, 5.00/3.0/094, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). GLARE/HALOES, PAIN PRESSURE, DROP OF VISION WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. EVENT DATE: UNK. IMPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12643685&pc=QCB | 2023826-2021-04015 | 2021-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Halo; Eye Pain | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, 6.00/1.0/090, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). GLARE/HALOES, PAIN PRESSURE, DROP OF VISION WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, 6.00/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED AT AN UNKNOWN DATE DUE TO EXCESSIVE VAULT, PAIN, PRESSURE AND DROP IN VISION. ON (B)(6) 2022 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED, AND THE PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8618161&pc=QCB | 2023826-2019-00798 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 15.5/+2.0/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO ICL WAS TOO BIG. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8850038&pc=QCB | 2023826-2019-01367 | 2019-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, TMICL 12.1, INTRAOCULAR LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2019 WAS REMOVED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP AND ANGLE CLOSURE. AT PATIENTS LAST VISIT ((B)(6) 2019) STATUS OF THE EYE WAS REPORTED AS DILATED AND A/C CELLS. ON (B)(6) 2019 IT WAS REPORTED THAT THE EXPLANT RESOLVED THE ISSUES AND PATIENT IS STILL RECOVERING. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE 3191: UNREACTIVE (FIXED) PUPIL, PIGMENT DISPERSION. DATE OF EVENT: CORRECTED TO (B)(6) 2019 IN INITIAL MDR. THE REPORTER INDICATED THAT A 12.1MM TMICL12.1 INTRAOCULAR LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, ANGLE CLOSURE, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, AND BLURRED VISION. IT WAS REPORTED THAT THE EXPLANT RESOLVED THE ISSUE. HOWEVER, THE POST-OP DATA STATES THAT THE PATIENT IS STILL EXPERIENCING DILATED PUPIL WITH A/C CELLS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9282596&pc=QCB | 2023826-2019-02091 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.5/1.0/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE) AND BLURRED VISION. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND DEBRIS ON THE LENS BUT NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454484&pc=QCB | 2023826-2019-02342 | 2018-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICH2.6, 6.50/1.5/159 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THAT THE LENS HAS BEEN EXPLANTED BY ANOTHER DOCTOR DUE TO ELEVATED IOP AND BLURRED VISION. CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION (LENS EXPLANT) H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454586&pc=QCB | 2023826-2019-02343 | 2018-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICH2.6, 5.00/1.0/031 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. IT WAS REPORTED THAT THE LENS HAS BEEN EXPLANTED BY ANOTHER DOCTOR DUE TO ELEVATED IOP AND BLURRED VISION. CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION (LENS EXPLANT). H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18210943&pc=QCB | 2023826-2023-05250 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2023 ELEVATED IOP (34MMHG), POOR VISUAL ACCUITY WAS OBSERVED. TAPERING TOPICAL STEROID AND ALPHAGAN IS BEING TAKEN. LENS REMAINS IMPLANTED. PROBLEM NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558794&pc=QCB | 2023826-2021-00700 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -15.5/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS IOP CONTROLLED WITH ALPHAGAM, HIGH VAULT, SHALLOW AC, "CRAPPY" VISION AT 20/200. PUPIL IS DILATED AND AR IS +2.75. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REFRACTIVE SURPRISE, SIGNIFICANT REDUCTION OF ICA, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, HYPEROPIC SURPRISE +2.00, GLARE/HALOS, BLURRED VISION. REPORTEDLY VISION IS NOW "OK" AT +1.00 HYPEROPIC. PILOCARPINE USED TO CONSTRICT PUPIL AND SET THE ICL IN THE SULCUS. H6-CLINICAL CODE 4581 (SIGNIFICANT REDUCTION OF ICA, PIGMENT DISPERSION, UNREACTIVE/FIXED PUPIL). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664235&pc=QCB | 2023826-2022-00658 | 2022-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -13.5/+2.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED. ADDITIONAL SYMPTOMS REPORTED: ELEVATED IOP, GLARE/HALOS, PIGMENT DISPERSION, AND BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. TRACE PIGMENT IS PRESENT ON THE ENDOTHELIUM AT 6 O'CLOCK AFTER EXCHANGE. HOWEVER, THE EVENT IS UNLIKELY TO BE CAUSED BY THE DEVICE AND NO INTERVENTION IS PLANNED. AN ENLARGEMENT OF PI-YAG WAS PERFORMED AFTER LENS EXPLANT AND PRIOR TO REPLACEMENT LENS IMPLANT. H6-CLINICAL CODE 4581 (PIGMENT DISPERSION). CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. IMPLANT DATE, EXPLANT DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18060314&pc=MTA | 2023826-2023-04944 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED. DOCTOR REQUESTED CONSULTATION FOR PATIENT WITH ICL ROTATION AND PIGMENT DISPERSION. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: 4581-PIGMENT DISPERSION CLAIM#: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18060314&pc=QCB | 2023826-2023-04944 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED. DOCTOR REQUESTED CONSULTATION FOR PATIENT WITH ICL ROTATION AND PIGMENT DISPERSION. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: 4581-PIGMENT DISPERSION CLAIM#: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112743&pc=QCB | 2023826-2019-01626 | 2019-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICH12.1, -8.00/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT, ELEVATED IOP, LENS DISLOCATION/SUBLUXATION, GLARE/HALOES, AND BLURRED VISION. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED, EYE IS QUIET. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). PRE-EXISTING CONDITION: NARROW ANTERIOR CHAMBER ANGLES. ADDITIONAL INFORMATION: PRE-EXISTING CORNEAL OR REFRACTIVE SURGERY. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140318&pc=QCB | 2023826-2021-02397 | 2021-06-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. GOOD POSITION AND VAULT, SYMPTOMS (GLARE) RESOLVED AND IOP NORMAL. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: PT INFO: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14416526&pc=QCB | 2023826-2022-01590 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: HE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2; +7.5/+6.0/161 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULT; REFRACTIVE SURPRISE; RESIDUAL ASTIGMATISM; ELEVATED IOP; GLARE/HALOES; BLURRED VISION. THE LENS WAS EXPLANTED AND LATER REPLACED WITH A SAME LENGTH LENS BUT DIFFERENT POWER/AXIS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND A PATIENT RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC SPLIT AND HAPTIC BROKEN. DIMENSIONAL/FUNCTION INSPECTION COULD NOT BE PEFORMED DUE TO SIGNIFICANTLY DAMAGED LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416136&pc=QCB | 2023826-2021-00506 | 2021-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.50/3.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP(44MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, GLARE/HALOES, BLURRED VISION, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, "UNUSUAL ANATOMY". REPORTEDLY, "ALL LOOKS GOOD, MUCH LESS GLARE IN NEW LENS PER PATIENT." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938959&pc=QCB | 2023825-2021-01652 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -14.50/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK WITH ELEVATED IOP(50MMHG), ANGLE CLOSURE WITH ELEVATED IOP, PIGMENT DISPERSION, UNREACTIVE(FIXED) PUPIL, GLARE/HALOES, BLURRED VISION, IRIS ATOROPHY, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, "UNFORTUNATELY DURING THE ICL EXCHANGE BECAUSE OF THE LIMITED DISTANCE BETWEEN THE ICL AND THE ENDOTHELIUM SHE SUSTAINED A 1MM TEAR IN DESCEMET'S MEMBRANE. AS OF NOW WE ARE WAITING TO SEE IF HER CORNEA WILL CONTINUE TO CLEAR. WHEN THE CORNEA CLEARS ENOUGH WE INTEND TO ASSESS THE STATUS OF THE ENDOTHELIUM TO SEE HOW MUCH DAMAGE SHE HAS SUSTAINED BY THE THREE PROCEDURES ON THIS EYE. CURRENTLY HER ANGLES ARE ADEQUATE IN BOTH EYES AND THE IOP IS ACCEPTABLE." Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US, (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14809062&pc=QCB | 2023826-2022-01953 | 2022-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00/150/66 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PATIENT HAS EXPERIENCED EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE, GLARE/HALOES AND BLURRED VISION. FOR PATIENTS CURRENT CONDITION THE REPORTER STATED " IRIS PERIPHERAL ANTERIOR SYNACHIAE WHICH REQUIRES REMOVAL OF CURRENT LENS SIZING OF 13.2MM TO 12.6MM SIZING. THE INITIAL LENS IS STILL IN PATIENT'S EYES." FOR ADDITIONAL INFORMATION THE REPORTER STATED " PATIENT HAS IRIS ATROPHY DUE TO HIGH LENS VAULTING BUT DR HAS BEEN MONITORING THIS CASE FOR 2 YEARS. NEEDS A SECONDARY INTERVENTION AS WELL DUE TO IRIS DISTORTION AND CAUSING PATIENT TO HAVE VISUAL DISTURBANCE ESPECIALLY AT NIGHTS." IN THE REPORTS OPINION THE CAUSE OF THIS EVENT IS USER ERROR. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON 28-JUN-2022 AND REPLACED WITH A SHORTER LENGTH LENS ON THE SAME DATE BUT DIFFERENT SURGERY AND THE PROBLEM RESOLVED. THE REPORTER STATED "ANTERIOR SYNECHIAE AND IRIS ATROPHY STILL PERSIST BUT VAULTING REDUCED AND EYE LOOKED BETTER, LESS GLARE." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#:(B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17060362&pc=QCB | 2023826-2023-02031 | 2023-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO TRANSIENT LOSS OF BCVA (ASSESSED ON (B)(6)2023, ELEVATED IOP (50MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2023). PIGMENT DISPERSION. GLARE, HALOES, AND BLURRED VISION. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PARACENTESIS DONE AND DIAMOX GIVEN TO REDUCE IOP. PATIENT DOES NOT WANT NEW LENS IMPLANTED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6: HEALTH EFFECT CLINICAL CODE 4581: PIGMENT DISPERSION. H6:WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT - IMPACT CODE: "4625- PARACENTESIS DONE." B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER-13.5/+1.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED TRANSIENT LOSS OF BCVA, ELEVATED IOP, PIGMENT DISPERSION, GLARE/HALOES, BLURRED VISION, DILATED PUPIL, MEDICATION WAS PRESCRIBED. REPORTEDLY "PARACENTESIS DONE AND GAVE DIAMOX TO REDUCE IOP." ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ALSO REPORTED, "IOP REDUCED AND PUPIL STILL DILATE AFTER REMOVAL. PATIENT NOT KEEN TO IMPLANT ANOTHER ICL ANYMORE." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393935&pc=QCB | 2023826-2015-00168 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.0/+3.5/174 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS HAS AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY, SIGNIFICANT REDUCTION OF THE IRIDO-CORNEAL ANGLES (ICA), UNREACTIVE PUPIL (FIXED), ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, GLARE/HALOS, BLURRED VISION AND IRIS ATROPHY. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. EXPLANTED DATE UNKNOWN. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED AND OBSERVED NO VISIBLE DAMAGE TO THE LENS. (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). RESUBMISSION OF INITIAL MDR REQUESTED BY FDA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528011&pc=QCB | 2023826-2019-02495 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 6.50/+2.0/068 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO HIGH INTRAOCULAR PRESSURE (IOP) AND EXCESSIVE VAULTING. THERE WAS NO PATIENT INJURY. THE LENS WAS NOT REPLACED WITH ANOTHER LENS AT THIS TIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. A YAG WAS PERFORMED TO ENLARGE THE PIS. CAUSE OF THE EVENT WAS UNKNOWN/OTHER (PI'S WIDE ENOUGH? NEEDS SMALLER LENS?) . THE PATIENT'S POST-OP CONDITION WAS SOME PIGMENT DISPERSION. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CONTACT CASE, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, YAG PERFORMED TO ENLARGE PIS H6 - PATIENT CODE: 3191 - PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - CORRECTED TO (B)(6) 2019 IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, THE "PIGMENT DISPERSION HAS IMPROVED BUT HAS NOT YET FULLY RESOLVED." THE PATIENT STATES "VISION HAS IMPROVED SINCE PROCEDURE OVERALL." CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528024&pc=QCB | 2023826-2019-02494 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 7.00/+3.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S EYE ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO HIGH INTRAOCULAR PRESSURE (IOP) AND EXCESSIVE VAULTING. THERE WAS NO PATIENT INJURY. THE LENS WAS NOT REPLACED WITH ANOTHER LENS AT THIS TIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. A YAG WAS PERFORMED TO ENLARGE THE PIS. CAUSE OF THE EVENT WAS UNKNOWN/OTHER (PI'S WIDE ENOUGH NEEDS SMALLER LENS) . THE PATIENT'S POST-OP CONDITION WAS SOME PIGMENT DISPERSION. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A CONTACT CASE, WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, YAG PERFORMED TO ENLARGE PIS. PATIENT CODE: 3191 - PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: IT WAS LATER INDICATED BY THE REPORTER THAT AS OF 25/FEB/2020 "THE PATIENT'S PIGMENT DISPERSION HAS IMPROVED BUT HAS NOT FULLY HEALED. THE PATIENT'S VISION HAS IMPROVED SINCE THE PROCEDURE OVERALL THOUGH" CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140634&pc=QCB | 2023826-2020-01200 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT REPORTS ELEVATED IOP WHICH HAS BEEN RESOLVED BUT PAIN PERSISTS. ALSO REPORTED: FIXED MYDRIASIS, NO RESPONSE TO LIGHT STIMULI, LASER IRIDOTOMY, PILOCARPINE ADMINISTRATION, CONTINUOUS GRAVITATIONAL PAIN-FEELING SOMETHING PUSHING ON THE EYE, HALOS, CIRCLES, AND BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14957809&pc=QCB | 2023826-2022-02108 | 2022-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.50/+3.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE AND DISCOMFORT. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14958057&pc=QCB | 2023826-2022-02109 | 2022-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.0/+4.0/65 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE AND DISCOMFORT. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908414&pc=QCB | 2023826-2019-01486 | 2019-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50/+3.0/106 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE PATIENT REPORTED DISCOMFORT. THE LENS WILL BE EXPLANTED BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8913315&pc=QCB | 2023826-2019-01487 | 2019-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50/+2.5/094 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE PATIENT REPORTED DISCOMFORT. THE LENS WILL BE EXPLANTED BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10915175&pc=QCB | 2023826-2020-02800 | 2020-10-21 | Injury | STAAR SUGICAL COMPANY | 2020-11-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.2, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE REPORTER STATED, " POM1 WENT TO ED AROUND MIDNIGHT FOUND TO HAVE IOP OF 50. GIVEN NARCOTICS, DIAMOX, AND TIMOLOL, AND SEEN IN CLINIC THE NEXT AM W/ VAULT-200 AND IOP OF 10. PT. REPORTED SX OF EYE PAIN/PRESSURE ONCE SINCE THAT EVENT. ON ALL OTHER F/U'S, VAULT- 700." IT WAS REPORTED THAT THE SURGEON, " OPTED TO ENLARGE THE LPIS .FOR PATIENT CONDITION (LAST VISIT DATE: (B)(6) 2020) THE REPORTER REPORTED BCVA 20/20-2; UCVA 20/25-2; VAULT 650. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS), "QUIET; PT ASYMPTOMATIC. ENLARGED LPIS AT THIS VISIT," WAS REPORTED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (EXPIRATION DATE): UNK, SERIAL NUMBER REPORTED MISSING A DIGIT. (DEVICE MANUFACTURING DATE): UNK, SERIAL NUMBERS REPORTED MISSING A DIGIT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17454729&pc=QCB | 2023826-2023-03241 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00/1.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO HIGH VAULT CAUSING PAIN AND HIGH PRESSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED ON THE SAME DATE AND THE PROBLEM RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271533&pc=QCB | 2023826-2023-05502 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (INTRAOCULAR PRESSURE) ANS PAIN. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278375&pc=QCB | 2023826-2021-03031 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/108 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS OTHER - ICL LENS LENGTH CALCULATION ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS WAS EXCHANGED ON (B)(6) 2021 WITH A DIFFERENT MODEL, SIMILAR POWER, SHORTER LENGTH LENS WHICH RESOLVED THE ISSUE. REPORTEDLY, DOCTOR AND PATIENT HAPPY WITH THE FINAL RESULT. THE REPORTER MARKED THE CAUSE OF THE EVENT AS PATIENT-RELATED FACTOR. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 721078.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15469467&pc=QCB | 2023826-2022-03140 | 2022-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Iritis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS OF -16.50/1.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO IRITIS/IOP ELEVATION. LENS WAS NOT REPLACED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15474202&pc=QCB | 2023826-2022-03244 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Iritis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS OF --17.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO IRITIS/HIGH IOP. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SEPIMENT CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16479533&pc=QCB | 2023826-2023-00722 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Intraocular Pressure Increased; Iritis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE EYE OF A PATIENT WITH PRE-EXISTING GLAUCOMA ON (B)(6) 2022. PIGMENT DISPERSION AND IRITIS WERE OBSERVED. CAUSE OF THE EVENT IS BOTH A UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY THE STATUS OF THE EYE IS, "MOSTLY QUIET" AND "TAPERING PREDNISOLONE INITITATED LATANOPROST IN MY OPINION BEST IRRESPECTIVE OF PIGMENT DISPERSION, IRITIS.". Manufacturer Narrative: B5 - "REPORTEDLY, 'BOTH ICLS WERE REMOVED TUESDAY (B)(6) 2023 FOR PERSISTENT IOP ISSUES FROM PIGMENT DISPERSION. MAXIMAL MEDICAL THERAPY HAD BEEN TRIED, AS WELL AS SLT LASER FOR THE LEFT.'" SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE: (B)(6) 2023 SHOULD BE ADDED TO THE INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE EYE OF A PATIENT WITH PRE-EXISTING GLAUCOMA ON (B)(6) 2022. PIGMENT DISPERSION, PRESSURE SPIKES, AND IRITIS ARE OBSERVED. REPORTEDLY THE STATUS OF THE EYE IS, "MOSTLY QUIET" AND "TAPERING PREDNISOLONE INITIATED LATANOPROST IN MY OPINION BEST IRRESPECTIVE OF PIGMENT DISPERSION, IRITIS." REPORTEDLY, 'BOTH ICLS WERE REMOVED TUESDAY (B)(6) 2023 FOR PERSISTENT IOP ISSUES FROM PIGMENT DISPERSION. MAXIMAL MEDICAL THERAPY HAD BEEN TRIED, AS WELL AS SLT LASER FOR THE LEFT. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. ADVERSE EVENT PROBLEM: PATIENT RELATED FACTOR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16585378&pc=QCB | 2023826-2023-00994 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Intraocular Pressure Increased; Iritis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE EYE OF A PATIENT WITH PRE-EXISTING GLAUCOMA ON (B)(6) 2022. PIGMENT DISPERSION AND IRITIS WERE OBSERVED. CAUSE OF THE EVENT IS BOTH A UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY, THE STATUS OF THE EYE IS, "MOSTLY QUIET" AND "TAPERING PREDNISOLONE INITIATED LATANOPROST IN MY OPINION BEST IRRESPECTIVE OF PIGMENT DISPERSION, IRITIS." Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. ADVERSE EVENT PROBLEM: PATIENT RELATED FACTOR. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: D6B: EXPLANTATION DATE: 28-MAR-2023 ADDED FOR CORRECTION. H1: SERIOUS INJURY SHOULD BE ADDED AND MALFUNCTION OMITTED FOR CORRECTION. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE EYE OF A PATIENT WITH PRE-EXISTING GLAUCOMA ON (B)(6) 2022. PIGMENT DISPERSION, PRESSURE SPIKES, AND IRITIS ARE OBSERVED. REPORTEDLY THE STATUS OF THE EYE IS, "MOSTLY QUIET" AND "TAPERING PREDNISOLONE INITIATED LATANOPROST IN MY OPINION BEST IRRESPECTIVE OF PIGMENT DISPERSION, IRITIS." REPORTEDLY, 'BOTH ICLS WERE REMOVED TUESDAY (B)(6) 2023 FOR PERSISTENT IOP ISSUES FROM PIGMENT DISPERSION. MAXIMAL MEDICAL THERAPY HAD BEEN TRIED, AS WELL AS SLT LASER FOR THE LEFT. THE ONLY THING WE DID AFTER SURGERY WAS PERFORM A GONIOTOMY ON THE LEFT EYE AT THE TIME OF THE REMOVAL.' CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10969886&pc=QCB | 2023826-2020-02884 | 2020-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Iritis; Pain; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/018 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. THE PATIENT DEVELOPED AQUEOUS MISDIRECTION AND EXPERIENCED PAIN. THE SURGEON ENLARGED THE PI BY SURGICAL PROCEDURE. ANOTHER ICL LENS WAS NOT IMPLANTED. THE PATIENT IS IMPROVING, WITH SOME RESIDUAL IRITIS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405559&pc=QCB | 2023826-2021-00562 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/092, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT AFTER LENS REPOSITIONING THE PATIENT'S IOP INCREASED, ALTHOUGH THE PROBLEM WAS STILL RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241308&pc=QCB | 2023826-2021-02895 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO INTRAOCULAR PRESSURE (IOP) WAS ELEVATED. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK, SEX: UNK, WEIGHT: UNK, RACE: UNK, ETHNICITY: UNK, DATE OF EVENT: UNK, EXPIRATION DATE UNK IMPLANT DATE: UNK, EXPLANT DATE: UNK, DEVICE MFR DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241986&pc=QCB | 2023826-2021-02896 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO INTRAOCULAR PRESSURE (IOP) WAS ELEVATED. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK, SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK, EXPIRATION DATE UNK, IMPLANT DATE: UNK, EXPLANT DATE: UNK, DEVICE MFR DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13144034&pc=QCB | 2023826-2022-00017 | 2021-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS A PLANNED EXCHANGE TO BE PERFORMED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE RIGHT PATIENT'S EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "TORIC LENS VAULT WAS TOO HIGH". CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182549&pc=QCB | 2023826-2022-00033 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON STATES THERE IS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON STATES THERE IS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 13.2MM VTICH13.2, +6.0/+5.0/099 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. H6-DEVICE CODE 1494: ACD<3.0MM. CLAIM# (B)(4). Manufacturer Narrative: SEX, AGE, WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4). Manufacturer Narrative: SEX, AGE, WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161442&pc=QCB | 2023826-2022-01084 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.5/2.0/046 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15227548&pc=QCB | 2023826-2022-02765 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO A PATIENT'S EYE IN (B)(6) 2022. THE SURGEON NOW REPORTS A PLANNED LENS EXCHANGE FOR A SHORTER LENGTH LENS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -3.50/+1.00/077 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15674217&pc=QCB | 2023826-2022-03632 | 2022-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS OF -18/3.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP, ANGLE CLOSURE WITH ELEVATED IOP, UNREACTIVE (FIXED) PUPIL, AND IRIS ATROPHY. CAUSE OF THE EVENT IS THE DEVICE. THE REPORTER STATED THAT THE LENS WAS, "OVERSIZED." THE PROBLEM IS NOT RESOLVED. REPORTEDLY, "EXPLANTED ICL, AWAITING PUPIL DILATION TO RESOLVE BEFORE RE-IMPLANTING." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "VTICM 12.6" SHOULD BE CORRECTED TO "VTICMO 12.6" CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4582- ANGLE CLOSURE; IRIS ATROPHY; UNREACTIVE (FIXED) PUPIL. ADDITIONAL DATA: B5: IT WAS LATER REPORTED "THE LENS WAS NOT REPLACED." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15763937&pc=QCB | 2023826-2022-04065 | 2022-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16465175&pc=MTA | 2023826-2023-00767 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6:1494-OFF LABEL: PROGRESSIVE MYOPIA. D9: RETURN DATE- 05-APR-2023 CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16465175&pc=QCB | 2023826-2023-00767 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6:1494-OFF LABEL: PROGRESSIVE MYOPIA. D9: RETURN DATE- 05-APR-2023 CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595289&pc=MTA | 2023826-2023-00985 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1494-OFF LABEL- PROGRESSIVE MYOPIA D9: RETURN DATE: 05-APR-2023 (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595289&pc=QCB | 2023826-2023-00985 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1494-OFF LABEL- PROGRESSIVE MYOPIA D9: RETURN DATE: 05-APR-2023 (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409310&pc=QCB | 2023826-2019-00378 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -16.0/+3.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, AND ANGLE CLOSURE. THE STATUS OF THE EYE IS "GOOD" AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS "SIZING." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(6). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475395&pc=QCB | 2023826-2019-00520 | 2019-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, PUPIL BLOCK, AND ANGLE CLOSURE. IN ADDITION, THE SURGEON PERFORMED ENLARGEMENT OF THE PI. NO ADDITIONAL INFORMATION CURRENTLY AVAILABLE REGARDING PATIENT'S CONDITION AFTER LENS EXCHANGE. Manufacturer Narrative: PREVIOUSLY STATED A LENS SPHERE/CYLINDER/AXIS VALUE OF -8.5/+2.5/103. ACTUAL VALUE IS -8.5/+2.5/108. THE LENS EXCHANGE RESOLVED THE PROBLEM. "EXCELLENT OUTCOME." CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, PUPIL BLOCK, ANGLE CLOSURE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8489348&pc=QCB | 2023826-2019-00555 | 2018-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/1.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4)..
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8501145&pc=QCB | 2023826-2019-00505 | 2019-03-12 | Injury | STAAR SURGICAL COMPANY | 2019-04-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -10.00/2.0/066 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (35MMHG) WITHOUT PUPIL BLOCK AND ANGEL CLOSURE AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8501252&pc=QCB | 2023826-2019-00506 | 2019-03-12 | Injury | STAAR SURGICAL COMPANY | 2019-04-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/1.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (26MMHG) WITHOUT PUPIL BLOCK AND ANGEL CLOSURE AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511516&pc=QCB | 2023826-2019-00597 | 2019-03-14 | Injury | STAAR SURGICAL COMPANY | 2019-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, - 6.0/+4.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE (IOP), ANGLE CLOSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. PRIOR TO EXPLANT, GAVE THE PATIENT DIAMOX, TOPICAL OCULAR ANTIHYPERTENSIVE MEDICATIONS, THEN IV MANNITOL ATTEMPTING TO LOWER IOP, THEN ATTEMPTED TAPPED ANTERIOR CHAMBER AT SLIT LAMP ATTEMPTING TO LOWER IOP. THESE MEASURES WERE UNSUCCESSFUL AND IOP CONTINUED TO RISE. GONIOSCOPY DEMONSTRATED ANGLE CLOSURE, AND THE DECISION WAS MADE TO EXPLANT. THE SURGEON PLANS TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO OVERSIZED LENS. THE PATIENT CONDITION AT ONE DAY POST-OP, MILD CORNEAL EDEMA AND ANGLE OPEN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519014&pc=QCB | 2023826-2019-00610 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.5/+4.0/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. IN ADDITION, THE SURGEON PERFORMED ENLARGEMENT OF THE PI. THE PROBLEM WAS RESOLVED. THE SURGEON NOTES "EXCHANGE WITHOUT COMPLICATION" AND "EXCELLENT 20/20 OD." Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND RESIDUE ON THE LENS SURFACE AND THE OPTIC WAS SCRATCHED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586217&pc=QCB | 2023826-2019-00735 | 2019-04-12 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, IN THE PATIENTS EYE ON (B)(6) 2019. THE REPORTER INDICATED THE SURGERY WENT WELL BUT AT ONE DAY POST-OP, THE INTRAOCULAR PRESSURE (IOP) WAS 50MMHG, WITH SHALLOW CHAMBER AND ANGLE CLOSURE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8620926&pc=QCB | 2023826-2019-00871 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICM5_13.2, -7.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXPLANT RESOLVED THE PROBLEM. THE SURGEON AND PATIENT ARE CONSIDERING IMPLANTING A NEW LENS AFTER THREE OR FOUR MONTHS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621026&pc=QCB | 2023826-2019-00872 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICM5_13.2, -7.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 27 (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXPLANT RESOLVED THE PROBLEM. THE SURGEON AND PATIENT ARE CONSIDERING IMPLANTING A NEW LENS AFTER THREE OR FOUR MONTHS. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 3189 SHOULD BE CORRECTED TO 2993 IN INITIAL MEDWATCH REPORT. DEVICE EVALUATED BY MFR: DEVICE EVALUATION: LENS RETURNED DRY WITH SURGICAL RESIDUE IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633170&pc=QCB | 2023826-2019-00886 | 2019-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.0/+1.0/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION, ELEVATED IOP, AND LENS DISLOCATION/SUBLUXATION INFERIORLY. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. THE REPORTER STATES, "LARGER SULCUS HORIZONTAL THAN ESTIMATED BY WTW CAUSING ICL TO DROP INFERIORLY (HOLE AT INFERIOR PUPIL EDGE)." REFERENCE CLAIM# (B)(4) FOR REPLACEMENT LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A ZIPLOC BAGGIE, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8634346&pc=QCB | 2023826-2019-00804 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/177 (SPHERE/CYLINDER/AXIS), FROM THE PATIENTS RIGHT EYE (OD) DURING THE SAME SURGERY. THE LENS WAS REMOVED DUE TO EXCESSIVE VAULTING, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE (IOP), UNREACITVE (FIXED) PUPIL AND IRIS ATROPHY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO LENS SIZE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652099&pc=QCB | 2023826-2019-00933 | 2019-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL DRY WITH SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8731290&pc=QCB | 2023826-2019-01084 | 2019-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -9.50/4.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON OBSERVED EXCESSIVE VAULT THE NEXT DAY FOLLOWING SURGERY, LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND ELEVATED IOP (31 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT, OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). DATE OF EVENT IS (B)(6) 2019 CONCOMITANT MEDICAL PRODUCT: INJECTOR LOT# 1431485. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 OF -9.5/4.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON OBSERVED EXCESSIVE VAULT WITH ROTATION THE NEXT DAY FOLLOWING SURGERY; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 AND REPLACED FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. DATE RECEIVED BY MANUFACTURER: DATE RECEIVED IN ORIGINAL MDR IS NOT APPLICABLE; ADDITIONAL INFORMATION WAS RECEIVED ON 12-JULY-2019. PATIENT CODE:3191 SECONDARY SURGERY (EXPLANT); SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 OF -9.5/4.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON OBSERVED EXCESSIVE VAULTING THE NEXT AT FOLLOWING SURGERY; LOSS OF BCVA; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP (31MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. PATIENT CODE: 3191 SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8731314&pc=QCB | 2023826-2019-01101 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/4.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON OBSERVED EXCESSIVE VAULT THE NEXT DAY FOLLOWING SURGERY, LOSS OF BCVA AND ELEVATED IOP (31 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 OF -9.0/4.5/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS). THE SURGEON OBSERVED EXCESSIVE VAULTING THE NEXT AT FOLLOWING SURGERY; LOSS OF BCVA; AND ELEVATED IOP (31MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. PATIENT CODE: 3191 SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT, OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 OF -9.0/4.5/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS). THE SURGEON OBSERVED EXCESSIVE VAULT WITH ROTATION THE NEXT DAY FOLLOWING SURGERY; AND ELEVATED IOP (31MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS WAS REMOVED AND LATER REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. DEVICE CODE 1401 IN ORIGINAL MDR IS NOT APPLICABLE. DEVICE CODE 2923 ADDED FOR ROTATION. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736583&pc=QCB | 2023826-2019-01126 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP. A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019 WITHIN THE SAME SURGERY AS THE REMOVAL OF THE ORIGINAL LENS. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K #: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8809460&pc=QCB | 2023826-2019-01259 | 2019-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED INTRAOCULAR PRESSURE. ANOTHER LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENTS CONDITION AT LAST VISIT WAS THE PUPIL WAS BIGGER THAN NORMAL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889701&pc=QCB | 2023826-2019-01418 | 2019-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -14.00/4.0/118 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (40 MMHG), AND EXCESSIVE VAULT WERE OBSERVED. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SMALLER LENGTH LENS. REPORTER NOTED "VAULT IS STILL GENEROUS BUT SAFE." THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS RECOMMENDED LENS TOO LARGE. PER REPORTER, PATIENTS CONDITION ON LAST VISIT, (B)(6) 2019, BCVA 20/50, "EYE IS SAFE, UNINFLAMED, GOOD IOP, STILL GENEROUS VAULT BUT SAFE EXAM & IOP, ANGLES OPEN." Manufacturer Narrative: SEX: FEMALE SHOULD BE CORRECTED TO MALE. RETURNED TO MANUFACTURER: 08-AUG-2019 SHOULD BE CORRECTED TO 30-JUL-2019. (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889704&pc=QCB | 2023826-2019-01419 | 2019-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (50 MMHG), AND EXCESSIVE VAULT WERE OBSERVED. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SMALLER LENGTH LENS. REPORTER NOTED "VAULT IS STILL GENEROUS BUT SAFE." THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "RECOMMENDED LENS TOO LARGE." PER REPORTER, PATIENTS CONDITION ON LAST VISIT, (B)(6) 2019, BCVA 20/50, "EYE IS SAFE, UNINFLAMED, GOOD IOP, STILL GENEROUS VAULT BUT SAFE EXAM & IOP, ANGLES OPEN." Manufacturer Narrative: EVENT DESCRIPTION: "PER REPORTER, PATIENTS CONDITION ON LAST VISIT, (B)(6) 2019, BCVA 20/50, "EYE IS SAFE, UNINFLAMED, GOOD IOP" SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT. INFORMATION IS NOT APPLICABLE. ON (B)(6) 2019 THE LENS WAS REMOVED SHOULD BE CORRECTED TO (B)(6) 2019 THE LENS WAS REMOVED, IN INITIAL MEDWATCH REPORT. DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894304&pc=QCB | 2023826-2019-01441 | 2019-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. ON (B)(6) 2019 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918840&pc=QCB | 2023826-2019-01497 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A PIECE OF HAPTIC TORN OFF. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8998364&pc=QCB | 2023826-2019-01665 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT LENS A 12.6MM, VTICM5_12.6, -13.00/+2.00/78 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AS A REPLACEMENT LENS BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (INTRAOCULAR PRESSURE). IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED, PATIENT CONDITION IS WELL. NO PLAN TO IMPLANT OTHER LENS. Manufacturer Narrative: REFERENCE MFR# 2023826-2019-01615 FOR INITIAL LENS. DEVICE EVALUATION: DIMENSIONAL INSPECTION DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. INVESTIGATION FOUND THAT IT IS LIKELY THAT THE CUSTOMERS SWITCHED THE INITIAL AND EXCHANGE LENSES WHEN PACKAGING THEM FOR RETURN. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. DEVICE CODE 1494: OFF-LABEL USE- AGE < 21 YEARS OLD. CLAIM#: (B)(4). Manufacturer Narrative: H6 - METHOD CODE 3331: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9070787&pc=QCB | 2023826-2019-01707 | 2019-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/1.0/091 SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: EVENT DESCRIPTION: INITIAL MDR SHOULD HAVE STATED: ELEVATED IOP WAS ALSO REPORTED. IT WAS REPORTED THAT AT NINE DAYS POST-OP, 'HIS IOP AFTER MEDICATIONS IS 27MMHG.' (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120432&pc=QCB | 2023826-2019-01735 | 2019-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -13.5/+2.5/127 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED WITH EXPLANT, NO REPLACEMENT LENS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120434&pc=QCB | 2023826-2019-01736 | 2019-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -13.5/+1.5/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED WITH EXPLANT, NO REPLACEMENT LENS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9194673&pc=QCB | 2023826-2019-01927 | 2019-07-30 | Injury | STAAR SURGICAL COMPANY | 2019-10-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). ANOTHER ICL LENS WAS NOT IMPLANTED. THE LENS WAS DISCARDED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9202984&pc=QCB | 2023826-2019-01911 | 2019-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/5.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE PROBLEM WAS RESOLVED. THE CAUSE WAS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: SEX, WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9216530&pc=QCB | 2023826-2019-01987 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO 13.2, -8.5/1.5/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE, AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: DESCRIPTION OF EVENT OR PROBLEM- "-8.5/1.5/069" SHOULD BE CORRECTED TO "-8.5/1.5/068" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279652&pc=QCB | 2023826-2019-02063 | 2018-08-01 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +10.0/+1.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP. A SECONDARY YAG WAS ALSO PERFORMED. ON (B)(6) 2019 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE SURGEON NOTES, "PERIPHERAL SYNAECHIAE TO WOUND AND PARASCENTESIS." THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279655&pc=QCB | 2023826-2019-02069 | 2019-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9292244&pc=QCB | 2023826-2019-02051 | 2019-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS AND EXCESSIVE VAULT WAS OBSERVED. ELEVATED IOP (38 MMHG) WAS ASSESSED ON (B)(6) 2019. LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDDOTHIAL CELL LOSS AND EXCESSIVE VAULT WAS OBSERVED." CORRECTED TO: "SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND EXCESSIVE VAULT WERE OBSERVED." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9299931&pc=QCB | 2023826-2019-02078 | 2019-09-07 | Injury | STAAR SURGICAL COMPANY | 2019-11-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/051 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS DISCARDED BY THE HOSPITAL AND WILL NOT BE RETURNED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306729&pc=QCB | 2023826-2019-02050 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -14.0/1.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306739&pc=QCB | 2023826-2019-02049 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -14.5/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312711&pc=QCB | 2023826-2019-02079 | 2019-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312732&pc=QCB | 2023826-2019-02080 | 2019-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317888&pc=QCB | 2023826-2019-02129 | 2019-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336755&pc=QCB | 2023826-2019-02173 | 2019-10-18 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -10.0/+3.0/175 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. IOP VALUE WENT FROM 45MMHG TO 15MMHG AFTER EXPLANT. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A CONTACT LENS CONTAINER. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336761&pc=QCB | 2023826-2019-02172 | 2019-10-18 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -10.0/+3.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. IOP VALUE WENT FROM 59MMHG TO 12MMHG AFTER EXPLANT. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A CONTACT LENS CONTAINER. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9335805&pc=QCB | 2023826-2019-02149 | 2019-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2, -11.00/1.0/085 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (23 MMHG) .THIS RESOLVED THE PROBLEM. AN ENLARGEMTNE OF PI AND YAG WAS PERFORME. PER THE REPORTER, CAUSE OF THE EVENT IS REPORTED AS "STEEP VAULT.". Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH CLEAR SURGICAL RESIDUE IN A LENS VIAL INSIDE A ZIPLOC BAG. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9338940&pc=QCB | 2023826-2019-02120 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+4.0/087 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND NARROW ANGLES. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6)2019. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS IN THE LEFT EYE. THE PATIENT IS DOING WELL. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: "A REPLACEMENT LENS OF A SHORTER LENGTH WAS IMPLANTED ON (B)(6) 2020. THE REPLACEMENT LENS RESOLVED THE ISSUE. PATIENT IS NOW FINE." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9338965&pc=QCB | 2023826-2019-02119 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/102 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND NARROW ANGLES. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, INCISION ENLARGEMENT. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND NARROW ANGLES. THE INCISION WAS ENLARGED TO REMOVE THE LENS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE VAULT IS VERY GOOD AND THE PATIENT CONDITION IS FINE. CLAIM # (B)(4). Manufacturer Narrative: (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019 IN THE INITIAL MDR. G4-(B)(6) -2019 SHOULD BE CORRECTED TO (B)(6) 2019 IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. WEIGHT UNK. AGE: UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498077&pc=QCB | 2023826-2019-02407 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.5/+6.0/084 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, IRIS ATROPHY, UNREACTIVE (FIXED) PUPIL AND ANGLE CLOSURE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. THE SURGEON DID NOT IMPLANT ANOTHER LENS. POST-OP THE PATIENT STILL HAS IRIS ATROPHY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512598&pc=QCB | 2023826-2019-02423 | 2019-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -15.5/+2.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP. IMMEDIATE POSTOP RESULTS WERE GOOD FOLLOWING THE EXCHANGE. PRIOR TO EXCHANGE THE SURGEON HAD PERFORMED A RELEASE OF FLUID VIA PARACENTESIS AND HAD THE PATIENT ON "MAXIMAL MEDICAL THERAPY." Manufacturer Narrative: CORRECTED DATA: A2: PATIENT'S AGE: ON (B)(6) 1983. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512611&pc=QCB | 2023826-2019-02422 | 2019-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -14.5/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP. IMMEDIATE POSTOP RESULTS WERE GOOD FOLLOWING THE EXCHANGE. PRIOR TO EXCHANGE THE SURGEON HAD PERFORMED A RELEASE OF FLUID VIA PARACENTESIS AND HAD THE PATIENT ON "MAXIMAL MEDICAL THERAPY." Manufacturer Narrative: PATIENT'S AGE: (B)(6) 1983. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# 430811
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516563&pc=QCB | 2023826-2019-02448 | 2019-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED. ELEVATED IOP (50MMHG) WAS ASSESSED ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559343&pc=QCB | 2023826-2020-00046 | 2019-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM A PATIENT'S EYE FOR AN UNKNOWN REASON. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +5.0/+2.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UNREACTIVE (FIXED) PUPIL, AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE SYMPTOMS RESOLVED AFTER EXPLANT. Manufacturer Narrative: B5-PLEASE DISREGARD PREVIOUS B5 STATEMENT. H6-ADDITIONAL PATIENT CODE 3191: NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF ICA, UNREACTIVE PUPIL). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE 1937 SHOULD BE IN THE PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE NAME: ESUBMITTER SOFTWARE DOES NOT ALLOW FOR AN UNKNOWN ENTRY. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565909&pc=QCB | 2023826-2020-00062 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -6.50/-1.50/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B3- DATE OF EVENT: (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. B5- THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -6.50/-1.50/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM. D6- IMPLANT DATE: (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -6.50/-1.50/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP AND UNREACTIVE (FIXED) PUPIL. THE EXCHANGE RESOLVED THE PROBLEM. H6- PATIENT CODE 3191- UNREACTIVE (FIXED) PUPIL. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9639933&pc=QCB | 2023826-2020-00201 | 2019-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, - 14.00/+2.0/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET. THE SURGEON AND PATIENT ARE HAPPY. Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED ELEVATED IOP. SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #2. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9656163&pc=QCB | 2023826-2020-00212 | 2020-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 14.00/+2.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT HAS NO PAIN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689251&pc=QCB | 2023826-2020-00250 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT AND ELEVATED IOP (20 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2020. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER STATES "PATIENT IS HAPPY". PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9690171&pc=QCB | 2023826-2020-00278 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). AN ADDITION OF A PI-YAG WAS ALSO PERFORMED. ON THE SAME DATE IN A SEPARATE SURGERY, A SHORTER LENGTH LENS WAS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS "SIZING CALCULATION ERROR." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED ON (B)(6) 2019 YAG WAS PERFORMED. POSTOP PATIENT RESULTS ARE "SO FAR EVERYTHING SEEMS FINE". CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721910&pc=QCB | 2023826-2020-00307 | 2019-09-10 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +4.5/+2.0/060 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE SYMPTOMS RESOLVED AFTER EXPLANT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9757629&pc=QCB | 2023826-2020-00388 | 2019-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REPORTEDLY, IOP OF 31 WAS OBSERVED ON (B)(6) 2019 THAT RESPONDED TO IOP LOWERING MEDICATION (COMBIGAN) GIVING AN IOP OF 20 ON (B)(6) 2019 ("NEAR THEIR BASELINE OF 19."). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (MEDICAL INTERVENTION: COMBIGAN MEDICATION). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781511&pc=QCB | 2023826-2020-00296 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT AND ELEVATED IOP (28 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT OTHER LENS DUE TO CONSTANT ELEVATED IOP. PER THE REPORTER ON (B)(6) 2020, "PATIENT'S CURRENT CONDITION IS GOOD, IOP WAS TREATED WITH MEDICATION AND THE SURGEON KEEP MONITORING THE PATIENT'S IOP. AFTER REMOVING THE LENS, ALL THE PROBLEMS WERE RESOLVED AND THE DOCTOR PRESCRIBED GLASSES." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781512&pc=QCB | 2023826-2020-00297 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -10.00/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT AND ELEVATED IOP (27 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT OTHER LENS DUE TO CONSTANT ELEVATED IOP. PER REPORTER ON (B)(6) 2020, "PATIENT IN GOOD CONDITION, IOP WAS TREATED WITH MEDICATION AND THE SURGEON KEEP MONITORING THE PATIENT'S IOP, AND AFTER REMOVING THE LENS, ALL THE PROBLEMS WERE RESOLVED AND THE DOCTOR DESCRIBED GLASSES." Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842477&pc=QCB | 2023826-2020-00547 | 2019-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/3.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THIS RESOLVED THE PROBLEM, PATIENT CURRENT CONDITION IS GOOD. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842582&pc=QCB | 2023826-2020-00548 | 2019-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.5/5.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. LENS EXPLANT RESOLVED THE PROBLEM, PATIENT CURRENT CONDITION IS GOOD. Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 1937 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842748&pc=QCB | 2023826-2020-00505 |  | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -9.5/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. THE SURGEON REPORTS "SIZE SEEMS TO BE BIG, WITH PERSISTENT ELEVATED IOP REFRACTIVE TO MEDICAL TREATMENT." THE SURGEON PLANS TO INCREASE THE SIZE OF THE PI AND WILL DETERMINE IF AN EXCHANGE IS NEEDED. IT IS UNKNOWN AT THIS TIME AS TO WHETHER THE PI WAS INCREASED OR THE EXCHANGE OCCURRED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, YAG- ENLARGEMENT OF PI). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842787&pc=QCB | 2023826-2020-00506 |  | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -9.5/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. THE SURGEON REPORTS "SIZE SEEMS TO BE BIG, WITH PERSISTENT ELEVATED IOP REFRACTIVE TO MEDICAL TREATMENT." THE SURGEON PLANS TO INCREASE THE SIZE OF THE PI AND WILL DETERMINE IF AN EXCHANGE IS NEEDED. IT IS UNKNOWN AT THIS TIME AS TO WHETHER THE PI WAS INCREASED OR THE EXCHANGE OCCURRED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, YAG- ENLARGEMENT OF PI). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9891498&pc=QCB | 2023826-2020-00597 | 2020-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -13.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PUPIL BLOCK WITH ELEVATED IOP (GREATER THAN 50MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (GREATER THAN 50MMHG) WERE ASSESSED ON (B)(6) 2020, AND EXCESSIVE VAULT WAS OBSERVED. ADDITION OF NEW PI AND YAG WERE PERFORMED. THE LENS WAS REMOVED ON (B)(6) 2020. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE (EXCESSIVE VAULT). REPORTEDLY, "PATIENT HAPPY." Manufacturer Narrative: PATIENT CODE- 3191- SECODARY YAG NEEDS BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913332&pc=QCB | 2023826-2020-00737 | 2020-01-14 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+0.5/152 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THREE DAYS LATER, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED AGAIN. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON IS NOT PLANNING TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953427&pc=QCB | 2023826-2020-00872 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/2.0/74 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953577&pc=QCB | 2023826-2020-00874 | 2018-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.5/3.0/96 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6- TYPE OF INVESTIGATION CODE 10 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953691&pc=QCB | 2023826-2020-00875 | 2018-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.5/3.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(60 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953821&pc=QCB | 2023826-2020-00882 | 2019-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957585&pc=QCB | 2023826-2020-00847 | 2020-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -3.00/0.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO ANGLE CLOSURE WITH ELEVATED IOP (20.5MMHG) BEING OBSERVED ON (B)(6) 2020 AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY - ALL IS FINE." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957586&pc=QCB | 2023826-2020-00846 | 2020-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -3.50/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO ANGLE CLOSURE WITH ELEVATED IOP (28.5MMHG) BEING OBSERVED ON (B)(6) 2020 AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY - ALL IS FINE." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CORRECTED DATA: H6-PATIENT CODE: 3191-ANGLE CLOSURE SHOULD BE INCLUDED IN INITIAL MDR. PATIENT CODE 1937- ELEVATED IOP SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960150&pc=QCB | 2023826-2020-00856 | 2019-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9991666&pc=QCB | 2023826-2020-00950 | 2019-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025642&pc=QCB | 2023826-2020-01002 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/045 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE PROBLEM WAS RESOLVED AFTER EXPLANT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AS WELL AS THE OPTIC AND HAPTIC DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3- DEVICE EVALUATION: "WITH MOISTURE" SHOULD BE REMOVED FROM PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10032664&pc=QCB | 2023826-2020-01022 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, -3.00/2.5/017 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ANGLE CLOSURE WITH ELEVATED IOP (20 MMHG) ON (B)(6) 2020 AND EXCESSIVE VAULT WERE OBSERVED. ADDITION OF NEW PI AND YAG WAS PERFORMED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS AND THIS RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - -3.00/2.5/107 SHOULD BE CORRECTED TO +3.00/2.5/107 FOR LENS (SPHERE/CYLINDER/AXIS). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. IT WAS ALSO REPORTED THE PATIENT WAS PRESCRIBED MEDICATION. PATIENT CODE 3191: NO CODE AVAILABLE (MEDICATION). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10044695&pc=QCB | 2023826-2020-01053 | 2019-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "..THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT." CORRECTED TO "..THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP OF 40MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ." IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10044899&pc=QCB | 2023826-2020-01054 | 2019-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-14.0/1.0/124 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE 40 MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10044986&pc=QCB | 2023826-2020-01055 | 2020-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.1,-12.0/1.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA : B5-THE REPORTER INDICATED THAT A 13.7MM, VTICMO 13.7, -12.0/1.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. CLAIM# : (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242445&pc=QCB | 2023826-2020-01423 | 2020-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE AND MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392733&pc=QCB | 2023826-2020-01666 | 2020-05-18 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +05.00/+1.0/034 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED 6 HOURS LATER ON THE SAME DAY DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS (PLANNED FOR A LATER DATE). THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE (LENS WAS TOO LARGE). THE STATUS OF THE EYE IS PALE (QUIET) AND NORMAL FINDINGS OF ENDOTHELIUM AND LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392777&pc=QCB | 2023826-2020-01667 | 2020-05-18 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +05.50/+1.0/147 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED 6 HOURS LATER ON THE SAME DAY DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS (PLANNED FOR A LATER DATE). THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE (LENS WAS TOO LARGE). THE STATUS OF THE EYE IS PALE (QUIET) AND NORMAL FINDINGS OF ENDOTHELIUM AND LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392960&pc=QCB | 2023826-2020-01672 | 2020-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE (FIXED) PUPIL. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND DEBRIS ON THE LENS. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398745&pc=QCB | 2023826-2020-01707 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS WAS TOO LARGE, CAUSING ELEVATED INTRAOCULAR PRESSURE (IOP). THE INCISION WAS ENLARGED TO REMOVE THE LENS AND A SUTURE WAS REQUIRED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. AN ADDITIONAL SURGERY WAS PERFORMED - ENLARGEMENT/ADDITION OF PI BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. H6: HEALTH EFFECT IMPACT CODE: 4625, ADDITIONAL SURGERY - ENLARGEMENT/ADDITION OF PI BY YAG. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416135&pc=QCB | 2023826-2020-01714 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS HAD EXCESSIVE VAULTING, WITH ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020 AND WAS EXCHANGED FOR A SHORTER LENS. THE EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED AC SHALLOWING. CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZING ISSUE". H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PARTIAL AC SHALLOWING, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10417803&pc=QCB | 2023826-2020-01752 | 2019-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -4.5/1.0/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP, IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER LENS EXPLANT. THE REPORTER STATED " THE BASIC INTRAOCULAR PRESSURE OF THE PATIENT WAS HIGH, AND POST-OP WAS HIGH. THE IOP WAS STILL GREATER THAN 26MMHG AFTER MEDICATION. CONSIDERING THE HIGH RISK OF GLAUCOMA, THE LENS WAS REMOVED, AND THEN THE CORNEAL LASER OPERATION WAS PERFORMED. AFTER LASER SURGERY THE PATIENTS UCVA REACHED OU1.2 AND THE IOP IS NORMAL (17MMHG)". Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6 - HEALTH IMPACT: 4644 - PATIENT WAS PRESCRIBED MEDICATION, SHOULD HAVE BEEN INCLUDED. H6 - HEALTH IMPLACT: 4625 - LASER REFRACTIVE SURGERY, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440601&pc=QCB | 2023826-2020-01816 | 2020-07-29 | Injury | STAAR SURGICAL COMPANY | 2020-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/174 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. A COUPLE OF WEEKS LATER, THE VAULT HAD COME DOWN, SO THE SURGEON DECIDED TO LEAVE THE LENS IN THE EYE AND MONITOR THE PATIENT. THE PATIENT WAS TAKING DROPS FOR PRESSURE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445835&pc=QCB | 2023826-2020-01815 | 2020-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/010 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476011&pc=QCB | 2023826-2020-01938 | 2020-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.00/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476114&pc=QCB | 2023826-2020-01934 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THAT SAME DAY ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR REPORT # 2023826-2020-01935 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS DUE TO EXCESSIVE VAULTING. Manufacturer Narrative: ADDITIONAL INFORMATION: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. IT IS LIKELY THAT THESE LENSES WERE INADVERTENTLY SWAPPED DURING RETURN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL WIDTH INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. DIMENSIONAL OVERALL LENGTH INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416144&pc=QCB | 2023826-2020-01715 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/064 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS HAD EXCESSIVE VAULTING, WITH ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WA EXPLANTED ON (B)(6) 2020 AND EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT WITH RESIDUE ON THE LENS. CORRECTED DATA: B5 - IT WAS ALSO REPORTED THE PATIENT EXPERIENCED PARTIAL AC SHALLOWING. CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZING ISSUE". H6 - HEALTH EFFECT CLINICAL CODE: 4581: PARTIAL AC SHALLOWING SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10498204&pc=QCB | 2023826-2020-01979 | 2020-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +5.00/+2.0/162 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND IRIDOCORNEAL TOUCH. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE PATIENT HAS INTERMITTENT ANGLE CLOSURE AND IS STABLE ON PILOCARPINE. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - WORK ORDER SEARCH RESULTS AND INFORMATION SHOULD BE REMOVED AS IT IS NO LONGER APPLICABLE TO THIS CLAIM. H6 - 213 AND 4310 CODES ADDED DUE TO PREVIOUS PRODUCT EVALUATION RESULTS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS,+5.00/+2.0/162 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND IRIDOCORNEAL TOUCH. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE VAULT MUCH IMPROVED, ANGLE OPEN, IOP NORMAL, CORNEAL EDEMA, WHICH CAUSED A REDUCTION IN VISION. AT FINAL ASSESSMENT, THE VAULT WAS SATISFACTORY AND VISION WAS GOOD. THE ISSUE WAS RESOLVED AND PATIENT HAS A GOOD OUTCOME. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH A PRE-EXISTING CONDITION - KERATOCONUS. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +4.50/+2.0/165 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020, DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND IRIDOCORNEAL TOUCH. THE PATIENT WAS ADMINISTERED PILOCARPINE AND WAS STABLE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE VAULT MUCH IMPROVED, ANGLE OPEN, IOP NORMAL, CORNEAL EDEMA, WHICH CAUSED A REDUCTION IN VISION. AT FINAL ASSESSMENT, THE VAULT WAS SATISFACTORY AND VISION WAS GOOD. THE ISSUE WAS RESOLVED AND PATIENT HAS A GOOD OUTCOME. CLAIM # 717114 Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12957740&pc=QCB | 2023826-2021-04643 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -17.00/+4.0/117 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED IOP. THE ANTERIOR CHAMBER WAS IRRIGATED/ EVACUATED OF VISCO/ FLUIDS AND THE WOUND WAS BURPED, CONFIRMING THERE WAS NO VISCO REMAINING. THE PIS WERE ENLARGED BY YAG. THE PROBLEM WAS RESOLVED AND PATIENT IS NO LONGER IN PAIN. THE IOP HAS DECREASED SIGNIFICANTLY AND ANGLES ARE OPEN. THE CAUSE OF THE EVENT WAS DUE TO THE ICL WAS TOO BIG, CAUSING PERIPHERAL ANGLE CLOSURE. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN A SPECIMEN CUP WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND STRETCHED OUT WITH RESIDUE THROUGHOUT THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12965310&pc=QCB | 2023826-2021-04644 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -17.50/+4.0/065 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED IOP. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/ FLUIDS AND THE WOUND WAS BURPED, CONFIRMING THERE WAS NO VISCO REMAINING. THE PIS WERE ENLARGED BY YAG. THE PROBLEM WAS RESOLVED AND PATIENT IS NO LONGER IN PAIN. THE IOP HAS DECREASED SIGNIFICANTLY AND ANGLES ARE OPEN. THE CAUSE OF THE EVENT WAS DUE TO THE ICL WAS TOO BIG, CAUSING PERIPHERAL ANGLE CLOSURE.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -17.50/+4.0/065 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED IOP. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/ FLUIDS AND THE WOUND WAS BURPED, CONFIRMING THERE WAS NO VISCO REMAINING. THE PIS WERE ENLARGED BY YAG. THE PROBLEM WAS RESOLVED AND PATIENT IS NO LONGER IN PAIN. THE IOP HAS DECREASED SIGNIFICANTLY AND ANGLES ARE OPEN. THE CAUSE OF THE EVENT WAS DUE TO THE ICL WAS TOO BIG, CAUSING PERIPHERAL ANGLE CLOSURE. Manufacturer Narrative: H3: DEVICE EVALUATION : LENS WAS RETURNED IN AA SPECIMEN CUP DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182702&pc=QCB | 2023826-2022-00022 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -13.5/+3.5/162 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SHALLOW AC AND ANGLES, PAIN, AND VOMITING. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. HEALTH EFFECT (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182711&pc=QCB | 2023826-2022-00023 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -14.0/+3.5/162 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SHALLOW AC AND ANGLES, PAIN, AND VOMITING. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11967251&pc=QCB | 2023826-2021-01840 | 2021-02-04 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE PATIENT HAD EYE PAIN AND THE IOP WAS CONTROLLED BY ANTI-GLAUCOMA DROPS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17970222&pc=QCB | 2023826-2023-04631 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) -2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE AND PIGMENT DISPERSION. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED (VAULT). Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE, PIGMENT DISPERSION. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, PIGMENT DISPERSION, PAIN, BLURRED VISION AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH THE SAME MODEL AND POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "IOP NORMAL WITH SINGLE DRUG GLAUCOMA THERAPY". ADDITIONAL INFORMATION PROVIDED - "WAITING FOR IOP TO LOWER FURTHER. THE ANGLES ARE OPEN NOW. PIGMENT DISPERSION STILL PRESENT". THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "RAISED IOP DUE TO ANGLE CLOSER AND HIGH VAULT AND PIGMENT DISPERSION". H6: HEALTH EFFECT- CLINICAL CODE: "2137" AND "1994" SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "4644" SHOULD BE ADDED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17378095&pc=QCB | 2023826-2023-03101 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Intraocular Pressure Increased; Pain; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6A: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS IMPLANTED ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023, DUE TO HIGH VAULTING, CAUSING PAIN AND HIGH PRESSURE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611930&pc=QCB | 2023826-2019-00782 | 2019-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -10.00/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SHALLOW ANTERIOR CHAMBER AND LOW VAULT WAS OBSERVED. PATIENT HAD ELEVATED IOP AND PAIN. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT IS REPORTED AS "CALCULATOR AND REP RECOMMENDED LENGTH OF LENS DUE TO AGE (13.7MM). 13.2MM IMPLANTED AFTER EXPLANT" OF LENS. Manufacturer Narrative: EVENT: THE REPORTER INDICATED THAT A 13.7MM TMICL 13.7 -10.00/2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE(OD) ON (B)(6) 2019. SHALLOW ANTERIOR CHAMBER AND EXCESSIVE VAULT WAS OBSERVED. PATIENT HAD ELEVATED IOP AND PAIN. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9034019&pc=QCB | 2023826-2019-01672 | 2019-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -12.50/3.5/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY THE "LENS WAS REMOVED DUE TO SEVERE PAIN, LENS WAS WRONG LENGTH. SURGEON REMOVED LENS AND DID NOT IMPLANT ANY OTHER LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: SUPPLEMENTAL #1 MEDWATCH REPORT IN THE EVENT DESCRIPTION SHOULD BE CORRECTED TO: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2, -12.5/3.5/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (50 MMHG) WITHOUT PUPIL BLOCK WAS OBSERVED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS OVERSIZED LENS. PER PATIENT'S LAST VISIT ON "(B)(6) 2019, BVCA 20/30, IOP 20 MMHG, EYE STATUS IS WHITE AND QUIET. IOP REMAINED ELEVATED 42 MMHG (B)(6) 2019 DESPITE MAX MEDICAL TREATMENT AND AQUEOUS RELEASE." CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE: 1937. PATIENT CODE 3191: NO CODE AVAILABLE (MEDICAL INTERVENTION). CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019 SHOULD BE REMOVED AND REPLACED WITH UNK AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (52 MMHG) WITHOUT PUPIL BLOCK WAS OBSERVED ON (B)(6) 2019. LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS OVERSIZED PER REPORTER, PER PATIENT'S LAST VISIT ON (B)(6) 2019, BCVA: 20/40, IOP: 18 MMHG, EYE STATUS WHITE AND QUIET." (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9034043&pc=QCB | 2023826-2019-01671 | 2019-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY THE "LENS WAS REMOVED DUE TO SEVERE PAIN, LENS WAS WRONG LENGTH. SURGEON REMOVED LENS AND DID NOT IMPLANT ANY OTHER LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (52 MMHG) WITHOUT PUPIL BLOCK WAS OBSERVED ON (B)(6) 2019. LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS OVERSIZED PER REPORTER, PER PATIENT'S LAST VISIT ON (B)(6) 2019, BCVA: 20/40, IOP: 18 MMHG, EYE STATUS WHITE AND QUIET." CLAIM#: (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER: 16-SEP-2019 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 MEDWATCH REPORT. PATIENT CODE: 1937. PATIENT CODE: 3191 NO CODE AVAILABLE (MEDICAL INTERVENTION). PATIENT RECEIVED MEDICAL INTERVENTION AND ELEVATED IOP STILL REMAINED. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. .WEIGHT:UNK. ETHNICITY- UNK. RACE: UNK. -DATE IF EVENT:UNK EXPLANTED: UNK. .IMPLANTED: UNK. (B)(4).WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13240884&pc=QCB | 2023826-2022-00156 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL_12.6; -10.50/+1.00/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE; ACUTE ANGLE CLOSURE SECONDARY TO PUPILLARY BLOCK FOLLOWING PHARMACOLOGIC MYDRIASIS; PAIN; DECREASED VISUAL ACUITY; NAUSEA; PUPIL DILATION. IT WAS REPORTED THAT "SUBJECT TREATED WITH ORAL ACETAZOLAMIDE AND TOPICAL TIMOLOL; PILOCARPINE; LATANOPROST AND BRINZOLAMIDE OVER THE COURSE OF APPROXIMATELY 2 HOURS. SUBJECT DISCHARGED WITH INSTRUCTIONS TO USE TIMOLOL; BRIMONIDINE; LATANOPROST AND ACETAZOLAMIDE AT HOME AND RETURN FOR FOLLOW-UP EXAM THE NEXT DAY". POST OP NOTES INDICATED THAT "THE SUBJECT RETURNED FEELING MUCH BETTER; WITHOUT PAIN OR HEADACHE. INTRAOCULAR PRESSURE MEASURED 8 MMHG; AND THE CORNEA WAS CLEAR OS. UNCORRECTED VISUAL ACUITY IMPROVED TO 20/16 OS. BRIMONIDINE WAS TAPERED; AND ALL OTHER MEDICATIONS DISCONTINUED". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS NOTED AS THE DEVICE. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: B5: "ON (B)(6) 2022 LENS OPACITY (CATARACT) ADC WAS OBSERVED. ANTERIOR SUBCAPSULAR OPACITY NOTED AS GRADE1/MILD. NO ACTION TAKEN, PATIENT REMAINS ON STUDY." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17200420&pc=QCB | 2023826-2023-02472 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pain; Pupillary Block; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE - THE LENS HAD NOT CENTRAL HOLE. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS (THE PERFORATING HOLE WAS NOT PRESENT). DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. H6: DEVICE HISTORY RECORD (DHR) REVIEW: THE INVESTIGATIVE RESULTS DETERMINED THE MILLING TOOL NOT BEING REPLACED AFTER ITS MAX LIFE WAS REACHED IS MOST LIKELY THE ROOT CAUSE TO WHY THE CENTRAL HOLE WAS NOT PRESENT ON THE LENS. AUGUST 2022, IT HAS BEEN MADE A COMMON PRACTICE ACROSS MANUFACTURING TO CHANGE THE PRIMARY TOOL PRIOR TO MAX LIFE BEING REACHED TO MINIMIZE CHIPPED EDGES/HOLES. THIS PRACTICE WILL ALSO HELP PREVENT FURTHER OCCURRENCES OF A MISSING CENTRAL HOLE. IN ADDITION, RETRAINING FOR SOP & WORK INSTRUCTION ON DETECTION OF MISSING CENTRAL HOLE AND/OR LENS DEFECTS HAVE BEEN PERFORMED. SPECIFIC OPERATORS RESPONSIBLE HAVE BEEN MADE AWARE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2023. PRODUCT NON-CONFORMITY WAS NOTED AS "THE ICL LENS HAVE NO CENTRAL HOLE".THE PATIENT EXPERIENCED ELEVATED IOP WITH PUPIL BLOCK, PAIN AND DISCOMFORT. REPORTEDLY, "THE PATIENT FELT PAIN AND UNCOMFORTABLE".ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME PARAMETERS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "THE PATIENT HAS JUST COMPLETED THE SURGERY AND IS STILL IN FOLLOW-UP. THE PATIENT'S EYE CONDITION IS GOOD". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "THE ICL LENS HAVE NO CENTRAL HOLE". H6: HEALTH EFFECT- CLINICAL CODE: "2330" AND "1994" SHOULD BE ADDED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478860&pc=QCB | 2023826-2021-00635 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS BEING TOO BIG, THE CORNERS OF THE ANTERIOR CHAMBER ARE CLOSED, INTRAOCULAR PRESSURE IS ELEVATED, DESPITE REGULAR USE OF ANTIHYPERTENSIVE DRUGS AND THE PATIENT COMPLAINS OF SQUEEZING PAIN FROM TIME TO TIME. THERE IS ALSO A RETINAL DETACHMENT. THE CAUSE OF THE EVENT IS UNKNOWN. THE PATIENTS VISION IS NORMAL AND THERE ARE NO STRONG COMPLAINTS FROM PATIENT. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - YAG PI, MEDICAL INTERVENTION OUTSIDE STANDARD PROTOCOL. H6 - HEALTH EFFECT CLINICAL CODE: 2047 SHOULD BE REMOVED. B5 - IT WAS LATER CONFIRMED BY THE REPORTER THAT THE PATIENT DID NOT EXPERIENCE RETINAL DETACHMENT. ALSO ADDITIONAL YAG SURGERIES WERE PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11349059&pc=QCB | 2023826-2021-00497 | 2020-12-01 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.50/1.0/023 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP . THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS "LENGTH ICL" Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942691&pc=QCB | 2023826-2021-01763 | 2021-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -16.0/+1.0/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP. THE PROBLEM IS RESOLVED. Manufacturer Narrative: B5-ADDITIONAL INFO: "DESPITE ALL CALCULATIONS INDICATING 12.6 OR EVEN 13.2 THE LENS WAS CLEARLY OVERSIZED IN THE EYE AND A 12.1 WAS NEEDED. THERE WAS NOTHING IN THE PREOPERATIVE EVALUATION THAT SHOULD HAVE INDICATED THAT RESULT." CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057192&pc=QCB | 2023826-2021-04911 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS SLOWLY RECOVERING AND THE PUPIL IS GETTING BETTER. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, THE PROBLEM IS RESOLVED. THE PATIENT IS HAPPY. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14654881&pc=QCB | 2023826-2022-01797 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO HIGH VAULT AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THE IOP WAS 12 MMHG, VAULT WAS 750UM AND BCVA 20/25. THE OVERSIZED LENS CAUSED THE PUPIL BLOCK AND HIGH PRESSURE. Manufacturer Narrative: H6- MEDICAL DEVICE PROBLEM CODE: 1583 SHOULD BE CORRECTED TO 2993. B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON 06-MAY-2022. THE LENS WAS EXPLANTED ON 13-MAY-2022 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) WITH PUPIL BLOCK. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THE IOP WAS 12 MMHG, VAULT WAS 750UM AND BCVA 20/25. THE OVERSIZED LENS CAUSED THE PUPIL BLOCK AND HIGH PRESSURE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14760468&pc=QCB | 2023826-2022-01928 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.00/+1.00/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.00/+1.00/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE STATUS OF THE EYE IS "ALL FINE AND PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11042998&pc=QCB | 2023826-2020-03085 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-9.5/1,5/109, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LOW VAULT AND PUPIL BLOCK WITH ELEVATED IOP 26MMHG (ASSESSED: (B)(6) 2020). ON (B)(6) 2020 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: N5- "ANGLE CLOSURE WITH ELEVATED IOP WAS REPORTED."- SHOULD BE IN INITIAL MDR. H6: CLINICAL CODE 4581 - ANGLE CLOSURE- SHOULD BE IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515299&pc=QCB | 2023826-2021-00653 | 2021-02-24 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7, -16.00/4.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO PUBIL BLOCK WITH ELEVATED IOP(34MMHG), ANGLE CLOSURE WITH ELEVATED IOP((34MMHG) ASSESSED ON (B)(6) 2021), AND EXCESSIVE VAULT, ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515298&pc=QCB | 2023826-2021-00652 | 2021-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7, -16.00/4.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO PUBIL BLOCK WITH ELEVATED IOP(39MMHG), ANGLE CLOSURE WITH ELEVATED IOP(39MMHG) ASSESSED ON (B)(6) 2021, AND EXCESSIVE VAULT, ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7, -16.00/4.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS ON (B)(6) 2021 DUE TO PUBIL BLOCK WITH ELEVATED IOP(39MMHG), ANGLE CLOSURE WITH ELEVATED IOP(39MMHG) ASSESSED ON (B)(6) 2021, AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS USER ERROR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7, -16.00/4.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS ON (B)(6) 2021 DUE TO PUBIL BLOCK WITH ELEVATED IOP(39MMHG), ANGLE CLOSURE WITH ELEVATED IOP(39MMHG) ASSESSED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12270523&pc=QCB | 2023826-2021-03056 | 2021-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TICM12.6; -14.0/1.0/133 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING; PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE; UNREACTIVE (FIXED) PUPIL; AND IT WAS INDICATED THAT ENLARGEMENT OF PI (YAG) WAS PERFORMED ON (B)(6) 2021. THE PROBLEM IS NOT RESOLVED BUT IT WAS NOTED A PLANNED LENS REPLACEMENT IS TAKING PLACE ON (B)(6) 2021. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12686036&pc=QCB | 2023826-2021-04108 | 2021-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK AND ANGLE CLOSURE WITH ELEVATED IOP. PROBLEM RESOLVED. AT A LATER UNKNOWN DATE "SOME IRIS ATROPHY + 3MM PUPIL" WAS REPORTED. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN, " PATIENT ANATOMY/ OVERSIZE MANUFACTURED ICL" WAS REPORTED FOR CAUSE DETAILS. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A DIFFERENT MODEL AND SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK AND ANGLE CLOSURE WITH ELEVATED IOP AND THE PROBLEM RESOLVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS ALSO REPORTED AS ADDITIONAL INTERVENTION/TREATMENT. "SOME IRIS ATROPHY + 3MM PUPIL, DEVELOPED DURING AC WASH" WAS REPORTED. FOR ADDITIONAL INFORMATION THE REPORTER STATED, "PUPIL STILL 3MM BUT BETTER FUNCTION AND REACTS TO LIGHT (SLOWER THAN OD)." CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN, "PATIENT ANATOMY/OVER MANUFACTURED ICL" WAS REPORTED FOR CAUSE DETAILS. H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286583&pc=QCB | 2023826-2022-00174 | 2021-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -7.00/+3.50/100 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED IOP; MEDICATION WAS PRESCRIBED (DIAMOX); ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS BUT IT WAS LATER REPORTED THAT THE ISSUE WAS NOT RESOLVED AND THE PATIENT HAS PUPIL BLOCK/FIXED PUPIL; PRESSURE SPIKES AND PRK SURGERY WAS PERFORMED DUE TO THE VISION. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER UNEXPECTED SHORT CILIARY SULCUS SPACE". SEE MDR # 2023826-2022-00176 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317692&pc=QCB | 2023826-2022-00202 | 2021-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +4.5/+2.0/009 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 AN ENLARGEMENT/ADDITION OF PI/YAG WAS PERFORMED. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, AND UNREACTIVE (FIXED) PUPIL AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +4.5/+2.0/009 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 AN ENLARGEMENT/ADDITION OF PI/YAG WAS PERFORMED. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, AND UNREACTIVE (FIXED) PUPIL AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15936540&pc=QCB | 2023826-2022-04499 | 2022-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -14.0/1.5/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE; FIXED (UNREACTIVE) PUPIL; ENLARGEMENT OF PI - YAG WHICH WAS PERFORMED ON (B)(6) 2022; AND ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF SHORTER LENGTH LENS WAS IMPLANTED BUT THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16142929&pc=QCB | 2023826-2023-00098 | 2022-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), PUPIL BLOCK, ANGLE CLOSURE, UNREACTIVE (FIXED) PUPIL AND IRIS ATROPHY. AN ADDITIONAL PI WAS PERFORMED ON (B)(6) 2022 AND THERE WAS ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2023-00099 FOR REPLACEMENT LENS. REMAINING PUPIL DISSIPATION AFTER AN ACUTE GLAUCOMA ATTACK. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNFICANT LENS HAPTIC DAMAGE. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14116256&pc=QCB | 2023826-2022-01185 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: VAULT WAS REPORTER IN THE RIGHT EYE (OD) OF A PATIENT WITH AN IMPLANTABLE COLLAMER LENS. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.0/128 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. PUPIL BLOCK WITH ELEVATED IOP, PIGMENT DISPERSION, FIXED PUPIL, IRIS ATROPHY, WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXPLANTED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS REPORTED AS OTHER, THE REPORTER STATED "FIXED DILATED PUPIL, TWO WEEKS POST ICL. IOP SLIGHTLY ELEVATED, THOUGHT MIGHT BE PUPILLARY BLOCK. PIGMENT DISPERSION NOTED, THOUGHT MIGHT BE DUE TO PRESSURE SPIKES NOT IN OFFICE. NOW THINK MOXI-DEX (5MG/CC) MIGHT BE CASE." ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT OF PI YAG, ADDITIONAL SUTURE AND ATTEMPT TO EXCHANGE) WAS PERFORMED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "PHAKIC WITHOUT ICL IMPLANTED. IRIS TRAUMA TEMPORALLY, SUTURE IN PLACE, NOW WITH MR OF -12.75 +1.75X119. THREE DIOPTER SHIFT. WAITING TO SETTLE TO SEE IF PROCEED WITH SECONDARY IMPLANT. FOR ADDITIONAL INFORMATION THE REPORTER STATED "FIRST CASE OF IRIS ATROPHY. WAS CONCERNED IT WAS IOP RELATED TO HIGH VAULT. ADDITIONAL PI ENLARGEMENT NO HELP. PROGRESSIVE TID DECIDED TO EXCHANGE ICL. VERY DIFFICULT CASE, WITH INDUCTION LEVEL ANESTHESIA NEEDED, STILL NOT ABLE TO KEEP PATIENT STILL. LUCKY TO GET LENS OUT. IRIS TRAUMA/PROLAPSE. ATTEMPTED TO PLACED, THEN PATIENT BEGAN TO ASPIRATE, FORTUNATE TO BE DONE WITH CASE. CONTEMPLATING SECONDARY ICL UNDER GENERAL ANESTHESIA BUT WANT REFRACTION TO STABILIZE. ATTEMPTED TO IMPLANT TMICL 12.6 -11.0/1.5/109 (ONLY ONE AVAIL URGENTLY) SN: (B)(6)." H6- HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION, UNREACTIVE/FIXED PUPIL, IRIS ATROPHY, TRANSILLUMINATION IRIS DEFECTS. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTION: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - DILATED PUPIL, IRIS TRAUMA/PROLAPSE. SHOULD HAVE BEEN INCLUDED. (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14505054&pc=QCB | 2023826-2022-01687 | 2021-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -11.0/1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; REFRACTIVE SURPRISE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PUPIL BLOCK; ANGLE CLOSURE; ELEVATED INTRAOCULAR PRESSURE; UNREACTIVE (FIXED) PUPIL; BLURRED VISION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. ON (B)(6) 2021 A REPLACEMENT LENS OF SHORTER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: PRODUCT EVALUATION-LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794474&pc=QCB | 2023826-2022-01950 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO TRANSIENT LOSS OF BCVA, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND IRIS ATROPHY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794499&pc=QCB | 2023826-2022-01951 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO TRANSIENT LOSS OF BCVA, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND IRIS ATROPHY. IT WAS REPORTED THAT THE PROBLEM DID NOT RESOLVE, THE REPORTER STATED "THE DOCTOR THINKS THAT THE PATIENT'S MYOPIA PROBLEM HAS NOT BEEN SOLVE." CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440465&pc=QCB | 2023826-2024-00046 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.00/1.0/173 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2023 THE LENS WAS IMPLANTED AND REMOVED INTRA-OP DUE TO EXCESSIVE VAULT; REFRACTIVE SURPRISE; PUPIL BLOCK WITH ELEVATED IOP; AND ANGLE CLOSURE WITH ELEVATED IOP. ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS DEVICE. SEE MFR# 2023826-2024-00047 FOR REPLACEMENT LENS. Manufacturer Narrative: H6-CLINICAL CODE 4581: ANGLE CLOSURE WITH ELEVATED IOP. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366544&pc=QCB | 2023826-2021-00477 | 2020-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method; Unintended Movement | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE UPSIDE DOWN. THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS. REPORTEDLY, DURING INSERTION THE ICL APPEARED TO UNFOLD PROPERLY AND IN THE PROPER ORIENTATION. AFTER EXAMINATION IT WAS DETERMINED THAT THE ICL WAS INSERTED INVERTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: A 13.2MM TMICL13.2 ICL WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL SYMPTOMS REPORTED-SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES (ICA), PUPIL BLOCK WITH ELEVATED IOP, PIGMENT DISPERSION, BLURRED VISION, AND ENLARGEMENT OF PI-YAG. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12662188&pc=QCB | 2023826-2021-04050 | 2021-09-12 | Injury | STAAR SURGICAL COMPANY | 2021-10-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.5/074 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND PUPIL BLOCK WITH INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10245039&pc=QCB | 2023826-2020-01403 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.50/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (34 MMHG) WAS OBSERVED ON (B)(6) 2020. SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS (ASSESSED ON (B)(6) 2020), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, UNREACTIVE (FIXED) PUPIL, LENS OPACITY AND EXCESSIVE VAULT WERE OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, INCISION ENLARGEMENT, ADDITIONAL SUTURE AND THE LENS WAS REMOVED ON (B)(6) 2020. REPORTEDLY THERE ARE NO PLANS TO REPLACE THE LENS. PER REPORTER ON (B)(6) 2020, PRESCRIPTION OF IOP LOWERING DROPS PROVIDED AND AS OF (B)(6) 2020 IOP WAS 18MMHG AND PUPILS ARE STILL UNREACTIVE. Manufacturer Narrative: CORRECTED DATA: B5 - 15.50/2.5/100 (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO -15.00/3.0/100 (SPHERE/CYLINDER/AXIS). H6 - HEALTH EFFECT - CLINICAL CODE: 4581 (UNREACTIVE FIXED PUPIL) AND (LENS OPACITY) AND (SIGNIFICANT ENDOTHELIAL CELL LOSS) SHOULD HAVE BEEN INCLUDED. H6 - HEALTH IMPACT - DEVICE CODE: 4625 ADDITIONAL SURGERY - INCISION ENLARGEMENT, ADDITIONAL SUTURE, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. SHOULD HAVE BEEN INCLUDED. ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN WITH FIBERS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10245040&pc=QCB | 2023826-2020-01402 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/3.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (35 MMHG) WAS OBSERVED ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, UNREACTIVE (FIXED) PUPIL, AND EXCESSIVE VAULT WAS OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, INCISION ENLARGEMENT, ADDITIONAL SUTURE AND THE LENS WAS REMOVED ON (B)(6) 2020. REPORTEDLY THERE ARE NO PLANS TO REPLACE THE LENS. PER REPORTER ON (B)(6) 2020, PRESCRIPTION OF IOP LOWERING DROPS PROVIDED AND AS OF (B)(6) 2020 IOP WAS 19MMHG AND PUPILS ARE STILL UNREACTIVE. Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT - CLINICAL CODE: 4581 (UNREACTIVE FIXED PUPIL). SHOULD HAVE BEEN INCLUDED. H6 - HEALTH IMPACT - DEVICE CODE: 4625 ADDITIONAL SURGERY - INCISION ENLARGEMENT, ADDITIONAL SUTURE, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. SHOULD HAVE BEEN INCLUDED. ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN WITH FIBER ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10279904&pc=QCB | 2023826-2020-01454 | 2020-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED THE SAME DAY ON (B)(6) 2020 DUE TO EXCESSIVE VAUTLING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON HAD ENLARGED THE PI BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE (HIGH VAULT). H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A PLASTIC BAG, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. H6: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6: HEALTH EFFECT IMPACT CODE: 4625 - SECONDARY SURGICAL INTERVENTION; PI, YAG. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15990641&pc=QCB | 2023826-2022-04529 | 2022-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO ICL IMPLANTABLE COLLAMER LENS. THE SURGEON REPORTED LATE POSTOPERATIVE PUPILLARY BLOCK WITH FIBRIN OCCLUDING THE PUPIL. THE SURGEON PERFORMED A PI, WHICH RESOLVED THE HIGH IOP. CULTURE OF THE OCCLUSION WAS NEGATIVE AND THE PATIENT DID WELL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - UNREACTIVE (FIXED) PUPIL. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 -PARS PLANE VITRECTOMY. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTED LATE POSTOPERATIVE PUPILLARY BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE (IOP), UNREACTIVE (FIXED) PUPIL AND FIBRIN OCCLUDING THE PUPIL. THE SURGEON ENLARGED THE PI, WHICH RESOLVED THE HIGH IOP. THE PATIENT HAD A RETINAL DETACHMENT MACULA ON (B)(6) 2022, WHICH REQUIRED A PARS PLANE VITRECTOMY SURGERY. CULTURE OF THE OCCLUSION WAS NEGATIVE AND THE PATIENT DID WELL. THE LENS REMAINED IMPLANTED. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM#: (B)(4). Manufacturer Narrative: AGE AT TIME OF EVENT, DATE OF BIRTH :UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. MODEL: EXPIRATION DATE UNK. IF IMPLANTED, GIVE DATE: UNK. IF EXPLANTED, GIVE DATE: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10956516&pc=QCB | 2023826-2020-02851 | 2020-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP, AND BLURRED VISION. ON (B)(6) 2020 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. ON (B)(6) 2020 (LAST VISIT) PATIENTS CURRENT CONDITION WAS BCVA 20/40, UCVA 20/70 AND IOP WAS 10MMHG. "BLURRY VISION AND MILD-MOD IRRITATION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.0/2.5/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON 06 NOV 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP, AND BLURRED VISION. ON 19 NOV 2020 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. ON (B)(6) 2020 (LAST VISIT) PATIENTS CURRENT CONDITION WAS BCVA 20/40, UCVA 20/70 AND IOP WAS 10MMHG. "BLURRY VISION AND MILD-MOD IRRITATION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). HOWEVER, SIGNS/SYMPTOMS WERE NOT BEYOND THAT OF TYPICAL SURGERY. IT WAS REPORTED THAT IRRITATION AND BLURRY VISION RESOLVED. PATIENT IS NOW 20/20. H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17432754&pc=QCB | 2023826-2023-03243 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Red Eye(s); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -13.50/+2.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, GLARES, REDNESS, MEDICATION USED AS WELL AS GONIOSCOPY PERFORMED TO OPEN ANGLES. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9039904&pc=QCB | 2023826-2019-01703 | 2019-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Retinal Detachment; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 VTICMO, -18.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH ROTATION, ELEVATED IOP, REFRACTIVE SURPRISE, AND RETINA DETACHMENT WAS OBSERVED. SURGICAL INTERVENTION (REPOSITIONING, RETINAL REPAIR AND GAS INFUSION) AND MEDICAL INTERVENTION (ANTIBIOTICS, STEROIDS, GLAUCOMA MEDICATION) WERE PERFORMED/ PRESCRIBED. LENS CURRENT STATUS WAS REPORTED TO BE IN POSITION BUT VAULT WAS APPROXIMATELY 100 MICRONS. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED VIA EMAIL. THE LENS WAS REPOSITIONED A SECOND TIME ON (B)(6) 2019. "PATIENT HAS GAS LEAKAGE IN THE EYE. SO AGAIN RETINA SURGEON HAS GAS FILLED UP AND THEY WILL WAIT FOR FOUR WEEK FOR VISUAL RECOVERY UPDATE." CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED TO STATE THAT THE LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE ON (B)(6) 2019. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271368&pc=QCB | 2023826-2021-03013 | 2021-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; UGH (Uveitis Glaucoma Hyphema) Syndrome; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO UGH SYNDROME CAUSED IRIS ATROPHY AND CONTINUOUS PIGMENT IN THE ANTERIOR CHAMBER. REMOVAL WENT WELL WITHOUT COMPLICATIONS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND UNREACTIVE (FIXED) PUPIL/IRIS. THE PIS WERE ENLARGED BY YAG ON 26-FEB-2021. THE CAUSE OF THE EVENT WAS THE DEVICE. RECENT POST-OP VISIT FOUND EYE STILL WITH PIGMENT IN AC AND FIXED IRIS. H3: DEVICE EVALUATION NOT REQUIRED. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - UNREACTIVE (FIXED) PUPIL/IRIS H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ENLARGE PIS BY YAG CORRECTED DATA: B3: 08-FEB-2021. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE PATIENT HAD A BCVA OF 20/20 AFTER EXPLANT. PER EMAIL RECEIVED ON (B)(6) 2021, " A REPLACEMENT LENS HAS NOT BEEN IMPLANTED YET, AS WE ARE STILL WAITING FOR PIGMENT/IRITIS TO RESOLVE." IN ADDITION THE REPORTER ALSO STATED THAT HIGH ICL VAULT CAUSED UGH SYNDROME. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18378904&pc=QCB | 2023826-2023-05715 | 2023-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Uveitis; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/1.50/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, BLURRED VISOIN AND PROLONGED POSTOPERATIVE ANTERIOR UVEITIS. IT WAS REPORTED THAT THE IOP WAS CONTROLLED WITH MEDICATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12686762&pc=QCB | 2023826-2021-04111 |  | Injury | STAAR SURGICAL COMPANY | 2021-10-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: ON 23 SEP 2021 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN JMED CASE REPORTS ENTITLED, 'UNILATERAL URRETS-ZAVALIA SYNDROME AFTER IMPLANTABLE COLLAMER LENS IMPLANTATION: A CASE REPORT AND REVIEW OF THE LITERATURE'. THE ARTICLE REPORTS ELEVATED IOP, DILATED PUPIL AND GLARE (SUSPECT URRETS-ZAVALIA SYNDROME) IN ONE EYE FOLLOWING ICL V4C IMPLANTATION. MEDICAL INTERVENTION (ACETAZOLAMIDE, PILOCARPINE) WAS PRESCRIBED. TWO MONTHS AFTER SURGERY THE DILATATION OF THE RIGHT PUPIL PARTIALLY REVERSED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112820&pc=QCB | 2023826-2019-01627 | 2019-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICH12.1, -8.00/1.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT, ELEVATED IOP, LENS DISLOCATION/SUBLUXATION AND GLARE/HALOES. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED, EYE IS QUIET. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS" INJECTOR MODEL-MSI-TF; LOT#-UNK SHOULD HAVE BEEN INCLUDED. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED. PRE-EXISTING CONDITION: NARROW ANTERIOR CHAMBER ANGLES SHOULD HAVE BEEN INCLUDED. PRE-EXISTING CONDITION: PREVIOUS CORNEAL OR REFRACTIVE SURGERY. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980474&pc=QCB | 2023826-2020-02945 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: REDDIT COMPLAINT: THE PATIENT HAD A 12.6MM VTICM5_12.6 IMPLANTABLE LENS, -12.5/+1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTED INTO HIS LEFT (OS) EYE. THE PATIENT REPORTS HALOS, DOUBLE VISION, LIGHT RINGS, POSTOP PAIN, PRESSURE, AND VOMITING. REPORTEDLY, THE VOMITING WAS POSSIBLY CAUSED BY THE PATIENT TAKING XANAX PRIOR TO SURGERY. POSTOP PAIN AND PRESSURE RELIEVED BY "PUSHING ONTO THE CUT TO FORCE FLUID OUT." Manufacturer Narrative: B5: ADDITIONAL INFORMATION: 6HRS POST-OP AC FLUID DRAINED VIA EXISTING PARACENTESIS. DIAMOX ADMINISTERED. GHOSTING OF VISION IN LOW LIGHT. ELEVATED IOP DUE TO RETAINED OCCUCOAT. POST-OP HALO/GLARE IS EXPECTED DUE TO SCOTOPIC PUPIL SIZE AND RELATIVELY HIGHT MYOPIC PRESCRIPTION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980487&pc=QCB | 2023826-2020-02946 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: REDDIT COMPLAINT: THE PATIENT HAD A 12.6MM VTICM5_12.6 IMPLANTABLE LENS, -13.0/+2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTED INTO HIS RIGHT (OD) EYE. THE PATIENT REPORTS HALOS, LIGHT RINGS, AND POSTOP VOMITING. REPORTEDLY, THE VOMITING WAS POSSIBLY CAUSED BY THE PATIENT TAKING XANAX PRIOR TO SURGERY. Manufacturer Narrative: B5-ADDITIONAL INFO: ELEVATED IOP 6HRS POST-OP DUE TO RETAINED OCCUCOAT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: CLINICAL CODE: 4581, FOR "GHOSTING VISION". 2227, HALO IS NOT APPLICABLE. DEVICE PROBLEM CODE: 3001, GLARE/HALO IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671674&pc=QCB | 2023826-2019-01020 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: REFERENCE MFR REPORT# 2023826-2019-00886 FOR INITIAL LENS IMPLANT. THE SURGEON EXCHANGED THE INITIAL IMPLANT WITH A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/013 (SPHERE/CYLINDER/AXIS). THE SURGEON REPORTS LENS ROTATION, HALOS, STARBURST, "COMA DOUBLE IMAGE" AND PATIENT IS ON "SYMBIZA+MONOPROST". Manufacturer Narrative: PRODUCT PROBLEM SHOULD BE CORRECTED TO ADVERSE EVENT IN INITIAL MEDWATCH REPORT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT DUE TO MEDICATION INTERVENTION. SURGEON REPORTED PER LAST VISIT DATE ON (B)(6) 2019, PATIENT STILL IS EXPERIENCING MILD HALOS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: MALFUNCTION SHOULD BE CORRECTED TO SERIOUS INJURY IN INITIAL MEDWATCH REPORT. ADDITIONAL INFORMATION: CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE: 1937 SHOULD BE CORRECTED TO PATIENT CODE: 1937, THE PATIENT'S IOP IS MAINTAINED ON TWO IOP LOWERING MEDICATIONS: SIMBRINZA (BRINZOLAMIDE/BRIMONIDINE TARTRATE OPHTHALMIC SUSPENSION) AND MONOPROST. THE ICL REMAINS IMPLANTED AS OF THIS TIME. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140593&pc=QCB | 2023826-2020-01201 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT REPORTS ELEVATED IOP WHICH HAS BEEN RESOLVED BUT PAIN PERSISTS. ALSO REPORTED: FIXED MYDRIASIS, NO RESPONSE TO LIGHT STIMULI, LASER IRIDOTOMY, PILOCARPINE ADMINISTRATION, HALOS AND CIRCLES. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403864&pc=QCB | 2023826-2020-01732 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN THE AMERICAN JOURNAL OF OPHTHALMOLOGY WAS RECEIVED ON (B)(6) 2020, ENTITLED 'COMPARISON OF TORIC IMPLANTABLE COLLAMER LENS AND TORIC ARTIFLEX PHAKIC IOLS IN TERMS OF VISUAL OUTCOME, A PAIRED CONTRALATERAL EYE STUDY.' THE ARTICLE STATES THAT AFTER IMPLANTATION OF THE VISIAN TORIC IMPLANTABLE COLLAMER LENS (TICL), 3 EYES DEVELOPED AN INCREASE OF IOP AFTER 3 WEEKS. THESE WERE TREATED WITH TIMOLOL FOR 1 MONTH AND THE PROBLEM WAS RESOLVED. ALSO 52% OF PATIENTS REPORTED SOME DEGREES OF GLARE AND HALOS WHICH IMPROVED OVER TIME. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6541708&pc=QCB | 2023826-2016-01557 | 2014-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.5/+2.0/97 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP AND REFRACTIVE CHANGE OVER TIME. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. PATIENT LIVES OUTSIDE OF COUNTRY AND WILL ARRANGE FOLLOW-UP WITH SURGEON. Manufacturer Narrative: PMA/510(K): NA, THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6549621&pc=QCB | 2023826-2016-01558 | 2014-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2017-05-05 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.5/+1.0/109 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT HAS EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP AND REFRACTIVE CHANGE OVER TIME. PATIENT LIVES OUTSIDE OF COUNTRY AND WILL ARRANGE FOLLOW-UP WITH SURGEON. Manufacturer Narrative: NA- THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552490&pc=QCB | 2023826-2019-00648 | 2019-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -15.5/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE VA OF THE EYE AFTER EXPLANT IS REPORTED AS "FOGGY, SMEARED." Manufacturer Narrative: PATIENT CODE 2140: VISUAL DISTURBANCE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10030329&pc=QCB | 2023826-2020-01045 | 2020-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), EXCESSIVE VAULT AND HALOS. MEDICAL INTERVENTION WAS PRESCRIBED TO TREAT IOP (INTRAOCULAR PRESSURE). THE SURGEON PLANS TO EXCHANGE THE LENS WITH A SHORTER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "REPORTEDLY, "EVERYTHING WAS OK" AFTER THE EXCHANGE." SHOULD BE REMOVED FROM SUPPLEMENTAL MDR#1. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS HAS BEEN EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. REPORTEDLY, "EVERYTHING WAS OK" AFTER THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10030336&pc=QCB | 2023826-2020-01044 | 2020-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -15.00/5.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), EXCESSIVE VAULT AND HALOS. MEDICAL INTERVENTION WAS PRESCRIBED TO TREAT IOP (INTRAOCULAR PRESSURE). THE SURGEON PLANS TO EXCHANGE THE LENS WITH A SHORTER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR." SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE 2682 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - INVESTIGATION CODE 4117 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS HAS BEEN EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. REPORTEDLY, "EVERYTHING WAS OK" AFTER THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17356616&pc=QCB | 2023826-2023-03045 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Iritis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.50/3.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2023. IT WAS INDICATED THAT THE PATIENT NOW HAS MILD INFLAMMATION AND SIGNIFICANT VAULT REDUCTION. IT WAS ADDITIONALLY INDICATED A POSSIBLE LENS EXPLANT/REPLACE IS PLANNED. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11168622&pc=QCB | 2023826-2021-00132 | 2021-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Iritis; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -16.00/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE REPORTER STATED THAT THE PATIENT IS VAULTED AT OVER 900 AND HAS A MILD ON AND OFF IRITIS. LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. ADDITIONAL INFORMATION: H6 - HEALTH EFFECT CLINICAL CODE : 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES,IRIDOCYCLITIS. B5 - IT WAS LATER REPORTED THAT THE PATIENT WAS CURRENTLY RECOVERING WELL WITH BETTER VAULT AND ANGLE. POSTOPERATIVELY THE PATIENT WAS GLASSES RX FOR IMPROVING HER DISTANCE BUT IS OVERALL FAIRLY PLEASED WITH THE OUTCOME. SHE IS USING MEDICATION FOR MILD POSTOPERATIVE IRIDOCYLITIS WHICH IS A RE-OCCURING PATIENT RELATED FACTOR. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "OVERSIZED ICL" AND A PATIENT RELATED FACTOR. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231194&pc=QCB | 2023826-2021-00198 | 2020-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Iritis; Keratitis | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE PATIENT DEVELOPED IRITIS CAUSED BY THE HIGH VAULT AND KERATITIIS CAUSED BY THE ICL REMOVAL. TO DATE THE IRITIS AND KERATITIS HAS IMPROVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334451&pc=QCB | 2023826-2021-00361 | 2021-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Iritis; UGH (Uveitis Glaucoma Hyphema) Syndrome | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING, WITH LENS ROTATION AND PATIENT EXPERIENCED IRITIS. THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE PATIENT HAD MILD IRITIS/POSSIBLE UGH (UVEITIS-GLAUCOMA-HYPHEMA), THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE EYE IS GOOD AND IOP NORMAL, ICL STABLE, WITH NO RECURRENCE OF IRITIS. ELECTED TO EXCHANGE ICL DESPITE ADEQUATE VAULT DUE TO ICL ROTATION (60 DEGREES) AND VARIABLE IOP AND POST-OP CELL/FLARE. INCREASING SIZE APPEARS TO HAVE SOLVED THE PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332958&pc=QCB | 2023826-2020-01569 |  | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Irritation; Pain; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, PAIN DUE TO ROTATION, SHIFTING VAULT AND IRIS IRRITATION. LENS WAS ROTATED, THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - RESULT CODE: 213 CHANGED TO RESULT CODE: 114 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973336&pc=QCB | 2023826-2021-04715 | 2021-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Macular Edema; Inflammation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO PIGMENT DISPERSION AND AC INFLAMMATION WITH CME. REPORTEDLY THE PROBLEM IS RESOLVED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE PATIENT WAS TREATED WITH ORAL AND TOPICAL STEROIDS. NO INFECTION OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973341&pc=QCB | 2023826-2021-04714 | 2021-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Macular Edema; Inflammation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/090 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO PIGMENT DISPERSION AND AC INFLAMMATION WITH CME. REPORTEDLY THE PROBLEM IS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY THE PATIENT WAS TREATED WITH ORAL AND TOPICAL STEROIDS. NO INFECTION OCCURRED AND THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10570818&pc=QCB | 2023826-2020-02258 | 2020-06-23 04:00 | Malfunction | STAAR SUGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: A 12.6MM, TMICL12.6, -15.0/2.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS RETURNED DAMAGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, INSIDE A VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. D9: CORRECTED RETURNED TO MANUFACTURER DATE FROM 14 AUG 2020 TO 10 AUG 2020. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10759304&pc=QCB | 2023826-2020-02537 |  | Injury | STAAR SUGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: IT WAS REPORTED THAT A TORIC ICL PATIENT EXPERIENCED LENS ROTATION. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS REPORTED THAT A TORIC ICL PATIENT EXPERIENCED LENS ROTATION. LENS EXPLANT HAS BEEN PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10759312&pc=QCB | 2023826-2020-02536 |  | Injury | STAAR SUGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: IT WAS REPORTED THAT A TORIC ICL PATIENT EXPERIENCED LENS ROTATION. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS REPORTED THAT A TORIC ICL PATIENT EXPERIENCED LENS ROTATION. LENS EXPLANT HAS BEEN PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4). Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBERS REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10875730&pc=QCB | 2023826-2020-02716 | 2020-09-18 04:00 | Malfunction | STAAR SUGICAL COMPANY | 2020-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A13.2MM, VTICM5_13.2, -10.5/+2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO INJURY. THE REPORTER STATED "TICL WAS BROKEN DUE TO THE JAM OF INJECTOR". ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10825379&pc=QCB | 2023826-2020-02643 | 2020-10-14 04:00 | Malfunction | STAAR SUGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON NOTED A SMALL TEAR ON A 12.6MM, VTICM5_12.6, -10.50/1.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WHILE INSERTING INTO THE PATIENTS EYE,LENS WAS REMOVED BEFORE DEPLOYMENT. THERE WAS PATIENTS CONTACT BUT NO PATIENT INJURY. ANOTHER LENS WAS SUCCESSFULLY IMPLANTED. THE REPORTER STATED " TEAR LIKELY OCCURRED DURING LOADING BUT NOT VISIBLE WHILE IN CARTRIDGE. (TEAR IS EXTREMELY SMALL)". Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10777844&pc=QCB | 2023826-2020-02573 | 2020-09-28 04:00 | Malfunction | STAAR SUGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.0/+2.50/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION INTO THE EYE (HAPTIC GOT STUCK IN THE PLUNGER AND GOT TORN). THERE WAS PATIENT CONTACT BUT NO INJURY. IT WAS REPORTED THAT THE SURGEON REMOVED THE LENS WITHIN THE SAME SURGERY. A REPLAEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. PATIENT IS HAPPY. CAUSE OF EVENT " LENS GOT STUCK IN THE PLUNGER AND GOT TORN." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# 717559.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779178&pc=QCB | 2023826-2020-02575 | 2020-10-05 04:00 | Malfunction | STAAR SUGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A 12.1MM, VTICM5_12.1, -9.50/+2.0/089, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT CONTACT. CAUSE OF EVENT WAS UNKNOWN. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED, PATIENT IS HAPPY. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A HAPTIC AND THE OPTIC TORN. RESIDUE ON LENS SURFACE WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - "THAT ON (B)(6) 2020" SHOULD BE CORRECTED TO "THAT ON (B)(6) 2020", ALSO "(B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED" SHOULD BE CORRECTED TO "(B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11122218&pc=QCB | 2023826-2021-00090 | 2020-12-03 | Malfunction | STAAR SUGICAL COMPANY | 2021-01-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -3.00/0.5/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING THE LOADING PROCESS. THE LENS HAD NO CONTACT WITH THE EYE. THE DOCTOR IMPLANTED A STAND BY LENS INSTEAD. PROBLEM RESOLVED, PATIENT IS HAPPY. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: 123830-2604, UNK, UNK, UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# 718336.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11122225&pc=QCB | 2023826-2021-00071 | 2020-09-17 04:00 | Malfunction | STAAR SUGICAL COMPANY | 2021-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.00/5.0/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THE REPORTER STATED " THE DOCTOR SOLVES THE SURGERY WITHOUT PUTTING A NEW ICL LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11241826&pc=QCB | 2023826-2021-00247 | 2020-12-28 05:00 | Injury | STAAR SUGICAL COMPANY | 2021-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.00/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). HIGH VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8/0.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. TYPE OF INVESTIGATION: 4110 - DISREGARD PREVIOUS WORK ORDER SEARCH, PREVIOUS LENS REPORTED IS NOT APPLICABLE FOR THIS CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11241828&pc=QCB | 2023826-2021-00246 | 2020-12-28 05:00 | Injury | STAAR SUGICAL COMPANY | 2021-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). HIGH VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM ,"PATIENT IS GREAT". CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "SIZING." CORRECTED DATA: TYPE OF INVESTIGATION: DISREGARD PREVIOUS 4110 - LENS WORK ORDER SEARCH WAS FOR LENS NOT APPLICABLE FOR CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8/+1/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER AND NOTED "SIZING". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11325804&pc=QCB | 2023826-2021-00425 |  | Malfunction | STAAR SUGICAL COMPANY | 2021-02-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.00/3.5/068, IMPLANTABLE COLLAMER LENS TOUCHED THE PATIENTS EYE. THE REPORTER STATED "DR SAID IT WOULD NOT FIT, NO OTHER STAAR LENS WAS USED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11373577&pc=QCB | 2023826-2021-00499 | 2021-01-21 05:00 | Injury | STAAR SUGICAL COMPANY | 2021-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -12.5/2.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS INSERTED AND REMOVED. LENS WAS TOO LARGE, REPLACED WITH ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, INSIDE MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC AND THE OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11382610&pc=QCB | 2023826-2021-00533 | 2020-11-12 | Injury | STAAR SUGICAL COMPANY | 2021-02-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -12.5/1.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11383350&pc=QCB | 2023826-2021-00539 | 2020-12-21 05:00 | Injury | STAAR SUGICAL COMPANY | 2021-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.0/2.0/89(SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS OTHER "THE LENS SIZE DID NOT FIT THE PATIENTS EYE." Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. RESIDUE ON LENS SURFACE WAS OBSERVED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10505357&pc=QCB | 2023826-2020-02053 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -14.0/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10505367&pc=QCB | 2023826-2020-02051 | 2020-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.0/71 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10505368&pc=QCB | 2023826-2020-02052 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -14.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10505369&pc=QCB | 2023826-2020-02054 | 2020-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10508376&pc=QCB | 2023826-2020-02062 | 2020-08-11 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 THE SURGEON NOTED A TEAR ON 13.2MM, VTICM5_13.2, -13.0/4.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WHEN THE CARTRIDGE WAS ALREADY IN THE PATIENTS LEFT EYE (OS). THE LENS LENS WAS NOT INJECTED INTO THE EYE. HE REMOVED THE LOADED CARTRIDGE FROM THE EYE AND ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED. THERE WAS NO PATIENT INJURY. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511714&pc=QCB | 2023826-2020-01981 | 2020-07-09 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511774&pc=QCB | 2023826-2020-01982 | 2020-07-15 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10508682&pc=QCB | 2023826-2020-02060 | 2020-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.50/4.00/18 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. OTHER, "LENS SIZE WAS ACCORDING TO OCOS NOMOGRAM CORRECT, BUT VAULT TOO HIGH AFTER IMPLANTATION" WAS REPORTED TO BE THE CAUSE OF THIS EVENT. LENS EXCHANGE IS PLANNED BUT NOT YET TAKEN PLACE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5:THE REPORTER INDICATED THAT A12.6MM, VTICMO12.6, -14.50/4.00/18 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). ON 11 SEP 2020 THE LENS WAS REMOVED AND LATER EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5:THE REPORTER INDICATED THAT A12.6MM, VTICMO12.6, -14.50/4.00/18 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511666&pc=QCB | 2023826-2020-01987 | 2019-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.50/+1.0/121 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511746&pc=QCB | 2023826-2020-01993 | 2020-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.50/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511754&pc=QCB | 2023826-2020-01995 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, 11.50/+2.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511760&pc=QCB | 2023826-2020-01988 | 2020-08-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/089 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND BENT, WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511786&pc=QCB | 2023826-2020-01983 | 2019-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511788&pc=QCB | 2023826-2020-01984 | 2019-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515397&pc=QCB | 2023826-2020-02019 | 2020-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -6.00/3.0/104 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED AND REPLACED DUE TO LOW VAULT WITH ROTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -6.00/3.0/104 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2020 DUE TO LOW VAULT. LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTER STATED THAT THERE WAS VISCOELASTING AFTER EXCHANGE IN THE EYE; SURGEON WILL CLOSE MONITOR THE PATIENT. CAUSE WAS REPORTED AS UNKNOWN. 6A. (B)(6) 2020. 6B. (B)(6) 2020. H6-IMPACT CODE 4629; LENS EXCHANGED. CORRECTED DATA B5- IN INITIAL MDR THE STATEMENT "THE LENS WAS REMOVED AND REPLACED DUE TO LOW VAULT WITH ROTATION." IS NOT APPLICABLE. ROTATION DOES NOT APPLY FOR THIS CASE. H6-TYPE OF INVESTIGATION 4110. NEEDED TO BE INCLUDED IN INITIAL MDR. H10-LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. NEEDED TO BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515408&pc=QCB | 2023826-2020-02031 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.00/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT, SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515420&pc=QCB | 2023826-2020-02029 | 2020-08-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICM5_13.2, -12.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION INTO PATIENTS' EYE AND WAS REMOVED. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "THE LENS WAS SANDWICHED BETWEEN THE FTP AND THE CARTRIDGE INNER CYLINDER." REPORTEDLY, "PATIENT EXPERIENCED POSTOPERATIVE MILD INFLAMMATION WHICH IS COMMONLY SEEN AFTER EYE SURGERY. THERE IS NO HEALTH DAMAGE." PLAN TO IMPLANT ANOTHER LENS IS SCHEDULED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT IMPACT CODE 2645 SHOULD BE CORRECTED TO 2199 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515453&pc=QCB | 2023826-2020-02027 | 2020-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.00/2.5/009 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10523555&pc=QCB | 2023826-2020-02032 | 2020-01-21 | Injury | STAAR SURGICAL COMPANY | 2020-09-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.5/2.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10524392&pc=QCB | 2023826-2020-02036 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.0/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: B5-PREVIOUSLY STATED A -5.0/+1.0/074 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED. CORRECT VALUE IS -5.0/+1.0/093. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10532418&pc=QCB | 2023826-2020-02035 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -13.5/3.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10529736&pc=QCB | 2023826-2020-02072 | 2020-07-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/090 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10529866&pc=QCB | 2023826-2020-02075 | 2019-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING WITH LENS ROTATION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUD. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK. UNK. UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT IMPACT CODE: (B)(4). HEALTH EFFECT IMPACT CODE: (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10530326&pc=QCB | 2023826-2020-02080 | 2020-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.00/+1.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH FOREIGN RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10530402&pc=QCB | 2023826-2020-02078 | 2019-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.00/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10530461&pc=QCB | 2023826-2020-02077 | 2020-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: UNK. UNK. UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT IMPACT CODE: (B)(4). DEVICE CODE: (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10530598&pc=QCB | 2023826-2020-02071 | 2020-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK. UNK. UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT IMPACT CODE: (B)(4). DEVICE CODE: (B)(4). CLAIM #:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10530645&pc=QCB | 2023826-2020-02070 | 2020-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/117 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUDN THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK. UNK. UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT IMPACT CODE: (B)(4). DEVICE CODE: (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10531804&pc=QCB | 2023826-2020-02159 | 2020-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/2.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10532956&pc=QCB | 2023826-2020-02042 | 2020-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/1.5/075 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10538711&pc=QCB | 2023826-2020-02045 | 2020-06-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-15 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.0/3.5/061 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), A TEAR ON THE HAPTIC WAS NOTED. THE DEFECT WAS NOTED IN THE LENS ONCE IT WAS LOADED. THERE WAS NO PATIENT CONTACT WITH THE LENS. AN ALTERNATE LENS WAS IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537771&pc=QCB | 2023826-2020-01992 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE SURGEON ENLARGED THE PI BY YAG PROCEDURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10538712&pc=QCB | 2023826-2020-02048 | 2020-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON ATTEMPTED TO IMPLANT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/1.50/002 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS WAS TORN WHILE REMOVING IT FROM THE VIAL. THERE WAS NO PATIENT CONTACT. THIS OCCURRED ON (B)(6) 2020. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON THE SAME DATE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND FOREIGN DEBRIS AND A SCRATCH ON THE OPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10538716&pc=QCB | 2023826-2020-02044 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -10.50/2.5/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548194&pc=QCB | 2023826-2020-02111 | 2020-08-26 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -7.00/4.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548209&pc=QCB | 2023826-2020-02110 | 2020-08-26 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.50/4.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548104&pc=QCB | 2023826-2020-02082 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548110&pc=QCB | 2023826-2020-02083 | 2020-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/2.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS STUCK IN CARTRIDGE/INJECTION SYSTEM DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED REPLACED INTRAOPERATIVELY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, "THE LENS WAS STUCK IN CARTRIDGE AND THE SURGEON FOUND OUT AFTER HE PUT THE LENS INTO PATIENT'S EYE, THEN HE TAKE IT OU, SO THERE IS NO DAMAGE TO PATIENT." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: B5 - LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, AND REPLACED WITH AN ALTERNATE LENS ON (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548118&pc=QCB | 2023826-2020-02088 | 2020-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.00/4.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548145&pc=QCB | 2023826-2020-02095 | 2020-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND A LONGER LENGTH LENS WAS LATER IMPLANTED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE PROBLEM WAS RESOLVED, SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - "ON (B)(6) 2020 THE LENS WAS REMOVED" SHOULD BE CORRECTED TO "ON (B)(6) 2020 THE LENS WAS REMOVED" WHICH SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH #2. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - DATE OF EVENT - SHOULD BE CORRECTED TO UNKNOWN. B5 - SEE MFR# 2023826-2020-02661 FOR FOLLOW UP CLAIM REGARDING EXCHANGED LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548152&pc=QCB | 2023826-2020-02096 | 2020-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND A LONGER LENGTH LENS WAS LATER IMPLANTED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548189&pc=QCB | 2023826-2020-02115 | 2020-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.50/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548195&pc=QCB | 2023826-2020-02107 | 2019-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.00/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548199&pc=QCB | 2023826-2020-02114 | 2020-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -5.50/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REFRACTIVE TREATMENT PERFORMED. THIS DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "VAULT IS STILL HIGH, THE SURGE ON WILL DECIDE WHETHER TO CHANGE A SMALLER SIZE AFTER A WEEK." Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED REFRACTIVE TREATMENT WAS INCORRECTLY REPORTED AND SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6/H10 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. - STATEMENT SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551347&pc=QCB | 2023826-2020-02143 | 2020-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551348&pc=QCB | 2023826-2020-02141 | 2020-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+2.0/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551349&pc=QCB | 2023826-2020-02129 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/130 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- INVESTIGATION CODE 4110 FOR TREND ANALYSIS SHOULD HAVE BEEN INCLUDED. RESULT CODE IN INITIAL MDR CORRECTED FROM 114 TO 213. H10- 4110- LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551352&pc=QCB | 2023826-2020-02121 | 2020-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/124 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551355&pc=QCB | 2023826-2020-02126 | 2020-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE LOADING A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE CARTRIDGE, THE LENS TORE. IT WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE TEAR OF THE LENS IS OBSERVED IN THE SUPERIOR TEMPORAL PUPILLARY BORDER. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B5-THE REPORTER INDICATED THAT, WHILE LOADING A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE CARTRIDGE, THE LENS TORE. IT WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE TEAR OF THE LENS IS OBSERVED IN THE SUPERIOR TEMPORAL PUPILLARY BORDER. THE LENS WAS EXCHANGED ON (B)(6)2020, FOR A SAME MODEL/ SIZE AND POWER LENS. THE EXCHANGE RESOLVED THE PROBLEM. D6B- (B)(6) 2020. H6-IMPACT CODE 4629; LENS EXCHANGED. CORRECTED DATA: B7- PROGRESSIVE MYOPIA. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION; H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND LENS RETURNED IN PIECES. RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551578&pc=QCB | 2023826-2020-02152 | 2020-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -10.50/5.50/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE BEFORE/DURING LOADING. THE REPORTER STATED " THE LENS WAS LOADED IN THE CARTRIDGE AND THE FOAM TIP WAS PLACED IN THE INJECTOR AND THEN CARTRIDGE PUT IN THE INJECTOR. AS THE DOCTOR WAS PUSHING THE LENS FORWARD WITH FOAM TIP TO CHECK THAT EVERYTHING WAS IN PLACE, HE NOTICED THAT THE ALIGNMENT OF THE LENS WAS NOT RIGHT AND DECIDED TO RELOAD THE LENS TO AVOID COMPLICATIONS ON INJECTING THE LENS UPSIDE DOWN. HE REMOVED THE ICL AND WHEN RELOADING IN THE INJECTOR, HE NOTICED A TEAR IN THE LENS AND DECIDED IT'S NOT SAFE TO IMPLANT." THERE WAS NO PATIENT CONTACT AND THE CAUSE OF THE EVENT IS UNKNOWN. A REPLACEMENT LENS WILL BE IMPLANTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN AND RESIDUE ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10556372&pc=QCB | 2023826-2020-02156 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.50/3.00/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED IN A SEPARATE SURGERY ON 4 AUG 2020 AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH DEBRIS ON PRODUCT. VISUAL INVESTIGATION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10557005&pc=QCB | 2023826-2020-02144 | 2020-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10557007&pc=QCB | 2023826-2020-02148 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.5/+4.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CORRECTED DATA: B5 - IT WAS CLARIFIED THAT THE PATIENT HAD EXPERIENCED LOW VAULT AND ON 25/JUL/2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10557017&pc=QCB | 2023826-2020-02145 | 2018-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10558066&pc=QCB | 2023826-2020-02165 | 2020-08-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.50/+0.5/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND REMAINS IMPLANTED. THE LENS WILL BE EXPLANTED/EXCHANGED AT A LATER DATE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5:THE REPORTER INDICATED THE LENS WILL REMAIN IMPLANTED, THE ANGLE IS OPEN AND THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10563698&pc=QCB | 2023826-2020-02220 | 2020-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.0/5.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10563700&pc=QCB | 2023826-2020-02215 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.5/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10563703&pc=QCB | 2023826-2020-02199 | 2020-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -14.0/2.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10563714&pc=QCB | 2023826-2020-02217 | 2020-04-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT WITHIN THE SAME SURGERY AND PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED WITH DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10563722&pc=QCB | 2023826-2020-02216 | 2020-04-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.5/1.5/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT WITHIN THE SAME SURGERY AND PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10563724&pc=QCB | 2023826-2020-02219 | 2020-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.50/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CORRECTED DATA: B4 - CORRECTED TO 28/08/2020. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10563725&pc=QCB | 2023826-2020-02204 | 2020-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -14.50/4.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10563732&pc=QCB | 2023826-2020-02218 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -12.0/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566518&pc=QCB | 2023826-2020-02181 | 2020-04-22 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566573&pc=QCB | 2023826-2020-02178 | 2020-08-13 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +6.00/+2.0/132 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: UNK, UNK, UNK, UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 717296.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566513&pc=QCB | 2023826-2020-02174 | 2020-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566529&pc=QCB | 2023826-2020-02173 | 2020-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+2.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # 717228.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566541&pc=QCB | 2023826-2020-02167 | 2020-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/160 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566554&pc=QCB | 2023826-2020-02187 | 2020-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566557&pc=QCB | 2023826-2020-02183 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/012 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. D4: SERIAL # (B)(6) D4: EXPIRATION DATE: 31-OCT-2021 H4: DEVICE MANUFACTURE DATE: 08-NOV-2018 CLAIM # (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 717301
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566584&pc=QCB | 2023826-2020-02172 | 2020-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.5/083 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10570629&pc=QCB | 2023826-2020-02253 | 2020-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/0.50/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT, THE REPORTER STATED " MICRO-CYST UNDER THE IRIS WHICH MAKES THE WTW VALUE NOT CORRESPOND TO REAL LENGTH OF THE WTW." THIS EXCHANGE RESOLVED THE PROBLEM, THE REPORTER STATED " PATIENT VISION IS OK FOR NOW, UNDER REGULAR OBSERVATION." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10577513&pc=QCB | 2023826-2020-02196 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN BMC OPHTHALMOLOGY WAS RECEIVED ON 31-AUG-2020, ENTITLED 'SPONTANEOUS ROTATION OF A TORIC IMPLANTABLE COLLAMER LENS RELATED TO ABNORMAL CILIARY BODY MORPHOLOGY: A CASE REPORT.' THE ARTICLE STATES THAT AFTER IMPLANTATION OF A TICM12.1 LENS -12.5/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE IT ROTATED. THE LENS WAS REPOSITIONED AND ROTATED AGAIN. THE LENS WAS EXCHANGED WITH A LONGER NOT-TORIC LENS AND THE PROBLEM WAS RESOLVED. A SHORT AND SMALL CILIARY PROCESS WITH SHALLOW CILIARY SULCUS AND POSTERIORLY POSITIONED CILIARY BODY WAS FOUND BY ULTRASOUND BIOMICROSCOPY (UBM). Manufacturer Narrative: ADDITIONAL INFORMATION: IT WAS DISCOVERED THAT THIS CLAIM IS A DUPLICATE OF FILE# (B)(4). PLEASE DISREGARD CURRENT REPORT. CLAIM# (B)(4). Manufacturer Narrative: G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10577514&pc=QCB | 2023826-2020-02195 | 2020-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10577518&pc=QCB | 2023826-2020-02191 | 2020-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEFORMED/STRETCHED OUT WITH BLOOD STAINS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10577519&pc=QCB | 2023826-2020-02192 | 2020-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. RE-HYDRATION PROCESS DETERMINED LENS HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10577520&pc=QCB | 2023826-2020-02190 | 2020-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. THE LENS WAS SUCCESSFULLY EXCHANGED INTRA-OPERATIVELY. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10585623&pc=QCB | 2023826-2020-02117 | 2020-08-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2, -11.00/1.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CORRECTED DATA: B5 - "THE LENS WAS IMPLANTED AND REMOVED AND REPLACED INTRAOPERATIVELY" SHOULD BE CORRECTED TO "THE LENS WAS REPLACED INTRAOPERATIVELY WITH ANOTHER LENS". THE LENS WAS NOT IMPLANTED. D6A- "31-AUG-2020" SHOULD BE REMOVED AS IT IS NOT APPLICABLE. D6B - "31-AUG-2020" SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10589271&pc=QCB | 2023826-2020-02257 | 2020-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -15.0/2.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, AND RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10600062&pc=QCB | 2023826-2020-02230 | 2020-08-11 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -6.00/+3.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10599982&pc=QCB | 2023826-2020-02223 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). APPROXIMATELY 2 WEEKS LATER, THE LENS HAD ROTATED 90 DEGREES AND THE PATIENT INDICATED "SHAKING HIS HEAD TOO MUCH". THE SURGEON ROTATED THE LENS BACK INTO POSITION. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.00/+4.0/072 (SPHERE/CYLINDER/AXIS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR/UNKNOWN. THE PATIENTS EYE POST-OP IS QUIET, WELL VAULTED, NO EDEMA, FORMED CHAMBER, PATENT PIS, GOOD VISION AND PLANO REFRACTION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10600109&pc=QCB | 2023826-2020-02229 | 2020-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+0.5/179 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND BLACK STAIN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. DEVICE EVALUATION FOR (T529650) PERFORMED FOR INVESTIGATIONAL PURPOSES. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331:THE SUBJECT LENS WAS DETERMINED TO BE INADVERTENTLY SWAPPED WITH ANOTHER SAME SIZE/MODEL LENS OF DIFFERENT POWER FROM A SEPARATE LOT DURING MANUFACTURING. THE SWAPPED LENS WAS STILL IN THE WAREHOUSE AND HAD NOT YET BEEN SHIPPED TO A CUSTOMER. REFRACTIVE EVALUATION OF BOTH LENSES CONFIRM THE SWAP. A LENS WORK ORDER SEARCH OF THE SWAPPED LENS LOT FOUND NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS. BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THE SWAPPED LENS IS LIKELY DUE TO A LINE CLEARANCE FAILURE DURING FINAL INSPECTION REWORK. CORRECTIVE ACTIONS FOR A SIMILAR SWAPPED LENS OCCURRENCE HAVE BEEN IDENTIFIED AND ARE BEING IMPLEMENTED UNDER CAPA M19-08. H6-INVESTIGATION CODE 4110: FOR (T529650)- NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. THE SUBJECT LENS WORK ORDER, WO# (B)(4), DID NOT HAVE A RE-WORK NOTED ON FORM-0185(ICL IN-PROCESS REWORK/REJECT FORM). HOWEVER AN EXTRA LABEL/TRAY/LID WERE REQUESTED DURING FINAL INSPECTION, AND THE SUBJECT LENS WAS NOTED TO BE REWORKED FOR THE TYVEK LID ON FORM-0400 (LABEL REPRINT LOG). IT APPEARS LIKELY THAT DURING REWORK ON S/N (B)(6) THE SUBJECT LENS WAS INADVERTENTLY SWAPPED WITH S/N (B)(6). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10600977&pc=QCB | 2023826-2020-02234 | 2020-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED WITH A LONGER LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL IN LIQUID WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10600981&pc=QCB | 2023826-2020-02233 | 2020-09-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 12.6MM VTICM5_12.6, -11.00/5.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND EXCHANGED DURING THE INITIAL SURGERY ON (B)(6) 2020 WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10600993&pc=QCB | 2023826-2020-02238 | 2020-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/2.5/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. A LENS CALCULATIONS ERROR OCCURRED WHEN THE SURGEON ENTERED A POSITIVE CYLINDER NUMBER INSTEAD OF A MINUS CYLINDER NUMBER. LENS REMAINS IMPLANTED. REPORTEDLY, "NEW LENS ORDERED." CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: G3 - SHOULD BE CORRECTED TO 07/OCT/2020 IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B1 - CHANGE FROM PRODUCT PROBLEM TO ADVERSE EVENT. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT /DAMAGE. B5 - LENS REMOVED AND EXCHANGED WITH 13.2MM, VTICM5_13.2, -12.00/2.5/089 (SPHERE/CYLINDER/AXIS) LENS ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "PATIENT IS SUPER HAPPY." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10609818&pc=QCB | 2023826-2020-02271 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM TORIC LENS WAS IMPLANTED, AND HIGH VAULT WAS OBSERVED. THE REPORTER STATED " THE PATIENT IS VERY HAPPY, AND PRESSURES ARE NORMAL AT 17. THIS SURGERY IS NOW 3 WEEKS POST-OP. THEY ARE UNSURE IF THEY SHOULD LEAVE THE LENS IN AS IS, OR EXCHANGE IT FOR A DIFFERENT SIZE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL RETURNED IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED. H6: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: T WAS LATER CLARIFIED THAT A 13.2MM TMICL13.2; -15.50/2.5/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 10/AUG/2020. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. CLAIM#: (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10617616&pc=QCB | 2023826-2020-02266 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10617635&pc=QCB | 2023826-2020-02268 | 2020-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10617656&pc=QCB | 2023826-2020-02265 | 2020-08-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT, BEFORE/DURING LOADING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/091 (SPHERE/CYLINDER/AXIS), IT TORE. THIS WAS DISCOVERED AFTER OPEINING THE VIAL. THIS OCCURRED ON (B)(6) 2020. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S.. INVESTIGATION TYPE: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10617660&pc=QCB | 2023826-2020-02267 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/078 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10619472&pc=QCB | 2023826-2020-02277 | 2020-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VTICM5_12.6, -10.50/3.50/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10619488&pc=QCB | 2023826-2020-02275 | 2020-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VTICM5_12.1, -8.50/5.00/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10624465&pc=QCB | 2023826-2020-02278 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10624467&pc=QCB | 2023826-2020-02279 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634690&pc=QCB | 2023826-2020-02291 | 2020-08-26 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/053 (SPHERE/CYLINDER/AXIS), INTO THE CARTRIDGE AND THE LENS TORE. THERE WAS NO PATIENT CONTACT. ANOTHER LENS WILL BE IMPLANTED AT A LATER DATE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635476&pc=QCB | 2023826-2020-02248 | 2020-09-15 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -8.50/1.5/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS DISLOCATION/SUBLUXATION WAS OBSERVED. ON (B)(6) 2019 AND (B)(6) 2020 THE LENS WAS REPOSITIONED. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634529&pc=QCB | 2023826-2020-02284 | 2019-01-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.00/+2.0/081 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LENS WAS TOO LONG. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. IT WAS THE SURGEONS PREFERENCE TO CHANGE THE LENS. Manufacturer Narrative: ADDITIONAL DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC STRETCHED OUT. THERE WAS RESIDUE ON LENS. CORRECTED DATA: H6: HEALTH EFFECT IMPACT CODE: 2199 - NO HEALTH CONSEQUENCES OR IMPACT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634621&pc=QCB | 2023826-2020-02288 | 2020-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # 717382.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634735&pc=QCB | 2023826-2020-02286 | 2020-07-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS), AND THE LENS TORE DURING DELIVERY. THERE WAS PATIENT CONTACT BUT NO INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTICS TORN AND RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635451&pc=QCB | 2023826-2020-02244 | 2019-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION- LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635453&pc=QCB | 2023826-2020-02243 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635454&pc=QCB | 2023826-2020-02239 | 2020-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/4.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, RESIDUE ON LENS. CORRECTION: A1 - REPORTED AS 90521, CORRECT TO 90251. A3 - REPORTED AS FEMALE, CORRECT TO MALE. D6A - REPORTED AS (B)(6) 2020, CORRECT TO (B)(6) 2020. D6B - REPORTED AS (B)(6) 2020, CORRECT TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: D6A IS (B)(6) 2020, SHOULD BE REMOVED FROM INITIAL MDR AS IT IS NOT APPLICABLE FOR THIS CLAIM. D6B IS (B)(6) 2020, SHOULD BE REMOVED FROM INITIAL MDR AS IT IS NOT APPLICABLE FOR THIS CLAIM. D9 - CORRECTED TO (B)(6) 2020 . CLAIM# (B)(4). Manufacturer Narrative: D6A - CORRECTED TO (B)(6) 2020. D6B - CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635456&pc=QCB | 2023826-2020-02240 | 2020-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.00/5.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE, RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, RESIDUE ON LENS. CORRECTION: A1 - REPORTED AS (B)(4), CORRECT TO (B)(4); A3 - REPORTED AS FEMALE, CORRECT TO MALE; D6A - REPORTED AS (B)(6) 2020, CORRECT TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 LENS OF -8.00/5.00/77 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020, FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN." D6A - CORRECTED TO (B)(6) 2020. D9 - D9 DATE IN SUPPLEMENTAL #1 MDR IS INCORRECT. D9 DATE IN THE INITIAL MDR IS ACCURATE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635458&pc=QCB | 2023826-2020-02241 | 2020-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635459&pc=QCB | 2023826-2020-02245 | 2019-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TOR AND DEFORMED WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635472&pc=QCB | 2023826-2020-02250 | 2020-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/4.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635473&pc=QCB | 2023826-2020-02252 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS UNFOLDED INAPPROPRIATELY AND EVENTUALLY ENDED UPSIDE DOWN. THE LENS WAS REMOVED AND WAS TORN DURING REMOVAL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE IN INITIAL MDR. B5 - CORRECTED TO: THE REPORTER INDICATED THAT A TORIC IMPLANTABLE COLLAMER LENS OF -13.0/ 3.5/82 (SPHERE/ CYLINDER/ AXIS) WAS OPENED AND PLACED IN THE INSERTER BUT ON EXAMINATION OF THE EYE, THE CRYSTALLINE LENS HAD A LARGE TEAR IN THE LENS CAPSULE (THERE WAS A SMALL IRREGULARITY IN THE CRYSTALLINE LENS UNDER THE SITE OF THE PI) WHICH WOULD REQUIRE A LENSECTOMY AND IOL THE FOLLOWING DAY. THE REPORTER ALSO STATED THAT THIS WAS NOT A PROBLEM WITH THE TICL, BUT RATHER AN ISSUE WITH THE CRYSTALLINE LENS/ LENS CAPSULE DURING THE LASER PI WHICH COMPLICATED THE SURGERY. THERE HAS BEEN NO DEFICIENCY STATED AGAINST THE LENS. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635474&pc=QCB | 2023826-2020-02251 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS UNFOLDED INAPPROPRIATELY AND EVENTUALLY ENDED UPSIDE DOWN. THE LENS WAS REMOVED AND WAS TORN DURING REMOVAL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE IN INITIAL MDR. B5 - CORRECTED TO: THE REPORTER INDICATED THAT A TORIC IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS UNFOLDED INAPPROPRIATELY AND EVENTUALLY ENDED UPSIDE DOWN. THE LENS WAS EXPLANTED AND WAS TORN DURING REMOVAL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10639549&pc=QCB | 2023826-2020-02309 | 2020-09-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -14.50/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME INITIAL SURGERY AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: H3: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, INSIDE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10639551&pc=QCB | 2023826-2020-02307 | 2020-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -16.0/6/52 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS DUE TO PATIENT-RELATED FACTOR, THE REPORTER STATED " THE PROBLEM IS WITH THE SULCUS, NARROW SULCUS". Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICO-CENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659330&pc=QCB | 2023826-2020-02318 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON(B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. CLAIM # : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659331&pc=QCB | 2023826-2020-02324 | 2020-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/000 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: REPORT WAS LATER CLARIFIED BY THE REPORTER THAT THIS CLAIM IS FOR PATIENT'S RIGHT EYE (OD) AND NOT PATIENT'S LEFT EYE (OS). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659332&pc=QCB | 2023826-2020-02316 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659336&pc=QCB | 2023826-2020-02320 | 2020-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.0/+4.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659340&pc=QCB | 2023826-2020-02314 | 2020-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/+4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT AXIS LENS DUE TO LENS ROTATION. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659482&pc=QCB | 2023826-2020-02321 | 2020-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.5/+4.0/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659486&pc=QCB | 2023826-2020-02323 | 2020-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2020 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE LENS." Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10668281&pc=QCB | 2023826-2020-02396 | 2020-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM,VTICMO12.6, -8.5/3.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # 717427.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10668282&pc=QCB | 2023826-2020-02399 | 2020-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/1.5/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10668283&pc=QCB | 2023826-2020-02393 | 2020-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.5MM, TICM125V4, -6.00/+2.50/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2008. ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE/DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10668284&pc=QCB | 2023826-2020-02392 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM,VTICMO12.6, -8.0/1.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10668285&pc=QCB | 2023826-2020-02391 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM,VTICMO12.6, -9.0/1.5/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10668286&pc=QCB | 2023826-2020-02398 | 2020-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.0/1.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10668287&pc=QCB | 2023826-2020-02395 | 2020-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.5/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669788&pc=QCB | 2023826-2020-02344 | 2020-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669848&pc=QCB | 2023826-2020-02339 | 2020-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669852&pc=QCB | 2023826-2020-02338 | 2020-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669922&pc=QCB | 2023826-2020-02335 | 2020-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669938&pc=QCB | 2023826-2020-02334 | 2020-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+4.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678729&pc=QCB | 2023826-2020-02366 | 2019-11-21 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/1.5/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678683&pc=QCB | 2023826-2020-02374 | 2020-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.00/2.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678686&pc=QCB | 2023826-2020-02372 | 2020-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/2.5/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SAME LENGTH LENS BUT DIFFERENT POWER (-13.50/2.5/092 (SPHERE/CYLINDER/AXIS)) DUE TO WRONG ICL CALCULATION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678688&pc=QCB | 2023826-2020-02352 | 2020-08-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/2.0/092 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED DURING INITIAL SURGERY ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE LENS GOT STUCK ON THE INJECTOR DURING THE IMPLANTATION PROCESS AND THE LENS WAS ACCIDENTALLY PULLED OUT WHILE ADJUSTING." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678692&pc=QCB | 2023826-2020-02351 | 2019-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678696&pc=QCB | 2023826-2020-02356 | 2020-09-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020, A VTICMO 12.6, -18.00/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. NO PATIENT CONTACT. DAMAGE WAS NOTED AFTER OPENING VIAL. A BACKUP LENS WAS IMPLANTED ON (B)(6) 2020, AND THIS RESOLVED THE PROBLEM. PER REPORTER ON 24/SEP/2020, "PATIENT'S CURRENT CONDITION IS PERFECT." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678720&pc=QCB | 2023826-2020-02362 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEFORMED AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678724&pc=QCB | 2023826-2020-02361 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/1.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678727&pc=QCB | 2023826-2020-02365 | 2020-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.50/1.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678728&pc=QCB | 2023826-2020-02367 | 2020-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/4.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VSUAL INSPECTION FOUND HAPTIC TORN/ BENT WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678751&pc=QCB | 2023826-2020-02369 | 2020-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678754&pc=QCB | 2023826-2020-02386 | 2020-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.05/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.05/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 IN PREVIOUS MDR- CORRECT THE B5 STATEMENT TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678755&pc=QCB | 2023826-2020-02371 | 2020-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.5/145 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678771&pc=QCB | 2023826-2020-02359 | 2020-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687532&pc=QCB | 2023826-2020-02397 | 2020-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM,VTICM5_13.2, -3.0/1.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687535&pc=QCB | 2023826-2020-02401 | 2020-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -13.50/3.00/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CORRECTED DATA: H1 - SHOULD HAVE BEEN SELECTED AS "SERIOUS" IN INITIAL MEDWATCH REPORT. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687536&pc=QCB | 2023826-2020-02394 | 2020-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A 12.6MM, VTICMO12.6, -18.0/+2.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT (OD) EYE . THE SURGEON REMOVED THE LENS WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. THE REPORTER STATED " LENS GOT STUCK IN THE PLUNGER AND GOT TORN ". IT WAS REPORTED THAT ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687537&pc=QCB | 2023826-2020-02403 | 2020-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A 12.6MM, VTICM5_12.6, -12.00/+2.00/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS) . THE LENS WAS REMOVED WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. IT WAS REPORTED THAT ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CORRECTED DATA: D10: CORRECTED TO CARTRIDGE MODEL-SFC-45; LOT #: UNK. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687539&pc=QCB | 2023826-2020-02402 | 2020-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/4.00/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR THE SAME LENGTH LENS. THE REPORTER STATED " SINCE THERE WAS NO VAULTING ISSUE AND ONLY ROTATED, THE LENS WAS REPLACED WITH THE SAME SIZE. AFTER REPLACING THE LENS, THE PROBLEM RESOLVED." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT IMPACT CODE: 4628 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687732&pc=QCB | 2023826-2020-02381 | 2020-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+5.0/002 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687733&pc=QCB | 2023826-2020-02390 | 2020-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687734&pc=QCB | 2023826-2020-02383 | 2020-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.5/2.0/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B5 - -9.5/2.0/180 SHOULD BE CORRECTED TO -9.5/2.0/093, FOR INITIAL LENS (SPHERE/CYLINDER/AXIS). ADDITIONAL INFORMATION: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687739&pc=QCB | 2023826-2020-02388 | 2020-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/176 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687745&pc=QCB | 2023826-2020-02380 | 2020-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+5.0/004 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BENT WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10693249&pc=QCB | 2023826-2020-02405 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.50/1.0/051 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS KEEPS ON ROTATING, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10693250&pc=QCB | 2023826-2020-02406 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.00/1.5/111(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS KEEPS ON ROTATING, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10701147&pc=QCB | 2023826-2020-02427 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. LENS ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10701118&pc=QCB | 2023826-2020-02425 | 2019-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.50/1.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10701156&pc=QCB | 2023826-2020-02424 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -6.00/4.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5: IT WAS LATER REPORTED THAT LENS DID NOT ROTATE AND ONLY HAD LOW VAULTING. H6: MEDICAL DEVICE PROBLEM CODE: 2923 SHOULD BE REMOVED AS IT'S NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH DEBRIS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN DEBRIS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708894&pc=QCB | 2023826-2020-02444 | 2020-07-30 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-20 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2020. REPORTEDLY, THERE WAS NO PATIENT INJURY AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706634&pc=QCB | 2023826-2020-02412 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706654&pc=QCB | 2023826-2020-02413 | 2020-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/089 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706715&pc=QCB | 2023826-2020-02416 | 2020-05-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/080 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH MOISTURE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706726&pc=QCB | 2023826-2020-02418 | 2020-05-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706803&pc=QCB | 2023826-2020-02422 | 2020-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708895&pc=QCB | 2023826-2020-02436 | 2019-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708896&pc=QCB | 2023826-2020-02438 | 2017-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, HAPTIC TORN AND BENT, AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708905&pc=QCB | 2023826-2020-02437 | 2017-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708925&pc=QCB | 2023826-2020-02457 | 2020-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.5/1.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708927&pc=QCB | 2023826-2020-02458 | 2020-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.5/1.0/82 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708928&pc=QCB | 2023826-2020-02451 | 2020-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.0/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708931&pc=QCB | 2023826-2020-02460 | 2020-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.0/2.0/142 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708933&pc=QCB | 2023826-2020-02455 | 2020-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.5/3.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708934&pc=QCB | 2023826-2020-02456 | 2020-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/3.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708939&pc=QCB | 2023826-2020-02445 | 2020-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -14.0/0.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10711586&pc=QCB | 2023826-2020-02493 | 2020-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.0/2/76 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10711821&pc=QCB | 2023826-2020-02496 | 2020-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/+1.5/68(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10711935&pc=QCB | 2023826-2020-02495 | 2020-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/+1.5/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED LOW VAULT, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT A LENS EXCHANGED IS PLANNED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5:THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/+1.5/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10722603&pc=QCB | 2023826-2020-02490 | 2020-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10722606&pc=QCB | 2023826-2020-02489 | 2019-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+0.5/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10731835&pc=QCB | 2023826-2020-02476 | 2020-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+3.0/101 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. ANOTHER LENS WILL BE IMPLANTED AT A LATER DATE.THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED, AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10731834&pc=QCB | 2023826-2020-02480 | 2020-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50/+0.5/035 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, THE HAPTIC TORN AND BENT WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740158&pc=QCB | 2023826-2020-02485 | 2019-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE, RESIDUE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740471&pc=QCB | 2023826-2020-02477 | 2020-09-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/104 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO SIZING ISSUE (TOO LARGE). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5: THE OVD IRRIGATION AND EVACUATION WAS PERFORMED INTRA-OPERATIVE DURING IMPLANTATION OF THE INITIAL LENS. POST-OP THE BCVA WAS 20/20 AND GOOD POST-OP APPEARANCE. THE LENS WAS NOT RETURNED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740519&pc=QCB | 2023826-2020-02486 | 2020-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), BUT THE LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT CONTACT . ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTICS TORN AND THERE WAS RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740541&pc=QCB | 2023826-2020-02484 | 2019-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE, DEBRIS AND RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THERE WAS DEBRIS AND RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740546&pc=QCB | 2023826-2020-02483 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/085 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS WITHIN THE SAME SURGERY, AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10741883&pc=QCB | 2023826-2020-02530 | 2020-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -9.0/1.5/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749975&pc=QCB | 2023826-2020-02522 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Patient-Device Incompatibility; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WANTS TO OBTAIN ADVICE AS TO HOW MUCH TO ROTATE A TMICL LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE SURGEON STATED THAT THERE IS NO ISSUE WITH THE IMPLANTED LENS. THIS IS A CORNEAL ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10750005&pc=QCB | 2023826-2020-02524 | 2020-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -16.5/1.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: B5 - PREVIOUSLY STATED A VTICM5_12.1 -16.5/+1.5/098 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED. CORRECTED INFORMATION: A VTICM5_12.1 -9.5/+1.5/098 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755729&pc=QCB | 2023826-2020-02473 | 2020-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -6.00/0.5/171 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -6.00/0.5/171 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755750&pc=QCB | 2023826-2020-02462 | 2020-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, -0.00/5.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-REPORTEDLY, "THE LENS IMPLANTED AND ALIGNED AS PER IOD TO 14 DEGREES CLOCKWISE (166 DEGREES) ROUTINE AND IN ABSENCE OF COMPLICATIONS. ON EXAMINATION POST-OP DAY 1 IT WA FOUND TO HAVE MOVED TO A POSITION OF 15 DEGREES (29 DEGREES FROM INTENDED AXIS). THIS RESULTED IN A SIGNIFICANT REFRACTION FOR WHICH WE HAVE UNDERTAKEN A SECONDARY ROTATION - WELL ALIGNED AND BACK IN POSITION. THE LENS HAD ROTATED 25 DEGREES FROM THE INTENDED AXIS TO A POSITION OF 10 DEGREES FOLLOWING REPOSITIONING." LENS WAS EXPLANTED ON (B)(6) 2020. IT WAS REPORTED THAT THE SURGEON CONCLUDED THAT THE AUTONOMY OF THE EYE WAS TOO LARGE FOR AN EXCHANGE THEREFORE DECIDED TO EXPLANT AND WILL CONTINUE LASER PROCEDURE FOR THIS EYE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755784&pc=QCB | 2023826-2020-02470 | 2020-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4) . Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755787&pc=QCB | 2023826-2020-02518 | 2020-08-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Ejection Problem | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A TMICL12.6, -8.50/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS "WAS HUNG UP IN PLUNGER. LENS NOT IMPLANTED." CAUSE OF THE EVENT IS REPORTED AS "LENS GOT HUNG UP IN THE PLUNGER." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK, UNK, UNK, UNK, UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# 431171.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755802&pc=QCB | 2023826-2020-02464 | 2020-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/4.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755805&pc=QCB | 2023826-2020-02472 | 2020-10-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WAS TO IMPLANT A 13.2MM VTICM5_13.2, -12.00/4.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS)ON (B)(6) 2020. THE LENS TORE DURING INJECTION AND PART OF THE LENS STUCK IN THE CARTRIDGE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. CAUSE OF THE EVENT IS REPORTED AS DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755806&pc=QCB | 2023826-2020-02471 | 2020-09-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 12.6MM VTICMO12.6, -11.00/2.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED DURING INITIAL SURGERY ON (B)(6) 2020. AN EXCHANGE LENS WAS IMPLANTED ON (B)(6) 2020. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11093447&pc=QCB | 2023826-2020-03269 | 2020-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.0/92 (SPHERE/CYLINDER/AXIS),IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11093450&pc=QCB | 2023826-2020-03270 | 2020-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.5/088 (SPHERE/CYLINDER/AXIS),IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # 718226
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11093453&pc=QCB | 2023826-2020-03274 | 2020-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -9.5/1.0/092 (SPHERE/CYLINDER/AXIS),IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098616&pc=QCB | 2023826-2020-03252 | 2019-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/101 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098619&pc=QCB | 2023826-2020-03253 | 2020-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "LENS SIZE BASED ON PARAMETER. 11.85 CALLIPER (IOL MASTER 12.3) IS 13.2 BUT SURPRISINGLY THE VAULT WAS 2.5 TIMES CORNEAL THICKNESS WITH POST-OP CHAMBER LESSER THAN 1MM." Manufacturer Narrative: B1-UPDATED. B5-UPDATED INFORMATION: ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM IS RESOLVED. H6-HEALTH IMPACT CODE UPDATED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098620&pc=QCB | 2023826-2020-03243 | 2020-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE REMOVING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/080 (SPHERE/CYLINDER/AXIS) FROM THE VIAL, IT WAS DISCOVERED TO BE TORN. THIS OCCURRED ON (B)(6) 2020. ON THE SAME DATE AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098624&pc=QCB | 2023826-2020-03247 | 2020-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098627&pc=QCB | 2023826-2020-03249 | 2020-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098628&pc=QCB | 2023826-2020-03240 | 2020-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098629&pc=QCB | 2023826-2020-03241 | 2020-11-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), IT WAS TORN. THIS OCCURRED ON (B)(6) 2020 AND THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. ON (B)(6) 2020 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "THIS OCCURED ON (B)(6) 2020". SHOULD BE "THIS OCCURED ON (B)(6) 2020" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098634&pc=QCB | 2023826-2020-03248 | 2020-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098636&pc=QCB | 2023826-2020-03239 | 2020-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098637&pc=QCB | 2023826-2020-03246 | 2020-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11093446&pc=QCB | 2023826-2020-03271 | 2020-08-07 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -13.0/0.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098626&pc=QCB | 2023826-2020-03237 | 2020-12-02 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-30 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A PIECE OF A TICL IMPLANTABLE COLLAMER LENS TORE DURING LOADING DUE TO EXTREME RESISTANCE WITHIN THE CARTRIDGE. A BACKUP LENS WAS THEN CHOSEN; NO PATIENT INVOLVEMENT. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087586&pc=QCB | 2023826-2020-03162 | 2020-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -14.00/4.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087587&pc=QCB | 2023826-2020-03161 | 2020-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087588&pc=QCB | 2023826-2020-03160 | 2020-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -7.00/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087589&pc=QCB | 2023826-2020-03163 | 2020-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/2.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE LENS WAS REMOVED AND REPLACED ON 26/NOV/2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# 718286.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087590&pc=QCB | 2023826-2020-03165 | 2020-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/3.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087592&pc=QCB | 2023826-2020-03164 | 2020-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REMOVED AND REPLACED ON 28/NOV/2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT T REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: UNK, UNK ,UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# 718287.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087609&pc=QCB | 2023826-2020-03168 | 2020-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/4.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087611&pc=QCB | 2023826-2020-03170 | 2020-09-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.00/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087614&pc=QCB | 2023826-2020-03171 | 2020-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.50/1.0/107 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087616&pc=QCB | 2023826-2020-03166 | 2020-10-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/1.5/134 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS STUCK IN CARTRIDGE/INJECTION SYSTEM DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. A SAME LENGTH REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -15.00/1.5/134 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE LENS STUCK IN CARTRIDGE/INJECTION SYSTEM DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. LATER THE LENS WAS REPLACED WITH A SAME LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087617&pc=QCB | 2023826-2020-03167 | 2020-10-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS STUCK IN CARTRIDGE/INJECTION SYSTEM DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. A SAME LENGTH REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087618&pc=QCB | 2023826-2020-03174 | 2020-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/4.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087620&pc=QCB | 2023826-2020-03172 | 2020-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091342&pc=QCB | 2023826-2020-03234 | 2020-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.50/+1.0/008 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE, RESIDUE, AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091425&pc=QCB | 2023826-2020-03235 | 2020-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091451&pc=QCB | 2023826-2020-03225 | 2020-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091525&pc=QCB | 2023826-2020-03231 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING LOW VAULTING AND LENS ROTATION. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/122 (SPHERE/CYLINDER/AXIS) ON (B)(6) 2016. CAUSE OF THE EVENT WAS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. E1: ESTABLISHMENT NAME - (B)(6). CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/122 (SPHERE/CYLINDER/AXIS) ON (B)(6) 2016. ON (B)(6) 2021; THE LENS WAS EXPLANTED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF EVENT IS UNKNOWN. #: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/122 (SPHERE/CYLINDER/AXIS) ON (B)(6) 2016. ON (B)(6) 2021; THE LENS WAS EXPLANTED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING; LENS ROTATION; AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - THE LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091550&pc=QCB | 2023826-2020-03224 | 2020-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091989&pc=QCB | 2023826-2020-03215 | 2020-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY AND DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087581&pc=QCB | 2023826-2020-03099 | 2020-12-04 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -5.00/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON 04/DEC/2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "IOP IS DOWN, STILL CORNEAL EDEMA AND PUPIL REACTION IS NOT AS GOOD AS PRE-OP." PER REPORTER ON (B)(6) 2020, CORNEAL OEDEMA GETTING BETTER AND BETTER, NEARLY AS GOOD AS PRE-OP. PATIENT IS OKAY AND FINE. SURGEON ANTS TO WAIT AND WILL DECIDE TO IMPLANT ANOTHER SHORTER LENS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087583&pc=QCB | 2023826-2020-03096 | 2020-12-01 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.50/4.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: ADVERSE EVENT: OUTCOME ATTRIBUTED TO ADVERTISEMENT: REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE . DESCRIBE EVENT OR PROBLEM: (B)(6) THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.50/4.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON AND UNKNOWN DATE. IT IS REPORTED THAT THE LENS ROTATION RESOLVED THE PROBLEM. LENS REMAINS IMPLANTED. TYPE OF REPORTABLE EVENT: SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087615&pc=QCB | 2023826-2020-03169 | 2020-11-02 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/2.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091361&pc=QCB | 2023826-2020-03229 | 2020-12-03 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/173 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091952&pc=QCB | 2023826-2020-03230 | 2020-12-02 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS HAPPY. FTHE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091976&pc=QCB | 2023826-2020-03232 | 2020-09-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+5.5/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073977&pc=QCB | 2023826-2020-03200 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.5/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073979&pc=QCB | 2023826-2020-03198 | 2020-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.00/3.0/72 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073980&pc=QCB | 2023826-2020-03202 | 2020-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/2.0/67 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B3: CORRECTED TO (B)(6) 2020 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073982&pc=QCB | 2023826-2020-03192 | 2020-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, BENT AND DEFORMED WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073987&pc=QCB | 2023826-2020-03199 | 2020-09-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073992&pc=QCB | 2023826-2020-03193 | 2020-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/2.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073299&pc=QCB | 2023826-2020-03185 | 2020-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.0/2/74 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073300&pc=QCB | 2023826-2020-03187 | 2020-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -8.5/2.0/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073303&pc=QCB | 2023826-2020-03188 | 2020-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -13.5/2.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065556&pc=QCB | 2023826-2020-03092 | 2020-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON 03/SEP/2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: UNK, UNK , UNK . THIS PRODUCT IS NOT MARKETED IN THE US DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT CLAIM# 718202.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065567&pc=QCB | 2023826-2020-03093 | 2020-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN AND REMOVED A 13.2MM VTICMO13.2, -12.00/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. A SAME LENGTH LENS WAS IMPLANTED ON (B)(6) 2020. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065596&pc=QCB | 2023826-2020-03105 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.50/3.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED WITH A LONGER LENGTH SPHERICAL LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN, WITH RESIDUE ON LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065607&pc=QCB | 2023826-2020-03101 | 2020-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.00/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065608&pc=QCB | 2023826-2020-03108 | 2020-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.00/1.0/076 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.00/1.0/076 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY" H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065613&pc=QCB | 2023826-2020-03107 | 2020-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.50/1.5/095 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -10.50/1.5/095 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTION OF INFORMATION: D6A - REPORTED AS (B)(6) 2021, CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM AND "PATIENT IS HAPPY." CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065561&pc=QCB | 2023826-2020-03087 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -16.00/3.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. PER THE REPORTER, LENS WILL BE REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11057963&pc=QCB | 2023826-2020-03139 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11058106&pc=QCB | 2023826-2020-03126 | 2020-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11058199&pc=QCB | 2023826-2020-03127 | 2020-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50/+4.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN RESIDUE/DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11058201&pc=QCB | 2023826-2020-03132 | 2020-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11058282&pc=QCB | 2023826-2020-03133 | 2020-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER (AXIS) LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4) Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11058443&pc=QCB | 2023826-2020-03121 | 2020-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11058446&pc=QCB | 2023826-2020-03118 | 2020-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00/+0.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11058462&pc=QCB | 2023826-2020-03124 | 2020-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060093&pc=QCB | 2023826-2020-03148 | 2020-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060107&pc=QCB | 2023826-2020-03147 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060131&pc=QCB | 2023826-2020-03141 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060136&pc=QCB | 2023826-2020-03152 | 2020-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON 1(B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060141&pc=QCB | 2023826-2020-03140 | 2020-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060157&pc=QCB | 2023826-2020-03145 | 2020-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem; Unintended Movement | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -16.0/+2.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS THE LENS WAS IMPLANTED UPSIDE DOWN. REPORTEDLY, THE LENS WAS MANIPULTATED INTO THE CORRECT POSITION AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6: IMPLACT CODE 4628 IS CORRECTED TO 4625 FOR "MANIPULATION OF LENS IN THE RIGHT POSITION." DEVICE PROBLEM CODE 2993 CORRECTED TO 3026 FOR "LENS UPSIDE DOWN." CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060182&pc=QCB | 2023826-2020-03150 | 2020-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT / DEFORMED WITH DEBRIS AND RESIDUE ON PRODUCT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060193&pc=QCB | 2023826-2020-03143 | 2020-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060196&pc=QCB | 2023826-2020-03149 | 2020-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/103 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060200&pc=QCB | 2023826-2020-03155 | 2020-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060204&pc=QCB | 2023826-2020-03144 | 2020-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/017 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060207&pc=QCB | 2023826-2020-03153 | 2020-11-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE VAULT PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: THE LENS IMPLANT WAS FOUND TO BE TOO LARGE IN THE PATIENT'S EYE AFTER IMPLANT. RIGHT INFERIOR CORNEAL EDEMA IS REPORTED AFTER EXCHANGE; REFERENCE MFR FILE# 718182 FOR REPLACEMENT LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A SPECIMEN CU; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: DATE RECD BY MFR: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060290&pc=QCB | 2023826-2020-03146 | 2020-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CORRECTED DATA: B5 - "THE LENS WAS EXCHANGED" SHOULD BE CORRECTED TO "THE LENS WAS EXPLANTED AND LATER THE SAME DATE REPLACED". CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060295&pc=QCB | 2023826-2020-03142 | 2020-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11025436&pc=QCB | 2023826-2020-03075 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1,-5.00/1.0/4 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND LATER REPLACED DUE TO EXCESSIVE VAULT. THE REPORTER STATED " THE 12.1 CRYSTAL OF THE PATIENTS RIGHT EYE WAS IMPLANTED FOR THE FIRST TIME. IT WAS TOO LARGE, WITH HIGH ARCH, SHALLOW ANTERIOR CHAMBER AND HIGH GLAUCOMA RISK. THEREFORE, A 12.1 CRYSTAL WITH VERTICAL IMPLANTATION WAS REPLACED, AND THE ARCH HEIGHT AND ANTERIOR CHAMBER WERE IN GOOD CONDITION" IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11025527&pc=QCB | 2023826-2020-03076 | 2019-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1,-9.0/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS LATER REMOVED AND REPLACED DUE TO EXCESSIVE VAULT. THE REPORTER STATED " THE FIRST ONE IS PLACED HORIZONTALLY AT 12.1. AND THE SECOND IS PLACED VERTICALLY, WHICH CAN LOWER THE ARCH". IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11032717&pc=QCB | 2023826-2020-03113 | 2020-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+3.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR ANOTHER SAME LENGTH BUT DIFFERENT MODEL AND DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11032921&pc=QCB | 2023826-2020-03111 | 2020-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11033206&pc=QCB | 2023826-2020-03074 | 2020-10-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-11.5/+6.0/148 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11033304&pc=QCB | 2023826-2020-03069 | 2020-07-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-16.0/2.0/92 (SPHERE/CYLINDER/SXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY AND WITH RESIDUE ON THE PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, DEFORMED, AND STRETCHED OUT WITH FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11023681&pc=QCB | 2023826-2020-03114 | 2020-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.00/+4.5/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024118&pc=QCB | 2023826-2020-03044 | 2020-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024119&pc=QCB | 2023826-2020-03054 | 2020-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024121&pc=QCB | 2023826-2020-03035 | 2020-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024124&pc=QCB | 2023826-2020-03052 | 2020-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEFORMED WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024125&pc=QCB | 2023826-2020-03037 | 2020-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/065 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024127&pc=QCB | 2023826-2020-03043 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024130&pc=QCB | 2023826-2020-03042 | 2020-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -6.0/+2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024132&pc=QCB | 2023826-2020-03045 | 2020-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.0/+2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024134&pc=QCB | 2023826-2020-03036 | 2020-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024135&pc=QCB | 2023826-2020-03034 | 2020-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024139&pc=QCB | 2023826-2020-03053 | 2020-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/126 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024146&pc=QCB | 2023826-2020-03033 | 2020-10-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/071 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE LEFT (OS) EYE. REPORTEDLY, THE LENS GOT STUCK IN THE INJECTOR AND WAS TORN. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2020. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010303&pc=QCB | 2023826-2020-03061 | 2020-07-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.1,-12.0/3.0/90 (SPHERE/CYLINDER/SXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5 VTICMO13.1 SHOULD BE CORRECTED TO VTICMO13.2. CLAIM (B)(4). Manufacturer Narrative: CORRECTED DATA/ADDITIONAL INFORMATION: B5 IT WAS LATER CLARIFIED THAT THE REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010308&pc=QCB | 2023826-2020-03077 | 2020-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6,-7.00/5.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010318&pc=QCB | 2023826-2020-03073 | 2020-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2,-6.00/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS LATER REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. "SURPRISE OF THE EXCESSIVE VAULT" WAS REPORTED AS CAUSE OF EVENT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(6)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010333&pc=QCB | 2023826-2020-02981 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.5/1.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010348&pc=QCB | 2023826-2020-03065 | 2020-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-10.0/1.0/69 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010426&pc=QCB | 2023826-2020-03064 | 2020-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-7.50/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010431&pc=QCB | 2023826-2020-03060 | 2020-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-11.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010452&pc=QCB | 2023826-2020-03062 | 2020-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-17.50/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012198&pc=QCB | 2023826-2020-03020 | 2020-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.50/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012229&pc=QCB | 2023826-2020-03011 | 2018-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14/2.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2020 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012230&pc=QCB | 2023826-2020-03028 | 2020-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012261&pc=QCB | 2023826-2020-03021 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.50/0.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INPSECTION FOUND A OPTIC TORN, HAPTIC TORN, AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012267&pc=QCB | 2023826-2020-03027 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012292&pc=QCB | 2023826-2020-03025 | 2020-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/3.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND WITH DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012312&pc=QCB | 2023826-2020-03009 | 2020-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.00/1.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "EVERYTHING IS GOOD BUT NOT COMFORTABLE TO KEEP THE LENS WHILE VAULT IS LOW." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: ADDITIONAL INFORMATION: B1 - ADVERSE EVENT B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.00/1.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. REPORTEDLY, "EVERYTHING IS GOOD AND PATIENT HAPPY AFTER EXCHANGING THE LENS." D6B - (B)(6)2020 H6 - HEALTH EFFECT REPORTED AS 2199 - UPDATE TO 4629 CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012317&pc=QCB | 2023826-2020-03014 | 2020-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.50/3.8/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11013628&pc=QCB | 2023826-2020-02997 | 2020-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HATIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014052&pc=QCB | 2023826-2020-03006 | 2020-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014315&pc=QCB | 2023826-2020-02986 | 2019-05-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/090 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014431&pc=QCB | 2023826-2020-03000 | 2020-08-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/089 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH A SAME MODEL ,BUT DIFFERENT LENGTH (12.1MM) LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS LOST, AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: ADDITIONAL INFORMATION H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. ,BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014441&pc=QCB | 2023826-2020-02990 | 2019-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014535&pc=QCB | 2023826-2020-02994 | 2020-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.26MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+4.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE PATIENT HAD LOOSE SUSPENSORY LIGAMENTS AND THE LENS DISLOCATED/SUBLUXED. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. B5 - 13.26MM SHOULD BE CORRECTED TO 13.2MM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010432&pc=QCB | 2023826-2020-03063 | 2020-09-25 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-17.00/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012290&pc=QCB | 2023826-2020-03010 | 2018-01-07 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.50/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO LOW VAULT. A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11013748&pc=QCB | 2023826-2020-02999 | 2020-04-03 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014348&pc=QCB | 2023826-2020-02987 | 2020-05-14 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/077 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014649&pc=QCB | 2023826-2020-03003 | 2020-11-14 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.50/+6.0/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND THE EYE HAS IMPROVED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014693&pc=QCB | 2023826-2020-02991 | 2020-05-26 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +4.00/+6.0/166 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE LENS REMAINED IMPLANTED AND THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10989501&pc=QCB | 2023826-2020-02973 | 2020-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.5/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10989502&pc=QCB | 2023826-2020-02965 | 2020-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.0/2.0/57 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10989503&pc=QCB | 2023826-2020-02969 | 2020-08-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.00/1.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS WITHIN THE SAME SURGERY DUE TO LOW VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10989504&pc=QCB | 2023826-2020-02959 | 2020-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/1.0/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10989499&pc=QCB | 2023826-2020-02975 | 2020-10-04 | Injury | STAAR SURGICAL COMPANY | 2020-12-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.50/1.5/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10990222&pc=QCB | 2023826-2020-02974 | 2020-08-31 | Injury | STAAR SURGICAL COMPANY | 2020-12-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -16.0/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE REPOSITIONING DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A NON TORIC LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984714&pc=QCB | 2023826-2020-02953 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE DOCTOR HAS DECIDED HE WANTS TO REMOVE AND EXCHANGE PATIENTS ICL(IMPLANTABLE COLLAMER LERNS) WITH A SMALLER SIZE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. ON AN UNKNOWN DATE THE LENS REMOVED AND A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. CLAIM# (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984746&pc=QCB | 2023826-2020-02852 | 2020-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/4.5/171 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984751&pc=QCB | 2023826-2020-02865 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984771&pc=QCB | 2023826-2020-02861 | 2020-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -13.00/3.0/006 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, BENT, AND DEFORMED WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984774&pc=QCB | 2023826-2020-02864 | 2020-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -7.00/1.5/091 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. DATE REC¿D BY MFR - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984783&pc=QCB | 2023826-2020-02854 | 2019-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.00/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984788&pc=QCB | 2023826-2020-02860 | 2020-11-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICMO12.6, -15.50/2.5/010 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2020. THE LENS TORE DURING INJECTION INTO PATIENT'S EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. LENS WAS REMOVED AND EXCHANGED DURING INITIAL SURGERY ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984790&pc=QCB | 2023826-2020-02856 | 2019-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, BENT, AND DEFORMED WITH DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984791&pc=QCB | 2023826-2020-02853 | 2020-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984795&pc=QCB | 2023826-2020-02863 | 2020-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.00/2.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. SEX: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984797&pc=QCB | 2023826-2020-02855 | 2019-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.00/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984800&pc=QCB | 2023826-2020-02862 | 2020-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/5.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984815&pc=QCB | 2023826-2020-02876 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984835&pc=QCB | 2023826-2020-02868 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/6.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTHL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SPECIMEN CUP, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/6.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, AND REPORTER STATES THE PROBLEM WAS RESOLVED. HOWEVER, ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS REPORTED TO RESOLVE AGAIN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984841&pc=QCB | 2023826-2020-02869 | 2020-10-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020, A VTICM5 12.6, -12.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING INTO CARTRIDGE. NO PATIENT CONTACT. REPORTEDLY, THE LENS TORE WHILE PULLING IN THE CARTRIDGE. ON (B)(6) 2020, A SIMILAR LENS WAS IMPLANTED, AND THIS RESOLVED THE PROBLEM. REPORTEDLY, "PATIENT AND DOCTOR HAPPY." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE, RESIDUE, AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, RESIDUE, AND DEBRIS ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4); CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984845&pc=QCB | 2023826-2020-02874 | 2020-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984767&pc=QCB | 2023826-2020-02857 | 2020-11-12 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -5.50/2.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "EXCHANGE IS NOT PLANNED. PATIENT IS HAPPY. PROBLEM RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984851&pc=QCB | 2023826-2020-02877 | 2020-08-04 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/4.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980476&pc=QCB | 2023826-2020-02943 | 2020-08-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. THE LENS WAS EXPLANTED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE OPTIC AND HAPTIC WERE TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980477&pc=QCB | 2023826-2020-02944 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020 AS A REPLACEMENT LENS FOR MFR REPORT# 2023826-2020-02943. REPORTEDLY, THE REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM THOUGH THE STATUS OF THE EYE IS REPORTED AS "GOOD." Manufacturer Narrative: ON 10-DEC-2020 THE REPORTER STATED THAT THE PROBLEM WAS RESOLVED WITH THE REPLACEMENT LENS. PLEASE DISREGARD THE CURRENT CLAIM AS IT WAS REPORTED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980479&pc=QCB | 2023826-2020-02908 | 2020-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "02-OCT-2020" SHOULD BE CORRECTED TO "02-NOV-2020" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980480&pc=QCB | 2023826-2020-02929 | 2020-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980482&pc=QCB | 2023826-2020-02936 | 2020-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEVRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980483&pc=QCB | 2023826-2020-02933 | 2020-11-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.5/+1.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS LENS STUCK IN CARTRIDGE/INJECTION SYSTEM AND LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY IS REPORTED. REPORTEDLY, THE LENS ROTATED AND IS DECENTERED. THE CAUSE OF THE EVENT IS REPORTED AS "INJECTOR OR LENS VERY SOFT, POSTERIOR PART CUT." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980489&pc=QCB | 2023826-2020-02928 | 2020-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC ON THE LENS WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980491&pc=QCB | 2023826-2020-02931 | 2020-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3: "VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS" SHOULD HAVE BEEN INCLUDED WITH PREVIOUS DEVICE EVALUATION INFORMATION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980492&pc=QCB | 2023826-2020-02932 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.1/+2.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT CYLINDER AND AXIS LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980508&pc=QCB | 2023826-2020-02919 | 2020-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFO: ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM IS RESOLVED. H6-IMPACT CODE UPDATED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980509&pc=QCB | 2023826-2020-02924 | 2020-07-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENS INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980512&pc=QCB | 2023826-2020-02922 | 2020-07-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRA-OPERATIVELY ON (B)(6) 2020. THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980513&pc=QCB | 2023826-2020-02913 | 2020-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/108 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980515&pc=QCB | 2023826-2020-02915 | 2020-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE THE LENS. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980519&pc=QCB | 2023826-2020-02917 | 2020-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.0/+4.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980520&pc=QCB | 2023826-2020-02907 | 2020-11-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), IT TORE. IT WAS REMOVED INTRAOPERATIVELY AND THERE WAS NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2020. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980523&pc=QCB | 2023826-2020-02912 | 2020-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "SMALL DIAMETER OF ICL RELATED TO SULCUS-SULCUS DIAMETER." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980507&pc=QCB | 2023826-2020-02923 | 2020-06-09 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980510&pc=QCB | 2023826-2020-02921 | 2020-09-28 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10969932&pc=QCB | 2023826-2020-02902 | 2020-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED ON A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10969990&pc=QCB | 2023826-2020-02901 | 2020-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+0.5/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10969999&pc=QCB | 2023826-2020-02900 | 2020-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970003&pc=QCB | 2023826-2020-02894 | 2020-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970044&pc=QCB | 2023826-2020-02895 | 2020-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970163&pc=QCB | 2023826-2020-02899 | 2020-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970183&pc=QCB | 2023826-2020-02897 | 2020-05-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/088 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970198&pc=QCB | 2023826-2020-02890 | 2020-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970246&pc=QCB | 2023826-2020-02896 | 2020-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. D6B: (B)(6) 2020. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970250&pc=QCB | 2023826-2020-02889 | 2020-10-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/091 (SPHERE/CYLINDER/AXIS), WITH FORCEPS AND THE LENS TORE. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970260&pc=QCB | 2023826-2020-02898 | 2020-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+4.5/116 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970210&pc=QCB | 2023826-2020-02904 | 2020-09-25 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10970213&pc=QCB | 2023826-2020-02903 | 2020-05-11 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10955162&pc=QCB | 2023826-2020-02883 | 2020-10-12 | Injury | STAAR SURGICAL COMPANY | 2020-12-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM,VTICM5 13.2, -5.50/+0.5/034 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE AND A REPLACEMENT LENS WAS LATER IMPLANTED. THERE WAS NO PATIENT INJURY. THE REPORTER STATED " THE DOCTOR THOUGHT THAT THE HAPTIC WOULD BE DAMAGED". Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10949817&pc=QCB | 2023826-2020-02849 | 2019-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/2.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC'D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937882&pc=QCB | 2023826-2020-02786 | 2020-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/4.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937944&pc=QCB | 2023826-2020-02783 | 2019-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019, FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937948&pc=QCB | 2023826-2020-02780 | 2020-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.50/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937962&pc=QCB | 2023826-2020-02771 | 2020-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/1.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION INTO THE PATIENT'S EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED SAME LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, "AFTER THE LENS WAS INJECTED INTO THE EYE, IT WAS NOTICED THAT THE LENS WAS BROKEN.THE SURGEON WANTED TO FOLLOW UP FOR A FEW DAYS TO SEE IF IT WOULD AFFECT THE LOCALIZATION OF THE LENS, SO HE DIDN'T REMOVE IT IN THE SAME SURGERY." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937965&pc=QCB | 2023826-2020-02781 | 2020-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4), Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937980&pc=QCB | 2023826-2020-02789 | 2020-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937986&pc=QCB | 2023826-2020-02784 | 2020-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10938011&pc=QCB | 2023826-2020-02790 | 2020-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON 22/OCT/2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10938013&pc=QCB | 2023826-2020-02795 | 2020-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50/2.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10938022&pc=QCB | 2023826-2020-02796 | 2020-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10938038&pc=QCB | 2023826-2020-02788 | 2020-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEFORMED, WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930490&pc=QCB | 2023826-2020-02832 | 2020-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND HE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930502&pc=QCB | 2023826-2020-02840 | 2020-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930524&pc=QCB | 2023826-2020-02838 | 2020-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, THE SURGEON REPORTING THAT THE LENS WAS TILTED TO THE INFERIOR TEMPORAL AND THE SUSPENSORY LIGAMENT WAS CONSIDERED TO BE SLACK OR FRACTURED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930525&pc=QCB | 2023826-2020-02835 | 2019-12-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/044 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. THE LENS WAS EXPLANTED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HATIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930541&pc=QCB | 2023826-2020-02831 | 2020-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.5/+2.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND HE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930904&pc=QCB | 2023826-2020-02810 | 2020-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930905&pc=QCB | 2023826-2020-02808 | 2020-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN VIAL IN LIQUID. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5- THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT POWER. THE EXCHANGE LENS RESOLVED THE PROBLEM. H1- SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930939&pc=QCB | 2023826-2020-02811 | 2020-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10931052&pc=QCB | 2023826-2020-02818 | 2020-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10931070&pc=QCB | 2023826-2020-02820 | 2020-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS TORE DURING FORCEPS LOADING. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10931151&pc=QCB | 2023826-2020-02822 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+6.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10931157&pc=QCB | 2023826-2020-02821 | 2020-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930844&pc=QCB | 2023826-2020-02807 | 2020-10-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/164 (SPHERE/CYLINDER/AXIS), AND THE LENS LOADING FORCEPS TORE THE LENS. THERE WAS NO PATIENT CONTACT AND NO INJURY. ANOTHER ICL LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR/UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930988&pc=QCB | 2023826-2020-02816 | 2019-12-10 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/128 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10931019&pc=QCB | 2023826-2020-02817 | 2020-01-13 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+0.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10915373&pc=QCB | 2023826-2020-02804 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/+5.00/76 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXCHANGED. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10915499&pc=QCB | 2023826-2020-02801 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -6.50/+3.0/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10911406&pc=QCB | 2023826-2020-02803 | 2020-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -13.00/+3.00/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10911407&pc=QCB | 2023826-2020-02802 | 2020-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -11.50/+2.00/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK.. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10909680&pc=QCB | 2023826-2020-02770 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/180 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) BUT THE LENS WAS TORN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10909952&pc=QCB | 2023826-2020-02769 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +08.50/+2.5/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING ROTATED AND WAS RE-POSITIONED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10899963&pc=QCB | 2023826-2020-02745 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10899974&pc=QCB | 2023826-2020-02744 | 2020-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10899992&pc=QCB | 2023826-2020-02748 | 2020-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/089 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND STABLE. THE SURGEON AND PATIENT ARE HAPPY. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10900087&pc=QCB | 2023826-2020-02753 | 2020-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND WITH RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10900105&pc=QCB | 2023826-2020-02752 | 2020-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10900146&pc=QCB | 2023826-2020-02754 | 2020-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905569&pc=QCB | 2023826-2020-02732 | 2019-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: HE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US .(B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905573&pc=QCB | 2023826-2020-02737 | 2020-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -14.00/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -14.00/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORENAL ANGLES . THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905600&pc=QCB | 2023826-2020-02729 | 2020-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.00/6.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905636&pc=QCB | 2023826-2020-02728 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS, SIMILAR POWER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, "THE LENS SIZE WAS SMALL.". Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905661&pc=QCB | 2023826-2020-02727 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/2.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS, SIMILAR POWER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, "THE LENS SIZE WAS SMALL." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905665&pc=QCB | 2023826-2020-02733 | 2020-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: HE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM; VTICMO12.6, -9.5/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED, AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10898461&pc=QCB | 2023826-2020-02763 | 2020-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CORRECTED DATA: B5 - IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. SHOULD HAVE PREVIOUSLY BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10898466&pc=QCB | 2023826-2020-02762 | 2020-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. SHOULD HAVE BEEN INCLUDED IN PREVIOUS SUPPLEMENTAL. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS OPTIC DEFORMED, THE HAPTIC TORN, BROKEN,BENT AND DEFORMED. CORRECTED DATA: B5 - IT WAS REPORTED THE PROBLEM WAS RESOLVED. SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10898521&pc=QCB | 2023826-2020-02761 | 2020-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10898556&pc=QCB | 2023826-2020-02760 | 2020-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+4.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEFORMED WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10881025&pc=QCB | 2023826-2020-02756 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -10.00/1.0/131 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN A VIAL. VISUAL INSPECTION FOUND A TORN OPTIC AND HAPTIC AND RESIDUE AND DEBRIS ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10875928&pc=QCB | 2023826-2020-02717 | 2020-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.0/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 22020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THIS EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIOANAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY. VISUAL INSPECTION FOUND OPTIC SPLIT AND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO LENS RETURNED SPLIT. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10876509&pc=QCB | 2023826-2020-02755 | 2020-09-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00\+4.00\80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS NOTED TO BE DAMAGED AFTER OPENING THE VIAL. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, HAPTIC TORN, AND BENT WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10868412&pc=QCB | 2023826-2020-02714 | 2020-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.50/+3.50/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A NON TORIC LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 717756.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10865705&pc=QCB | 2023826-2020-02711 | 2020-08-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.5/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AND REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2020 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851784&pc=QCB | 2023826-2020-02666 | 2020-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.50/4.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WITH ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1 - REPORTED AS PRODUCT PROBLEM - UPDATED TO ADVERSE EVENT. B2 - UPDATE TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5 - REPORTED AS REMAINS IMPLANTED. LENS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2020. RESOLVED THE PROBLEM. D6B - UPDATED TO (B)(6) 2020. H6 - HEALTH EFFECT REPORTED AS 2199, UPDATED TO 4629 (EXCHANGE). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851789&pc=QCB | 2023826-2020-02667 | 2020-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/2.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT WITH FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852491&pc=QCB | 2023826-2020-02671 | 2020-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852495&pc=QCB | 2023826-2020-02672 | 2020-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852501&pc=QCB | 2023826-2020-02674 | 2020-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS, AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852525&pc=QCB | 2023826-2020-02678 | 2020-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING, WITH LENS ROTATION AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852601&pc=QCB | 2023826-2020-02685 | 2020-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+3.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THIS WAS THE REPLACEMENT LENS FOR MFR REPORT # 2023826-2020-02684, THE PROBLEM WAS RESOLVED AND THE EYE WAS STABLE WITH THE IMPLANT OF THE REPLACEMENT LENS, BUT THERE WAS STILL A HIGH VAULT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H10: PREVIOUS SUPPLEMENTAL WAS SUBMITTED IN ERROR AND DOES NOT APPLY TO THIS MFR#. DISREGARD SUPPLEMENTAL #1. CLAIM#: (B)(4). Manufacturer Narrative: H10 - CORRECTION TO SUPPLEMENTAL MDR #2 - "CLAIM#: (B)(4). IN H10 SECTION SHOULD BE CORRECTED TO "CLAIM#: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852619&pc=QCB | 2023826-2020-02675 | 2020-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS. THE LENS WAS IMPLANTED IN A VERTICAL POSITION AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852635&pc=QCB | 2023826-2020-02676 | 2020-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS. THE LENS WAS IMPLANTED IN A VERTICAL POSITION AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED, WITH RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856343&pc=QCB | 2023826-2020-02690 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS A DECREASE IN VAULT AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: B5 - "THE SURGEON REPORTS A DECREASE IN VAULT" SHOULD BE CORRECTED TO: IT WAS REPORTED THAT THE PATIENT HAD HIGH VAULT AND INDICATED, "I SHOULD EXCHANGE THE ICL DOWN ONE SIZE ALSO; THOUGH THE VAULTS FROM HONG KONG SUGGEST THE LENS MAY HAVE DROPPED BACK TO AN ACCEPTABLE LEVEL". CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - PREVIOUS SUPPLEMENTAL #1 SHOULD BE DISREGARDED AND INFORMATION SHOULD BE REVERTED BACK TO INITIAL MEDWATCH REPORT. SUPPLEMENTAL #1 INFORMATION IS N/A. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856404&pc=QCB | 2023826-2020-02691 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS A DECREASE IN VAULT AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITIH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: UNKNOWN. DATE OF EVENT-UNKNOWN. SERIAL #: UNKNOWN, IMPLANT DATE: UNKNOWN. DEVICE MANUFACTURE DATE: UNKNOWN. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856408&pc=QCB | 2023826-2020-02694 | 2020-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY STATED THE LENS WAS EXCHANGED WITH A SAME SIZE LENS. CORRECTED STATEMENT: LENS WAS EXCHANGED WITH A SHORTER LENS. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851817&pc=QCB | 2023826-2020-02668 | 2020-09-04 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.50/2.5/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON 04/SEP/2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856353&pc=QCB | 2023826-2020-02689 | 2020-09-24 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/055 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED, THE SURGEON REPORTING LOW VAULT AND LENS ROTATION; THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXCHANGED WITH A LONGER LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE AND PATIENT-RELATED FACTOR, "ICL ROTATED PROB TOO SMALL, PLACED ICL WITH HIGHER VAULT." Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10843080&pc=QCB | 2023826-2020-02708 | 2020-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/1/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10843162&pc=QCB | 2023826-2020-02707 | 2020-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.50/2.0/154 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10843233&pc=QCB | 2023826-2020-02706 | 2020-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITIH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10843403&pc=QCB | 2023826-2020-02705 | 2020-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -7.5/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10843536&pc=QCB | 2023826-2020-02709 | 2020-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.0/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10845097&pc=QCB | 2023826-2020-02712 | 2019-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10846944&pc=QCB | 2023826-2020-02713 | 2020-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.00/+3.00/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10843021&pc=QCB | 2023826-2020-02703 | 2020-08-19 | Injury | STAAR SURGICAL COMPANY | 2020-11-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.5/3.5/78 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10822737&pc=QCB | 2023826-2020-02654 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/6.00/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10822837&pc=QCB | 2023826-2020-02655 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.50/3.00/036 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10822878&pc=QCB | 2023826-2020-02660 | 2020-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.00/0.5/040 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823658&pc=QCB | 2023826-2020-02589 | 2020-05-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/2.0/015 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED DURING THE INITIAL SURGERY ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: SEX, WEIGHT, RACE- UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823663&pc=QCB | 2023826-2020-02597 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: HE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -14.50/3.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823665&pc=QCB | 2023826-2020-02594 | 2020-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO EXCESSIVE VAULT.. A SHORTER LENGTH REPLACEMENT LENS WAS THEN IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823666&pc=QCB | 2023826-2020-02593 | 2019-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823668&pc=QCB | 2023826-2020-02591 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/0.5/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823669&pc=QCB | 2023826-2020-02590 | 2019-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823672&pc=QCB | 2023826-2020-02596 | 2020-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823674&pc=QCB | 2023826-2020-02600 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH FOREIGN RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823678&pc=QCB | 2023826-2020-02601 | 2020-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823889&pc=QCB | 2023826-2020-02658 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.0/2.5/140 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: PREVIOUS SUPPLEMENTAL MDR IS NOT APPLICABLE TO THIS CLAIM. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED INSIDE A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE. MDR SUPPLEMENTAL #1 WAS WRONGFULLY SUBMITTED UNDER MFR# 2020-02685 CLAIM# (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10825051&pc=QCB | 2023826-2020-02659 | 2020-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -16.5/2.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823656&pc=QCB | 2023823-2020-02588 | 2019-12-13 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823661&pc=QCB | 2023823-2020-02587 | 2019-12-12 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823930&pc=QCB | 2023826-2020-02653 | 2020-05-15 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/2.0/82 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823964&pc=QCB | 2023826-2020-02663 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. REPORTEDLY, REFRACTIVE ERROR WAS OBSERVED IN PATIENT WITH A SMALL AMOUNT OF RESIDUAL ASTIGMATISM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10815914&pc=QCB | 2023826-2020-02564 | 2020-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10815920&pc=QCB | 2023826-2020-02566 | 2020-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE AND DEBRIS ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10815926&pc=QCB | 2023826-2020-02567 | 2020-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER (CYLINDER/AXIS) LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE AND DEBRIS ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10815931&pc=QCB | 2023826-2020-02568 | 2020-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS DEBRIS AND RESIDUE ON LENS. CLAIM # 717673 Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10815936&pc=QCB | 2023826-2020-02569 | 2020-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE AND DEBRIS ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10815955&pc=QCB | 2023826-2020-02570 | 2020-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.00/+2.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER (AXIS) LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE AND DEBRIS ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10815957&pc=QCB | 2023826-2020-02554 | 2020-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # 717616.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10815974&pc=QCB | 2023826-2020-02552 | 2020-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10816025&pc=QCB | 2023826-2020-02563 | 2020-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+1.0/024 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTICS TORN, WITH A PORTION OF THE LENS MISSING. THERE WAS RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10816028&pc=QCB | 2023826-2020-02558 | 2020-08-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812675&pc=QCB | 2023826-2020-02613 | 2020-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE PATIENT DID NOT WANT TO UNDERGO AN ADDITIONAL OPERATION. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812677&pc=QCB | 2023826-2020-02619 | 2019-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED; REFERENCE MFR FILE#: (B)(4) FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRSI ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812683&pc=QCB | 2023826-2020-02631 | 2019-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/077 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812700&pc=QCB | 2023826-2020-02626 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/119 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812709&pc=QCB | 2023826-2020-02614 | 2020-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/160 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812721&pc=QCB | 2023826-2020-02612 | 2020-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020, THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE PATIENT DID NOT WANT TO UNDERGO AN ADDITIONAL OPERATION. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812743&pc=QCB | 2023826-2020-02622 | 2020-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +2.0/+5.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812754&pc=QCB | 2023826-2020-02634 | 2020-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.0/+4.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALAUTION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THE HAPTIC IS SCRUNCHED TOGETHER. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812757&pc=QCB | 2023826-2020-02630 | 2020-04-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+3.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENS INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812758&pc=QCB | 2023826-2020-02629 | 2019-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/163 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812765&pc=QCB | 2023826-2020-02620 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/2.0/004 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B1-CORRECTED WITH UPDATED INFORMATION. B5-UPDATED INFORMATION: THE LENS WAS EXCHANGED WITH A LONGER LENS. H6-HEALTH IMPACT CODE UPDATED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812771&pc=QCB | 2023826-2020-02618 | 2019-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/014 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND FOREIGN RESIDUE ON THE OPTIC OF THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812788&pc=QCB | 2023826-2020-02610 | 2020-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+2.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812800&pc=QCB | 2023826-2020-02608 | 2020-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812889&pc=QCB | 2023826-2020-02621 | 2020-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812779&pc=QCB | 2023826-2020-02633 | 2020-07-11 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. WITHIN THE SAME SURGERY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10804684&pc=QCB | 2023826-2020-02551 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10792881&pc=QCB | 2023826-2020-02557 | 2020-10-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/059 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS HAD EXCESSIVE VAULTING. A SHORTER LENS WAS IMPLANTED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10796094&pc=QCB | 2023826-2020-02555 | 2020-09-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00/+1.50/101 (SPHERE/CYLINDER/AXIS), AND THE LENS WAS TORN. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS STUCK IN LIQUID IN CARTRIDGE INSIDE A SYRINGE. VISUAL INSPECTION FOUND THE LENS RETURNED IN CARTRIDGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786142&pc=QCB | 2023826-2020-02576 | 2015-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM, TICM12V4, -5.50/+2.50/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A HAPTIC AND THE OPTIC TORN. DEBRIS ON LENS SURFACE WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786224&pc=QCB | 2023826-2020-02577 | 2015-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A TICM115V4, -5.50/+4.00/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMGE TO LENS. DEBRIS ON LENS SURFACE WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786700&pc=QCB | 2023826-2020-02503 | 2020-03-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE, RESIDUE, AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786719&pc=QCB | 2023826-2020-02510 | 2020-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.00/3.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786722&pc=QCB | 2023826-2020-02501 | 2020-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2, -10.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2018, A SECOND SURGERY, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT, WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: UNK. UNK. UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786735&pc=QCB | 2023826-2020-02502 | 2020-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/2.5/121 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND STRETCHED OUT WITH DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786754&pc=QCB | 2023826-2020-02498 | 2020-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786761&pc=QCB | 2023826-2020-02506 | 2020-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786766&pc=QCB | 2023826-2020-02509 | 2020-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.00/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND LENS HAPTIC DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786774&pc=QCB | 2023826-2020-02511 | 2020-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A 13.2MM, VTICMO13.2, -16.50/5.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS EXCHANGED FOR A SAME LENGTH LENS ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CORRECTED DATA: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE . H1- SERIOUS . CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT ON (B)(6) 2020 THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE. THE LENS WAS LATER EXPLANTED ON AN UNKNOWN DATE FOR AN UNKNOWN REASON. A NEW LENS OF SAME LENGTH BUT DIFFERENT DIOPTER WAS IMPLANTED AND THE PROBLEM WAS REPORTEDLY RESOLVED. (B)(6) 2020. H6- HEALTH IMPACT CODE CORRECTED FROM 2645 TO 4627. MEDICAL DEVICE PROBLEM CODE CORRECTED FROM 3189 TO 2993. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786778&pc=QCB | 2023826-2020-02517 | 2020-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -14.00/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786782&pc=QCB | 2023826-2020-02507 | 2020-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786725&pc=QCB | 2023826-2020-02500 | 2020-08-11 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2, -9.50/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10789131&pc=QCB | 2023826-2020-02605 | 2019-01-06 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS; -15.00/5.0/099 (SPHERE/CYLINDER/AXIS); INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THERE WAS NO PROBLEM WITH THE LENS BUT THE PATIENT HAS RESIDUAL REFRACTION ERRORS THAT EFFECT THE VISUAL ACUITY. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/SAME LENGHT LENS. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. THE ERROE OCCURED DURING MEASUREMENT OF THE REFRACTION ERROR BY THE OPTOMETRIST. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779832&pc=QCB | 2023826-2020-02549 | 2020-09-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+1.5/109 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE LEFT (OS) EYE. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2020. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. NO REPLACEMENT LENS HAS YET BEEN IMPLANTED AS THEY ARE WAITING FOR AN ALTERNATE LENS TO BE MADE AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779840&pc=QCB | 2023826-2020-02540 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/075 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. REPORTEDLY, THE LENS ROTATED. IT WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AS IT ROTATED AGAIN. THE LENS WAS THEN EXPLANTED. Manufacturer Narrative: ON 06-NOV-2020 CONFIRMATION WAS RECEIVED THAT THE CURRENT CLAIM IS A DUPLICATE OF MFR FILE# (B)(4). PLEASE DISREGARD THE CURRENT CLAIM. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779854&pc=QCB | 2023826-2020-02548 | 2020-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779859&pc=QCB | 2023826-2020-02545 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/074 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779870&pc=QCB | 2023826-2020-02543 | 2020-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/076 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT.THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# 717579.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779872&pc=QCB | 2023826-2020-02539 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, (-8.0/+2.0/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. REPORTEDLY, THE LENS ROTATED. IT WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AS IT ROTATED AGAIN. THE LENS WAS THEN EXPLANTED. Manufacturer Narrative: ON 06-NOV-2020 CONFIRMATION THAT THE CURRENT CLAIM IS A DUPLICATE OF MFR FILE# 431155 WAS CONFIRMED. PLEASE DISREGARD THE CURRENT REPORT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779889&pc=QCB | 2023826-2020-02538 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WILL POSSIBLY REMOVE AND REPLACE A TMICL LENS FROM THE PATIENT'S RIGHT (OD) EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: INFORMATION RECEIVED FOR THE CURRENT PATIENT ON (B)(6) 2020. CURRENT REPORT IS A DUPLICATE OF MFR FILE#: (B)(4). PLEASE DISREGARD CURRENT REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 - MFR# FILE (B)(4) SHOULD BE CORRECTED TO MFR# 2023826-2020-02223 PLEASE REFERENCE THAT REPORT FOR COMPLETE CLAIM DETAILS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11103588&pc=QCB | 2023826-2021-00025 | 2020-11-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INCORRECTLY IMPLANTED A 13.2 VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/1.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LEFT EYE (OS) LENS WAS INCORRECTLY IMPLANTED IN THE OD. THE LENS WAS REMOVED AND REPLACED WITH THE CORRECT OD LENS DURING THE INITIAL SURGERY. THE PROBLEM WAS RESOLVED AND THE "PATIENT IS HAPPY." THE CAUSE OF THIS EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116954&pc=QCB | 2023826-2021-00026 | 2020-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.00/1.5/048 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116955&pc=QCB | 2023826-2021-00029 | 2020-11-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICMO13.2, -11.50/1.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS TORE DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED DURING THE INITIAL SURGERY ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. PER THE REPORTER ON (B)(6) 2020, "DOCTOR DOES NOT PLANT TO IMPLANT AGAIN." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116957&pc=QCB | 2023826-2021-00032 | 2020-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.50/3.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116958&pc=QCB | 2023826-2021-00035 | 2020-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -6.50/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE RECEIVED BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116961&pc=QCB | 2023826-2021-00027 | 2020-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.00/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116963&pc=QCB | 2023826-2021-00028 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -10.50/6.0/110 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "ZONULAR ABNORMALITY". REPORTEDLY, "PARTIAL RESOLUTION OF THE PROBLEM, MAYBE MOTIVATED BY ABNORMALITY ON THE ZONULAE LIGAMENT SYSTEM. LOW VAULT, CLOSE TO 100 UM." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116967&pc=QCB | 2023826-2021-00033 | 2020-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON 10/NOV/2020 THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, DEBRES ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11116970&pc=QCB | 2023826-2021-00034 | 2020-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -6.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11117661&pc=QCB | 2023826-2021-00036 | 2020-11-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2020 A 13.2MM, VTICM5 13.2, -13.50/2.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. DAMAGE WAS NOTED AFTER OPENING VIAL. NO PATIENT CONTACT. ANOTHER SIMILAR LENS WAS IMPLANTED ON (B)(6) 2020 AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN AND WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11119918&pc=QCB | 2023826-2021-00008 | 2020-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11119940&pc=QCB | 2023826-2021-00023 | 2020-12-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON OPENED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.00/+3.0/090 (SPHERE/CYLINDER/AXIS), AND DETECTED A MARGINAL DEFECT UNDER THE MICROSCOPE. THE SURGEON DECIDED TO IMPLANT THE LENS AND WHEN EJECTING THE LENS FROM THE CARTRIDGE, A LENS CRACK WAS FOUND. THE LENS WAS INSERTED AND REMOVED DURING THE SAME SURGERY. ANOTHER LENS WAS IMPLANTED AND THE RESULT IS GOOD AND PATIENT IS SATISFIED. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON OPENED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.00/+3.0/093 (SPHERE/CYLINDER/AXIS), AND DETECTED A MARGINAL DEFECT UNDER THE MICROSCOPE. THE SURGEON DECIDED TO IMPLANT THE LENS AND WHEN EJECTING THE LENS FROM THE CARTRIDGE, A LENS CRACK WAS FOUND. THE LENS WAS INSERTED AND REMOVED DURING THE SAME SURGERY. ANOTHER LENS WAS IMPLANTED AND THE RESULT IS GOOD AND PATIENT IS SATISFIED. D4: SERIAL # (B)(6). CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11120003&pc=QCB | 2023826-2021-00020 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND A HAPTIC TORN AND DEFORMED. THERE WAS RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11120072&pc=QCB | 2023826-2021-00015 | 2020-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/065 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS (12.6MM) AND THE PROBLEM WAS RESOLVED. THIS WAS THE REPLACEMENT LENS FOR MFR. # 2023826-2021-00014. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11120074&pc=QCB | 2023826-2021-00019 | 2020-11-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11120087&pc=QCB | 2023826-2021-00018 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: H3 - LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11120088&pc=QCB | 2023826-2021-00014 | 2020-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Mechanical Jam; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/088 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS STUCK IN CARTRIDGE OR INJECTION SYSTEM. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS NOT RESOLVED. SEE MFR. # 2023826-2021-00015 FOR REPLACEMENT LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11120130&pc=QCB | 2023826-2021-00016 | 2020-04-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/090 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER ICL LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11120137&pc=QCB | 2023826-2021-00004 | 2020-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/ +1.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020, DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11120169&pc=QCB | 2023826-2021-00002 | 2020-06-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/083 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11120206&pc=QCB | 2023826-2021-00024 | 2020-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND WAS DISSATISFIED WITH POSTOPERATIVE VISUAL ACUITY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. VISUAL ACUITY IS STABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11122312&pc=QCB | 2023826-2021-00076 | 2020-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-15.00/2.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11122374&pc=QCB | 2023826-2021-00080 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-6.0/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11122375&pc=QCB | 2023826-2021-00081 | 2020-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-8.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFY EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129612&pc=QCB | 2023826-2021-00075 | 2020-12-09 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -12.0/1.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE WHILE LOADING. THERE WAS NO PATIENT CONTACT, NO PATIENT INJURY AND A BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN CARTRIDGE. VISUAL INSPECTION FOUND RESIDUE ON LENS AND CARTRIDGE WING DAMAGED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129005&pc=QCB | 2023826-2021-00088 | 2020-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/3.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129007&pc=QCB | 2023826-2021-00087 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-10.0/2.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129037&pc=QCB | 2023826-2021-00089 | 2020-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. IT WAS REPORTED THAT ON (B)(6) 2020 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LNEGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. D6B- DATE CORRECTED TO (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135884&pc=QCB | 2023826-2021-00052 | 2020-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/121 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135889&pc=QCB | 2023826-2021-00060 |  | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +9.5/+3.5/040 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135903&pc=QCB | 2023826-2021-00045 |  | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+0.5/022 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A LONGER LENGTH LENS. AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE LENS WAS IMPLANTED ON (B)(6) 2020 - SHOULD BE IN PREVIOUS MDR. D6A: (B)(6) 2020 SHOULD BE IN THE PREVIOUS MDR . CLAIM# (B)(4). Manufacturer Narrative: CORRECTION: B5 - "THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED." CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135912&pc=QCB | 2023826-2021-00048 | 2020-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/143 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135921&pc=QCB | 2023826-2021-00041 | 2020-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/057 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135956&pc=QCB | 2023826-2021-00051 | 2020-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: D9: RETURNED DATE: (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135960&pc=QCB | 2023826-2021-00058 | 2019-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.0/+4.5/009 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135964&pc=QCB | 2023826-2021-00037 | 2020-12-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Integrity Problem | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -13.0/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS REMOVED AND REPLACED WITH A BACKUP LENS, THE SURGEON STATING HE COULD NOT SEE THE TORIC MARKS ON THE LENS. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. H6-HEALTH IMPACT CODE CORRECTED H6-METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUIED, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS; TORIC MARKS ARE LIGHT. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11156378&pc=QCB | 2023826-2021-00063 | 2020-11-13 | Injury | STAAR SURGICAL COMPANY | 2021-01-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/3.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11156343&pc=QCB | 2023826-2021-00066 | 2020-12-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Unintended Movement | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN AND REMOVED A 13.2MM VTICMO13.2, -12.50/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE ON 01/DEC/2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. ON 04/DEC/2020 A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, PER THE REPORTER ON 24/DEC/2020, "DOCTOR MISTAKENLY TOUCHED LENS WITH UNSTERILIZED SURFACE." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11156347&pc=QCB | 2023826-2021-00067 | 2020-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/1.5/151 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11158550&pc=QCB | 2023826-2021-00093 | 2020-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020 DUE TO BEING TOO LONG. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND GRADE 2 ANGLE NASALLLY AND TEMPORALY SUPERIORLY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THE PATEINT HAD BLURRY VISION. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11158589&pc=QCB | 2023826-2021-00096 | 2020-12-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.00/+1.5/083 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11158592&pc=QCB | 2023826-2021-00097 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/116, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/116 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS NOT RESOLVED. THE PATIENT WAS NOT SATISFIED BECAUSE OF HALOS. SEE MFR. REPORT # 2023826-2021-00468 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. D6A: (B)(6) 2020. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11158730&pc=QCB | 2023826-2021-00102 | 2019-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/173 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159174&pc=QCB | 2023826-2021-00126 | 2020-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+5.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159186&pc=QCB | 2023826-2021-00116 | 2020-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S EYE ON (B)(6) 2020. REPORTEDLY, THE LENS KEPT ROTATING AND WAS EXPLANTED ON (B)(6) 2020. IT WAS NOT REPLACED DUE TO PATIENT ANATOMY. NO PATIENT INJURY REPORTED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: APPARENTLY THIS OCCURRED DUE TO THE PATIENT'S ANATOMY. DURING THE SURGERY THE PATIENT'S EYE WOULDN'T ADHERE TO THE DROPS IN ORDER TO PLACE THE ICL INTO THE PATIENT'S EYES. CLAIM# (B)(6). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159194&pc=QCB | 2023826-2021-00125 | 2020-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+5.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159198&pc=QCB | 2023826-2021-00118 | 2020-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/078 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISUAL DAMAGE WITH FOREIGN RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159252&pc=QCB | 2023826-2021-00120 | 2020-12-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.0/+2.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS THE LENS WAS IMPLANTED UPSIDE DOWN. INTRAOPERATIVELY THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT WITH DEBRIS AND RESIDUE ON THE LENS. CLAIM #(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159269&pc=QCB | 2023826-2021-00119 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/001 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'SEY. REPORTEDLY, THE LENS WAS EXPLANTED AND REPLACED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1-CORRECTED. B5-PLEASE DISREGARD PREVIOUS B5 STATEMENT. H6-HEALTH IMPACT CODE CORRECTED. MEDICAL DEVICE CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: B5. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/001 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY, THE LENS BECAME UNSTERILIZED WHILE FOLDING AND AN ALTERNATE LENS WAS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159271&pc=QCB | 2023826-2021-00115 | 2020-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S EYE ON (B)(6) 2020. REPORTEDLY, THE LENS KEPT ROTATING AND WAS EXPLANTED ON (B)(6) 2020. IT WAS NOT REPLACED DUE TO PATIENT ANATOMY. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: APPARENTLY THIS OCCURRED DUE TO THE PATIENT'S ANATOMY. DURING THE SURGERY THE PATIENT'S EYE WOULDN'T ADHERE TO THE DROPS IN ORDER TO PLACE THE ICL INTO THE PATIENT'S EYE. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY, THE ORIGINAL LENS WAS DISLOCATED. AFTER REPOSITIONING ON (B)(6) 2020 THE LENS WAS REPLACED WITH A LONGER LENS. SEE MFR FILE# (B)(4), FOR REPLACEMENT LENS. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11160963&pc=QCB | 2023826-2021-00135 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.00/30 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFO ON 27 JAN 2021. B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.00/30 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. ADDITIONAL INFO ON 24 FEB 2021. B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.00/30 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A LONGER LENGTH LENS WAS LATER IMPLANTED ON THE SAME DAY AS REMOVAL. IT HAS BEEN REPORTED THAT LOW VAULT IS STILL REMAINING, PROBLEM DID NOT RESOLVE. FINAL EXPLANT IS PLANNED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11161069&pc=QCB | 2023826-2021-00139 | 2020-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.0/1.0/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION NOT REQUIRED. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A LONGER LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11161081&pc=QCB | 2023826-2021-00134 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/1.00/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFO ON (B)(6) 2021. B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/1.00/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. ADDITIONAL INFO ON (B)(6) 2021. B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/1.00/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A LONGER LENGTH LENS WAS LATER IMPLANTED ON THE SAME DAY AS REMOVAL. IT HAS BEEN REPORTED THAT LOW VAULT IS STILL REMAINING, PROBLEM DID NOT RESOLVE. FINAL EXPLANT IS PLANNED CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11161137&pc=QCB | 2023826-2021-00140 | 2020-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/1.0/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION NOT REQUIRED. ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A LONGER LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11162949&pc=QCB | 2023826-2021-00068 | 2020-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.0/139 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SAME LENGTH, DIFFERENT CYLINDER AXIS LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "AS THE LENS WAS IMPLANTED AT VERTICAL DIRECTION, THE VAULT DECREASED. (VERTICAL STS IS GREATER THAN HORIZONTAL STS)" Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND WITH RESIDUE ON LENS. CLAIM #(B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H3 - DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE RECEIVED BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11182460&pc=QCB | 2023826-2021-00077 | 2020-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2,-10.5/1.0/74 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11189746&pc=QCB | 2023826-2021-00147 | 2020-11-25 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/104 (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11189698&pc=QCB | 2023826-2021-00157 | 2019-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11189707&pc=QCB | 2023826-2021-00154 | 2020-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THAT ONLY A SMALL PIECE OF THE LENS WAS RETAINED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11189730&pc=QCB | 2023826-2021-00151 | 2020-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION; H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11189766&pc=QCB | 2023826-2021-00143 | 2020-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REPORTED AS LENS DISLOCATION/SUBLUXATION, ONE HAPTIC WASN'T PLACED IN THE SULUS. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY. H6: ADVERSE EVENT PROBLEM - MEDICAL DEVICE PROBLEM CODE - CHANGE CODE FROM 1583 TO 2923. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3: DATE OF EVENT: ON (B)(6) 2020. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, DISLOCATION/SUBLUXATION, AND ONE HAPTIC WASN'T PLACED IN THE SULUS. THE LENS WAS REPOSITIONED (THE HAPTIC WAS PLACED IN THE SULCUS) AND THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY. CLAIM # 718425 Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11196995&pc=QCB | 2023826-2021-00182 | 2020-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -4.5/+2.5/061 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. REPORTEDLY, IT WILL POSSIBLY BE EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1-UPDATED B5-UPDATED INFORMATION: ON 13-JAN-2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "DESPITE ACCURATE WTW MEASUREMENTS AND APPROPRIATE ICL SIZE, EXCESSIVE VAULT WAS PRESENT WITH SLIT AC DEPTH BUT NORMAL IOP." THE STATUS OF THE EYE IS REPORTED AS MYOPIA AND ASTIGMATISM. H6-HEALTH IMPACT CODE UPDATED. DEVICE CODE UPDATED DEVICE CODE 1494 (PREVIOUS CORNEAL OR REFRACTIVE SURGERY) CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11197015&pc=QCB | 2023826-2021-00181 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. REPORTEDLY, IT WILL POSSIBLY BE EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1-UPDATED. B5-ADDITIONAL INFORMATION: REPORTEDLY, THE LENS WAS EXPLANTED ON (B)(6) 2021, DUE TO EXCESSIVE VAULT. THE STATUS OF THE EYE IS MYOPIA AND ASTIMATISM. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "DESPITE ACCURATE WTW MEASUREMENTS AND APPROPRIATE ICL SIZE, EXCESSIVE VAULT WAS PRESENT WITH SLIT AC DEPTH BUT NORMAL IOP." H6-HEALTH IMPACT CODE: UPDATED. DEVICE CODE: UPDATED. DEVICE CODE 1494: OFF-LABEL USE (HISTORY OF PREVIOUS CORNEAL OR REFRACTIVE SURGERY). CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11205156&pc=QCB | 2023826-2021-00186 | 2020-12-17 | Injury | STAAR SURGICAL COMPANY | 2021-01-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.0/1.00/017 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2020 THE LENS WAS EXPLANTED, THE REPORTER STATED " THEY EXPLANT THE LENS, SWITCH IT AND IMPLANT CORRECTLY." PROBLEM RESOLVED, PATIENT IS HAPPY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11200454&pc=QCB | 2023826-2021-00189 | 2020-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.00/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED WITH MOISTURE AND RESIDUE/ DEBRIS, INSIDE MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. RESIDUE AND DEBRIS ON LENS WAS OBSERVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206964&pc=QCB | 2023826-2021-00107 | 2020-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.50/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED IN A SECOND SURGERY ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTION: H1 IN FOLLOW-UP #1 SHOULD BE "SERIOUS INJURY" NOT "MALFUNCTION" CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231112&pc=QCB | 2023826-2021-00209 | 2020-12-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/109 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231146&pc=QCB | 2023826-2021-00207 | 2020-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231178&pc=QCB | 2023826-2021-00202 | 2020-11-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/013 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS AS IMPLANTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231227&pc=QCB | 2023826-2021-00194 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. THE PATIENT IS IN GOOD CONDITION. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231257&pc=QCB | 2023826-2021-00195 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/063 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231270&pc=QCB | 2023826-2021-00193 | 2020-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+4.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT MODEL (VICM5) LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: A THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231309&pc=QCB | 2023826-2021-00205 | 2020-11-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED FOR ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231338&pc=QCB | 2023826-2021-00204 | 2020-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON OPENED THE VIAL AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/078 (SPHERE/CYLINDER/AXIS), AND NOTED WHITE SEDIMENT ON THE SURFACE OF THE ICL LENS. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE SURGEON IMPLANTED ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238047&pc=QCB | 2023826-2021-00169 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -16.00/2.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS FOR PATIENT LEFT EYE (OS) AND RETURNED WITH THE HAPTIC TORN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238017&pc=QCB | 2023826-2021-00160 | 2020-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/2.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238023&pc=QCB | 2023826-2021-00162 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238025&pc=QCB | 2023826-2021-00165 | 2020-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.00/2.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238033&pc=QCB | 2023826-2021-00167 | 2020-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISURE, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238036&pc=QCB | 2023826-2021-00161 | 2020-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.00/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238069&pc=QCB | 2023826-2021-00163 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4) .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238096&pc=QCB | 2023826-2021-00170 | 2020-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.50/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND REPLACED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "LENS SIZE TOO BIG, CAUSING CROWDING OF ANGLE AND HIGH VAULT. NEED ONE SIZE SMALLER." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE, AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN STRETCHED OUT, WITH RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11239117&pc=QCB | 2023826-2021-00220 | 2020-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11239154&pc=QCB | 2023826-2021-00223 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11239172&pc=QCB | 2023826-2021-00216 | 2020-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11239266&pc=QCB | 2023826-2021-00221 | 2020-09-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/+1.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. INTRAOPERATIVELY THE LENS WAS REMOVED, THE SURGEON STATING THAT WHILE FLUSHING OUT THE VISCOELASTIC THE LENS ROTATED. SINCE IT WAS A TICL THE DOCTOR WAS CONCERNED THAT THE LENS WOULD LEAD TO LOW VAULT AND LENS ROTATION. THEREFORE ON (B)(6) 2020 A LONGER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11241958&pc=QCB | 2023826-2021-00248 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11264199&pc=QCB | 2023826-2021-00285 | 2020-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.5/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11264225&pc=QCB | 2023826-2021-00286 | 2020-09-24 | Injury | STAAR SURGICAL COMPANY | 2021-02-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11264411&pc=QCB | 2023826-2021-00289 | 2020-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/1.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11270224&pc=QCB | 2023826-2021-00274 | 2020-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+6.0/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5 - "LENS ROTATION ALSO FACTORED INTO WHY THE LENS WAS EXCHANGED FOR A LONGER LENS." SHOULD BE ADDED TO THE INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE 2923 ADDED TO INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11270278&pc=QCB | 2023826-2021-00268 | 2020-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2020. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11270349&pc=QCB | 2023826-2021-00279 | 2020-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/125 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273774&pc=QCB | 2023826-2021-00227 | 2020-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273777&pc=QCB | 2023826-2021-00234 | 2021-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.5/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273778&pc=QCB | 2023826-2021-00237 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -10.50/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. REPORTEDLY, "LENS TOO BIG FOR PATIENT'S EYE. NEEDED REPLACEMENT." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273780&pc=QCB | 2023826-2021-00240 | 2020-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/1.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6)2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273788&pc=QCB | 2023826-2021-00236 | 2021-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -10.50/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)6) 2021. THIS RESOLVED THE PROBLEM. REPORTEDLY, "LENS TOO BIG FOR PATIENT'S EYE. NEEDED REPLACEMENT." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(6) 2021. THE REPORTER INITIALLY STATED THAT THE PATIENT'S PUPIL WAS MID-DILATED BUT IMPROVING. THEY LATER STATED THAT THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273789&pc=QCB | 2023826-2021-00229 | 2020-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND WITH DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273790&pc=QCB | 2023826-2021-00238 | 2020-11-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273793&pc=QCB | 2023826-2021-00243 | 2020-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/1.5/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DFORMED AND WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273795&pc=QCB | 2023826-2021-00239 | 2020-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273797&pc=QCB | 2023826-2021-00244 | 2020-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.50/1.5/163 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE US DEVICE PROBLEM CODE: (B)(4). CLAIM# 718551
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11280258&pc=QCB | 2023826-2021-00254 | 2020-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11280302&pc=QCB | 2023826-2021-00203 | 2020-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11280327&pc=QCB | 2023826-2021-00255 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11280333&pc=QCB | 2023826-2021-00261 | 2021-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11280363&pc=QCB | 2023826-2021-00210 | 2020-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00/+5.5/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11282820&pc=QCB | 2023826-2021-00293 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.5/1.50/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND WITH DEBRIS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH DEBRIS AND RESIDUE ON THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11282841&pc=QCB | 2023826-2021-00294 | 2020-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -12.50/ 0.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH DEBRIS ON THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11282966&pc=QCB | 2023826-2021-00292 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -11.0/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND WITH DEBRIS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS AND RESIDUE ON THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11290094&pc=QCB | 2023826-2021-00299 | 2020-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.00/3.50/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A -9.50 DIOPTER SPHERICAL LENS OF THE SAME LENGTH DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11290110&pc=QCB | 2023826-2021-00296 | 2020-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -12.5/3.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11290160&pc=QCB | 2023826-2021-00300 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 12.6 LENS WAS GOING TO BE EXPLANTED DUE TO SIZING. IT WAS LATER REPORTED THAT THE PATIENT DID NOT REQUIRE EXPLANT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, 4.0/1.0/169, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AND LOW VAULT WAS OBSERVED. IT WAS LATER REPORTED THAT AT 2-DAY POST- OP IT WAS DETERMINED THAT EXPLANTATION WAS NOT NECESSARY. PROBLEM RESOLVED. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11290168&pc=QCB | 2023826-2021-00301 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 12.6 LENS WAS GOING TO BE EXPLANTED DUE TO SIZING. IT WAS LATER REPORTED THAT THE PATIENT DID NOT REQUIRE EXPLANT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, 4.5/1.5/005, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AND LOW VAULT WAS OBSERVED. IT WAS LATER REPORTED THAT AT 2-DAY POST- OP IT WAS DETERMINED THAT EXPLANTATION WAS NOT NECESSARY. PROBLEM RESOLVED. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11290259&pc=QCB | 2023826-2021-00307 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.50/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED HIGH VAULT, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11290280&pc=QCB | 2023826-2021-00302 | 2021-01-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -16.00/4.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS GOING TO BE IMPLANTED AS A REPLACEMENT LENS TO CORRECT EXCESSIVE VAULT, SINIFICANT RESUCTION OF ICA AND MILD A/C CELL BUT THE LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS REMOVED AND REPLACED WITH THE BACKUP LENS WITHIN THE SAME SURGERY. THE REPORTER STATED " PLUNGER RODE OVER THE TRAILING HAPTICS AND TORE THEM OFF AT INJECTION." FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER REPORTED " RECOVERING WELL". Manufacturer Narrative: B5-IT WAS LATER REPORTED THAT AFTER THE BACKUP LENS WAS IMPLANTED THE PATIENTS CURRENT CONDITION WAS GOOD WITH BETTER VA AND VAULT. IT WAS ALSO REPORTED THAT THE BACKUP LENS RESOLVED THE PROBLEM (INITIAL REPORTED ADVERSE EVENTS). H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299009&pc=QCB | 2023826-2021-00339 | 2020-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299010&pc=QCB | 2023826-2021-00340 | 2020-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299013&pc=QCB | 2023826-2021-00334 | 2021-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTEDLY IS CONSIDERING EXCHANGING THE LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON B)(6) 2020. ON B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "LOW VAULT RESULTED IN SIGNIFICANT ICL ROTATION. LENS EXCHANGED FOR A LARGER SIZE AND NO ROTATION OCCURRED." AFTER EXCHANGE THE PATIENT WILL REPORTEDLY UNDERGO PRK FOR RESIDUAL REFRACTIVE ERROR WHICH WAS AN EXPECTED OUTCOME. Manufacturer Narrative: B1-UPDATED. B5-UPDATED. PLEASE DISREGARD PREVIOUS B5 STATEMENT. H6-HEALTH IMPACT CODE UPDATED. DEVICE CODE UPDATED. (B)(4) Manufacturer Narrative: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR UNK OR BLANK ENTRY. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299045&pc=QCB | 2023826-2021-00341 | 2021-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299086&pc=QCB | 2023826-2021-00335 | 2021-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTEDLY IS CONSIDERING EXCHANGING THE LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED THEN REPLACED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM IS RESOLVED. THE CAUSE IS REPORTED AS DEVICE: "LOW VAULT RESULTED IN SIGNIFICANT ICL ROTATION. THE LENS WAS EXCHANGED FOR A LARGER SIZE AND NO ROTATION OCCURRED." THE PATIENT IS DOING WELL AFTER EXCHANGE WILL UNDERGO PRK FOR RESIDUAL REFRACTIVE ERROR WHICH WAS AN EXPECTED OUTCOME. Manufacturer Narrative: ESUBMITTER SOFTWARE DOES NOT ALLOW FOR UNK OR BLANK ENTRY. (B)(4). Manufacturer Narrative: B1-UPDATED. B5-UPDATED. PLEASE DISREGARD PREVIOUS B5 STATEMENT. H6-HEALTH IMPACT CODE UPDATED. MEDICAL DEVICE CODE UPDATED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299089&pc=QCB | 2023826-2021-00351 | 2021-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299134&pc=QCB | 2023826-2021-00342 | 2020-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - SHOULD BE CORRECTED TO "VISUAL INSPECTION FOUND THE LENS HAPTIC TORN." CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299136&pc=QCB | 2023826-2021-00337 | 2020-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/110 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299162&pc=QCB | 2023826-2021-00344 | 2020-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11304417&pc=QCB | 2023826-2021-00343 | 2020-09-27 | Injury | STAAR SURGICAL COMPANY | 2021-02-09 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/164 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11305337&pc=QCB | 2023826-2021-00368 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM512.1, -11.50/2.50//93, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11305374&pc=QCB | 2023826-2021-00369 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -11.50/3.0/076, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11305615&pc=QCB | 2023826-2021-00372 | 2020-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.5/4.0/085, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC STRETCHED OUT AND FOREIGN RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11305782&pc=QCB | 2023826-2021-00373 | 2020-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM512.6, -11.00/1.0/091, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308770&pc=QCB | 2023826-2021-00315 | 2020-12-10 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308892&pc=QCB | 2023826-2021-00329 | 2020-10-15 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER REPLACING THE LENS, THE ARCH HEIGHT IS WITHIN THE NORMAL RANGE." Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE - UNK. DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308902&pc=QCB | 2023826-2021-00321 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: AN ARTICLE PUBLISHED IN JOURNAL OF REFRACTIVE SURGERY WAS RECEIVED ON (B)(6) 2021, ENTITLED 'VISUAL AND REFRACTIVE OUTCOMES WITH THE EYECRYL PHAKIC TORIC IOL VERSUS THE VISIOAN TORIC IMPLANTABLE COLLAMER LNS: RESULTS OF A 2-YEAR PROSPECTIVE COMPARATIVE STUDY.' THE ARTICLE STATES THAT TICL GROUP: 1 EYE REQUIRED REALIGNMENT FOR SIGNIFICANT ROTATION AT ROTATED AGAIN AFTER 2 MONTHS. TICL WAS SLIGHTLY UNDERSIZED WITH A VAULT OF 180. EXCHANGE TO ON-SIZE LARGER WAS RECOMMENDED, HOWEVER PATIENT OPTED TO WEAR GLASSES FOR RESIDUAL RX. BOTH EYES OF ANTHER PATIENT REQUIRED LENS EXCHANGE FOR A SMALLER SIZE DUE TO HEGH VAULT AT 3 MONTHS POST-OP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. MODEL #: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE : UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308908&pc=QCB | 2023826-2021-00326 | 2021-01-08 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308933&pc=QCB | 2023826-2021-00330 | 2020-12-15 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/4.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11305873&pc=QCB | 2023826-2021-00376 | 2020-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/1.5/74, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11305981&pc=QCB | 2023826-2021-00374 | 2020-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, DEBRIS ON LENS AND A TORN HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11306034&pc=QCB | 2023826-2021-00370 | 2020-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.50/2.00/91, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME LENGTH DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND DEBRIS ON PRODUCT, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308733&pc=QCB | 2023826-2021-00311 | 2021-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE RECEIVED BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308780&pc=QCB | 2023826-2021-00314 | 2019-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/1.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/2.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308814&pc=QCB | 2023826-2021-00313 | 2020-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308849&pc=QCB | 2023826-2021-00319 | 2021-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2, -12.50/2.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS TO BE EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2, -12.50/2.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT IS REPORTED THAT ON THE SAME DAY THE LENS WAS EXPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(6) 2021. MALFUNCTION IS CORRECTED TO SERIOUS INJURY. IMPACT CODE 2199 IS CORRECTED TO 4627. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308882&pc=QCB | 2023826-2021-00333 | 2021-01-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -3.50/1.0/175 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. "A PLACK DOT WAS FOUND IN LENS OPTICAL AREA AND IT TURNED WHITE AFTER RINSE AND IMPLANTATION. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, "SURGEON DOES NOT PLAN TO REMOVE THE LENS AS THE PATIENT SHOWS NO FURTHER SIGNS OR SYMPTOMS." Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308904&pc=QCB | 2023826-2021-00323 | 2020-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308923&pc=QCB | 2023826-2021-00322 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308930&pc=QCB | 2023826-2021-00325 | 2021-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND "RESIDUE"/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND DEBRIS ON LENS AND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11322655&pc=QCB | 2023826-2021-00420 | 2020-12-02 | Injury | STAAR SURGICAL COMPANY | 2021-02-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/3.0/82, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11322262&pc=QCB | 2023826-2021-00417 | 2021-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.5/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11322585&pc=QCB | 2023826-2021-00419 | 2020-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/3.0/91, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334397&pc=QCB | 2023826-2021-00263 | 2020-12-20 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+3.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334454&pc=QCB | 2023826-2021-00365 | 2021-01-18 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/163, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334314&pc=QCB | 2023826-2021-00258 | 2020-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334464&pc=QCB | 2023826-2021-00360 | 2020-12-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -06.50/+3.0/079 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN IN THE PATIENTS LEFT EYE (OS). THE LENS WAS TORN UPON REMOVAL FROM THE EYE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP, PATIENT HAD SOME MILD CORNEAL SWELLING BUT IS HEALING WELL (SURGEON INDICATED WAS THE RESULT OF REMOVAL AND INSERTION OF NEW ICL). THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL, WITH DEBRIS AND RESIDUE ON LENS. VISUAL INSPECTION FOUND A HAPTIC TORN, WITH DEBRIS AND RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334465&pc=QCB | 2023826-2021-00367 | 2021-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/082, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE /DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11340671&pc=QCB | 2023826-2021-00396 | 2020-10-27 | Injury | STAAR SURGICAL COMPANY | 2021-02-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+2.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11340677&pc=QCB | 2023826-2021-00388 | 2019-05-27 | Injury | STAAR SURGICAL COMPANY | 2021-02-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11340684&pc=QCB | 2023826-2021-00387 | 2019-09-19 | Injury | STAAR SURGICAL COMPANY | 2021-02-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11340668&pc=QCB | 2023826-2021-00393 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: "THE CORRECT DATE FOR THE LENS EXCHANGE IS (B)(6) 2020." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11340670&pc=QCB | 2023826-2021-00391 | 2021-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS ON THE LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11340675&pc=QCB | 2023826-2021-00390 | 2021-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343288&pc=QCB | 2023826-2021-00428 | 2020-11-07 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.5/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. IN A SUBSEQUENT PROCEDURE THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE RECD BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343396&pc=QCB | 2023826-2021-00427 | 2020-09-19 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.0/4.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11346262&pc=QCB | 2023826-2021-00426 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.0/0.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AS A REPLACMENT LENS TO CORRECT LOW VAULT BUT THE PROBLEM DID NOT RESOLVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11342783&pc=QCB | 2023826-2021-00433 | 2020-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11342915&pc=QCB | 2023826-2021-00432 | 2020-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.5/1.0/83 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH SPHERICAL LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. FOR ADDITIONAL INFORMATION THE REPORTER STATED THAT "THE PATIENT'S POSTOPERATIVE SUPPLY IS TOO HIGH. CONSIDERING THE PATIENT'S OPERATION SCHEDULE AND TICL DELIVERY TIME, ICL CRYSTALS WERE IMPLANTED AFTER COMMUNICATION WITH THE PATIENT". CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343204&pc=QCB | 2023826-2021-00409 | 2020-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -14.00/1.5/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FOR PATIENT'S LEFT EYE (OS), TORE DURING INJECTION/DELIVERY ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED INTRAOPERATIVELY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED, THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/119 (SPHERE/CYLINDER AXIS) WAS FOUND DAMAGED, DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS NOT IMPLANTED, BUT TOUCHED THE PATIENT'S EYE. A REPLACEMENT LENS OF THE SAME MODEL, SIZE AND SIMILAR POWER WAS LATER IMPLANTED. AND THIS RESOLVED THE PROBLEM. H6: INVESTIGATION FINDINGS: CODE 114, SHOULD HAVE BEEN REPORTED IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343207&pc=QCB | 2023826-2021-00413 | 2020-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -13.00/3.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343211&pc=QCB | 2023826-2021-00416 | 2020-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AFTER OPENING VIAL OF A 13.2MM VTICM5_13.2, -7.00/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS FOUND TO HAVE A SCRATCH ON LENS. AFTER REVIEWING LENS UNDER MICROSCOPE, "SCRATCH WAS FOUND ON JUNCTION OF OPTIC/HAPIC SURFACE." LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. SCRATCHES WERE ALSO NOTED ON THE JUNTION OF THE HAPTIC AND OPTIC. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11346066&pc=QCB | 2023826-2021-00434 | 2020-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.0/0.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, INSIDE MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC'D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11348310&pc=QCB | 2023826-2021-00488 | 2021-01-18 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/2.0/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11348991&pc=QCB | 2023826-2021-00496 | 2020-12-18 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -6.00/2.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH SPHERICAL LENS DUE TO LOW VAULT AND ROTATION . THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11353489&pc=QCB | 2023826-2021-00495 | 2021-01-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. NO PATIENT INJURY WAS REPORTED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11348717&pc=QCB | 2023826-2021-00489 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/2.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366576&pc=QCB | 2023826-2021-00484 | 2020-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXISLENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE REPLACEMENT LENS WAS POSITIONED IN THE VERTICAL POSITION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368716&pc=QCB | 2023826-2021-00455 | 2020-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+4.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368731&pc=QCB | 2023826-2021-00461 | 2021-01-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -06.00/+0.5/025 (SPHERE/CYLINDER/AXIS) DURING THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. ANOTHER LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE REPLACEMENT LENS WAS IMPLANTED ON (B)(60 2021 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. THERE WAS DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368742&pc=QCB | 2023826-2021-00456 | 2020-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+4.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368750&pc=QCB | 2023826-2021-00470 | 2020-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+4.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368773&pc=QCB | 2023826-2021-00454 | 2020-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC HAS A PINCH. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368850&pc=QCB | 2023826-2021-00453 | 2020-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368882&pc=QCB | 2023826-2021-00462 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING ROTATED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 - "THE LENS WAS REPORTED AS HAVING ROTATED" SHOULD BE CORRECTED TO "THE LENS WAS REPORTED AS TOO SMALL." H6: MEDICAL DEVICE PROBLEM CODE: 1583. CLAIM #: (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE OF EVENT: UNKNOWN. EXPIRATION DATE: UNKNOWN. IMPLANT DATE: UNKNOWN. EXPLANT DATE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE : UNKNOWN. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372189&pc=QCB | 2023826-2021-00445 | 2020-08-22 | Injury | STAAR SURGICAL COMPANY | 2021-02-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.00/1.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372196&pc=QCB | 2023826-2021-00446 | 2020-08-22 | Injury | STAAR SURGICAL COMPANY | 2021-02-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.00/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372203&pc=QCB | 2023826-2021-00448 | 2020-12-30 | Injury | STAAR SURGICAL COMPANY | 2021-02-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON(B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372209&pc=QCB | 2023826-2021-00440 | 2021-01-15 | Injury | STAAR SURGICAL COMPANY | 2021-02-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 12.6MM VTICMO12.6, -10.50/0.5/054 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO BEING IMPLANED UPSIDE DOWN AND LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "DOCTOR PLANNED TO IMPLANT THE BACK-UP LENS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372187&pc=QCB | 2023826-2021-00437 | 2020-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/2.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "THIS PATIENT WILL NOT BE IMPLANTED AGAIN." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372197&pc=QCB | 2023826-2021-00451 | 2020-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -6.0/1.0/176 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372198&pc=QCB | 2023826-2021-00439 | 2021-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.00/3.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "EXCHANGE IS PLANNED IN 3-4 WEEKS." Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.00/3.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372211&pc=QCB | 2023826-2021-00452 | 2021-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "EXCHANGE IS PLANNED IN 3-4 WEEKS." Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372212&pc=QCB | 2023826-2021-00447 | 2021-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.00/2.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS "LOW VAULT, SO ROTATION OF THE LENS WITH TIME." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11375610&pc=QCB | 2023826-2021-00472 | 2020-12-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/081, (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND THERE WAS DEBRIS ON LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11375624&pc=QCB | 2023826-2021-00464 | 2020-12-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/093, (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11385938&pc=QCB | 2023826-2021-00541 | 2020-10-15 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.0/72 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11386100&pc=QCB | 2023826-2021-00542 | 2020-08-23 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.5/3.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11386623&pc=QCB | 2023826-2021-00540 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM, TMICL12.6, -10.50/3.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FLIPPED DURING INSERTION AND WAS DAMAGED WHEN IMMEDIATELY REMOVED. DOCTOR ENDED UP USING BACKUP LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE-RACE: UNK. DATE OF EVENT: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11383334&pc=QCB | 2023826-2021-00537 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.5/3.0/71 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.5/3.0/71 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH DIFFERENT AXIS LENS AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11383336&pc=QCB | 2023826-2021-00536 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.5/3.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.5/3.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH DIFFERENT AXIS LENS AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11386152&pc=QCB | 2023826-2021-00543 | 2020-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/2.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400357&pc=QCB | 2023826-2021-00522 | 2020-12-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-02 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2020. THE LENS WAS REMOVED INTRAOPERATIVELY. ON THE SAME DATE IN A SEPARATE SURGERY AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400366&pc=QCB | 2023826-2021-00518 | 2020-12-28 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400367&pc=QCB | 2023826-2021-00521 | 2020-11-26 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+0.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400361&pc=QCB | 2023826-2021-00520 | 2021-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), IT WAS TORN. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY, THE TORN HAPTIC OF THE LENS WAS FOUND DURING THE FOLLOW UP. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400364&pc=QCB | 2023826-2021-00519 | 2020-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE; DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400365&pc=QCB | 2023826-2021-00516 | 2020-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400372&pc=QCB | 2023826-2021-00528 | 2021-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, BEFORE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+0.5/064 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), IT WAS DISCOVERED THAT THERE WAS WHITE SEDIMENT ON THE LENS. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS NOT IMPLANTED AND THE CAUSE OF THE PROBLEM IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: G4: UNABLE TO DESELECT COMBINATION PRODUCT-PLEASE DISREGARD. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. NO PARTICULATE WAS NOTED ON THE LENS OR IN THE VIAL. H6: METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400373&pc=QCB | 2023826-2021-00530 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/111 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11401975&pc=QCB | 2023826-2021-00567 | 2021-01-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -10.0/2.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED AND REPLACED WITH THE STANBY LENS WITHIN THE SAME SURGERY. NO INJURY TO THE PATIENT WAS REPORTED AND CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: DATE OF EVENT: (B)(6) 2021. DATE OF REPORT: 29-JAN-2021. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405357&pc=QCB | 2023826-2021-00550 | 2020-10-30 | Injury | STAAR SURGICAL COMPANY | 2021-03-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/174 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405471&pc=QCB | 2023826-2021-00556 | 2020-11-14 | Injury | STAAR SURGICAL COMPANY | 2021-03-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405547&pc=QCB | 2023826-2021-00561 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-03-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD). AT ONE YEAR POST-OP, THE LENS WAS FOUND TO HAVE ROTATED. THE LENS REMAINED IMPLANTED. THE SURGEON DOES NOT PLAN TO ROTATE OR EXPLANT/EXCHANGE THE LENS. PATIENT IS ABLE TO ACHIEVE 20/20 TO 20/25 UNCORRECTED DISTANCE VA EACH EYE. WITH AN ASTIGMATIC CORRECTION IN EACH EYE; PATIENT ACHIEVES 20/20. PATIENT IS CONTENT WITH THE VISUAL OUTCOME BUT SEES BETTER WITH ADDITIONAL ASTIGMATIC CORRECTIVE LENSES. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405604&pc=QCB | 2023826-2021-00571 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-03-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/095 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS). AT ONE YEAR POST-OP, THE LENS WAS FOUND TO HAVE ROTATED. THE LENS REMAINED IMPLANTED. THE SURGEON DOES NOT PLAN TO ROTATE OR EXPLANT/EXCHANGE THE LENS. PATIENT IS ABLE TO ACHIEVE 20/20 TO 20/25 UNCORRECTED DISTANCE VA EACH EYE. WITH AN ASTIGMATIC CORRECTION IN EACH EYE; PATIENT ACHIEVES 20/20. PATIENT IS CONTENT WITH THE VISUAL OUTCOME BUT SEES BETTER WITH ADDITIONAL ASTIGMATIC CORRECTIVE LENSES. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE:UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405330&pc=QCB | 2023826-2021-00549 | 2021-01-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/099 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. A SUTURE WAS REQUIRED. THE PATIENT IS DOING WELL. THE REPORTER INDICATED THEY FELT THE LENS WAS LOADED TOO EARLY AND CAUSED THE LENS TO ROTATE IN THE CARTRIDGE. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WRAPPED IN GAUZE, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405389&pc=QCB | 2023826-2021-00551 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+5.5/093, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405397&pc=QCB | 2023826-2021-00552 | 2020-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/093, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2020 AND THE LENS WAS STABLE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THE LENS WAS REPOSITIONED BUT THE RESULTS WERE NOT GOOD. THE LENS MAY BE EXPLANTED/EXCHANGED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS ON ((B)(6) 2021, DUE TO REFRACTIVE SURPRISE. THSI RESOLVED THE PROBLEM." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405481&pc=QCB | 2023826-2021-00557 | 2020-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11412245&pc=QCB | 2023826-2021-00569 | 2021-01-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.00/2.00/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 LOW VAULT WITHOUT ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. THE REPORTER STATED " DOC HAS DECIDED NOT TO MAKE CHANGES TO OD AND MONITOR. VA AND EYE PRESSURE IS GOOD." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11415994&pc=QCB | 2023826-2021-00503 | 2021-01-03 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416011&pc=QCB | 2023826-2021-00510 | 2020-09-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/5.5/013 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS WAS REPOSITIONED ON (B)(6) 2020, THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS "LENS SIZE TOO SMALL." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416131&pc=QCB | 2023826-2021-00504 | 2021-01-03 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11418646&pc=QCB | 2023826-2021-00572 | 2020-10-07 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.50/1.5/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416112&pc=QCB | 2023826-2021-00512 | 2020-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -6.50/2.0/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. REFRACTIVE TREATMENT (LASIK, PRK.ETC) WAS PERFORMED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -6.50/2.0/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. REFRACTIVE TREATMENT ((B)(4)) WAS PERFORMED. THE LENS WAS REPLACED WITH A SAME LENGTH SIMILARE (SPHERE/CYLINDER/AXIS) LENS ON (B)(6) 2021. REPORTEDLY, "PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -6.50/2.0/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. REFRACTIVE TREATMENT ((B)(4)) WAS PERFORMED. THE LENS WAS REPLACED WITH A SAME LENGTH SIMILARE (SPHERE/CYLINDER/AXIS) LENS ON (B)(6) 2021. REPORTEDLY, "PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11418952&pc=QCB | 2023826-2021-00581 | 2020-07-10 | Injury | STAAR SURGICAL COMPANY | 2021-03-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -10.5/2.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11419358&pc=QCB | 2023826-2021-00582 | 2021-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.5/4.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITHOUT ROTATION WAS OBSERVED, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_ 12.6, -8.5/4.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGED RESOLVED THE PROBLEMS. CAUSE OF EVENT WAS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11419399&pc=QCB | 2023826-2021-00583 | 2021-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5 2.6, 9.5/4.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WITHOUT ROTATION WAS OBSERVED, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_ 12.6, -9.5/4.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(4) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC'D BY MFR THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11441198&pc=QCB | 2023826-2021-00617 | 2021-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11441202&pc=QCB | 2023826-2021-00614 | 2020-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY AND WITH DEBRIS ON THE PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND STRETCHED OUT, WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11448875&pc=QCB | 2023826-2021-00611 | 2021-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-09 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2021. AN ALTERNATE WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS THE "LENS GOT STUCK IN THE PLUNGER AND GOT TORN WHILE PUSHING COMPLETELY IN AC." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11448798&pc=QCB | 2023826-2021-00608 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.5/+1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT IS REPORTED. THE LENS REMAINS IMPLANTED WITH A PLANNED EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2021 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. REPORTER STATES "PATIENT IS HAPPY." CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11448899&pc=QCB | 2023826-2021-00607 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.5/+1.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT IS REPORTED. THE LENS REMAINS IMPLANTED WITH A PLANNED EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2021 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. REPORTER STATES "PATIENT IS HAPPY." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453109&pc=QCB | 2023826-2021-00597 | 2020-12-29 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453112&pc=QCB | 2023826-2021-00600 | 2020-12-22 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453126&pc=QCB | 2023826-2021-00604 | 2021-02-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.50/3.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, "THE SURGEON HAS DECIDED TO MONITOR THE OD FOR THE NEXT 2 MONTHS. FOR NOW, NO ABNORMALITY IS FOUND." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453110&pc=QCB | 2023826-2021-00592 | 2020-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PLANNED LENS EXCHANGE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, ALL GOOD. PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11466899&pc=QCB | 2023826-2020-02588 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478829&pc=QCB | 2023826-2021-00642 | 2021-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+4.0/094 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478853&pc=QCB | 2023826-2021-00644 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-03-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478858&pc=QCB | 2023826-2021-00641 | 2021-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+4.5/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. EXPLANT DATE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478836&pc=QCB | 2023826-2021-00640 | 2021-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/092, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMS_ 13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/092, (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SAME SIZE LENS WITH A DIFFERENT DIOPTER ON (B)(6) 2021 DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE MANIFEST REFRACTION IS 0.0/-1.5/175. REPORTEDLY, "THE CLINIC HAS PLANNED A TRANS -PRK TO CORRECT THE ASTIGMATISM." THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478847&pc=QCB | 2023826-2021-00643 | 2021-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507473&pc=QCB | 2023826-2021-00666 | 2021-01-16 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507477&pc=QCB | 2023826-2021-00667 | 2021-01-16 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/177 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11506592&pc=QCB | 2023826-2021-00680 | 2021-02-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON DID NOT LIKE HOW A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/070 (SPHERE/CYLINDER/AXIS), LOOKED. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507469&pc=QCB | 2023826-2021-00672 | 2019-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.0/+4.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507472&pc=QCB | 2023826-2021-00665 | 2020-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT AXIS LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507476&pc=QCB | 2023826-2021-00679 | 2021-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507483&pc=QCB | 2023826-2021-00670 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, AFTER REMOVING FROM THE VIAL, BEFORE LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE INJECTOR, IT WAS TORN. THIS OCCURRED ON (B)(6) 2021. THERE WAS NO PATIENT CONTACT WITH THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "THE LENS WAS DAMAGED SINCE I FIRST SAW IT." AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: AGE/DATE OF BIRTH,WEIGHT, ETHNICITY, RACE -UNKNOWN. DATE REC¿D BY MFR -PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507488&pc=QCB | 2023826-2021-00671 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515295&pc=QCB | 2023826-2021-00649 | 2020-10-08 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/126 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515296&pc=QCB | 2023826-2021-00650 | 2021-01-16 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515307&pc=QCB | 2023826-2021-00658 | 2021-02-03 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515290&pc=QCB | 2023826-2021-00646 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS OBSERVED. SURGEON "BELIEVES IT MAY BE TRAPPED VISCO AT THIS ARLY STAGE." THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, SURGEON DECIDED TO MONITOR PATIENT AND NOT DO ANY EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/4.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS WRONG SIZE. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS THE LENS. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNKNOWN. SEX: UNKNOWN. WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE OF EVENT: UNKNOWN. PRODUCT CODE: UNKNOWN. MODEL #, SERIAL #, EXPIRATION DATE: UNKNOWN. IMPLANT DATE: UNKNOWN. EXPLANT DATE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE : UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515294&pc=QCB | 2023826-2021-00647 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED. SURGEON "BELIEVES IT MAY BE TRAPPED VISCO AT THIS ARLY STAGE." THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6)2021, SURGEON DECIDED TO MONITOR PATIENT AND NOT DO ANY EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON 26/MAR/2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS WRONG SIZE. SEE MFR # XXXXXX FOR CLAIM AGAINST THE INITIAL LENS. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION - THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS THE LENS. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH - UNK. SEX - UNK. WEIGHT - UNK. ETHNICITY- UNK. RACE - UNK. DATE OF EVENT - UNK. PRODUCT CODE: UNK. SERIAL # - UNK. IMPLANT DATE - UNK. IMPLANT DATE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US DEVICE MANUFACTURE DATE - UNK CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515302&pc=QCB | 2023826-2021-00651 | 2020-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -4.50/6.0/007 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY "PATIENT GOT AN IPCL (ANOTHER BRAND LENS) AFTER EXPLANTATION OF THE LENS." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND WITH SPOTS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B7- KERATOCONUS. H3- DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- DEVICE CODE: 1494-OFF LABEL USE FOR KERATOCONUS. CONCLUSION CODE IN PREVIOUS MDR SHOULD BE 4310. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515305&pc=QCB | 2023826-2021-00654 | 2021-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.50/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH SPHERICAL LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "REPLACING THE LENS SOLVED THE PROBLEM OF ROTATING LENS. RESULT: EXCELLENT VISION OF THE PATIENT. SATISFIED DOCTOR." Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.50/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON THE SAME DAY THE LENS WAS EXCHANGED FOR A LONGER LENGTH SPHERICAL LENS. REPORTEDLY, ADDITIONAL CORRECTION OF ASTIGMATISM WAS CARRIED OUT WITH A LASER'. THE PROBLEM WAS RESOLVED. HEALTH IMPACT CODE- 4625- LASER REFRACTIVE SURGERY. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11522369&pc=QCB | 2023826-2021-00682 | 2021-01-21 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/2.50/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY. NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2021 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11522372&pc=QCB | 2023826-2021-00683 | 2021-02-08 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/3.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. AT CHECK UP DOCTOR FOUND THE LENS TORN AND ROTATED. 1 CT VAULT VALUE WAS REPORTED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE. THE REPORTER STATED " WE NEED A REPLACEMENT BECAUSE IT IS KEEP ROTATING AS IT TORN AND NOT STABLE IN THE EYE, SUBLUXATED FROM THE TORN." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11522545&pc=QCB | 2023826-2021-00686 | 2021-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.5/2.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A STRETCHED OUT HAPTIC AND DEBRIS/RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(6). PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536796&pc=QCB | 2023826-2021-00725 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2021-03-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.50/1.5/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536797&pc=QCB | 2023826-2021-00724 | 2020-05-30 | Injury | STAAR SURGICAL COMPANY | 2021-03-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -9.00/1.0/025 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536799&pc=QCB | 2023826-2021-00726 | 2021-01-23 | Injury | STAAR SURGICAL COMPANY | 2021-03-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.00/1.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536801&pc=QCB | 2023826-2021-00731 | 2020-09-14 | Injury | STAAR SURGICAL COMPANY | 2021-03-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -5.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536802&pc=QCB | 2023826-2021-00733 | 2020-04-25 | Injury | STAAR SURGICAL COMPANY | 2021-03-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.5/1.0/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536798&pc=QCB | 2023826-2021-00730 | 2021-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.00/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536800&pc=QCB | 2023826-2021-00732 | 2020-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.00/1.5/175 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED.. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: B5: THE EXCHANGED LENS WAS IMPLANTED VERTICALLY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11547890&pc=QCB | 2023826-2021-00744 | 2019-11-18 | Injury | STAAR SURGICAL COMPANY | 2021-03-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11547906&pc=QCB | 2023826-2021-00742 | 2020-05-15 | Injury | STAAR SURGICAL COMPANY | 2021-03-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11547928&pc=QCB | 2023826-2021-00741 | 2021-02-05 | Injury | STAAR SURGICAL COMPANY | 2021-03-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11547872&pc=QCB | 2023826-2021-00743 | 2020-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11557384&pc=QCB | 2023826-2021-00748 | 2020-08-22 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+1.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558787&pc=QCB | 2023826-2021-00708 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/110 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558790&pc=QCB | 2023826-2021-00721 | 2021-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-23 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, BEFORE/DURING LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/091 (SPHERE/CYLINDER/AXIS), IT TORE. THIS WAS DISCOVERED AFTER OPENING THE VIAL AND THERE WAS NO PATIENT CONTACT WITH THE LENS. THIS OCCURRED ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558796&pc=QCB | 2023826-2021-00713 | 2021-02-19 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558797&pc=QCB | 2023826-2021-00711 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-03-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.0/+2.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WAS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558801&pc=QCB | 2023826-2021-00709 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11556880&pc=QCB | 2023826-2021-00747 | 2020-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH DEBRIS ON THE PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558782&pc=QCB | 2023826-2021-00722 | 2020-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON STATES THERE IS A HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: ADDITIONAL/UPDATED INFO: A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM IS RESOLVED. CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. EVENT DATE:UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MFR DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558784&pc=QCB | 2023826-2021-00717 | 2021-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.5/+4.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "SIZING." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY AND STUCK ON A PIECE OF GAUZE WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558788&pc=QCB | 2023826-2021-00716 | 2020-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +2.0/+5.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME SIZE DIFFERENT AXIS LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558791&pc=QCB | 2023826-2021-00710 | 2021-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/065 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REFERENCE MFR REP 2023826-2021-00711 FOR REPLACEMENT LENS. Manufacturer Narrative: CORRECTED DATA: H3 - "DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS". SHOULD HAVE BEEN INCLUDED WITH PREVIOUS SUPPLEMENTAL. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563241&pc=QCB | 2023826-2021-00687 | 2020-12-30 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563277&pc=QCB | 2023826-2021-00761 | 2020-11-06 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563246&pc=QCB | 2023826-2021-00689 | 2020-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.50/2.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS REPOSITIONED, BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563249&pc=QCB | 2023826-2021-00587 | 2021-02-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH13.2, -4.50/4.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. PARTIAL LENS STUCK IN CARTRIDG. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. "INJECTOR/CARTRIDGE NOT SUITABLE FOR LENS THICKNESS." Manufacturer Narrative: ADDITIONAL INFORMATION: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563250&pc=QCB | 2023826-2021-00694 | 2021-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.50/4.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021.THE LENS TORE DURING INJECTION/DELIVERY. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -14.50/4.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS TORE DURING INJECTION/DELIVERY. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED. REPORTEDLY INCISION ENLARGEMENT AND ADDITIONAL SUTURE WERE USED TO REMOVE THE LENS. IT WAS REPLACED ON (B)(6) 2021 WITH A SAME LENGTH/MODEL LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563256&pc=QCB | 2023826-2021-00697 | 2020-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.00/1.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563258&pc=QCB | 2023826-2021-00695 | 2021-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ROTATED A TORIC LENS 12.6MM AND WILL EXCHANGE IT WITH A LONGER LENGTH LENS 13.2MM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2019. THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS "LENS ROTATION DUE TO LOW VAULT." REPORTEDLY, LENS REPLACEMENT IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563259&pc=QCB | 2023826-2021-00698 | 2020-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.50/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH DEBRIS ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: ETHNICITY - UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563273&pc=QCB | 2023826-2021-00696 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE (OD) WAS OBSERVED WITH MISSALIGNMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -18.00/1.5/141 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (ODS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 AND (B)(6) 2021 THE LENS WAS REPOSITIONED, THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, A LENS REPLACEMENT IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -18.00/1.5/141 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (ODS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 AND ON (B)(6) 2021 THE LENS WAS REPOSITIONED, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021, THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -18.00/1.5/141 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 AND (B)(6) 2021 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. EPIRETINAL GLIOSIS REPORTED. CLAIM # (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11579927&pc=QCB | 2023826-2021-00767 | 2021-02-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -16.50/0.5/175(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WITH ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11579961&pc=QCB | 2023826-2021-00775 | 2020-12-08 | Injury | STAAR SURGICAL COMPANY | 2021-03-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -4.5/0.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11579971&pc=QCB | 2023826-2021-00773 | 2020-10-23 | Injury | STAAR SURGICAL COMPANY | 2021-03-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.0/1.0/029 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11579975&pc=QCB | 2023826-2021-00774 | 2020-12-08 | Injury | STAAR SURGICAL COMPANY | 2021-03-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -4.0/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11579976&pc=QCB | 2023826-2021-00772 | 2020-10-23 | Injury | STAAR SURGICAL COMPANY | 2021-03-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.5/1.0/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11579936&pc=QCB | 2023826-2021-00768 | 2021-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.00/1.00/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WITH ROTATION WAS OBSERVED. IT WAS ALSO REPORTED THAT MONOVISION IS NOT BEING TOLERATED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.1 MM, VTICMO12.1, -16.50/0.5/175(SPHERE/CYLINDER/AXIS ), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS LATER EXPLANTED ON AN UNKNOWN DATE DUE TO LOW VAULT WITH ROTATION. ON (B)(6) 2021 A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT POWER WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11585662&pc=QCB | 2023826-2021-00805 | 2021-01-21 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/086 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE RIGHT EYE (OD). THERE WAS NO INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11585663&pc=QCB | 2023826-2021-00803 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-03-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL IMPLANTABLE COLLAMER LENS AND THE LENS HAD ROTATED WITH A SHALLOW VAULT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11586687&pc=QCB | 2023826-2021-00811 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-03-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11586783&pc=QCB | 2023826-2021-00812 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-03-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11585657&pc=QCB | 2023826-2021-00802 | 2021-03-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -06.50/+2.5/091 (SPHERE/CYLINDER/AXIS), AND THE SURGEON NOTED A CRINKLE IN THE LENS AND DECIDED NOT TO IMPLANT THE LENS. THERE WAS PATIENT CONTACT BUT NO INJURY, NO INCISION ENLARGEMENT OR SUTURES. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11585671&pc=QCB | 2023826-2021-00804 | 2019-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED; THE REPORTER STATES THE PROBLEM WAS RESOLVED. ON THE SAME DAY IT IS ALSO REPORTED THAT THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED; DUE TO LENS ROTAION AND THE REPORTER STATES THE PROBLEM WAS RESOLVED. ON THE SAME DAY IT IS ALSO REPORTED THAT THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11586785&pc=QCB | 2023826-2021-00813 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11586787&pc=QCB | 2023826-2021-00814 | 2021-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11586788&pc=QCB | 2023826-2021-00815 | 2020-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11586790&pc=QCB | 2023826-2021-00819 | 2021-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BENT. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11586793&pc=QCB | 2023826-2021-00820 | 2021-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICROCENTRIFUGE/ VIAL, WITH RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11592159&pc=QCB | 2023826-2020-02587 | 2019-12-12 | Injury | STAAR SURGICAL COMPANY | 2021-03-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11593764&pc=QCB | 2023826-2021-00825 | 2020-12-28 | Injury | STAAR SURGICAL COMPANY | 2021-03-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -10.0/2.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11598785&pc=QCB | 2023826-2021-00778 | 2020-11-26 | Injury | STAAR SURGICAL COMPANY | 2021-03-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/1.0/058 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11598799&pc=QCB | 2023826-2021-00781 | 2021-02-24 | Injury | STAAR SURGICAL COMPANY | 2021-03-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11598815&pc=QCB | 2023826-2021-00798 | 2021-01-08 | Injury | STAAR SURGICAL COMPANY | 2021-03-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.50/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11598830&pc=QCB | 2023826-2021-00793 | 2020-09-11 | Injury | STAAR SURGICAL COMPANY | 2021-03-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11598856&pc=QCB | 2023826-2021-00792 | 2020-07-22 | Injury | STAAR SURGICAL COMPANY | 2021-03-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/4.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11606452&pc=QCB | 2023826-2021-00830 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM513.2, -8.5/5.00/075 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11606455&pc=QCB | 2023826-2021-00829 | 2021-02-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, 2.0/4.0/175, IMPLANTABLE COLLAMER LENS GOT STUCK IN CARTRIDGE OR INJECTION SYSTEM ON (B)(6) 2021. THERE WAS PATIENT CONTACT BUT NO INJURY TO THE PATIENT. LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE SAME SURGERY, AND THE PROBLEM RESOLVED. CAUSE IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11606741&pc=QCB | 2023826-2021-00831 | 2019-07-02 | Injury | STAAR SURGICAL COMPANY | 2021-04-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM012.1, -13.5/1.5/092 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11606761&pc=QCB | 2023826-2021-00832 | 2020-10-26 | Injury | STAAR SURGICAL COMPANY | 2021-04-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, 8.00/5.0/090 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11606784&pc=QCB | 2023826-2021-00833 | 2021-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, 8.00/2.5/118 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 IN A SECOND SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616795&pc=QCB | 2023826-2021-00842 | 2021-01-20 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED, THE ICL IS IN A GOOD POSITION. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616831&pc=QCB | 2023826-2021-00846 | 2021-03-10 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/055 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616840&pc=QCB | 2023826-2021-00848 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+4.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616999&pc=QCB | 2023826-2021-00853 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+4.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11617006&pc=QCB | 2023826-2021-00852 | 2021-01-13 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616834&pc=QCB | 2023826-2021-00847 | 2021-03-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +01.00/+6.0/103 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE (BY PULLING IT TOWARDS THE FRONT OF THE CARTRIDGE). THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN (MAY BE DUE TO THE THICKNESS OF THIS ICL). Manufacturer Narrative: ADDITIONAL DATA: B5: THERE WAS NO PATIENT CONTACT. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616842&pc=QCB | 2023826-2021-00849 | 2021-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616901&pc=QCB | 2023826-2021-00927 | 2020-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11617002&pc=QCB | 2023826-2021-00851 | 2021-03-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS WITH IN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11617003&pc=QCB | 2023826-2021-00850 | 2021-03-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/089 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11617010&pc=QCB | 2023826-2021-00855 | 2020-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -05.00/+1.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO VAULT WAS NOT PREDICTABLE WITH WTW MEASUREMENTS. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620597&pc=QCB | 2023826-2021-00872 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620602&pc=QCB | 2023826-2021-00869 | 2021-02-15 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-06 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.5/+1.5/007 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS THE LENS WAS IMPLANTED UPSIDE DOWN. INTRAOPERATIVELY THE LENS WAS REMOVED. ON (B)(6) 2021 AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: THE LENS GOT SUCK BETWEEN THE PLUNGER AND CARTRIDGE WALL, THE LENS WAS DAMAGED DURING REMOVAL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620603&pc=QCB | 2023826-2021-00879 | 2021-03-13 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/145 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND REPLACED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620606&pc=QCB | 2023826-2021-00871 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/065 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620607&pc=QCB | 2023826-2021-00866 | 2021-03-03 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620615&pc=QCB | 2023826-2021-00868 | 2021-03-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-06 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE ATTEMPTING TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), IT WAS TORN. REPORTEDLY THE LENS GOT STUCK BETWEEN THE PLUNGER AND CARTRIDGE WALL WHICH WAS THE CAUSE OF THE LENS DAMAGE. THIS OCCURRED ON (B)(6) 2021 AND THERE WAS PATIENT CONTACT WITH THE LENS BUT IT WAS NOT FULLY IMPLANTED. ON (B)(6) 2021 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. NO PATIENT INJURY IS REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620601&pc=QCB | 2023826-2021-00873 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620608&pc=QCB | 2023826-2021-00863 | 2021-01-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. IT WAS REPLACED WITH AN ALTERNATE LENS AT A LATER DATE AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11625144&pc=QCB | 2023826-2021-00935 | 2020-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/3.5/124(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITHOUT ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS REPORTED AS OTHER "OUT OF FIT". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/3.5/124 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER EXCHANGED ON (B)(6) 2021 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11625360&pc=QCB | 2023826-2021-00937 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.00/3.0/057 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE PATIENT DECIDED NO TO REPLACE THE LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11628511&pc=QCB | 2023826-2021-00858 | 2020-07-26 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/114 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11627536&pc=QCB | 2023826-2021-00936 | 2020-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE PATIENT DECIDED NO TO REPLACE THE LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11632378&pc=QCB | 2023826-2021-00939 | 2021-01-08 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON DISCOVERED A CRACK ON A 12.6MM, VTICMO12.6, -5.5/1.0/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS AFTER IT HAD BEEN INJECTED. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. NO PATIENT INJURY WAS REPORTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11632520&pc=QCB | 2023826-2021-00942 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P990013 | 1 | Event Description: IT WAS REPORTED THAT A TORIC ICL IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND ROTATION WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11648932&pc=QCB | 2023826-2021-00946 | 2020-12-25 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE LENS SIZE DID NOT FIT THE PATIENT'S EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11649007&pc=QCB | 2023826-2021-00950 | 2021-03-05 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/135 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651533&pc=QCB | 2023826-2021-00890 | 2020-11-28 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.50/2.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS LENS SIZE DID NOT FIT PATIENT'S EYE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651556&pc=QCB | 2023826-2021-00891 | 2020-11-25 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.00/2.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020 AND (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "LENS EXCHANGED BECAUSE SURGEON CONFIRMED LENS ROTATION AGAIN ON (B)(6). THE INITIAL LENS SIZE MAY HAVE BEEN SMALL FOR PATIENT." Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651559&pc=QCB | 2023826-2021-00889 | 2021-01-17 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.00/2.5/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS LENS SIZE DID NOT FIT PATIENT'S EYE. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. REPORT SOURCE- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651565&pc=QCB | 2023826-2021-00881 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -10.00/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS TOO SMALL AND WAS REPLACED WITH A LONGER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651571&pc=QCB | 2023826-2021-00888 | 2021-02-13 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.50/2.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS LENS SIZE DID NOT FIT THE PATIENT'S EYES. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651608&pc=QCB | 2023826-2021-00908 | 2021-01-12 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651626&pc=QCB | 2023826-2021-00906 | 2019-03-12 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/2.5/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT DURING A SECOND SURGERY ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651636&pc=QCB | 2023826-2021-00907 | 2021-03-02 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. REPORT SOURCE- THIS PRODUCT IS NOT MARKETED IN THE US.(B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651639&pc=QCB | 2023826-2021-00912 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.50/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651640&pc=QCB | 2023826-2021-00909 | 2020-10-30 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651642&pc=QCB | 2023826-2021-00902 | 2019-06-14 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/3.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2019. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4) OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651643&pc=QCB | 2023826-2021-00916 | 2021-03-05 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/2.5/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651649&pc=QCB | 2023826-2021-00919 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/3.0/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. REPORT SOURCE- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651650&pc=QCB | 2023826-2021-00903 | 2019-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651654&pc=QCB | 2023826-2021-00911 | 2020-10-30 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO13.2, -17.00/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.00/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE. COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11653732&pc=QCB | 2023826-2021-00990 | 2021-03-11 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11653867&pc=QCB | 2023826-2021-00991 | 2021-03-11 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/068 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(4) 2020. THE LENS WAS EXPLANTED ON (B)(4) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11653869&pc=QCB | 2023826-2021-00995 | 2021-03-15 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654096&pc=QCB | 2023826-2021-00964 | 2021-01-11 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/065 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS "SMALL SIZE LENS." Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654101&pc=QCB | 2023826-2021-00984 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654102&pc=QCB | 2023826-2021-00983 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654119&pc=QCB | 2023826-2021-00961 | 2020-11-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654128&pc=QCB | 2023826-2021-00963 | 2021-01-09 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. SEE MFR FILE# 719543 FOR REPLACEMENT LENS. CAUSE OF THE EVENT IS REPORTED AS "SMALL SIZE LENS." Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651567&pc=QCB | 2023826-2021-00882 | 2021-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.50/3.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651573&pc=QCB | 2023826-2021-00892 | 2020-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -16.50/3.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020 AND (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTHL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE SURGEON EXCHANGED THE LENS DUE LENS SIZE MAY HAVE BEEN SMALL FOR PATIENT." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11653718&pc=QCB | 2023826-2021-00989 | 2021-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00/+1.0/009 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5:THE REPORTER INDICATED THE LENS WAS EXPLANTED AND A CLEAR-LENS SURGERY WAS PERFORMED. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654099&pc=QCB | 2023826-2021-00956 | 2021-03-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/055 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S EYE. REPORTEDLY, THE LENS WAS IMPLANTED AND FLIPPED UPSIDE DOWN. THE LENS WAS EXPLANTED AND TORE UPON REMOVAL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-PREVIOUSLY STATED THE LENS FLIPPED UPSIDE DOWN. CORRECTED INFO: THE LENS TORE DURING INJECTION INTO THE PATIENT'S LEFT (OS) EYE. H6-DEVICE CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: "FLIPPED. TORE ON REMOVAL." H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/055 (SPHERE/CYLINDER/AXIS) TORE DURING DELIVERY INTO THE PATIENT'S LEFT EYE (OS). SINCE THIS LENS WAS A BACK UP LENS, THE PATIENT HAD TO WAIT TWO DAYS FOR AN ALTERNATE LENS TO BE IMPLANTED. THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654103&pc=QCB | 2023826-2021-00980 | 2021-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/124 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654104&pc=QCB | 2023826-2021-00981 | 2021-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654105&pc=QCB | 2023826-2021-00976 | 2021-02-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.0/+3.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654106&pc=QCB | 2023826-2021-00977 | 2021-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654110&pc=QCB | 2023826-2021-00966 | 2021-03-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/098 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE LEFT (OS) EYE. REPORTEDLY, THE LENS GOT STUCK IN THE INJECTOR AND WAS TORN. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2021.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/098 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE LEFT (OS) EYE. REPORTEDLY, THE LENS GOT STUCK IN THE INJECTOR AND WAS TORN. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2021. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5, D4:MODEL TYPE CORRECTION; VTICMO 12.6. SUPP1 NOT APPLICABLE. CLAIM (B)(4). Manufacturer Narrative: B5, D4 : MODEL TYPE CORRECTION; VTICMO12.6 . CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654125&pc=QCB | 2023826-2021-00951 | 2021-03-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/068 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S EYE ON (B)(6) 2021. REPORTEDLY, THE LENS FLIPPED AND WAS TORN DURING EXPLANT. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY, IN AN ATTEMPT TO REMOVE THE ICL (DUE TO NOT UNFOLDING CORRECTLY) FROM THE LEFT (OS) EYE, THE ICL "TORE." THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THE PROBLEM IS RESOLVED. H6-HEALTH IMPACT CODE CORRECTED. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654129&pc=QCB | 2023826-2021-00955 | 2021-03-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Material Integrity Problem | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/055 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S EYE. REPORTEDLY, THE LENS WAS ABSENT OF THE MARKINGS ON THE HAPTICS TO VERIFY THE CORRECT ORIENTATION. THE LENS WAS EXPLANTED AND TORE UPON REMOVAL. SEE MFR REP# 2023826-2021-00956 FOR REPLACEMENT LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1-CORRECTED B5-PREVIOUSLY STATED THE LENS WAS IMPLANTED. THE DOCTOR "FELT IFFY ABOUT IMPLANTING IT" SO THE BACK-UP LENS WAS IMPLANTED INSTEAD. H6-HEALTH IMPACT CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. HAPTIC MARKINGS ARE NOTED. (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11669695&pc=QCB | 2023826-2021-01052 | 2021-03-12 | Injury | STAAR SURGICAL COMPANY | 2021-04-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -9.5/5.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11676162&pc=QCB | 2023826-2021-01058 | 2021-02-18 | Injury | STAAR SURGICAL COMPANY | 2021-04-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -9.5/0.5/107 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11669637&pc=QCB | 2023826-2021-01051 | 2021-03-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -1.50/6.0/97 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2021. THE LENS TORE DURING INJECTION. LENS REMAINS IMPLANTED. NO PATIENT INJURY WAS REPORTED. CAUSE IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THA A 13.7MM, VTICM5_13.7, -1.50/6.0/97(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 18 FEB 2021. THE LENS TORE DURING INJECTION. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL AND LENGTH LENS AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11669715&pc=QCB | 2023826-2021-01053 | 2021-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.00/3.0/053 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2021. IT WAS REPORTED THAT 1 DAY POST OP LENS ROTATION (30 DEGREES) WAS OBSERVED. A VAULT VALUE OF 100 WAS REPORTED. PRE-OP AND POST-OP BCVA 20/20 WAS REPORTED. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. PATIENT IS VERY UNHAPPY WITH RESULT. FOR CAUSE OF EVENT " ICL LENGTH (12.6) TOO SMALL FOR THIS PATIENT" WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM, VTICM5_ 12.6, -7.00/3.0/053 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2021. IT WAS REPORTED THAT AT 1 DAY POST OP LENS ROTATION (30 DEGREES) WAS OBSERVED. A VAULT VALUE OF 100 WAS REPORTED. PRE-OP AND POST-OP BCVA 20/20 WAS REPORTED. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. PATIENT IS VERY UNHAPPY WITH RESULT. FOR CAUSE OF EVENT " ICL LENGTH (12.6) TOO SMALL FOR THIS PATIENT" WAS REPORTED. (B)(6) 2021 THE LENS WAS EXPLANTED, A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED, AND THE PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN, AND DEBRIS THROUGHOUT THE LENS. CLAIM #(B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683137&pc=QCB | 2023826-2021-01063 | 2020-05-04 | Injury | STAAR SURGICAL COMPANY | 2021-04-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -14.0/1.5/175 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683138&pc=QCB | 2023826-2021-01067 | 2020-06-04 | Injury | STAAR SURGICAL COMPANY | 2021-04-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -7.0/1.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683139&pc=QCB | 2023826-2021-01062 | 2020-08-13 | Injury | STAAR SURGICAL COMPANY | 2021-04-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683141&pc=QCB | 2023826-2021-01061 | 2021-02-25 | Injury | STAAR SURGICAL COMPANY | 2021-04-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.0/2.5/82 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTER AS OTHER : THE LENS SIZE DID NOT FIT THE PATIENTS EYE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683145&pc=QCB | 2023826-2021-01060 | 2021-01-24 | Injury | STAAR SURGICAL COMPANY | 2021-04-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -6.0/2.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTER AS OTHER : THE LENS SIZE DID NOT FIT THE PATIENTS EYE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683151&pc=QCB | 2023826-2021-01059 | 2020-12-16 | Injury | STAAR SURGICAL COMPANY | 2021-04-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -6.0/2.0/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTER AS OTHER : THE LENS SIZE DID NOT FIT THE PATIENTS EYE. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683176&pc=QCB | 2023826-2021-01075 | 2021-03-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.00/6.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683135&pc=QCB | 2023826-2021-01064 | 2020-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.0/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITHOUT ROTATION AND THE PROBLEM RESOLVED. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. PATIENT REFUSED TO CHANGE THE LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. IT WAS RETUNED DRY WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND A HAPTIC TORN, BROKEN, BENT, DEFORMED, AND STRETCHED OUT LENS. UNABLE TO PERFORM WIDTH FOR THE DIMENSIONAL INSPECTION DUE TO HAPTIC DAMAGE. CLAIM (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683203&pc=QCB | 2023826-2021-01076 | 2020-11-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -14.0/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING SURGERY, PRE INSERTING. AN ADDITIONAL LENS WAS USED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN THE ASSOCIATED LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697098&pc=QCB | 2023826-2021-01100 | 2021-03-05 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/076 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697107&pc=QCB | 2023826-2021-01078 | 2021-03-25 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/093 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. THERE WAS NO PATIENT CONTACT WITH THE LENS. REPORTEDLY THE SURGEON THOUGHT HE HAD ADEQUATE PURCHASE ON THE LENS WHILE PULLING IT THROUGH THE CARTRIDGE AND WAS SURPRISED IT TORE. THIS OCCURRED DURING THE SURGEON'S CERTIFICATION PROCESS. Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697128&pc=QCB | 2023826-2021-01098 | 2021-03-11 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697141&pc=QCB | 2023826-2021-01095 | 2020-12-18 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697143&pc=QCB | 2023826-2021-01085 | 2020-11-24 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.5/+0.5/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON THE SAME DATE, IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697157&pc=QCB | 2023826-2021-01082 | 2021-02-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+3.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. A SHORTER LENS WAS IMPLANTED ON A LATER DATE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697137&pc=QCB | 2023826-2021-01092 | 2021-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "THE LENS SIZE DID NOT FIT THE PATIENT." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: METHOD CODE: 10. RESULT CODE: 213. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697147&pc=QCB | 2023826-2021-01091 | 2021-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.5/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "THE LENS SIZE DID NOT FIT THE PATIENT." Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: METHOD CODE - 10; RESULT CODE - 213. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697155&pc=QCB | 2023826-2021-01093 | 2021-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/+2.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS REPOSITION DATE (B)(6) 2021 CORRECTED TO (B)(6) 2021. THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/+2.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11704874&pc=QCB | 2023826-2021-01105 | 2020-11-21 | Injury | STAAR SURGICAL COMPANY | 2021-04-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11704883&pc=QCB | 2023826-2021-01106 | 2020-11-21 | Injury | STAAR SURGICAL COMPANY | 2021-04-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.50/0.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11704904&pc=QCB | 2023826-2021-01109 | 2020-10-12 | Injury | STAAR SURGICAL COMPANY | 2021-04-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.0/2.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 LENS WAS LATER EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11713122&pc=QCB | 2023826-2021-01151 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4): THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11713149&pc=QCB | 2023826-2021-01118 | 2021-03-13 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/2.5/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11713150&pc=QCB | 2023826-2021-01115 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/2.0/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11713164&pc=QCB | 2023826-2021-01132 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.0/1.5/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11713165&pc=QCB | 2023826-2021-01119 | 2021-01-20 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: HOLD FOR AB 6.14 THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: WEIGHT, UNK. ETHNICITY: UNK,.RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM (B)(4).