https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8887948&pc=MTA | 2023826-2019-01416 | 2019-07-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 20/JUN/2019. PIGMENT DISPERSION, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. PER REPORTER, ON PATIENTS (B)(6) 2019 VISIT, "BCVA WAS 20/20, VAULT 867, IOP 17 MMHG, AND STATUS OF THE EYE PIGMENT DISPERSION. I WOULD LIKE TO EXCHANGE THE ICL." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT CODE 114 NOT APPLICABLE IN INITIAL MDR. RESULT CODE 213 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE PATIENT IS 'ASYMPTOMATIC'. THE BCVA WAS RECORDED AS 20/20 WITH A VAULT OF 801 UM. THE SURGEON HAS DECIDED NOT TO REPLACE THE LENS. THE LENS REMAINS IMPLANTED. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | OS |  | 07/15/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194491&pc=MTA | 2023826-2018-01904 | 2018-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances; No Code Available | N/A | Event Description: REPORTEDLY, THE PATIENT IS SUFFERING FROM DIPLOPIA IN THE RIGHT (OD) EYE AND LENS DISLOCATION. THE LENS REMAINS IMPLANTED. THIS IS A SPHERICAL 12.6MM ICL, -12.0 DIOPTER LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0 WAS IMPLANTED INTO THE PATIENT'S LEFT (OD) EYE. ON (B)(6) 2018 THE LENS WAS EXPLANTED, THE SURGEON NOTING PIGMENT DISPERSION AND LENS DISLOCATION WHICH CAUSED DIPLOPIA. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. THE PROBLEM WAS RESOLVED WITH THE LENS EXPLANT. Manufacturer Narrative: PT INFO: UNK. EVENT DATE: UNK. IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MFR DATE: UNK. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DESCRIBE EVENT OR PROBLEM-PLEASE DISREGARD PREVIOUS DESCRIBE EVENT OR PROBLEM STATEMENT. ADDITIONAL PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (PIGMENT DISPERSION, LENS EXPLANT, SECONDARY SURGERY). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | OS |  | 12/01/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140763&pc=MTA | 2023826-2020-01231 | 2019-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED, REPORTER STATED " THE PATIENT FELT THAT THE GLARE WAS OBVIOUS AFTER SURGERY, AND IT WAS DIFFICULT TO ADAPT, SO HE WAS REQUIRED TO REMOVE IT". EXPLANTING THE LENS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | OD |  | 05/12/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7329320&pc=MTA | 2023826-2018-00357 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-09 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER ON (B)(6) 2017 IN THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTED THAT "THE PATIENT IS SUFFERING FROM SEVERE GLARE AND HE BELIEVES THAT IT IS FROM THE CENTRAL PORTS OF THE LENSES IN HER EYES". THE REPORTER STATED THAT THE PATIENT STILL HAS "GLARES/ DECENTRATION". THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE REPORTER BUT NONE HAVE BEEN FORTHCOMING. A SUPPLEMENTAL MDR WILL BE SUBMITTED IF NEW INFORMATION IS RECEIVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | OS | Physician | 02/07/2018 | 2018
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8447848&pc=MTA | 2023826-2019-00480 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem; No Apparent Adverse Event | Visual Disturbances; No Known Impact Or Consequence To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO PATIENT'S EYE ON (B)(6) 2019. REPORTEDLY THE SURGEON "WILL DETERMINE IF THE LENS NEEDS TO BE EXCHANGED DUE TO VAULT". LENS REMAINS IMPLANTED. "THERE ARE NO COMPLICATIONS WITH PATIENT AS OF NOW." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMAITON; SERIAL# UPDATED FROM SERIAL NUMBER (B)(4) TO SERIAL NUMBER (B)(4). EXPIRATION DATE UPDATED FROM 31-AUG-2020 TO 31-JUL-2019. UNIQUE IDENTIFIER (UDI)# UPDATED FROM (B)(4) TO (B)(4). DEVICE MANUFACTURE DATE UPDATED FROM 14-SEP-2018 TO 11-AUG-2017. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THERE IS CONCERN FOR ASSESSMENT OF VAULTING AND THE SURGEON WILL MONITOR TO DETERMINE IF AN EXCHANGE IS NECESSARY. THE PATIENT IS DOING WELL WITH MILD COMPLAINTS OF GLARE. THE LENS REMAINS IMPLANTED. MICL 13.7 SHOULD BE ADDED TO PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | OS |  | 02/26/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9886889&pc=MTA | 2023826-2020-00696 | 2020-03-05 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Known Impact Or Consequence To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICM_13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019 AS A REPLACEMENT LENS. IT IS REPORTED THAT " THE PATIENT STILL HAS GLARE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. REFERENCE MFR# 2023826-2019-02020 FOR INITIAL LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | OD | Physician | 03/05/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4960431&pc=MTA | 2023826-2015-00853 | 2014-12-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contamination with Chemical or Other Material | Visual Impairment | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND ON (B)(6) 2014 THE SURGEON OBSERVED A WHITE MATERIAL ADHERED ON THE LENS. A DECREASE IN UVCA WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | OS | Other | 07/09/2015 | 2015 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11058380&pc=MTA | 2023826-2020-03119 | 2020-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Visual Impairment | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT MODEL AND DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT COMPLAINED OF UNSATISFACTORY VISUAL ACUITY AND ASKED TO BE REPLACED WITH A TICL LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. THE LENS WAS RETURNED SIGNIFICANTLY DAMAGED WHICH CAN CAUSE DIFFICULTY WHEN CENTERING THE LENS DURING FUNCTIONAL INSPECTION. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE LENS WIDTH AND LENGTH WAS IN SPECIFICATION BUT REFRACTION (FUNCTIONAL) VERIFICATION FAILED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | OD |  | 11/24/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5040390&pc=MTA | 2023826-2015-01107 | 2015-07-11 | Injury | STAAR SURGICAL COMPANY | 2015-08-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | Visual Impairment | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS TORE WHILE BEING INSERTED AND THE SURGEON DECIDED TO LEAVE THE LENS IN THE EYE. THE LENS HAS SINCE DECENTERED AND THE EDGE OF THE LENS IS IN THE PATIENT'S VISION. THE SURGEON'S PLAN IS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). | OD | Physician | 07/29/2015 | 2015 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4966256&pc=MTA | 2023826-2015-00943 |  | Injury | STAAR SURGICAL COMPANY | 2015-08-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Visual Impairment; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT EXPERIENCED A LOSS OF UCVA AND THE LENS WAS EXCHANGED FOR TICM. THIS RESOLVED THE PROBLEM. PATIENT'S LAST UCVA WAS 20/25. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS IN SPECIFICATION. (B)(4). | OD | Physician | 07/14/2015 | 2015 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5206084&pc=MTA | 2023826-2015-01470 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Visual Impairment; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT WAS UNDER-CORRECTED. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE PATIENT IS NOW 20/20. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, VISUAL INSPECTION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - REPORTEDLY, ICL (-5.5 D;12.6 MM) WAS EXPLANTED/EXCHANGED, THREE WEEKS POSTOPERATIVELY, FOR A DIFFERENT POWER ICL OF TO ADDRESS PATIENT RESIDUAL REFRACTIVE ERROR ("UNDERCORRECTION"). NO REPORTED PROBLEM WITH THE LENS POSITION PRIOR TO THE LENS EXCHANGE. ACCORDING TO THE REPORT, BY A DOCTOR, DATED ON (B)(6) 2015, THE REPLACEMENT LENS RESOLVED THE ISSUE AND UVA WAS 20/20. GIVEN THE INFORMATION THAT A DIFFERENT POWER LENS WAS IMPLANTED AS A REPLACEMENT AND THE REFRACTIVE ERROR WAS CORRECTED IT IS VERY LIKELY THAT USER ERROR (MISCALCULATION) WAS THE CAUSE OF THE EVENT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, THE MOST LIKELY ROOT CAUSE OF THE EVENT WAS USER ERROR (MISCALCULATION). (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION, WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE SURGERY FACILITY REPORTED THAT THERE WAS NO PROBLEM WITH THE LENS. AS REPORTED ON THE MEDICAL REVIEW, SINCE A DIFFERENT POWER LENS WAS IMPLANTED AS A REPLACEMENT AND THE REFRACTIVE ERROR WAS CORRECTED, IT IS VERY LIKELY THE USER ERROR (MISCALCULATION) WAS THE CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). | OD | Physician | 10/06/2015 | 2015 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5131893&pc=MTA | 2023826-2015-01231 | 2015-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-10-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Impairment; No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THE PATIENT HAD A HYPEROPIC REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP UCVA WAS 20/25 ON (B)(6) 2015. SEE MFR # 2023826-2015-01311 FOR LEFT EYE. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS CODE(S): (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. EVALUATION CONCLUSION CODE(S) BASED ON THE COMPLAINT HISTORY, DEVICE HISTORY RECORD REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT HAD A HYPEROPIC REFRACTIVE SURPRISE. TWO MONTHS POSTOPERATIVELY, THE SURGEON EXCHANGED THE LENS FOR ANOTHER OF THE SAME MODEL BUT DIFFERENT DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. AFTER THE EXCHANGE, PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | OD | Physician | 09/09/2015 | 2015 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4509422&pc=MTA | 2023826-2015-00115 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Visual Impairment; No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT'S VISION WAS HYPEROPIC. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH NO COMPLICATIONS.Event Description: ADDITIONAL INFORMATION RECEIVED - THERE WERE NO RELATED COMPLICATIONS TO THE EVENT AND THE ICL WAS EXPLANTED WITH NO REPORTED INJURY. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. AS REPORTED ON THE MEDICAL REVIEW, THE MEDICAL DEVICE WAS UNLIKELY THE CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF LENS HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | OD | Other | 01/12/2015 | 2015 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5395179&pc=MTA | 2023826-2016-00072 |  | Injury | STAAR SURGICAL COMPANY | 2016-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Visual Impairment; No Code Available | P030016 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO THE PATIENT COULD NOT SEE WELL WITH THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE ICL LENS WAS EXPLANTED DUE TO TOO SMALL/FLAT SO IT WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT DID WELL AFTER THE EXCHANGE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. PER MEDICAL REVIEW: REPORTEDLY, MICL (-15D; 12.1 MM) WAS EXPLANTED POSTOPERATIVELY TO ADDRESS PATIENT RESIDUAL REFRACTIVE ERROR ("COULD NOT SEE WELL WITH THE LENS"). NO REPORTED POSTOPERATIVE SEQUELAE. DESPITE MULTIPLE ATTEMPTS NO ADDITIONAL INFORMATION HAS BEEN RECEIVED TO DATE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | NA | Other | 01/05/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6359091&pc=MTA | 2023826-2017-00250 | 2017-01-13 | Injury | STAAR SURGICAL COMPANY | 2017-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Vitrectomy; No Code Available | P030016 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS), WHICH WAS TREATED LIKE ENDOPHTHALMITIS. A VITRECTOMY WAS PERFORMED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN, BUT WAS NOT PRODUCT RELATED. THE REPORTER INDICATED CULTURES WERE TAKEN AND WERE CONSISTENTLY NEGATIVE. THE RETINA SPECIALIST RECOMMENDED A VITRECTOMY, WHICH WAS PERFORMED ON (B)(6) 2017. THE PATIENTS POST-OP BCVA WAS 20/25. Manufacturer Narrative: DEVICE HISTORY RECORD REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECT DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: CARTRIDGE MODEL SFC-45 FP - LOT NUMBER UNK. CORRECTED DATA: DATE RETURNED TO MFR PREVIOUS DATE IS INCORRECT, CORRECT DATE IS 02/08/2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF WHITE AND CLEAR RESIDUE. METHOD: LENS WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | OS | Other | 01/27/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513720&pc=MTA | 2023826-2015-00119 | 2015-01-15 | Injury | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Vitrectomy; No Code Available | P030016 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS AND AS THE SURGEON WAS ROTATING THE ICL IN THE EYE, THE ICL HIT AND TORE THE NATURAL CAPSULE. THE INCISION WAS ENLARGED TO REMOVE THE LENS WITHIN THE SAME SURGERY AND AN IOL WAS IMPLANTED. A VITRECTOMY WAS PERFORMED BUT SUTURES WERE NOT REQUIRED. THE REPORTER STATED THE SURGEON DID NOT FEEL THERE WAS A PROBLEM WITH THE ICL OR INJECTION SYSTEM. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS INSERTED INTO THE PATIENT'S RIGHT EYE (OD). Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW. RESULTS: PER MEDICAL REVIEW - THIS EVENT WAS CAUSED BY SURGICAL TECHNIQUE RATHER THAN THE MEDICAL DEVICE. ACCORDING TO THE DIRECTION FOR USE (DFU), THE LEARNING CURVE OF ICL IMPLANTATION SURGICAL TECHNIQUE IS DEEP. A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING PROCESS OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. THE MOST LIKELY ROOT CAUSE OF THE EVENT IS SURGEON TECHNIQUE. (B)(4). Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THERE WAS BLACK PARTICLE RESIDUE FLOATING IN THE LIQUID INSIDE THE VIAL. CHECKED BOTH SIDES OF THE LENS OPTIC AND HAPTICS FOR ROUGH/SHARP EDGES. THE EDGES WERE FOUND TO BE ACCEPTABLE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | OD | Other | 01/21/2015 | 2015 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6789687&pc=MTA | 2023826-2017-01302 | 2017-02-27 | Injury | STAAR SURGICAL COMPANY | 2017-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Vitreous Detachment; No Code Available | P030016 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED POSTERIOR VITREAL DETACHMENT AND LATTICING. THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE FACILITY REPORTED THEY AGREED WITH THE PATIENT REPORT OF POSTERIOR VITREAL DETACHMENT AND LATTICING. THE DATE OF THE LAST VISIT WAS (B)(6) 2017 AND THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/20. THE FACILITY REPORTED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE; NO MEDICATIONS AND/OR TREATMENTS WERE REQUIRED AND THERE WERE NO OTHER HEALTH EFFECTS AS A RESULT OF THE OCULAR CONDITION. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | OD | Patient | 07/14/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6789693&pc=MTA | 2023826-2017-01303 | 2017-02-27 | Injury | STAAR SURGICAL COMPANY | 2017-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Vitreous Detachment; No Code Available | P030016 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED POSTERIOR VITREAL DETACHMENT AND LATTICING. THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE FACILITY REPORTED THEY AGREED WITH THE PATIENT REPORT OF POSTERIOR VITREAL DETACHMENT AND LATTICING. THE DATE OF THE LAST VISIT WAS (B)(6) 2017 AND THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/20. THE FACILITY REPORTED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE; NO MEDICATIONS AND/OR TREATMENTS WERE REQUIRED AND THERE WERE NO OTHER HEALTH EFFECTS AS A RESULT OF THE OCULAR CONDITION. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | OS | Patient | 07/14/2017 | 2017 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17200362&pc=MTA | 2023826-2023-02518 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Vitreous Floaters | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT COMPLAINED OF VITREOUS FLOATERS POST-OP AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED (PRE-EXISTING CONDITION OF FLOATERS). THE LENS REMAINED IMPLANTED AND THE EYE IS HEALTHY. CLAIM # (B)(4). | OD | Other Health Care Professional | 05/31/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540585&pc=MTA | 2023826-2017-00741 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Vitreous Floaters; Headache; Blurred Vision; Visual Disturbances; Halo; Discomfort; No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED IRIS DISTORTION, DISCOMFORT, HEADACHES, GLARE/HALOES, BLURRED VISION, DOUBLE VISION OR "GHOSTING", "VITREOUS FLOATERS", PUPIL REMAINS DILATED AND NONRESPONSIVE, AND OTHER MISCELLANEOUS VISUAL DISTURBANCES. THE PATIENT IS NOT HAPPY WITH THE RESULT OF THE IMPLANT. ON (B)(6) 2017 THE PATIENT BEGAN PHYLOCARPINE DROPS TREATMENT AND THE DR. IS MONITORING THE OCCURENCE. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, AND IRIS ATROPHY. (B)(6). (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. "PHYLOCARPINE" PILOCARPINE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | OS | Physician | 04/05/2017 | 2017 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6341778&pc=MTA | 2023826-2017-00244 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Vitreous Floaters; Visual Disturbances | P030016 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THEIR RIGHT EYE (OD). THE PATIENT REPORTED HAD FLOATERS AND WAS BLURRY SOMETIME. THE LENS WAS IMPLANTED ON (B)(6) 2011. THE FACILITY REPORTED THEY CONFIRMED THE PATIENT REPORT AND THE DATE OF LAST VISIT WAS (B)(6) 2014. IT WAS UNKNOWN IF THE EVENT WAS CAUSED BY THE DEVICE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP PROGNOSIS WAS GOOD AND VISUAL ACUITY WAS 20/30-2. Manufacturer Narrative: CONCLUSION: IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) AND PER DFU "INDICATIONS THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE". THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. (B)(4). Manufacturer Narrative: (B)(4). | OD | Patient | 07/25/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483346&pc=QCB | 2023826-2020-01947 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: REPORTER INDICATED SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO PATIENTS EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.00 DIOPTER INTO THE PATIENT RIGHT EYE ON AN UNKNOWN DATE. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON AN UNKNOWN DATE FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON. D4- MODEL#: VICMO 13.2. SERIAL #: (B)(6). EXPIRATION DATE: 30-APR-2023 D6A- UNK. D6B-UNK. H1- SERIOUS INJURY. H6-IMPACT CODE 4629: LENS WAS EXCHANGED. H6-METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.00 DIOPTER INTO THE PATIENT RIGHT EYE ON (B)(6) 2020. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON D6A: ON (B)(6) 2020, D6B: ON (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY. VISUAL INSPECTION FOUND HAPTIC TORN WITH THREADS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. H6 -TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | OD |  | 08/11/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483351&pc=QCB | 2023826-2020-01948 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: REPORTER INDICATED SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION B1- ADVERSE EVENT B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.50 DIOPTER INTO THE PATIENT LEFT EYE ON AN UNKNOWN DATE. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON AN UNKNOWN DATE FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON. D4- MODEL#: VICMO 13.2. SERIAL #: (B)(6),. EXPIRATION DATE: 30-APR-2023 H1- SERIOUS INJURY H6-IMPACT CODE 4629: LENS WAS EXCHANGED. H6-METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.50 DIOPTER INTO THE PATIENT LEFT EYE ON (B)(6) 2020. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON D6A:10-AUG-2020. D6B: 31-AUG-2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL WIDTH VERIFICATION INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. H6 -TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS (B)(4): DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | OS |  | 08/11/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778982&pc=QCB | 2023826-2021-01323 | 2021-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6 -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WERE OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PRESCRIBED DISTANCE GLASSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PATIENT WAS PRESCRIBED DISTANCE GLASSES. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6 - THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION ON THE OPTOCRAFT FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING; THE SPHERE WAS OUT OF TOLERANCE BY 0.19D. H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #719586. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). CLAIM# (B)(4). | OS | Physician | 04/13/2021 | 2021 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779013&pc=QCB | 2023826-2021-01322 | 2021-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6 -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, "PRESCRIBED DISTANCE GLASSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PATIENT WAS PRESCRIBED DISTANCE GLASSES. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6 - LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION ON THE OPTOCRAFT FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VAULES MEASURED AT THE TIME OF MANUFACTURING; THE SPHERE WAS OUT OF TOLERANCE BY 0.08D. H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: H3 - RESULT CODE: 114 IN SUPPLEMENTAL MDR #3 CLAIM # (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK. CLAIM# (B)(4). | OD | Physician | 04/13/2021 | 2021 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140273&pc=QCB | 2023826-2021-02413 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 TH LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED " ON THE FIRST DAY AFTER OPERATION, THE VISUAL ACUITY WAS 0.6 AND THE CORNEAL EDEMA WAS MILD." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM,VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED " ON THE FIRST DAY AFTER OPERATION, THE VISUAL ACUITY WAS 0.6 AND THE CORNEAL EDEMA WAS MILD." NO SPECIAL TREATMENT WAS NEEDED FOR THE EDEMA. THE EDEMA DISAPPEARED THE NEXT DAY, AND THERE WAS NO SIGNIFICANT CHANGE BETWEEN CORNEAL ENDOTHELIUM AND PREOPERATIVE EXAMINATION. CLAIM#: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | OS | Physician | 06/08/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13160463&pc=QCB | 2023826-2022-00077 | 2021-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7; -14.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME WERE ASSESSED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | OD | Physician | 12/30/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17542184&pc=QCB | 2023826-2023-03515 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -2.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION (PT. RIGHT EYE FIELD IS BLURRED) AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45 S OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | OD | Physician | 07/17/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882069&pc=QCB | 2023826-2023-04269 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF-4.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2023. BLURRED VISION; SEVERE DISTENDING SENSATION;PAIN; DIZZINESS WAS REPORTED AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS REPOSITIONED THEN ON SAME DAY REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: H6-CLINICAL CODE 4581- SEVERE DISTENDING SENSATION; DIZZINESS. CLAIM# (B)(4). | OD | Physician | 09/08/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882376&pc=QCB | 2023826-2023-04270 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF-4.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE(OS) ON 07-JUL-2023. BLURRED VISION; SEVERE DISTENDING SENSATION;PAIN; DIZZINESS WAS REPORTED AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS REPOSITIONED THEN ON SAME DAY REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLINICAL CODE 4581- SEVERE DISTENDING SENSATION; DIZZINESS. CLAIM# (B)(4). | OS | Physician | 09/08/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18430037&pc=MTA | 2023826-2024-00047 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON 11-SEPT-2023. REPORTER STATES THE WRONG LENS WAS IMPLANTED. ON 13-SEPT-2023 REPORTER STATES THE "GOOD LENS IMPLANTED." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2024-00046 FOR INITIAL LENS. Manufacturer Narrative: CLAIM# (B)(4). | OD | Physician | 12/05/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16981458&pc=QCB | 2023826-2023-01954 | 2023-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Optical Problem | Inflammation; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2;-5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTES THE LENS HAS SOME DIFFUSE WHITISH PLAQUES/FILM ON THE LENS, WHICH LOOKS LIKE IT'S ON THE POSTERIOR SURFACE OF THE EVO LENS AND NOT THE LENS OR CORNEA. THE SURGEON ALSO INDICATED THERE IS A BIT MORE INFLAMMATION THEN TYPICAL, THE WHITISH FILM/PLAQUES APPEAR STABLE, NEITHER INCREASING OR DECREASING AND THERE SEEMS TO BE NO CONTINUED INFLAMMATION. THE SURGEON ALSO REPORTS GLARE/HALOES, BLURRED VISION, AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY, A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT MODEL/POWER WAS IMPLANTED. IT WAS NOTED "VISION IMPROVED SIGNIFICANTLY WITH EXCHANGE". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "SUSPECT THIS WAS INFLAMMATORY DEBRIS/FIBRING BUILDING UP ON THE LENS". Manufacturer Narrative: ADDITIONAL INFORMATION: INITIAL MDR IS NOT APPLICABLE AS THERE WAS NO DEVICE PROBLEM, REPORTER LATER CLARIFIED THE ISSUE WAS DIRECTLY PATIENT RELATED. CLAIM#: (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT CLINICAL CODE: 4581- WHITISH PLAQUE/FILM. H6- TYPE OF INVESTIGATION: 4110- LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | OS | Physician | 04/21/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13362153&pc=QCB | 2023826-2022-00348 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7;-11.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; AND ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | OD | Physician | 01/20/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16465175&pc=MTA | 2023826-2023-00767 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6:1494-OFF LABEL: PROGRESSIVE MYOPIA. D9: RETURN DATE- 05-APR-2023 CLAIM# (B)(4). | OD | Other Health Care Professional | 01/31/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595289&pc=MTA | 2023826-2023-00985 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1494-OFF LABEL- PROGRESSIVE MYOPIA D9: RETURN DATE: 05-APR-2023 (B)(4) | OS | Other Health Care Professional | 02/21/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548152&pc=QCB | 2023826-2020-02096 | 2020-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND A LONGER LENGTH LENS WAS LATER IMPLANTED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | OS |  | 08/24/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021855&pc=QCB | 2023826-2021-01996 | 2020-02-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "ICL WAS TOO LONG DESPITE STAAR SIZING RECOMMENDATION." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | OS | Physician | 05/26/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148266&pc=QCB | 2023826-2021-02478 | 2021-05-13 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER GOT CAUGHT IN THE CARTRIDGE AND WAS DAMAGED DURING LOADING. THIS OCCURRED ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. AN ALTERNATE LENS WAS IMPLANTED ON THE SAME DATE AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | NA | Physician | 06/18/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178444&pc=QCB | 2023826-2021-02709 | 2021-04-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER GOT CAUGHT IN THE CARTRIDGE AND WAS DAMAGED DURING LOADING. THIS OCCURRED ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. AN ALTERNATE LENS WAS IMPLANTED ON THE SAME DATE AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | NA | Physician | 06/21/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14707154&pc=QCB | 2023826-2022-01885 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS NOTED THAT "DESPITE DR. ISKANDER RECOMMENDING A TORIC ICL; SHE ELECTED AN SPHERICAL ICL. SHE NOW WANTS A TORIC ICL AND HAVE -1.00 TARGET IN BOTH EYES". THERE IS A PLANNED LENS EXCHANGE. LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL_12.6; -10.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IT WAS NOTED THAT "DESPITE DR. ISKANDER RECOMMENDING A TORIC ICL; SHE ELECTED AN SPHERICAL ICL. SHE NOW WANTS A TORIC ICL AND HAVE -1.00 TARGET IN BOTH EYES". ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL (TORIC) LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR WITH NO DEVICE ISSUE. CLAIM# (B)(4). | OD | Other Health Care Professional | 05/16/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15345336&pc=QCB | 2023826-2022-03008 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: REPORTEDLY, A 12.6MM VICM5_12.6-7.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 AND REPLACED WITH A SHORTER LENGTH LENS ON AN UNKNOWN DATE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: EXPIRATION DATE: 07/31/2023. (B)(4). CLAIM#: (B)(4). | NA | Other Health Care Professional | 08/03/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989738&pc=QCB | 2023826-2023-01962 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-12.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | OD | Physician | 05/19/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313132&pc=QCB | 2023826-2023-02811 | 2022-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-7.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | OD | Physician | 06/15/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733288&pc=QCB | 2023826-2023-04026 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER-9.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CAIM # (B)(4). | OS | Physician | 09/11/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770216&pc=QCB | 2023826-2023-04056 | 2023-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-9.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | OS | Physician | 09/13/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17795902&pc=QCB | 2023826-2023-04161 | 2023-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-9.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | OD | Physician | 08/23/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17969244&pc=QCB | 2023826-2023-04661 | 2023-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER-12.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | OD | Physician | 10/18/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422405&pc=QCB | 2023826-2023-05758 | 2023-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER-8.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | OS | Physician | 12/19/2023 | 2023 | 28 | Female